https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16281080&pc=QCB | 2023826-2023-00394 | 2022-09-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -9.5/1.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -9.5/1.0/090 | OS | Physician | 01/30/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16281402&pc=QCB | 2023826-2023-00316 | 2022-08-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -3.5/1.0/097 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULTING WAS OBSERVED. Manufacturer Narrative: PATIENT DEMOGRAPHICS UNK. CLAIM#: (B)(4). Manufacturer Narrative: B5: LENS WAS REPLACED ON AN UNKNOWN DATE. CLAIM# (B)(4). Manufacturer Narrative: B5 - LENS WAS NOT REMOVED OR REPLACED. PROBLEM WAS RESOLVED. PATIENT IS HAPPY. HE WILL BE CLOSELY MONITORED. CLAIM# (B)(4). | -3.5/1.0/097 | OD | Physician | 01/04/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16281403&pc=QCB | 2023826-2023-00315 | 2023-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030017 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF A -16.5/2.0/112 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A REPLACEMENT LENS. REPORTEDLY, "HAPTIC WAS TORN DURING IMPLANTATION. THE DOCTOR DID NOT EXPLANT IT RIGHT AWAY." CAUSE OF THE EVENT WAS USER ERROR. THE PROBLEM WAS RESOLVED. "PATIENT IS HAPPY." Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -16.5/2.0/112 | OD | Physician | 01/04/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16284125&pc=QCB | 2023826-2023-00328 | 2022-10-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6 MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -9.50/3.5/096 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS BUT DIFFERENT POWER DUE TO LENS ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - YELLOW BAG RETURNED EMPTY, NO LENS INSIDE VIAL. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -9.50/3.5/096 | OS | Physician | 01/03/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16284301&pc=QCB | 2023826-2023-00329 | 2022-11-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1 MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -11.0/2.0/084 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL LONGER LENGTH LENS AND DIFFERENT POWER DUE TO LENS ROTATION. THE EXCHANGE LENS WAS IMPLANTED VERTICALLY AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -11.0/2.0/084 | OD | Physician | 01/03/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16284380&pc=QCB | 2023826-2023-00330 | 2022-11-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1 MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -12.50/2.0/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL LONGER LENGTH LENS AND DIFFERENT POWER DUE TO LENS ROTATION. THE EXCHANGE LENS WAS IMPLANTED VERTICALLY AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN WITH FIBERS ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -12.50/2.0/087 | OS | Physician | 01/03/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16285221&pc=QCB | 2023826-2023-00360 | 2022-12-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+2.0/075 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -8.5/+2.0/075 | OD | Physician | 01/03/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16285233&pc=QCB | 2023826-2023-00361 | 2022-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+3.0/099 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. DUE TO LOW VAULT THE LENS WAS REMOVED ON (B)(6) 2022. LATER ON (B)(6) 2022 A LONGER LENGTH LENS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -10.0/+3.0/099 | OD | Physician | 01/03/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16287273&pc=QCB | 2023826-2023-00366 | 2022-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+1.0/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -12.0/+1.0/090 | OS | Physician | 01/04/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16287274&pc=QCB | 2023826-2023-00362 | 2022-10-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.5/081 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: A2 : CORRECTED TO: (B)(6) 1999 IN THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | -7.0/+1.5/081 | OD | Physician | 01/04/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16287276&pc=QCB | 2023826-2023-00365 | 2022-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.0/080 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5 - CORRECTED TO: "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.0/085 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT." IN THE INITIAL MDR. CLAIM # (B)(4). | -11.5/+1.0/080 | OD | Physician | 01/04/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16287277&pc=QCB | 2023826-2023-00364 | 2022-11-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+6.0/082 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: (B)(4). | -12.5/+6.0/082 | OS | Physician | 01/04/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16287724&pc=QCB | 2023826-2023-00363 | 2022-11-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+3.0/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: (B)(4). | -9.0/+3.0/095 | OD | Physician | 01/04/2023 | 2023 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16287875&pc=QCB | 2023826-2023-00415 | 2022-11-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -9/+1.5/57 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE DUE TO THE LENS HAD BECAME STUCK IN THE CARTRIDGE/INJECTION SYSTEM. INTRAOPERATIVELY THE LENS WAS REMOVED WITH NO PATIENT INJURY. ON (B)(6) 2022 A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -9/+1.5/57 | OS | Physician | 01/25/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16300829&pc=QCB | 2023826-2023-00441 | 2023-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -14.50/2.5/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS LOW VAULTING WITH ROTATION. ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON THIS SAME DATE IN A SEPARATE SURGERY A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED "USER DECIDED TO USE 13.2 SMALLER SIZE THAN OCOS RECOMMENDATION OF 13.7". Manufacturer Narrative: (B)(4). | -14.50/2.5/095 | OD | Physician | 01/23/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301004&pc=QCB | 2023826-2023-00437 | 2022-04-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.0/4.5/87 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). | -14.0/4.5/87 | OS | Physician | 01/05/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301034&pc=QCB | 2023826-2023-00439 | 2022-04-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.5/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -9.5/2.0/089 | OD | Physician | 01/05/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301372&pc=QCB | 2023826-2023-00378 | 2022-10-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -14.5/2.5/073 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6)2022, THE LENS WAS EXCHANGED FOR A SAME SIZE LENS, IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT. REPORTEDLY, "THE LENS WAS PLACED VERTICALLY." THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: LENS RETURNED 20-FEB-2023. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND THE HAPTIC DEFORMED, BENT, AND BROKEN WITH RESIDUE/DEBRIS ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNK. (B)(4). | -14.5/2.5/073 | OD | Physician | 01/09/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301410&pc=QCB | 2023826-2023-00438 | 2022-12-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -5.50/2.0/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVEN WAS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -5.50/2.0/098 | OD | Physician | 01/06/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301481&pc=QCB | 2023826-2023-00408 | 2022-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+2.0/109 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -12.5/+2.0/109 | OD | Physician | 01/05/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301496&pc=QCB | 2023826-2023-00407 | 2022-07-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+2.5/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE LENS WAS IMPLANTED VERTICALLY AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | -11.5/+2.5/091 | OS | Physician | 01/05/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301497&pc=QCB | 2023826-2023-00409 | 2022-11-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+1.5/107 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -14.0/+1.5/107 | OS | Physician | 01/05/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301518&pc=QCB | 2023826-2023-00410 | 2022-11-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.5/+1.0/101 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -15.5/+1.0/101 | OD | Physician | 01/05/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301773&pc=QCB | 2023826-2023-00406 | 2022-12-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.0/+2.5/075 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -16.0/+2.5/075 | OD | Physician | 01/05/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16301928&pc=QCB | 2023826-2023-00389 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -10.50/2.5/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. LENS ROTATION WAS REPORTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH. NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -10.50/2.5/090 | OS | Physician | 01/06/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16303732&pc=QCB | 2023826-2023-00340 | 2023-01-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.5/1.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. CLAIM # (B)(4). | -7.5/1.0/086 | OD | Physician | 01/16/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16303733&pc=QCB | 2023826-2023-00341 | 2023-01-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/1.0/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. CLAIM # (B)(4). | -11.0/1.0/108 | OS | Physician | 01/16/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16303739&pc=QCB | 2023826-2023-00342 | 2023-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.0/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. CLAIM# (B)(4). | -9.0/2.5/091 | OS | Physician | 01/16/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16303743&pc=QCB | 2023826-2023-00343 | 2022-12-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 -11.0/0.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, UNDER 21 YRS OF AGE AT DATE OF IMPLANT. CLAIM# (B)(4). | -11.0/0.5/086 | OD | Physician | 01/16/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16303745&pc=QCB | 2023826-2023-00349 | 2023-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.0/1.0/013 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM # (B)(4). | -10.0/1.0/013 | OS | Physician | 01/17/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16303761&pc=QCB | 2023826-2023-00351 | 2023-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.0/1.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. CLAIM# (B)(4). | -13.0/1.5/094 | OD | Physician | 01/20/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16303763&pc=QCB | 2023826-2023-00348 | 2023-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.5/2.5/167 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT: UNK. PATIENT ETHNICITY: UNK. PATIENT RACE: UNK. (B)(4). | -10.5/2.5/167 | OD | Physician | 01/17/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16326924&pc=QCB | 2023826-2023-00268 | 2022-11-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.00/+2.5/180 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED SO THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH A LONGER LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT: UNK. PATIENT ETHNICITY: UNK. PATIENT RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). Manufacturer Narrative: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT, WITH FIBERS ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -13.00/+2.5/180 | OS | Physician | 01/09/2023 | 2023 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16326974&pc=QCB | 2023826-2023-00267 | 2022-11-23 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+2.5/075 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: CLAIM#(B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.50/+2.5/075 | OS | Physician | 01/12/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16326993&pc=QCB | 2023826-2023-00264 | 2022-10-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.50/+1.5/116 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS REPORTED AS HIGH VAULT (MORE THAN 2CT) AND REMAINED IMPLANTED. THE PATIENT IS STABLE (VISION AND IOP WISE) AND IS UNDER CLOSE FOLLOW UP. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.50/+1.5/116 | OD | Physician | 01/12/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16329541&pc=QCB | 2023826-2023-00463 | 2022-10-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS -9.00/2.0/081 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.00/2.0/081 | OS | Physician | 01/09/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16329546&pc=QCB | 2023826-2023-00462 | 2022-12-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -11.00/1.5/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN/SPLIT AND HAPTIC BROKEN/BENT. LENS RETURNED DAMAGED. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO DAMAGED STATE OF LENS. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -11.00/1.5/094 | OS | Physician | 01/09/2023 | 2023 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16329914&pc=QCB | 2023826-2023-00270 | 2022-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+1.0/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. H6 - INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: PATIENT WEIGHT : UNK. PATIENT ETHNICITY: UNK. PATIENT RACE: UNK. DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). | -09.00/+1.0/074 | OS | Physician | 01/10/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16330097&pc=QCB | 2023826-2023-00397 |  | Injury | STAAR SURGICAL COMPANY | 2023-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY, "ICL OS WAS TO BE EXCHANGED FOR A SMALLER LENS, HOWEVER AFTER EXPLANTATION OF ORIGINAL LENS, NEW LENS WAS UNABLE TO BE IMPLANTED DUE TO FLOPPY IRIS AND POSSIBLE ZONULAR INSUFFICIENCY." CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. REPORTEDLY, "ICL OS WAS TO BE EXCHANGED FOR A SMALLER LENS, HOWEVER AFTER EXPLANTATION OF ORIGINAL LENS, NEW LENS WAS UNABLE TO BE IMPLANTED DUE TO FLOPPY IRIS AND POSSIBLE ZONULAR INSUFFICIENCY." H6: 1494-OFF LABEL: AGE<21 YEARS OLD AT THE TIME OF IMPLANTATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -9.0/1.0/095 | OS | Physician | 02/06/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16330342&pc=QCB | 2023826-2023-00450 | 2022-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.0/+2.0/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -17.0/+2.0/091 | OS | Physician | 01/09/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16330343&pc=QCB | 2023826-2023-00449 | 2022-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/+1.0/073 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -6.0/+1.0/073 | OS | Physician | 01/09/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16330344&pc=QCB | 2023826-2023-00451 | 2022-12-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.0/076 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.0/+1.0/076 | OS | Physician | 01/09/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16330428&pc=QCB | 2023826-2023-00398 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/080 (SPHERE/CYLINDER/AXIS) WAS INTRAOPERATIVELY IMPLANTED AND REMOVED AS REPLACEMENT LENS FOR A PATIENT'S LEFT EYE (OS) ON (B)(6)2022. REPORTEDLY, "ICL OS WAS TO BE EXCHANGED FOR A SMALLER LENS, HOWEVER AFTER EXPLANTATION OF ORIGINAL LENS, NEW LENS WAS UNABLE TO BE IMPLANTED DUE TO FLOPPY IRIS AND POSSIBLE ZONULAR INSUFFICIENCY." THE CAUSE OF THE EVENT IS UNKNOWN. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: B1 - CORRECTED TO: PRODUCT PROBLEM (E.G., DEFECTS/MALFUNCTION) IN THE INITIAL MDR B2 - REQUIRED INTERVENTION TO PREVENT PERMANENT IMAPIRMENT/DAMAGE SHOULD BE REMOVED FROM THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -9.0/1.0/080 | OS | Physician | 02/06/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16330430&pc=QCB | 2023826-2023-00452 | 2022-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+2.5/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: (B)(4). | -8.0/+2.5/091 | OS | Physician | 01/09/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16334682&pc=QCB | 2023826-2023-00466 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -13.00/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). SURGEON OBSERVED A LOW VAULT ONE DAY POST OP. PER UPDATED INFORMATION THE PATIENT RETURNED FOR REVIEW/FOLLOW-UP AND THE VAULT WAS 285MICRONS. THE REPORTER STATES "WE DO NOT HAVE ANY COMPLAINTS FOR THIS PATIENT ANYMORE. BOTH PATIENT AND SURGEON ARE HAPPY". Manufacturer Narrative: IMPLANT DATE:UNK. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.00/2.0/089 | OD | Physician | 01/10/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16342928&pc=QCB | 2023826-2023-00454 | 2022-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.5/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -10.0/+2.5/090 | OS | Physician | 01/10/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16343059&pc=QCB | 2023826-2023-00453 | 2022-08-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.5/078 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+1.5/078 | OS | Physician | 01/10/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16344320&pc=QCB | 2023826-2023-00401 | 2022-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -9.5/4.0/79 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THE REPOSITION DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022, THE LENS WAS EXCHANGED WITH A REPLACEMENT LENS, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THE LENS MODEL SIZE RETURNED IS NOT LENS MODEL SIZE STATED IN CLAIM. DIMENSIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. LENS MODEL SIZE RETURNED IS NOT LENS MODEL SIZE STATED IN CLAIM. DIMENSIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | -9.5/4.0/79 | OS | Physician | 01/11/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16344321&pc=QCB | 2023826-2023-00400 | 2022-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -13.0/3.5/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THE REPOSITION DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022, THE LENS WAS REMOVED AND IN A SUBSEQUENT SURGERY LATER THAT DAY, A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. DIMENSIONAL INSPECTION FOUND THE LENS MODEL SIZE RETURNED IS NOT THE LENS MODEL SIZE STATE IN THE CLAIM. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: OFF LABEL - AGE <21. (B)(4). | -13.0/3.5/90 | OS | Physician | 01/11/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16344745&pc=QCB | 2023826-2023-00455 | 2023-01-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/+0.5/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) AND (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -6.0/+0.5/094 | OS | Physician | 01/12/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16344746&pc=QCB | 2023826-2023-00456 | 2022-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.0/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -10.0/+1.0/092 | OS | Physician | 01/13/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16345206&pc=QCB | 2023826-2023-00481 | 2022-10-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Appropriate Term/Code Not Available | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -10.0/1.5/074 (SPHERE/CYLINDER/AXIS) TORE/ BROKE DURING INJECTION/ DELIVERY INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS DEVICE. DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -10.0/1.5/074 | OS | Other Health Care Professional | 01/10/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16345231&pc=QCB | 2023826-2023-00482 | 2022-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/149 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LOW VAULT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS WITH DIFFERENT AXIS AND THE REPLACEMENT LENS RESOLVED THE PROBLEM. IT WAS NOTED THAT THE INITIAL LENS WAS VERTICAL IMPLANTED. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) SHOULD BE ADDED TO THE "INITIAL" MDR. CLAIM #(B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC BROKEN/BENT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -12.0/1.5/149 | OS | Physician | 01/10/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16351465&pc=QCB | 2023826-2023-00471 | 2022-09-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -15.00/1.00/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS WITHIN THE SAME SURGERY DUE TO OBSERVATION OF EXCESSIVE VAULT AND THE PROBLEM RESOLVED. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0/+1.0/92 (SPHERE/CYLINDER/AXIS) INTO THE RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT. THE LENS WAS IMPLANTED; REMOVED; AND REPLACED WITH THE SAME MODEL/ POWER, SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -15.00/1.00/092 | OD | Physician | 01/13/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16351488&pc=QCB | 2023826-2023-00468 | 2023-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -12.5/2.0/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). | -12.5/2.0/095 | OD | Physician | 01/11/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16351662&pc=QCB | 2023826-2023-00470 | 2022-05-31 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -7.50/2.5/108 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED AND REPLACED WITH A SAME LENGTH LENS WITHIN THE SAME SURGERY DUE TO OBSERVATION OF EXCESSIVE VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+2.5/108 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT. THE LENS WAS IMPLANTED; REMOVED; AND REPLACED WITH A DIFFERENT TORIC MODEL; SAME POWER AND LENGTH AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5-THE REPORTER INDICATED THAT THE SURGEON NOTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+2.5/108 (SPHERE/CYLINDER/AXIS) REPORTEDLY "LENS OPENED BUT DID NOT PUT IT IN" ON (B)(6) 2023. THERE WAS NO PATIENT CONTACT. THE LENS WAS NOT IMPLANTED. A REPLACEMENT LENS OF A DIFFERENT TORIC MODEL, SAME POWER AND LENGTH WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM IS NOT MDR REPORTABLE BUT INITIAL REPORT WAS ALREADY SUBMITTED. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.50/2.5/108 | OD | Physician | 01/13/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16358875&pc=QCB | 2023826-2023-00493 | 2022-12-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -12.00/+2.0/081 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON REPORTS LENS WAS IMPLANTED UPSIDE DOWN. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS AND THIS RESOLVED THE PROBLEM. THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/+2.0/081 | OD | Physician | 02/09/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16358877&pc=QCB | 2023826-2023-00495 | 2023-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -15.0/+3.5/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON REPORTS THE LENS TORE DURING INJECTION/DELIVERY INTO THE EYE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS AND THIS RESOLVED THE PROBLEM. THERE WAS NO PATIENT INJURY. Manufacturer Narrative: (B)(4). | -15.0/+3.5/095 | OD | Physician | 02/09/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16358880&pc=QCB | 2023826-2023-00494 | 2022-12-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -11.50/+1.5/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS LENS WAS IMPLANTED UPSIDE DOWN. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS AND THIS RESOLVED THE PROBLEM. THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.50/+1.5/093 | OS | Physician | 02/09/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16359065&pc=QCB | 2023826-2023-00486 | 2022-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/171 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LOW VAULT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS WITH DIFFERENT AXIS AND THE REPLACEMENT LENS RESOLVED THE PROBLEM. IT WAS NOTED THAT THE INITIAL LENS WAS VERTICAL IMPLANTED. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -12.0/1.5/171 | OS | Physician | 01/12/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16359304&pc=QCB | 2023826-2023-00460 | 2022-05-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+3.0/128 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LENS TWO SIZES LONGER IN LENGTH DUE TO LOW VAULT AND LENS ROTATION. THE REPLACEMENT LENS WAS IMPLANTED VERTICALLY. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.5/+3.0/128 | OS | Physician | 01/15/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16359305&pc=QCB | 2023826-2023-00458 | 2022-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.0/102 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -11.0/+1.0/102 | OS | Physician | 01/13/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16359307&pc=QCB | 2023826-2023-00459 | 2022-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.0/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+1.0/093 | OS | Physician | 01/13/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16364164&pc=QCB | 2023826-2023-00417 | 2022-10-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.0/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PT INFORMATION: UNK. CLAIM# (B)(4). | -12.0/2.5/093 | OD | Physician | 01/26/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16364171&pc=QCB | 2023826-2023-00418 | 2022-10-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6, -11.0/3.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). | -11.0/3.0/087 | OS | Physician | 01/26/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16364172&pc=QCB | 2023826-2023-00419 | 2023-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -10.5/1.5/064 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2023 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). | -10.5/1.5/064 | OS | Physician | 01/26/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16364185&pc=QCB | 2023826-2023-00425 | 2022-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.0/3.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.0/3.0/087 | OD | Physician | 01/29/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16364186&pc=QCB | 2023826-2023-00427 | 2022-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -15.0/5.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. CLAIM # (B)(4). | -15.0/5.5/083 | OS | Physician | 01/29/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16364192&pc=QCB | 2023826-2023-00424 | 2022-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.0/3.0/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFORMATION: UNK. CLAIM # (B)(4). | -11.0/3.0/070 | OS | Physician | 01/29/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16365725&pc=QCB | 2023826-2023-00525 |  | Injury | STAAR SURGICAL COMPANY | 2023-02-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. IT WAS REPORTED THAT THE LENS HAD ROTATED; SURGEON REPOSITIONED THE LENS; BUT THE LENS HAD ROTATED AGAIN. LENS EXCHANGE IS PLANNED. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER CLARIFIED THAT THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -7.50/1.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. IT WAS REPORTED THAT THE LENS HAD ROTATED; SURGEON REPOSITIONED THE LENS; BUT THE LENS HAD ROTATED AGAIN. LENS EXCHANGE IS PLANNED. LENS REMAINS IMPLANTED. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -7.50/1.5/090 | NA | Physician | 12/29/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16368749&pc=QCB | 2023826-2023-00477 | 2022-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -9.5/1.5/085 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "BOTH RESULTS OF THE TYPES OF CALCULATIONS IN CASIA2 WERE 12.6, BUT IN PRACTICE 13.2 WAS APPROPRIATE." Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 15-FEB-2023. H3-DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH FOREIGN MATERIAL ON THE LENS SURFACE, SPECIFICALLY, FIBERS. VISUAL INSPECTION FOUND THE OPTIC SPLIT. DIMENSIONAL INSPECTIONS WERE UNABLE TO BE PERFORMED DUE TO SIGNIFICANT LENS DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE:UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | -9.5/1.5/085 | OS | Physician | 01/17/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16368759&pc=QCB | 2023826-2023-00478 | 2022-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF A -4.0/1.0/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, AND DISCREPANCY BETWEEN WTW AND MONOGRAM. THE PROBLEM WAS RESOLVED. REPORTEDLY, "LENS OPACITY (MYDRIASIS WAS NOT MAXIMAL AT THE TIME OF ICL REMOVAL + ICL SIZE WAS 13.2, WHICH WAS LARGE AND TOOK TIME AND EFFORT TO REMOVE. ALSO, THE LENS BECAME CLOUDY DUE TO THE HIGH VISCOSITY OF THE LENS WHEN USING THE SHELL GUN.) THE PATIENT EYE IS STILL UNSTABLE POSTOPERATIVELY AND WE ARE STILL DISCUSSING THE FUTURE WITH THE PATIENT." THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: WEIGHT:UNK, ETHNICITY:UNK, RACE:UNK. CLAIM# (B)(4). | -4.0/1.0/93 | OD | Physician | 01/17/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16368902&pc=QCB | 2023826-2023-00540 | 2023-02-01 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -8.5/2.0/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THE LENS BECAME STUCK IN THE CARTRIDGE/INJECTION SYSTEM. INTRAOPERATIVELY THE LENS WAS REMOVED FROM THE PATIENT'S EYE WITH NO INJURY. ON (B)(6) 2023 A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -8.5/2.0/094 | OD | Physician | 02/09/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16369388&pc=QCB | 2023826-2023-00528 | 2022-10-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+3.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.0/+3.0/091 | OD | Physician | 01/16/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16369389&pc=QCB | 2023826-2023-00527 | 2023-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.5/088 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.0/+2.5/088 | OD | Physician | 01/16/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16369390&pc=QCB | 2023826-2023-00526 | 2022-05-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+2.5/069 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.0/+2.5/069 | OD | Physician | 01/15/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16375378&pc=QCB | 2023826-2023-00544 | 2022-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/029 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | -10.00/+1.0/029 | OS | Physician | 01/17/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16376168&pc=QCB | 2023826-2023-00552 | 2022-10-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.0/4.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS THE LENS WAS IMPLANTED UPSIDE DOWN. INTRAOPERATIVELY THE LENS WAS REMOVED WITH NO PATIENT INJURY. ON (B)(6) 2022 A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/4.5/089 | OS | Physician | 02/12/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16376451&pc=QCB | 2023826-2023-00533 | 2022-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+1.5/086 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -7.5/+1.5/086 | OS | Physician | 01/16/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16376452&pc=QCB | 2023826-2023-00530 | 2022-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+2.0/099 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A DIFFERENT LENS OF LONGER LENGTH DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: (B)(4). | -14.0/+2.0/099 | OD | Physician | 01/16/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16376453&pc=QCB | 2023826-2023-00531 | 2022-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+2.0/084 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A DIFFERENT LENS OF LONGER LENGTH DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -13.5/+2.0/084 | OS | Physician | 01/16/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16376454&pc=QCB | 2023826-2023-00532 | 2022-11-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+2.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A DIFFERENT LENS OF LONGER LENGTH DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: (B)(4). | -11.5/+2.0/091 | OD | Physician | 01/16/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16379720&pc=QCB | 2023826-2023-00513 | 2023-01-18 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -5.0/2.5/101 (SPHERE/CYLINDER/AXIS) WAS PARTIALLY INSERTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS PLANNED TO RELOAD, "BUT IT DROPPED SOMEWHERE AND WAS NEVER FOUND." THE BACKUP LENS WAS IMPLANTED SUCCESSFULLY. THERE WAS NO PATIENT HARM. Manufacturer Narrative: AGE, WEIGHT AND ETHNICITY:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS. CLAIM# (B)(4). | -5.0/2.5/101 | OS | Physician | 01/18/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16380510&pc=QCB | 2023826-2023-00514 | 2022-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.5/3.5/133 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME SIZE LENS IMPLANTED VERTICALLY DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, THE STATUS OF THE EYE IS, "OK." Manufacturer Narrative: PT INFORMATION: UNK. OFF-LABEL ACD<3.0. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -10.5/3.5/133 | OS | Physician | 01/17/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16380512&pc=QCB | 2023826-2023-00515 | 2022-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5/1.5/125 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2022, THE LENS WAS EXCHANGED FOR A SAME SIZE LENS IMPLANTED VERTICALLY DUE TO LOW VAULT WITHOUT ROATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, THE STATUS OF THE EYE IS, "OK." Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -11.5/1.5/125 | OS | Physician | 01/17/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16383018&pc=QCB | 2023826-2023-00546 | 2022-12-31 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00/+2.0/071 (SPHERE/CYLINDER/AXIS) DURING THE SAME SURGERY DUE TO THE LENS WAS IMPLANTED UPSIDE DOWN IN THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS DAMAGED WHEN REMOVED. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AT A LATER DATE AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK RACE: UNK WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -15.00/+2.0/071 | OS | Physician | 01/17/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16383021&pc=QCB | 2023826-2023-00545 | 2022-10-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.5/015 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -09.50/+1.5/015 | OD | Physician | 01/17/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16383258&pc=QCB | 2023826-2023-00590 | 2023-01-11 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -7.5/+2.5/80 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS TORE DURING INJECTION/DELIVERY INTO THE EYE. INTRAOPERATIVELY THE LENS WAS REMOVED WITH NO PATIENT INJURY. ON (B)(6) 2023 A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED "THE ICL WAS BROKEN WHEN THE DOCTOR IMPLANTED INTO THE EYE". Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(6). | -7.5/+2.5/80 | OS | Physician | 02/14/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16383725&pc=QCB | 2023826-2023-00534 | 2022-12-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.5/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+1.5/091 | OD | Physician | 01/17/2023 | 2023 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16383795&pc=QCB | 2023826-2023-00490 | 2023-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -11.0/4.5/102 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS ROTATION WAS OBSERVED. THE LENS WAS "ADJUSTED ONE TIME AND ROTATION HAPPENED AGAIN." THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -11.0/4.5/102 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS ROTATION WAS OBSERVED. THE LENS WAS "ADJUSTED ONE TIME AND ROTATION HAPPENED AGAIN." ON (B)(6) 2023 THE EXCHANGE LENS OF LARGER SIZE WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. D6B- (B)(6) 2023. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -11.0/4.5/102 | OD | Physician | 01/16/2023 | 2023 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16383815&pc=QCB | 2023826-2023-00491 | 2023-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF -11.0/6.0/001 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON 03-SEPT-2022. LENS ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: D4-SERIAL NUMBER: SHOULD BE CORRECT TO STATE (B)(6). CLAIM# (B)(6). Manufacturer Narrative: H10 - THIS SUPPLEMENTAL IS FOR N/A SUPPLEMENTAL# 2 REPORT NUMBER 2023826-2023-00491 AS WRONG PE WAS SUBMITTED DUE TO TYPO. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -11.0/6.0/001 | OS | Physician | 01/16/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16385279&pc=QCB | 2023826-2023-00556 | 2023-01-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -14.00/4.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS OF SHORTER LENGTH AND DIFFERENT POWER DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATED " THE DOCTOR DECIDED TO IMPLANT SPHERICAL LENS INSTEAD OF TORIC LENS AND CORRECT ASTIG (ASTIGMATISM) WITH PRK ON TOP". CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED INFORMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+4.0/92 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT, LENS ROTATION AND PRK WAS PERFORMED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SPHERICAL MODEL AND POWER; SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION PROVIDED: "THE DOCTOR DECIDED TO IMPLANTED SPHERICAL LENS INSTEND TORIC LENS AND CORRECT ASTIG WITH PRK ON TOP". THE CAUSE OF THE EVENT AS UNKNOWN AND "EXCESSIVE VAULT AND LENS ROTATION 90 DEGREES". CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -14.00/4.0/092 | OS | Physician | 01/17/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16385551&pc=QCB | 2023826-2023-00562 | 2023-01-19 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.0/3.5/100 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS IMPLANTED,REMOVED AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE REPORTER STATES THE CAUSE OF THIS EVENT IS DUE TO PATIENT RELATED FACTORS. THE REPORTER ALSO STATES THAT THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/3.5/100 | OD | Physician | 01/20/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16385864&pc=QCB | 2023826-2023-00539 | 2022-12-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.5/088 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+1.5/088 | OD | Physician | 01/20/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16385865&pc=QCB | 2023826-2023-00536 | 2023-01-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.0/068 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.5/+1.0/068 | OD | Physician | 01/18/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16385868&pc=QCB | 2023826-2023-00538 | 2022-11-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+3.5/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21 YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -7.5/+3.5/091 | OD | Physician | 01/20/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16390848&pc=QCB | 2023826-2023-00517 | 2022-12-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -12.5/2.0/95 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED, AND REPLACED WITH A SAME SIZE LENS DUE TO A FOREIGN BODY ADHESION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: B1: PRODUCT PROBLEM ADDED AND ADVERSE EVENT OMITTED FOR CORRECTION. B2: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE OMITTED FOR CORRECTION. D4: MODEL#VTICM512.6. H6: 2993 OMITTED FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH FOREIGN MATERIAL ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN MATERIAL ON THE LENS SURFACE, SPECIFICALLY THICK RESIDUE ON HAPTIC AND FOOTPLATE. CLAIM# (B)(4). Manufacturer Narrative: AGE, WEIGHT, AND ETHNICITY: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -12.5/2.0/95 | OD | Physician | 01/19/2023 | 2023 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16400259&pc=QCB | 2023826-2023-00548 | 2022-10-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -03.50/+4.0/104 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -03.50/+4.0/104 | OS | Physician | 01/19/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16400667&pc=QCB | 2023826-2023-00547 | 2022-10-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -01.00/+2.0/070 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -1.00/2.0/070 (SPHERE/ CYLINDER/ AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT, REFRACTIVE SURPRISE AND LENS ROTATION. REPORTEDLY "THE ICLS IMPLANTED DON'T CORRECT PATIENT'S ASTIGMATISM WHICH WAS THE MAIN GOAL FOR THE SURGERY. MOST LIKELY THE ICLS NEED TO BE EXCHANGED." THE LENS REMAINS IMPLANTED. THE PROBLEM WAS NOT RESOLVED. THE CAUSE OF THE EVENT IS THE DEVICE. CAUSE DETAILS: "RESIDUAL ASTIGMATISM". CLAIM# (B)(4) Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -01.00/+2.0/070 | OD | Physician | 01/19/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16404447&pc=QCB | 2023826-2023-00499 | 2022-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.5/2.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -16.5/2.5/090 | OS | Physician | 01/30/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16404454&pc=QCB | 2023826-2023-00502 | 2023-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -9.50/2.5/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFORMATION: UNK. (B)(4). | -9.50/2.5/098 | OS | Physician | 01/30/2023 | 2023 | 41 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16404457&pc=QCB | 2023826-2023-00504 | 2022-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -16.5/3.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, AFTER OPERATION, THE VISUAL ACUITY OF THE LEFT EYE WAS 0.8, AND THE INTRAOCULAR PRESSURE OF THE LEFT EYE WAS 20MM HG. THE ARCH IS 600UM HIGH. Manufacturer Narrative: (B)(4). | -16.5/3.0/090 | OS | Physician | 01/31/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16404463&pc=QCB | 2023826-2023-00505 | 2022-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -14.0/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, AFTER OPERATION, THE VISUAL ACUITY WAS 1.0 IN THE RIGHT EYE, AND THE INTRAOCULAR PRESSURE IN THE RIGHT EYE WAS 13MM HG. THE ARCH IS 550UM HIGH. Manufacturer Narrative: PT INFO: UNK. (B)(4). | -14.0/1.5/093 | OD | Physician | 01/31/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16404467&pc=QCB | 2023826-2023-00510 | 2022-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.0/1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFORMATION: UNK. (B)(4). | -12.0/1.0/092 | OS | Physician | 02/01/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16404480&pc=QCB | 2023826-2023-00511 | 2022-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1, -11.5/2.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY FROM THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. (B)(4). | -11.5/2.5/088 | OS | Physician | 02/02/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16404486&pc=QCB | 2023826-2023-00509 | 2022-10-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/1.5/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. (B)(4). | -10.5/1.5/099 | OD | Physician | 02/01/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16406189&pc=QCB | 2023826-2023-00653 | 2023-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -8/+2.5/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2023 THE LENS WAS EXPLANTED; ON THIS SAME DATE IN A SEPARATE SURGERY A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED BUT THE PROBLEM WAS NOT RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED "LOW VAULT" Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5 - CORRECTED TO: " THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -8/+2.5/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2023 THE LENS WAS EXPLANTED; ON THIS SAME DATE IN A SEPARATE SURGERY A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED BUT THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED "LOW VAULT"" IN THE INITIAL MDR. CLAIM # (B)(4). | -8/+2.5/95 | OD | Physician | 02/14/2023 | 2023 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16406192&pc=QCB | 2023826-2023-00654 | 2023-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -7/+2.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING. ON 09-FEB-2023 THE LENS WAS EXPLANTED; ON THIS SAME DATE IN A SEPARATE SURGERY A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED BUT THE PROBLEM WAS NOT RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED "LOW VAULT". Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5: CORRECTED TO: "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -7/+2.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2023, THE LENS WAS EXPLANTED; ON THIS SAME DATE IN A SEPARATE SURGERY A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED BUT THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED "LOW VAULT"" IN THE INITIAL MDR. CLAIM # (B)(4). | -7/+2.5/091 | OS | Physician | 02/14/2023 | 2023 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16406264&pc=QCB | 2023826-2023-00598 | 2022-06-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -6.50/1.0/108 (SPHERE/CYLINDER/AXIS) LENS TEAR/BREAK BEFORE LOADING WHILE REMOVING FROM VIAL ON (B)(6) 2022. THE LENS WAS NOT IMPLANTED INTO THE PATIENT'S EYE. DURING THE SAME SURGERY A BACKUP LENS WAS USED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS DEVICE. "LENS HAD A TEAR ALREADY IN THE VIAL. HENCE BACKUP USED." Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A CARTRIDGE CASE WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS RETURNED STUCK IN CARTRIDGE WITH RESIDUE ON LENS. THE CARTRIDGE WAS RETURNED IN A DEVICE TRAY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE DEVICE WITH RESIDUE IN CARTRIDGE. CLAIM# (B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -6.50/1.0/108 | NA | Physician | 01/24/2023 | 2023 | 42 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16406265&pc=QCB | 2023826-2023-00596 | 2023-01-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -11.50/1.50/019 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON 02-JAN-2023. ON 30-JAN-2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM # (B)(4). | -11.50/1.50/019 | OD | Physician | 01/23/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16406337&pc=QCB | 2023826-2023-00600 |  | Injury | STAAR SURGICAL COMPANY | 2023-02-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -8.0/1.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED. ON THE SAME DAY DIFFERENT SURGERY A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). | -8.0/1.0/094 | OD | Physician | 01/24/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16408656&pc=QCB | 2023826-2023-00500 | 2023-01-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.5/3.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.5/3.0/075 | OD | Physician | 01/30/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16411423&pc=QCB | 2023826-2023-00565 | 2023-01-26 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.00/+2.5/089 (SPHERE/CYLINDER/AXIS), DID NOT LOAD PROPERLY AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE PATIENT IS GOOD. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: AGE OR DATE OF BIRTH, SEX, WEIGHT, ETHNICITY, AND RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -07.00/+2.5/089 | NA | Physician | 01/26/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16411586&pc=QCB | 2023826-2023-00570 | 2023-01-20 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00/+1.5/075 (SPHERE/CYLINDER/AXIS), TORE DURING LOADING AND THERE WAS NO PATIENT CONTACT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT INFO: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA; B5: THE REPORTER INDICATED THAT 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -18.0/1.5/075 (SPHERE/CYLINDER/AXIS) TORE/BROKE DURING LOADING ON (B)(6) 2023. THE LENS WAS NOT IMPLANTED. A REPLACEMENT LENS OF THE SAME MODEL/LENGTH WAS IMPLANTED INTO THE LEFT EYE (OS) AND THE PROBLEM WAS RESOLVED. REPORTEDLY "I AM AFRAID WE MANAGED TO TAKE A BITE OUT OF ONE OF THE ICLS WITH THE LOADING FORCEPS LAST WEEK". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM#: (B)(4). | -18.00/+1.5/075 | OS | Physician | 01/24/2023 | 2023 | 42 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16412022&pc=QCB | 2023826-2023-00639 | 2023-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/+2.5/103 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -5.5/+2.5/103 | OD | Physician | 01/27/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16412067&pc=QCB | 2023826-2023-00638 | 2023-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+1.5/166 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.0/+1.5/166 | OD | Physician | 01/27/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16412083&pc=QCB | 2023826-2023-00643 | 2023-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.0/+4.0/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -16.0/+4.0/090 | OS | Physician | 01/28/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16412146&pc=QCB | 2023826-2023-00641 | 2022-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.5/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -8.5/+1.5/091 | OD | Physician | 01/27/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16413690&pc=QCB | 2023826-2023-00684 | 2023-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -11.0/1.5/047 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2023 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED "LONGER LENGTH NEEDED"; THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. IT WAS ADDITIONALLY NOTED "EXPLANTED ICL UNTIL LONGER LENGTH RECEIVED". Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A SPECIMEN CUP WITH MOISTURE. VISUAL INSPECTION FOUND OPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.00/+1.50/047 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED A LOW VAULT. REPORTEDLY "HE WILL BE IMPLANTING A LARGER ICL." ON (B)(6) 2023 THE LENS WAS EXPLANTED. A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. REPORTEDLY "THE EXCHANGE WENT WELL AND ALL IS WITHIN NORMAL LIMITS". THE REPORTER INDICATED THE CAUSE OF THE EVENT AS OTHER AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. CAUSE DETAILS- "LONGER LENGTH NEEDED". CLAIM# (B)(4). | -11.0/1.5/047 | OS | Physician | 02/10/2023 | 2023 | 45 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16413764&pc=QCB | 2023826-2023-00683 |  | Injury | STAAR SURGICAL COMPANY | 2023-02-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM_12.6; -9.0/1.0/069 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2023 THE LENS WAS EXPLANTED; ON THIS SAME DATE IN A SEPARATE SURGERY A REPLACEMENT LENS OF A LONGER LENGTH LENS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). | -9.0/1.0/069 | OD | Physician | 02/17/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16414014&pc=QCB | 2023826-2023-00604 |  | Injury | STAAR SURGICAL COMPANY | 2023-02-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -9.0/2.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). REPORTEDLY THE LENS WAS REPOSITIONED DUE TO ROTATION AND THE LENS "WILL NOT STAY IN PLACE." "EXCHANGING DUE TO SIZING. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -9.0/2.0/090 | OS | Physician | 01/27/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16417958&pc=QCB | 2023826-2023-00609 | 2020-11-27 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -5.0/0.5/77 (SPHERE/CYLINDER/AXIS) LENS TEAR/BREAK BEFORE LOADING BEFORE OPENING VIAL WAS OBSERVED. THE LENS WAS NOT IMPLANTED. A BACKUP LENS IMPLANTED OF THE SAME MODEL/SIZE AND DIFFERENT DIOPTER. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: B5-THE SURGEON NOTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.0/0.5/77 (SPHERE/CYLINDER/AXIS) TORE/BROKE BEFORE LOADING (B)(6) 2020. IT'S UNKNOWN WHEN THE DAMAGE OCCURRED AND THE DAMAGE WAS NOTED BEFORE OPENING THE VIAL. THE LENS WAS NOT IMPLANTED. A BACK UP LENS OF THE SAME MODEL/LENGTH WAS IMPLANTED INTO THE RIGHT EYE (OD) AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 TYPE OF INVESTIGATION CODE: 10, DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE, WITH DEBRIS ON THE OPTIC. (B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM #(B)(4). | -5.0/0.5/77 | OD | Physician | 01/29/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16421538&pc=QCB | 2023826-2023-00612 | 2022-08-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -6.00/1.0/80 (SPHERE/CYLINDER/AXIS) LENS TEAR/BREAK DURING LOADING. LENS TEAR DUE TO SNAGGING ON (B)(6) 2022. THE LENS WAS NOT IMPLANTED INTO THE PATIENTS LEFT EYE (OS). A REPLACEMENT LENS OF THE SAME MODEL/SIZE AND SIMILAR POWER WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -6.00/1.0/80 | OS | Physician | 01/31/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16422658&pc=QCB | 2023826-2023-00615 | 2022-09-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -6.00/2.50/125 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE REPLACEMENT LENS RESOLVED THE PROBLEM. SEE MFR # 2023826-2023-00614 FOR INITIAL LENS. Manufacturer Narrative: (B)(4). | -6.00/2.50/125 | OS | Physician | 01/31/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16422659&pc=QCB | 2023826-2023-00614 | 2022-09-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -6.00/2.50/125 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). ON (B)(6) 2022 THE LENS WAS IMPLANTED,REMOVED AND REPLACED INTRAOPERATIVELY FOR A LONGER LENGTH LENS AND THE PROBLEM WA NOT RESOLVED. SEE MFR # 2023826-2023-00615 FOR EXCHANGE LENS. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/+2.5/125 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED A LOW VAULT. THE LENS WAS IMPLANTED; REMOVED AND REPLACED INTRAOPERATIVELY WITH THE SAME MODEL AND POWER, LONGER LENGTH BUT THE PROBLEM WAS NOT RESOLVED. ADDITIONAL INFORMATION PROVIDED: "REPLACEMENT LENS WAS ALSO LOW VAULT, SO WILL BE REPLACED AGAIN". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -6.00/2.50/125 | OS | Physician | 01/31/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16422790&pc=QCB | 2023826-2023-00637 | 2022-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.5/+1.0/176 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE SURGEON REPORTED THAT THE "ANTERIOR CHAMBER DEPTH HAS NORMALIZED." CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -3.5/+1.0/176 | OD | Physician | 01/24/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16422955&pc=QCB | 2023826-2023-00521 | 2023-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -8.5/1.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. ON (B)(6)2023, THE LENS WAS REMOVED AND REPLACED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR. REPORTEDLY, "STILL WITH LOW VAULT, BUT GOOD VISION," AND "MONITORING VAULT. HARD TO IMAGINE AN EYE WITH 12.0 WTW WILL NEED 13.7 LENS." Manufacturer Narrative: WEIGHT-RACE:UNK. PATIENT RELATED FACTOR. CLAIM# (B)(4). | -8.5/1.5/093 | OD | Physician | 01/26/2023 | 2023 | 45 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16423023&pc=QCB | 2023826-2023-00616 | 2022-11-18 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -6.00/1.00/89 SPHERE/CYLINDER/AXIS) LENS TEAR/BREAK DURING LOADING. LENS TEAR DUE TO SNAGGING ON (B)(6)2022. THE LENS WAS NOT IMPLANTED INTO THE PATIENTS RIGHT EYE (OD). A REPLACEMENT LENS OF THE SAME MODEL/SIZE AND SIMILAR POWER WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | -6.00/1.00/89 | OD | Physician | 01/31/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16423066&pc=QCB | 2023826-2023-00522 | 2023-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -8.5/1.5/074 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED AND REPLACED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR. REPORTEDLY, "SURPRISING LOW VAULT WITH 13.2 LENS" AND "MONITORING FOR NOW, GOOD VISION." Manufacturer Narrative: H3: DEVICE EVALUATION - NO LENS WAS RETURNED IN BOX. LENS WAS RETURNED DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE THROUGHOUT THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFO: UNK. PATIENT RELATED FACTOR. OFF-LABEL USE - ACD <3.0. (B)(4). | -8.5/1.5/074 | OS | Physician | 01/26/2023 | 2023 | 45 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16423522&pc=QCB | 2023826-2023-00617 | 2023-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM CTICMO 13.7 IMPLANTABLE COLLAMER LENS OF -16.00/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6)2023 THE LENS WAS REMOVED BECAUSE IT WAS "TOO LARGE." THE LENS WAS REPLACED WITH A SHORTER LENGTH. NO PATIENT INJURY REPORTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN VIAL IN LIQUID. VISUAL INSPECTION FOUND OPTIC AND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -16.00/2.0/089 | OD | Other Health Care Professional | 02/01/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16423523&pc=QCB | 2023826-2023-00618 | 2023-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF -16.00/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS REMOVED BECAUSE IT WAS "TOO LARGE." THE LENS WAS REPLACED WITH A SHORTER LENGTH. NO PATIENT INJURY REPORTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN VIAL IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: B5 - LENS EXCHANGED, DUE TO EXCESSIVE VAULT, ON (B)(6) 2023 WITH A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. STATUS OF EYE "S/P EVO ICL EXCHANGE OS 6 WEEKS. IOP'S OS DECREASED FORM LAT VISIT (HIGH 20'S TO LOW TEENS TODAY). PATIENT HAS BEEN GETTING SEEN EVERY WEEK TO FOLLOW UP ON IOP OS. H6: DEVICE CODE: 1494: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -16.00/2.0/089 | OS | Other Health Care Professional | 02/01/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16428877&pc=QCB | 2023826-2023-00623 | 2023-02-07 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11.00/1.5/090 (SPHERE/CYLINDER/AXIS) LENS TORE DURING INJECTION AND ALSO GOT STUCK IN CARTRIDGE ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS IMPLANTED REMOVED AND REPLACED INTRAOPERATIVELY FOR SPHERICAL LENS OF THE SAME SIZE AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS DEVICE. "THE LENS WOULD NOT STAY FORWARD IN THE SHOOTER AND IT CONTINUED TO SLIP BACKWARDS AND GOT CAUGHT DURING INSERTION." Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -11.00/1.5/090 | NA | Other Health Care Professional | 02/07/2023 | 2023 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16431927&pc=QCB | 2023826-2023-00581 | 2022-08-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.00/+3.5/131 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -07.00/+3.5/131 | OS | Physician | 01/30/2023 | 2023 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16433880&pc=QCB | 2023826-2023-00690 |  | Injury | STAAR SURGICAL COMPANY | 2023-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -14.0/2.0/112 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION WITH SHALLOW VAULT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H6- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND A HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | -14.0/2.0/112 | OD | Physician | 01/25/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16433881&pc=QCB | 2023826-2023-00691 |  | Injury | STAAR SURGICAL COMPANY | 2023-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -14.0/2.5/074 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION WITH SHALLOW VAULT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H6- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND THE HAPTIC AND OPTIC TORN. CLAIM # (B)(4). | -14.0/2.5/074 | OS | Physician | 01/25/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16433882&pc=QCB | 2023826-2023-00692 | 2023-01-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 13.7MM VTICM5_13.7, -7.50/2.0/079 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS FLIPPED IN PATIENT'S LEFT EYE (OS). THE LENS WAS REMOVED AND THE BACKUP LENS WAS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6: HEALTH EFFECT- IMPACT CODE: 4627 SHOULD BE ADDED FOR CORRECTION. ADDITIONAL INFORMATION: D9: RETURN DATE: 20-FEB-2023. CLAIM#(B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -7.50/2.0/079 | OS | Physician | 01/27/2023 | 2023 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16433991&pc=QCB | 2023826-2023-00695 | 2023-01-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 THE SURGEON NOTED THAT A 12.6MM VTICMO12.6, -15.00/2.0/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS BROKEN/TORN AFTER INJECTION INTO THE LEFT EYE (OS). THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY WITH NO PATIENT INJURY. ON (B)(6) 2023 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED.CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.00/2.0/107 | OS | Physician | 01/27/2023 | 2023 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16439814&pc=QCB | 2023826-2023-00597 | 2023-01-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -11.00/1.0/092 (SPHERE/CYLINDER/AXIS) LENS WAS NOTICED TO BE TORN WHILE IN INJECTOR, NO PATIENT CONTACT. THE LENS WAS NOT IMPLANTED. A BACKUP LENS WAS SUCCESSFULLY IMPLANTED. CAUSE REPORTED AS DEVICE. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3: TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM#: (B)(4). | -11.00/1.0/092 | NA | Other Health Care Professional | 01/24/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16440080&pc=QCB | 2023826-2023-00710 | 2022-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -12.5/+2.0/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS HAD TORE/BREAK DURING INJECTION/DELIVERY INTO THE EYE. ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS. THERE WAS NO PATIENT INJURY. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# : (B)(4). | -12.5/+2.0/91 | OD | Physician | 02/21/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16444137&pc=QCB | 2023826-2023-00657 | 2023-01-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.0/0.5/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT IN A SECOND SURGERY ON (B)(6) 2023. A SHORTER LENGTH LENS WAS IMPLANTED IN A THIRD SURGERY ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT: UNK. PATIENT ETHNICITY: UNK. PATIENT RACE: UNK. CLAIM #: (B)(4). | -8.0/0.5/111 | OD | Physician | 02/06/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16444138&pc=QCB | 2023826-2023-00658 | 2022-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/1.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM, UNDER 21 YRS OF AGE AT DATE OF IMPLANT. CLAIM # (B)(4). | -11.0/1.5/085 | OD | Physician | 02/06/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16444144&pc=QCB | 2023826-2023-00659 | 2023-01-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.5/2.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, OVER 45 YRS OF AGE AT DATE OF IMPLANT. CLAIM # (B)(4). | -15.5/2.5/090 | OS | Physician | 02/06/2023 | 2023 | 46 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16444146&pc=QCB | 2023826-2023-00664 | 2022-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -16.0/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED VERTICALLY WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM# (B)(4). | -16.0/2.0/091 | OD | Physician | 02/08/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16444147&pc=QCB | 2023826-2023-00662 | 2023-01-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5 12.1, -11.00/2.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: UNKNOWN. CLAIM# (B)(4). | -11.00/2.0/094 | OD | Physician | 02/07/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16444154&pc=QCB | 2023826-2023-00667 | 2022-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.0/2.0/170 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO ELEVATED IOP (36MMHG) WITHOUT PUPILBLOCK AND ANGLE CLOSURE (ASSESSED ON (B)(6) 2022) AND EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, EYE IS QUIET IOP NORMAL 14 MMHG. PATIENT IS HAPPY. Manufacturer Narrative: CLAIM# (B)(4). | -15.0/2.0/170 | OD | Physician | 02/09/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16444155&pc=QCB | 2023826-2023-00669 | 2023-01-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.5/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: UNKNOWN. CLAIM # (B)(4). | -11.5/1.0/089 | OD | Physician | 02/08/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16444163&pc=QCB | 2023826-2023-00666 | 2022-12-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.0/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION UNKNOWN. CLAIM # (B)(4). | -11.0/2.0/089 | OS | Physician | 02/08/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16444173&pc=QCB | 2023826-2023-00663 | 2023-01-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5 12.1, -11.50/1.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4). | -11.50/1.5/088 | OS | Physician | 02/07/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16448776&pc=QCB | 2023826-2023-00645 | 2022-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.5/130 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.0/+1.5/130 | OS | Physician | 01/29/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16448870&pc=QCB | 2023826-2023-00646 | 2022-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/+4.0/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -14.5/+4.0/094 | OS | Physician | 01/29/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16448888&pc=QCB | 2023826-2023-00644 | 2023-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0/+3.0/072 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -15.0/+3.0/072 | OD | Physician | 01/28/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16449741&pc=QCB | 2023826-2023-00648 | 2022-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-02-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+1.5/087 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.0/+1.5/087 | OS | Physician | 01/29/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16461342&pc=QCB | 2023826-2023-00698 | 2022-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -13.00/2.0/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS AND THE PROBLEM RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.00/2.0/90 | OD | Physician | 01/30/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16461535&pc=QCB | 2023826-2023-00699 | 2022-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -8.50/2.0/94 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2022 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PROBLEM RESOLVED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND THE OPTIC TORN. DIMENSIONAL INSPECTIONS FOUND THE DEVICE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -8.50/2.0/94 | OS | Physician | 01/30/2023 | 2023 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16464623&pc=QCB | 2023826-2023-00762 | 2022-12-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.5/2.5/069 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED DUE TO LOW VAULTING. ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON THIS SAME DATE A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 17-MAY-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.5/2.5/069 | OS | Physician | 02/27/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16464624&pc=QCB | 2023826-2023-00763 | 2023-01-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.0/1.5/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE SURGEON REPORTERS LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT DIFFERENT AXIS; THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9-RETURN DATE: 04/17/2023. H3:DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). | -10.0/1.5/088 | OD | Physician | 02/28/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16464625&pc=QCB | 2023826-2023-00766 | 2022-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -7.5/1.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS. BUT THE PROBLEM WAS NOT RESOLVED. ADDITIONALLY IT WAS NOTED "AFTER THE DOCTOR REPLACED THE LENS; THERE WAS NO SIGNIFICANT CHANGE IN THE PATIENT'S ARCH HEIGHT". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: D9: RETURN DATE: 17-MAY-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH RESIDUE/DEBRIS ON THE LENS. AND THE RETURNED LENS MODEL IS NOT THE MODEL STATED IN THE CLAIM. DIMENSIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#(B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -7.5/1.5/090 | OS | Physician | 02/28/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16464654&pc=QCB | 2023826-2023-00765 | 2022-11-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.5/1.0/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS THERE WAS EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXPLANTED; AND THE PATIENT DOES NOT WANT TO UNDERGO A SECONDARY LENS IMPLANTATION PROCEDURE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9-RETURN DATE: 04/17/2023. H3:DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.5/1.0/094 | OS | Physician | 02/28/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16464853&pc=QCB | 2023826-2023-00742 | 2022-09-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.0/029 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -11.5/+1.0/029 | OD | Physician | 01/31/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16464855&pc=QCB | 2023826-2023-00743 | 2022-09-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+1.0/149 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -12.0/+1.0/149 | OS | Physician | 01/31/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16465099&pc=QCB | 2023826-2023-00704 | 2022-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -14.0/3.5/085 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023. THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PROBLEM RESOLVED. THE REPORTER STATED " DURING THE FOLLOW-UP VISIT ON (B)(6), IT WAS FOUND THAT THE TICL LENS OF THE PATIENT'S LEFT EYE ROTATED 10 CLOCKWISE, AND ITS ARCH HEIGHT WAS 400 MICRONS. AFTER CONTINUOUS OBSERVATION FOR 4 MONTHS, IT WAS FOUND THAT THE LENS POSITION WAS RELATIVELY FIXED AND THE LEFT EYE VISION WAS 0.8" PER REPORTER AFTER PATIENT UNDERWENT LENS REPLACEMENT THE VISUAL ACUITY OF THE LEFT EYE WAS 1.0, THE INTRAOCULAR PRESSURE OF THE LEFT EYE WAS 22 MM HG, AND THE ARCH HEIGHT WAS 400 MICRONS.CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A BENT HAPTIC. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD <3.0MM). CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -14.0/3.5/085 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND REPLACED IN A SEPARATE SURGERY THAT DAY FOR A SAME LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PROBLEM RESOLVED. THE REPORTER STATED " DURING THE FOLLOW-UP VISIT ON (B)(6), IT WAS FOUND THAT THE TICL LENS OF THE PATIENT'S LEFT EYE ROTATED 10 CLOCKWISE, AND ITS ARCH HEIGHT WAS 400 MICRONS. AFTER CONTINUOUS OBSERVATION FOR 4 MONTHS, IT WAS FOUND THAT THE LENS POSITION WAS RELATIVELY FIXED AND THE LEFT EYE VISION WAS 0.8" PER REPORTER AFTER PATIENT UNDERWENT LENS REPLACEMENT THE VISUAL ACUITY OF THE LEFT EYE WAS 1.0, THE INTRAOCULAR PRESSURE OF THE LEFT EYE WAS 22 MM HG, AND THE ARCH HEIGHT WAS 400 MICRONS.CAUSE OF EVENT WAS UNKNOWN. CLAIM# (B)(4). | -14.0/3.5/085 | OS | Physician | 01/31/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16465225&pc=QCB | 2023826-2023-00771 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 -17.00/2.0/155 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). ON 07-FEB-2023 THE LENS WAS REMOVED AND REPLACED WITH A SPHERICAL LENS OF LONGER LENGTH DUE TO LENS DISLOCATION. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -17.00/2.0/155 | OS | Other Health Care Professional | 02/02/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16473884&pc=QCB | 2023826-2023-00705 | 2023-01-03 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 12.6MM, VTICMO12.6, -9.5/1.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE RIGHT EYE (OD). THE LENS WAS IMPLANTED, REMOVED AND REPLACED INTRAOPERATIVELY WITH NO PATIENT INJURY AND THE PROBLEM RESOLVED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. THE REPORTER ALSO STATES THAT THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: CORRECTED DATA: H6 - MEDICAL DEVICE PROBLEM CODE 1494: OFF-LABEL USE (UNDER 21 YRS OF AGE AT DATE OF IMPLANTATION). CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -9.5/1.0/094 | OD | Physician | 01/31/2023 | 2023 | 4 MO | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16473945&pc=QCB | 2023826-2023-00703 | 2022-12-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 12.6MM, VTICM512.6, -4.50/4.5/014 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE LEFT EYE (OS). PATIENT CONTACT WITH NO PATIENT INJURY WAS REPORTED. ON (B)(6) 2023, A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -4.50/4.5/014 | OS | Physician | 01/31/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16480116&pc=QCB | 2023826-2023-00729 | 2022-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5/2.5/084 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/2.5/084 | OS | Physician | 02/06/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16480117&pc=QCB | 2023826-2023-00728 | 2022-05-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.0/4.5/082 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH SPHERIC MODEL LENS DUE TO LENS ROTATION NOT ASSOCIATED WITH LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR. REPORTEDLY, "DUE TO PATIENT REASONS, VICMO WERE EVENTUALLY IMPLANTED." Manufacturer Narrative: PATIENT RELATED FACTOR. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -12.0/4.5/082 | OD | Physician | 02/06/2023 | 2023 | 44 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16483011&pc=QCB | 2023826-2023-00778 | 2022-12-26 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2; -11.50/1.0/176 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS HAD TORE DURING LOADING; AND NOTED "LENS RUPTURE DUE TO SNAGGING". THIS OCCURRED ON (B)(6) 2022. THE PROBLEM WAS REPORTED AS RESOLVED AND NOTED "NEXT OPERATION DATE UNDECIDED". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.50/1.0/176 | NA | Physician | 03/02/2023 | 2023 | 44 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16485545&pc=QCB | 2023826-2023-00677 | 2023-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2, -14.50/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4). | -14.50/2.0/095 | OS | Physician | 02/14/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16485556&pc=QCB | 2023826-2023-00680 | 2022-12-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2M VTICM5 13.2, -09.00/2.5/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4). | -09.00/2.5/096 | OD | Physician | 02/17/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16485558&pc=QCB | 2023826-2023-00679 | 2023-02-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -13.0/1.0/003 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT:UNK ETHNICITY: UNK RACE: UNK CLAIM# (B)(4). | -13.0/1.0/003 | OS | Physician | 02/16/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16485559&pc=QCB | 2023826-2023-00678 | 2023-01-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2, -13.0/2.5/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY:UNKNOWN RACE:UNKNOWN. CLAIM# (B)(4). | -13.0/2.5/079 | OS | Physician | 02/14/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16486793&pc=QCB | 2023826-2023-00774 | 2023-02-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -11.50/1.50/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON NOTES THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE; AND THE LENS WAS STUCK IN THE CARTRIDGE/INJECTION SYSTEM. INTRAOPERATIVELY THE LENS WAS REMOVED WITH NO PATIENT INJURY. ON (B)(6) 2023 A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT REPORTED AS USER ERROR AND NOTED "LENS GOT STUCK IN PLUNGER WHILE IMPLANTATION OF LENS AND GOT TORN". Manufacturer Narrative: CLAIM#: (B)(4). | -11.50/1.50/90 | OD | Physician | 03/02/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16488018&pc=QCB | 2023826-2023-00674 | 2023-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.0/1.5/173 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED VERTICALLY WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/1.5/173 | OS | Physician | 02/13/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16488021&pc=QCB | 2023826-2023-00675 | 2021-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2M VTICMO 13.2, -15.50/1.5/113 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER SHORTER LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. CLAIM# (B)(4). | -15.50/1.5/113 | OD | Physician | 02/14/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16492174&pc=QCB | 2023826-2023-00731 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -13.5/1.5/113 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). THE LENS WAS REMOVED AND LATER REPLACED FOR A LONGER LENGTH LENS. Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -13.5/1.5/113 | OD | Physician | 02/07/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16492442&pc=QCB | 2023826-2023-00799 | 2022-12-26 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2; -11.5/1.0/22 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS HAD TORE DURING LOADING; AND NOTED "LENS RUPTURE DUE TO SNAGGING". THIS OCCURRED ON (B)(6) 2022. PROBLEM WAS REPORTED AS RESOLVED AND NOTED "NEXT OPERATION DATE UNDECIDED". CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.5/1.0/22 | NA | Physician | 03/02/2023 | 2023 | 44 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16492513&pc=QCB | 2023826-2023-00632 | 2023-02-03 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Particulates; Off-Label Use | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO_13.2; -8.0/1.0/073 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS THERE WAS A FOREIGN PARTICULATE FOUND ON A SFC45 CARTRIDGE DURING THE LOADING PROCESS. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. PROBLEM IS NOT RESOLVED. Manufacturer Narrative: MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -8.0/1.0/073 | OS | Physician | 02/09/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16494457&pc=QCB | 2023826-2023-00732 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -14.5/2.0/065 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS REMOVED AND LATER REPLACED FOR A LONGER LENGTH LENS. Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -14.5/2.0/065 | OS | Physician | 02/07/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16494753&pc=QCB | 2023826-2023-00672 | 2023-01-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.5/1.0/014 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM, OVER 45 YRS OF AGE AT DATE OF IMPLANT. CLAIM # (B)(4). | -8.5/1.0/014 | OS | Physician | 02/12/2023 | 2023 | 47 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16499736&pc=QCB | 2023826-2023-00802 | 2023-02-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Particulates; Off-Label Use | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO_13.2; -10/1.5/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE SURGEON REPORTS THERE WAS A FOREIGN PARTICULATE FOUND ON A SFC45 CARTRIDGE DURING THE LOADING PROCESS. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. PROBLEM IS NOT RESOLVED. Manufacturer Narrative: MEDICAL DEVICE PROBLEM CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). | -10/1.5/092 | OS | Physician | 02/09/2023 | 2023 | 17 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16499737&pc=QCB | 2023826-2023-00630 | 2022-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8.0/+3.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. THE SURGEON ADDITIONALLY NOTED "PATIENT HAS TO WEAR GLASSES AGAIN; BUT SHE IS HAPPY". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -8.0/+3.0/093 | OS | Physician | 02/09/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16499738&pc=QCB | 2023826-2023-00800 | 2023-01-20 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6; -8.5/1.5/87 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS HAD TORE DURING LOADING; AND NOTED "LENS RUPTURE DUE TO SNAGGING." THIS OCCURRED ON (B)(6) 2023. ON THIS SAME DATE A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -8.5/1.5/87 | NA | Physician | 03/02/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16499740&pc=QCB | 2023826-2023-00833 | 2023-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -7.5/1.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022 AS A REPLACEMENT LENS FOR LOW VAULTING. IT WAS REPORTED THAT "AFTER THE DOCTOR REPLACED THE LENS; THERE WAS NO SIGNIFICANT CHANGE IN THE PATIENT'S ARCH HEIGHT". PROBLEM IS NOT RESOLVED. LENS REMAINS IMPLANTED. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). | -7.5/1.5/091 | OS | Physician | 02/28/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16499867&pc=QCB | 2023826-2023-00785 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -16.5/1.5/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE. EXCESSIVE VAULTING IS OBSERVED AND A EXCHANGE IS PLANNED. Manufacturer Narrative: CORRECTED DATA: D4: EXPIRATION DATE SHOULD BE CORRECTED TO (B)(6) 2024. D4: MANUFACTURE DATE SHOULD BE CORRECTED TO 06-JUL-2022. H1: SHOULD BE CORRECTED TO SERIOUS. CLAIM# (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -16.5/1.5/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE ON AN UNKNOWN DATE. EXCESSIVE VAULTING WAS OBSERVED AND THE LENS WAS EXPLANTED ON AN UNKNOWN DATE. A SHORTER LENS WAS USED FOR REPLACEMENT LENS. IT IS UNCLEAR IF LENS REMOVAL AND REPLACEMENT OCCURRED INTRAOPERATIVELY. D9: RETURN DATE: 22-MAR-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -16.5/1.5/087 | NA | Physician | 02/09/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16499868&pc=QCB | 2023826-2023-00786 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -15.00/2.0/102 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE. EXCESSIVE VAULTING IS OBSERVED AND A EXCHANGE IS PLANNED. Manufacturer Narrative: CORRECTED DATA: H1: SHOULD BE CORRECTED TO SERIOUS INJURY. H3 - DEVICE EVALUATION: SHOULD BE CORRECTED TO: THE LENS WAS RETURNED IN LIQUID IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -15.00/2.0/102 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE ON AN UNKNOWN DATE. EXCESSIVE VAULTING WAS OBSERVED AND THE LENS WAS EXPLANTED ON AN UNKNOWN DATE.A SHORTER LENS WAS USED FOR REPLACEMENT LENS. IT IS UNCLEAR IF LENS REMOVAL AND REPLACEMENT OCCURRED INTRAOPERATIVELY. ADDITIONAL INFORMATION: D9: RETURN DATE: 22/MAR/2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4). | -15.00/2.0/102 | NA | Physician | 02/09/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16500967&pc=QCB | 2023826-2023-00822 | 2022-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -11.5/+1.0/091 | OD | Physician | 02/06/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16500968&pc=QCB | 2023826-2023-00821 | 2023-01-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+1.0/069 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -8.0/+1.0/069 | OS | Physician | 02/06/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16501005&pc=QCB | 2023826-2023-00824 | 2023-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+2.0/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -13.0/+2.0/094 | OS | Physician | 02/07/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16506839&pc=QCB | 2023826-2023-00629 | 2022-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.00/+2.5/081 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO LOW VAULTING AND THE PROBLEM WAS RESOLVED. THE PATIENT HAS TO WEAR GLASSES AGAIN, BUT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND THE HAPTIC DEFORMED AND STRETCHED OUT WITH RESIDUE/DEBRIS ON THE LENS. DIMENSIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -07.00/+2.5/081 | OD | Physician | 02/09/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16507964&pc=QCB | 2023826-2023-00841 | 2022-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.0/1.5/155 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING WITHOUT LENS ROTATION. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT IMPLANTED HORIZONTALLY; AS OPPOSED TO THE VERTICAL IMPLANT. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -9.0/1.5/155 | OS | Physician | 03/08/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16509343&pc=QCB | 2023826-2023-00626 | 2023-02-07 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 13.2MM VTICM5_13.2 -6.50/1.0/111 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS BROKE IN HALF IN THE INJECTOR. THERE WAS NO PATIENT CONTACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: D9: RETURN DATE: 20-FEB-2023. H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -6.50/1.0/111 | NA | Physician | 02/09/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16509699&pc=QCB | 2023826-2023-00842 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 -10.50/1.5/011 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS REMOVED AND REPLACED WITH A LENS OF LONGER LENGTH AND DIFFERENT POWER ON (B)(6) 2023. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND A HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: IMPLANT DATE: UNKNOWN. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -10.50/1.5/011 | NA | Other Health Care Professional | 02/08/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16519587&pc=QCB | 2023826-2023-00807 | 2022-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.0/4.50/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 AND (B)(6) 2022 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND LATER REPLACED WITH A LONGER LENGTH LENS, PROBLEM RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.0/4.50/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. ON (B)(6) 2022 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS, PROBLEM RESOLVED. CAUSE IS REPORTED AS UNKNOWN. H6- HEALTH EFFECT-IMPACT CODE: 4627 TO 4629 - PREVIOUS CODE NOT APPLICABLE, THE LENS WAS REPLACED NOT EXPLANTED. H6-DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTICS BENT/DEFORMED . DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/4.50/087 | OS | Physician | 02/09/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16520002&pc=QCB | 2023826-2023-00828 | 2022-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+3.5/087 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+3.5/087 | OD | Physician | 02/08/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16523648&pc=QCB | 2023826-2023-00636 | 2022-12-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF A -6.5/1.0/98 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS A REPLACEMENT INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE INITIAL LENS WAS EXCHANGED DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS NOT RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "AFTER THE DOCTOR REPLACED THE CRYSTAL IN THE PATIENT, IT WAS STILL LOW ARCH HIGH, BUT THE DOCTOR DECIDED NOT TO REPLACE THE LENS AGAIN." Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. ADDITIONAL DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+1.0/98 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022 AS A REPLACEMENT LENS. THE PATIENT EXPERIENCED A LOW VAULT. THE LENS WAS EXPLANTED. REPORTEDLY " AFTER THE DOCTOR REPLACED THE CRYSTAL IN THE PATIENT, IT WAS STILL LOW ARCH HIGH, BUT THE DOCTOR DECIDED NOT TO REPLACE THE LENS AGAIN". H6-HEALTH EFFECT- IMPACT CODE: "2199" SHOULD BE REMOVED AND CODE "4627" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE:"2993" SHOULD BE ADDED. H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | -6.5/1.0/98 | OD | Physician | 02/10/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16523649&pc=QCB | 2023826-2023-00633 | 2023-02-09 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF A -9.5/1.5/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED, AND REPLACED DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT, BUT NO PATIENT INJURY. REPORTEDLY, "EYE IS QUIET," AND "LOADING AND SURGERY WENT OFF WELL WITH NO OUTWARDS INCIDENT." THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.5/91 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. REPORTEDLY THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. ON (B)(4) 2023 THE SAME MODEL AND LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS REPORTED AS "EYE QUIET' AND IT WAS REPORTED THAT "LOADING AND SURGERY WENT OFF WELL WITH NO UNTOWARDS INCIDENT'. REPORTEDLY "LENS GOT STUCK IN THE PLUNGER WHILE IMPLANTING AND GOT TORN'. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -9.5/+1.5/91 | OD | Physician | 02/09/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16523651&pc=QCB | 2023826-2023-00635 | 2022-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -6.5/1.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS NOT RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. ADDITIONAL DATA: H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -6.5/1.0/093 | OD | Physician | 02/10/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16524851&pc=QCB | 2023826-2023-00816 | 2023-01-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 13.2MM, VTICMO13.2, -9.00/5.0/163 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE LEFT EYE (OS). THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY, WITH NO PATIENT INJURY. A REPLACEMENT LENS WAS IMPLANTED AT A LATER DATE AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -9.00/5.0/163 | OS | Physician | 02/10/2023 | 2023 | 50 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16529974&pc=QCB | 2023826-2023-00871 | 2023-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -14.50/5.0/119 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED LOW VAULTING WITH ROTATION. ON (B)(6) 2023 THE LENS WAS EXPLANTED; ON THIS SAME DATE A REPLACEMENT LENS OF A LONGER LENGTH LENS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM; "EYE IS WELL; PATIENT IS SATISFIED". THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND UNKNOWN AND NOTED "LOW VAULT CAUSING ROTATION". Manufacturer Narrative: (B)(4). | -14.50/5.0/119 | OD | Physician | 03/06/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16529982&pc=QCB | 2023826-2023-00870 | 2023-02-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -13.5/1.5/126 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED. ON (B)(6) 2023 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND UNKNOWN AND NOTED "HIGH VAULT; INOPERATIVENESS REMOVAL DUE TO NO MYOSIS AFTER THE LENS IMPLANTATION". Manufacturer Narrative: CLAIM#: (B)(4). | -13.5/1.5/126 | OD | Physician | 03/06/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16530286&pc=QCB | 2023826-2023-00746 | 2022-12-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -9.0/0.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION UNKNOWN. DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, UNDER 21 YRS OF AGE AT DATE OF IMPLANT. CLAIM# (B)(4). | -9.0/0.5/082 | OD | Physician | 02/20/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16530291&pc=QCB | 2023826-2023-00750 | 2023-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.0/1.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION UNKNOWN. DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM #(B)(4). | -11.0/1.0/094 | OD | Physician | 02/20/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16530293&pc=QCB | 2023826-2023-00747 | 2022-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.5/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION UNKNOWN. DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM, 21 YRS OF AGE AT DATE OF IMPLANT. CLAIM# (B)(4). | -11.5/1.0/091 | OS | Physician | 02/20/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16530296&pc=QCB | 2023826-2023-00749 | 2023-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.0/2.5/062 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION UNKNOWN. DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM # (B)(4). | -11.0/2.5/062 | OS | Physician | 02/20/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16530297&pc=QCB | 2023826-2023-00753 | 2022-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.0/1.0/151 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION UNKNOWN. CLAIM #(B)(4). | -10.0/1.0/151 | OS | Physician | 02/20/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16530301&pc=QCB | 2023826-2023-00756 | 2022-12-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.0/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION UNKNOWN. CLAIM# (B)(4). | -8.0/1.0/087 | OD | Physician | 02/22/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16530303&pc=QCB | 2023826-2023-00758 | 2022-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.0/2.0/152 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. CLAIM# (B)(4). | -10.0/2.0/152 | OD | Physician | 02/24/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16530309&pc=QCB | 2023826-2023-00759 | 2023-02-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.0/1.5/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION UNKNOWN. CLAIM # (B)(4). | -18.0/1.5/107 | OD | Physician | 02/24/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16530584&pc=QCB | 2023826-2023-00754 | 2023-01-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.5/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -7.5/1.0/087 | OD | Physician | 02/20/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16530589&pc=QCB | 2023826-2023-00755 | 2022-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.0/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM #:(B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.0/2.0/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) /2023. THIS RESOLVED THE PROBLEM. D4: MODEL # VTICMO_12.1. EXPIRATION DATE: (B)(6) 2024. UNIQUE IDENTIFIER (UDI) # SHOULD BE OMITTED FOR CORRECTION. SERIAL # (B)(6). H4: DEVICE MANUFACTURE DATE: 04-DEC-2021. CLAIM# (B)(4). | -10.0/2.0/088 | OS | Physician | 02/20/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16532949&pc=QCB | 2023826-2023-00888 | 2023-02-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -8.50/3.0/085 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE SURGEON NOTES THE LENS WAS IMPLANTED UPSIDE DOWN; INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND "OTHER"; AND NOTED "THE EVO LENS IS MORE FLIMSY AND EASY TO PUT IN WRONG IF THE PATIENT MOVES; HARD TO CONTROL". Manufacturer Narrative: (B)(4). | -8.50/3.0/085 | OS | Physician | 02/24/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16533267&pc=QCB | 2023826-2023-00894 | 2023-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+0.5/127 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -6.5/+0.5/127 | OD | Physician | 02/12/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16533278&pc=QCB | 2023826-2023-00896 | 2023-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.5/+4.5/083 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.5/+4.5/083 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). Manufacturer Narrative: B5- CORRECTED. CLAIM# (B)(4). | -15.5/+4.5/083 | OD | Physician | 02/12/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16533340&pc=QCB | 2023826-2023-00895 | 2023-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+1.0/171 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 . THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -6.5/+1.0/171 | OS | Physician | 02/12/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16534443&pc=QCB | 2023826-2023-00875 | 2023-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -11.50/4.0/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULT. LENS WAS EXPLANTED ON AN UNKNOWN DATE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EXPLANT DATE: UNKNOWN. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -12.00/4.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. H6- HEALTH EFFECT-IMPACT CODE: 4627 TO 4629 - PREVIOUS CODE NOT APPLICABLE, THE LENS WAS REPLACED NOT EXPLANTED. CLAIM # (B)(4). | -11.50/4.0/094 | OD | Physician | 02/11/2023 | 2023 | 41 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16534502&pc=QCB | 2023826-2023-00876 | 2023-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023 AND THE PATIENT EXPERIENCED EXCESSIVE VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -9.00/2.0/076 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. H6- HEALTH EFFECT-IMPACT CODE: 4627 TO 4629 - PREVIOUS CODE NOT APPLICABLE, THE LENS WAS REPLACED NOT EXPLANTED. CLAIM # (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -9.00/2.0/079 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. PATIENT IS HAPPY. H6: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -9.00/2.0/079 | OS | Physician | 02/11/2023 | 2023 | 41 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16538740&pc=QCB | 2023826-2023-00793 | 2023-02-01 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -15.50/3.0/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS INTRAOPERATIVELY IMPLANTED AND REMOVED DUE TO A LENS TEAR/BREAK DURING LOADING. THERE WAS PATIENT CONTACT, BUT NO PATIENT INJURY. ON (B)(6) 2023, THE REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE REPORTER STATED, "EYE IS WELL, BEST CORRECTED IS ACHIEVED." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "THE LENS WAS LOADED CORRECTLY BUT IT SEEMS THAT THE FOAM TIP PLUNGER CAUGHT A HAPTIC AND TEAR IT DURING THE INJECTION OF THE LENS." Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | -15.50/3.0/095 | OS | Physician | 02/12/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16542490&pc=QCB | 2023826-2023-00898 | 2023-01-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.5/071 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -11.5/+1.5/071 | OS | Physician | 02/13/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16542491&pc=QCB | 2023826-2023-00897 | 2023-01-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.5/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -11.0/+1.5/091 | OD | Physician | 02/13/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16542692&pc=QCB | 2023826-2023-00899 | 2022-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+2.0/095 INTO THE PATIENT'S LEFT EYE (OS) ON 24-OCT-2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.0/+2.0/095 | OS | Physician | 02/14/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16545492&pc=QCB | 2023826-2023-00849 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS). THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO AN OVERSIZED LENS. THE LENS WAS REPLACED WITH A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PATIENT INFORMATION UNKNOWN. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.00/+1.5/093 | OS | Physician | 02/17/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16545525&pc=QCB | 2023826-2023-00848 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.00/+2.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD). THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO AN OVERSIZED LENS. THE LENS WAS REPLACED WITH A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PATIENT INFORMATION UNKNOWN. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.00/+2.0/089 | OD | Physician | 02/17/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16545537&pc=QCB | 2023826-2023-00858 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -07.50/+2.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE LENS POSITION IS WHERE IT SHOULD BE. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WIDTH WAS FOUND TO BE IN SPECIFICATION BUT THE LENGTH (DIAMETER) VERIFICATION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT INFORMATION AND EVENT DATE UNKNOWN. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -07.50/+2.5/094 | OD | Physician | 02/13/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16545676&pc=QCB | 2023826-2023-00854 | 2023-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS LEFT EYE (OS). THE LENS WAS REPORTED AS LOW VAULTING, WITH LENS ROTATION. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -03.50/+2.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. CORRECTED DATA: D2: COMMON DEVICE NAME: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB. E1: ADDRESS: (B)(6). CLAIM # (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH, SEX, WEIGHT, ETHNICITY, RACE, DATE OF EVENT, EXPIRATION DATE, IMPLANT DATE, EXPLANT DATE, AND DEVICE MANUFACTURE DATE: UNK. (B)(4). | -03.50/+2.0/088 | OS | Physician | 02/13/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16545773&pc=QCB | 2023826-2023-00850 | 2022-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+1.5/107 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. POST-OP THERE WAS MINOR CORNEAL EDEMA AT NASAL SIDE. SEE MFR. REPORT # 2023826-2023-00851 FOR REPLACEMENT LENS. POST-OP THE PATIENT HAS STABILIZED AND IS FEELING BETTER. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION UNKNOWN. CLAIM # (B)(4). | -16.50/+1.5/107 | OS | Physician | 02/13/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16545817&pc=QCB | 2023826-2023-00857 | 2023-01-31 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -13.00/+1.5/086 (SPHERE/CYLINDER/AXIS) WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED ON (B)(6) 2023 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.00/+1.5/086 | OS | Physician | 02/13/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16545830&pc=QCB | 2023826-2023-00853 | 2022-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -04.00/+4.0/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIUDE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION UNKNOWN. DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | -04.00/+4.0/082 | OS | Physician | 02/13/2023 | 2023 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16550263&pc=QCB | 2023826-2023-00957 | 2022-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -14.0/3.0/180 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT IMPLANTED VERTICALLY; THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN/BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -14.0/3.0/180 | OS | Physician | 03/14/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16550785&pc=QCB | 2023826-2023-00906 | 2023-02-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.5/1.5/134 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS INTRAOPERATIVELY IMPLANTED AND REMOVED DUE TO EXCESSIVE VAULT. CAUSE OF THE EVENT IS A PATIENT RELATED FACTOR. REPORTEDLY, "EXCESSIVE VAULT AT THE IMPLANTATION WITH THE IRIS WHICH PASSED UNDER THE ICL IN SPITE OF CLEANING OF VISCOELASTIC SIGNIFICANT RISK OF PUPILLARY BLOCK. THE ICL HAS BEEN EXPLANTED DIRECTLY." Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -10.5/1.5/134 | OD | Physician | 02/15/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16554754&pc=QCB | 2023826-2023-00866 | 2023-02-09 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.00/+1.0/084 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPORTED AS LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED AT THIS TIME NO EXCHANGE IS PLANNED, THE SURGEON WILL CLOSELY MONITOR THE PATIENT. THE PROBLEM WAS RESOLVED AND THE PATIENT IS HAPPY. THE LENS REMAINED IMPLANTED. CLAIM # (B)(4). | -09.00/+1.0/084 | OD | Physician | 02/16/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16554755&pc=QCB | 2023826-2023-00865 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50/+3.0/109 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2023. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 11-JUL-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND RESIDUE/DEBRIS ON THE LENS SPECIFICALLY FIBRE AND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). | -17.50/+3.0/109 | OD | Physician | 02/15/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16559752&pc=QCB | 2023826-2023-00882 | 2022-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -14.00/1.5/88 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE EXPLANT RESOLVED THE PROBLEM. Manufacturer Narrative: H6: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE DEBRIS ON PRODUCT IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -14.00/1.5/88 | OS | Physician | 02/14/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16559999&pc=QCB | 2023826-2023-00881 | 2022-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -13.50/1.5/82 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE EXPLANT RESOLVED THE PROBLEM. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D4: (B)(4) D6A:IMPLANT DATE: ON (B)(6) 2022. (B)(4) . Manufacturer Narrative: H6- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | -13.50/1.5/82 | OD | Physician | 02/14/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16560156&pc=QCB | 2023826-2023-00903 | 2022-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.00/+3.5/086 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -9.00/+3.5/086 | OD | Physician | 02/15/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16560351&pc=QCB | 2023826-2023-00900 | 2022-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.5/099 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+1.5/099 | OD | Physician | 02/15/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16560360&pc=QCB | 2023826-2023-00959 | 2023-02-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6; -13.0/+2.5/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTENDED FOR THE PATIENT'S LEFT EYE (OS); HAD TORE DURING LOADING. LENS WAS NOT IMPLANTED. THIS OCCURRED ON (B)(6) 2023. ON THIS SAME DATE A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -13.0/+2.5/92 | OS | Physician | 03/15/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16560555&pc=QCB | 2023826-2023-00910 | 2022-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM513.7 IMPLANTABLE COLLAMER LENS OF -9.5/3.0/110 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS ROTATION NOT ASSOCIATED TO LOW VAULT WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2022 AND ON (B)(6) 2022. THE REPOSITIONS DID NOT RESOLVED THE PROBLEM. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SPHERIC LENS AND THE PROBLEM WAS RESOLVED. REPORTEDLY, "LRI WILL BE PERFORMED FOR ASTIMATISM." Manufacturer Narrative: WEIGHT: UNKNOWN; ETHNICITY: UNKNOWN; RACE:UNKNOWN. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+3.0/110 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED WITH A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2022 AND (B)(6) 2022 BUT THAT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SPHERICAL MODEL AND POWER, SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION PROVIDED WAS: "LRI PERFORMED FOR ASTIGMATISM." CLAIM# (B)(4). | -9.5/3.0/110 | OD | Physician | 02/16/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16560630&pc=QCB | 2023826-2023-00909 | 2022-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -14.50/2.074 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SPHERIC LENS DUE TO LENS ROTATION NO ASSOCIATED TO LOW VAULT AND THE PROBLEM WAS RESOLVED. REPORTEDLY, "LRI WILL BE PERFORMED FOR ASTIMATISM." Manufacturer Narrative: CORRECTED DATA: B5: IMPLANTATION DATE: (B)(6) 2020. D4: MODEL: VTICMO12.1. ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -14.50/2.074 | OS | Physician | 12/09/2020 | 2020 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16561004&pc=QCB | 2023826-2023-00904 | 2022-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.50/+3.0/104 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A DIFFERENT LENS OF LONGER LENGTH DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: (B)(4). | -8.50/+3.0/104 | OS | Physician | 02/15/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16564633&pc=QCB | 2023826-2023-00887 | 2023-02-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 13.2MM, VTICM5_13.2, -9.00/1.5/115 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE RIGHT EYE (OD). THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. THERE WAS NO PATIENT INJURY. ON (B)(6) 2023 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -9.00/1.5/115 | OD | Physician | 02/16/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16564636&pc=QCB | 2023826-2023-00886 | 2020-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS -18.00/2.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. CAUSE OF EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). Manufacturer Narrative: B5- PER UPDATED INFORMATION THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: D4: MODEL# CORRECTED TO VTICMO 12.1 IN INITIAL MDR. CLAIM#: (B)(4). | -18.00/2.5/092 | OD | Physician | 12/13/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16566642&pc=QCB | 2023826-2023-00913 | 2023-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS OF -8.5/1.5/75 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS A REPLACEMENT INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023, DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS NOT RESOLVED. CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR, AND THE DEVICE FAILED TO PERFORM AS INTENDED. REPORTEDLY, "LOW VAULT WITH ROTATION." THE STATUS OF THE EYE IS REPORTED AS, "SUPRISING LOW VAULT WITH 13.2 LENS," AND "MONITORING FOR NOW, GOOD VISION." Manufacturer Narrative: PATIENT INFORMATION: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). Manufacturer Narrative: B5: LENS WAS EXPLANTED ON AN UNKNOWN DATE. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN GAUZE DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT AND RESIDUE THROUGHOUT THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -8.5/1.5/75 | OS | Physician | 02/17/2023 | 2023 | 45 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16567729&pc=QCB | 2023826-2023-00867 | 2023-02-09 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.50/+1.5/101 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPORTED AS LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE REPORTER INDICATED AT THIS TIME NO EXCHANGE IS PLANNED, THE SURGEON WILL CLOSELY MONITOR THE PATIENT. THE PROBLEM WAS RESOLVED AND THE PATIENT IS HAPPY. THE LENS REMAINED IMPLANTED. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -08.50/+1.5/101 | OS | Physician | 02/16/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16570033&pc=QCB | 2023826-2023-00973 | 2022-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -15.5/3.5/077 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND IMPLANTED VERTICALLY; AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9-RETURN DATE: 04/17/2023. H3:DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -15.5/3.5/077 | OS | Physician | 03/16/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16579040&pc=QCB | 2023826-2023-00961 | 2023-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.0/072 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (OVER 45 YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -11.0/+1.0/072 | OD | Physician | 02/20/2023 | 2023 | 48 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16579098&pc=QCB | 2023826-2023-00962 | 2022-12-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.5/083 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+1.5/083 | OS | Physician | 02/18/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16580093&pc=QCB | 2023826-2023-00963 | 2023-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.5/110 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -9.5/+1.5/110 | OS | Physician | 02/18/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16585359&pc=QCB | 2023826-2023-00990 | 2022-12-31 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -9.5/3.5/107 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS INTRAOPERATIVELY IMPLANTED AND REMOVED DUE TO A LENS TEAR/BREAK DURING LOADING/INJECTION INTO THE EYE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "THE LENS WAS BROKEN DURING IMPLANTATION, AND THE PATIENT WAS RELUCTANT TO UNDERGO A SECOND OPERATION." Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -9.5/3.5/107 | OS | Physician | 02/20/2023 | 2023 | 50 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16585360&pc=QCB | 2023826-2023-00993 | 2022-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.0/2.0/001 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME SIZE LENS IMPLANTED HORIZTONALLY. REPORTEDLY, "THE FIRST TICL IS VERTICAL AND THE SECOND TICL IS HORIZONTAL" THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT INFORMATION UNKNOWN. ADVERSE EVENT PROBLEM: OFF-LABEL - ACD <3.0. CLAIM# (B)(4). | -9.0/2.0/001 | OD | Physician | 02/20/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16587349&pc=QCB | 2023826-2023-00964 | 2022-12-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+2.0/088 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.0/+2.0/088 | OD | Physician | 02/20/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16587625&pc=QCB | 2023826-2023-00966 | 2023-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+1.5/129 INTO THE PATIENT'S LEFT EYE (OS) ON 18-AUG-2022. THE LENS WAS EXCHANGED ON 11-FEB-2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE (B)(4): OFF-LABEL USE (ACD < 3.0MM) CLAIM# (B)(4). | -7.5/+1.5/129 | OS | Physician | 02/21/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16587632&pc=QCB | 2023826-2023-00941 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.50/+1.5/073 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS). THE LENS WAS REMOVED ON 13-FEB-2023. THE LENS WAS DISCARDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: H6: 2923 SHOULD BE ADDED FOR CORRECTION. B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.50/+1.5/073 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS). THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LENS ROTATION. THE LENS WAS DISCARDED. INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. CLAIM# (B)(4). Manufacturer Narrative: SEX: UNK WEIGHT: UNK ETHNICITY: UNK RACE: UNK EVET DATE: UNK IMPLANT DATE: UNK H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -07.50/+1.5/073 | OS | Physician | 02/21/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16588740&pc=QCB | 2023826-2023-00979 | 2022-11-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -12.0/1.0/013 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATED THE INITIAL LENS WAS IMPLANTED VERTICALLY, SO THE VAULT WAS LOW. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). | -12.0/1.0/013 | OS | Physician | 02/20/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16588823&pc=QCB | 2023826-2023-00983 | 2023-02-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -6.0/1.0/061 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON 01-SEP-2022. ON 20-FEB-2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS THAT WAS IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTICS BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: OFF-LABEL USE (ACD < 3.0MM) (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -6.0/1.0/061 | OD | Physician | 02/24/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16588927&pc=QCB | 2023826-2023-00967 | 2023-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+2.0/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -8.0/+2.0/090 | OD | Physician | 02/21/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16588928&pc=QCB | 2023826-2023-00968 | 2023-01-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+1.0/073 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -12.0/+1.0/073 | OS | Physician | 02/21/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16594725&pc=QCB | 2023826-2023-01018 | 2023-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.5/2.0/098 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS AND THIS RESOLVED THE PROBLEM. THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -11.5/2.0/098 | OD | Physician | 03/21/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16594739&pc=QCB | 2023826-2023-01017 | 2023-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -14.0/3.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT NOTED "AFTER THE CLINICIANS JUDGEMENT; THE TRANSVERSAL IMPLANT WAS CHANGED TO OBLIQUE IMPLANT". PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 17-MAY-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -14.0/3.0/089 | OD | Physician | 03/17/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16595138&pc=QCB | 2023826-2023-00954 | 2023-02-06 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.50/+1.5/083 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PATIENT WEIGHT: UNK. PATIENT ETHNICITY: UNK. PATIENT RACE: UNK. EXPLANT DATE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # 728400 | -06.50/+1.5/083 | OD | Physician | 02/23/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16595252&pc=QCB | 2023826-2023-00995 | 2023-02-21 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method | No Clinical Signs, Symptoms or Conditions | P030017 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023, A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -1.50/1.5/85 (SPHERE/CYLINDER/AXIS) LENS TORE/BREAK INJECTION/DELIVERY INTO THE RIGHT EYE (OD). THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. THE LENS WAS INTRAOPERATIVELY IMPLANTED AND REMOVED. CAUSE OF THE EVENT WAS REPORTED AS USER ERROR. Manufacturer Narrative: PATIENT INFO: UNK. USER ERROR. CLAIM# (B)(4). | -1.50/1.5/85 | OD | Physician | 02/21/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16595583&pc=QCB | 2023826-2023-00998 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-03-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MMVTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -10.00/2.0/74 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LOW VAULT WAS OBSERVED. CAUSE OF THE EVENT WAS REPORTED AS OTHER. Manufacturer Narrative: ADDITONAL INFORMATION: B5- "THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.0/74 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON 14/MARCH/2022. THE PATIENT EXPERIENCED A LOW VAULT. THE LENS REMAINS IMPLANTED. THE PROBLEM WAS RESOLVED. REPORTEDLY "THE CASE IS PERFECTLY FINE, PATIENT IS DOING WELL SO FAR, AND THERE IS NO PLAN TO CHANGE OR REPLACE THE LENS NOW". THE CAUSE OF THE EVENT IS OTHER." CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -10.00/2.0/74 | OD | Physician | 02/22/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16596848&pc=QCB | 2023826-2023-00916 | 2022-12-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6M VTICM5 12.6, -10.50/2.5/002 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK. CLAIM# (B)(4). | -10.50/2.5/002 | OS | Physician | 02/28/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16596851&pc=QCB | 2023826-2023-00915 | 2022-12-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM512.6, -11.0/3.0/179 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS REPLACED VERTICALLY WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK. CLAIM# (B)(4). | -11.0/3.0/179 | OD | Physician | 02/28/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16596856&pc=QCB | 2023826-2023-00922 | 2023-02-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.5/0.5/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.5/0.5/109 | OD | Physician | 02/28/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16596857&pc=QCB | 2023826-2023-00924 | 2022-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.5/1.0/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM # (B)(4). | -9.5/1.0/071 | OD | Physician | 03/01/2022 | 2022 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16596858&pc=QCB | 2023826-2023-00923 | 2022-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -6.5/1.5/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED VERTICALLY WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM # (B)(4). | -6.5/1.5/078 | OD | Physician | 02/28/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16596859&pc=QCB | 2023826-2023-00926 | 2023-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -8.50/3.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. REPORTEDLY, EYE QUIET LENS IN POSITION Manufacturer Narrative: CLAIM# (B)(4). | -8.50/3.0/092 | OD | Physician | 03/02/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16596862&pc=QCB | 2023826-2023-00925 | 2022-11-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.0/1.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, EVALUATION CODES - DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM # (B)(4). | -9.0/1.0/095 | OS | Physician | 03/01/2023 | 2023 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16596865&pc=QCB | 2023826-2023-00927 | 2023-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -9.5/3.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION ON (B)(6)2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK. CLAIM# (B)(4). | -9.5/3.5/089 | OS | Physician | 03/02/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16596868&pc=QCB | 2023826-2023-00933 | 2023-02-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -5.5/1.5/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: = UNK. DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM, UNDER 21 YRS OF AGE AT DATE OF IMPLANT. (B)(4). | -5.5/1.5/112 | OD | Physician | 03/03/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16596872&pc=QCB | 2023826-2023-00931 | 2023-01-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.0/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM# (B)(4). | -9.0/1.5/093 | OS | Physician | 03/03/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16597323&pc=QCB | 2023826-2023-00936 | 2023-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.0/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION: UNKNOWN. DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM # (B)(4). | -12.0/1.0/091 | OS | Physician | 03/03/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16597325&pc=QCB | 2023826-2023-00935 | 2023-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.0/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, UNDER 21 YRS OF AGE AT DATE OF IMPLANT. CLAIM # (B)(4). | -12.0/1.5/091 | OS | Physician | 03/03/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16600879&pc=QCB | 2023826-2023-01063 | 2023-03-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -13.0/+3.0/99 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS EXCESSIVE VAULTING. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.0/+3.0/99 | OD | Physician | 03/02/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16601914&pc=QCB | 2023826-2023-00970 | 2023-01-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.50/+3.50/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS NOT RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -6.50/+3.50/093 | OD | Physician | 02/23/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16601952&pc=QCB | 2023826-2023-00971 | 2023-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.50/+3.0/081 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023 AS A REPLACEMENT LENS. THIS REPLACEMENT LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. INITIAL LENS INFORMATION CAN BE SEEN IN MFR REPORT. Manufacturer Narrative: CLAIM#: (B)(4). | -6.50/+3.0/081 | OD | Physician | 02/23/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16606536&pc=QCB | 2023826-2023-01001 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MMVTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -9.0/1.5/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LOW VAULT WAS OBSERVED. CAUSE OF THE EVENT WAS REPORTED AS OTHER. Manufacturer Narrative: PATIENT INFORMATION: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -9.0/1.5/088 | OS | Physician | 02/22/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16606537&pc=QCB | 2023826-2023-01002 | 2023-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -16.50/3.0/068 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. CAUSED OF THE EVENT WAS A PATIENT RELATED FACTOR. REPORTEDLY, "NARROW SULCUS." THE STATUS OF THE EYE WAS STATED AS, "EYE IS WELL." PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT & RACE :UNK. PATIENT RELATED FACTOR. CLAIM# (B)(4). | -16.50/3.0/068 | OD | Physician | 02/23/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16609603&pc=QCB | 2023826-2023-00984 | 2023-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.50/2.00/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. IT WAS REPORTED THAT A REPLACEMENT LENS WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. NO PATIENT INJURY WAS REPORTED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.50/2.00/103 | OS | Physician | 02/24/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16613558&pc=QCB | 2023826-2023-01086 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2 -10.00/1.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE DURING THE LOADING PROCESS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -10.00/1.0/089 | NA | Other Health Care Professional | 02/22/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16615097&pc=QCB | 2023826-2023-01085 | 2023-02-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2 -16.50/1.0/033 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REMOVED DUE TO IT BEING IMPLANTED UPSIDE DOWN. NO PATIENT INJURY. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: D4. MODEL#: "VICMO 13.2" SHOULD BE REMOVED AND "VTICMO13.2" SHOULD BE ADDED. H6-HEALTH EFFECT- IMPACT CODE: "2645" SHOULD BE REMOVED AND CODE "2199" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "3026" SHOULD BE ADDED. ADDITONAL INFORMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.50/1.0/033 (SPHERE/CYLINDER/AXIS) INTO THE LEFT EYE (OS) ON (B)(6)2023. THE LENS WAS IMPLANTED UPSIDE DOWN. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED AND REPLACED INTRAOPERATIVELY WITH A LENS OF THE SAME PARAMETERS AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION PROVIDED: "ORIGINAL LENS REMOVED AND REPLACED WITH BACK UP LENS. ORIGINAL LENS WENT IN UPSIDE DOWN. NO COMPLICATION, PATIENT DOING WELL". THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS THE DEVICE. CAUSE DETAILS PROVIDED: " LENS FLIPPED DURING DELIVERY INTO EYE. LENS WAS UPSIDE DOWN. LENS CUT IN ½". CLAIM# (B)(4). | -16.50/1.0/033 | OS | Other Health Care Professional | 02/24/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16620652&pc=QCB | 2023826-2023-01019 | 2023-01-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.5/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION UNKNOWN. ADVERSE EVENT DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM # (B)(4). | -10.5/1.5/090 | OD | Physician | 03/06/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16620681&pc=QCB | 2023826-2023-01027 | 2023-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.0/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE PROBLEM CODE: (B)(4) - OFF-LABEL USE, UNDER 21 YRS OF AGE AT DATE OF IMPLANT. CLAIM# (B)(4). | -11.0/1.5/090 | OD | Physician | 03/08/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16620682&pc=QCB | 2023826-2023-01024 | 2022-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/4.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.0/4.5/088 | OD | Physician | 03/06/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16620701&pc=QCB | 2023826-2023-01028 | 2023-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.0/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE PROBLEM CODE: (B)(4) - OFF-LABEL USE, UNDER 21 YRS OF AGE AT DATE OF IMPLANT. CLAIM# (B)(4). | -10.0/1.5/090 | OS | Physician | 03/08/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16622851&pc=QCB | 2023826-2023-01089 | 2023-03-07 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT WHILE ATTEMPTING TO IMPLANT A 12.6MM VTICM5_12.6; -11.50/+1.0/110 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS); THE LENS HAD TORE DURING INJECTION. THE LENS WAS NOT IMPLANTED. THERE WAS PATIENT CONTACT BUT NO INJURY. THIS OCCURRED ON (B)(6) 2023. ON (B)(6) 2023 A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.50/+1.0/110 | OS | Physician | 03/15/2023 | 2023 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16625158&pc=QCB | 2023826-2023-01091 | 2023-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+2.5/088 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -12.5/+2.5/088 | OS | Physician | 02/27/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16625159&pc=QCB | 2023826-2023-01093 | 2023-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+1.0/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023 . THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -6.5/+1.0/095 | OD | Physician | 02/27/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16625169&pc=QCB | 2023826-2023-01098 | 2023-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+1.5/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -13.0/+1.5/092 | OD | Physician | 02/27/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16625170&pc=QCB | 2023826-2023-01097 | 2023-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+1.5/073 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: ADVERSE EVENT PROBLEM DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#(B)(4). | -13.5/+1.5/073 | OS | Physician | 02/27/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16626826&pc=QCB | 2023826-2023-01023 | 2023-01-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.0/4.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.0/4.5/083 | OS | Physician | 03/07/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16633012&pc=QCB | 2023826-2023-01100 | 2023-02-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+3.5/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -14.0/+3.5/093 | OD | Physician | 02/28/2023 | 2023 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16633014&pc=QCB | 2023826-2023-01101 | 2023-01-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+1.0/109 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023 . THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -14.0/+1.0/109 | OD | Physician | 03/01/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16633212&pc=QCB | 2023826-2023-01011 | 2023-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -11.0/1.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED AND THE LONGER LEGNTH LENS WAS IMPLANTED. CAUSE OF THE EVENT IS UNKNOWN. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: B5: LENS WAS REMOVED AND REPLACED DUE TO LOW VAULT WITHOUT ROTATION. CLAIM# (B)(4). Manufacturer Narrative: OFF-LABEL - ACD <3.0. CLAIM# (B)(4). | -11.0/1.0/094 | OS | Physician | 02/27/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16636086&pc=QCB | 2023826-2023-01072 | 2023-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00/+2.5/150 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_ 13.2 IMPLANTABLE COLLAMER LENS, -12.00/+2.5/150 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPORTED AS EXCESSIVE VAULTING. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.00/+2.5/150 | OS | Physician | 03/02/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16636132&pc=QCB | 2023826-2023-01070 | 2023-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.50/+1.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LOW VAULTING AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.50/+1.0/093 | OD | Physician | 03/01/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16638141&pc=QCB | 2023826-2023-01074 | 2023-02-28 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.50/+1.5/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE REPORTER INDICATED THE TORIC LENS DID NOT HAVE POSITIONING MARKS ON THE LENS. THE LENS REMAINED IMPLANTED BECAUSE THE PATIENT DID NOT HAVE A PROBLEM WITH THE LENS. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: PATIENTS INFORMATION: UNK. EXPLANT DATE: NA. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -16.50/+1.5/090 | OS | Physician | 03/02/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16638183&pc=QCB | 2023826-2023-01071 | 2023-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -14.00/+3.5/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LOW VAULTING AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC STRETCHED OUT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -14.00/+3.5/082 | OS | Physician | 03/01/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16639377&pc=QCB | 2023826-2023-01118 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -13.00/2.5/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6), 2023. LENS ROTATION AFTER IMPLANT WAS OBSERVED, LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -13.00/2.5/095 | OS | Physician | 02/27/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16639378&pc=QCB | 2023826-2023-01119 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 -14.50/1.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LENS ROTATION AFTER IMPLANT WAS OBSERVED, LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -14.50/1.5/089 | OD | Physician | 02/27/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16640304&pc=QCB | 2023826-2023-01120 | 2023-03-09 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT WHILE ATTEMPTING TO IMPLANT A 13.2MM VTICM5_13.2; -9.50/+1/93 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD); THE LENS HAD TORE. THIS OCCURRED ON (B)(6) 2023. THERE WAS PATIENT CONTACT WITH NO INJURY. THE LENS WAS NOT IMPLANTED. ON (B)(6) 2023 A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: B5 - PROBLEM WAS RESOLVED, PATIENT VERY SATISFIED. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -9.50/+1/93 | OD | Physician | 03/10/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16640595&pc=QCB | 2023826-2023-01102 | 2023-01-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+1.0/074 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -13.5/+1.0/074 | OS | Physician | 03/01/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16644571&pc=QCB | 2023826-2023-01128 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -11.0/1.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED AND LATER REPLACED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND HAPTIC TORN TO THE LENS CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION UNKNOWN. CLAIM# (B)(4). | -11.0/1.5/092 | OS | Physician | 03/01/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16644900&pc=QCB | 2023826-2023-01127 |  | Injury | STAAR SURGICAL COMPANY | 2023-03-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -12.0/1.0/078 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) (B)(6) 2023. THE LENS WAS REMOVED AND LATER REPLACED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. Manufacturer Narrative: DEVICE EVALUATON: H3- TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. ADDITIONAL INFORMATION: H6- TYPE OF INVESTIGATION CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CORRECTED DATA: H6- HEALTH EFFECT- IMPACT CODE: "4629" SHOULD BE REMOVED AND CODE "4627" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. CLAIM# (B)(4). | -12.0/1.0/078 | OD | Physician | 03/01/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16646877&pc=QCB | 2023826-2023-01136 | 2023-02-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -10.5/2.5/112 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/2.5/112 | OD | Physician | 02/28/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16648640&pc=QCB | 2023826-2023-01138 | 2022-12-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -13.0/2.0/013 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) UPSIDE DOWN ON (B)(6) 2022. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY WITH NO PATIENT INJURY. ON (B)(6) 2022 A REPLACEMENT LENS WAS IMPLANTED VERTICALLY AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.0/2.0/013 | OD | Physician | 02/28/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16649544&pc=QCB | 2023826-2023-01139 | 2023-01-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -12.0/2.5/072 (SPHERE/CYLINDER/AXIS) TORE/BROKE DURING LOADING, NO PATIENT CONTACT. THE REPORTER STATES THAT THE SURGEON IS NEW AND WAS NERVOUS DURING OPERATION. ON (B)(6) 2023 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -12.0/2.5/072 | NA | Physician | 02/28/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16649708&pc=QCB | 2023826-2023-01079 | 2022-12-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50/+5.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTICS BENT. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN; ETHNICITY: UNKNOWN; RACE: UNKNOWN; WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # 728506 | -14.50/+5.0/090 | OD | Physician | 03/03/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16650145&pc=QCB | 2023826-2023-01080 | 2023-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+2.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS BUT THE PROBLEM WAS NOT RESOLVED. SEE MFR. REPORT # 2023826-2023-01081 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.50/+2.5/093 | OD | Physician | 03/03/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16650149&pc=QCB | 2023826-2023-01081 | 2023-01-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50/+2.5/097 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS REPORTED AS LOW VAULTING. THIS LENS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-01080. THE REPORTER INDICATED POST-OP THE PATIENTS CURRENT ARCH HEIGHT IS 200UM. THE SURGEON DECIDED TO OBSERVE AND NOT REPLACE IT TEMPORARILY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.50/+2.5/097 | OD | Physician | 03/03/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16652193&pc=QCB | 2023826-2023-01109 | 2023-02-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+1.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.0/+1.0/093 | OD | Physician | 03/02/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16652195&pc=QCB | 2023826-2023-01104 | 2022-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-03-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.0/066 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -10.5/+1.0/066 | OD | Physician | 03/02/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16652196&pc=QCB | 2023826-2023-01108 | 2023-01-09 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+2.0/002 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.0/+2.0/002 | OD | Physician | 01/09/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16652198&pc=QCB | 2023826-2023-01110 | 2023-02-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-03-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+1.5/095 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.5/+1.5/095 | OS | Physician | 03/02/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16669120&pc=QCB | 2023826-2023-01196 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-04-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 -7.00/1.0/072 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE. EXCESSIVE VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: IMPLANT DATE: UNK. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -7.00/1.0/072 | NA | Physician | 03/06/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16673866&pc=QCB | 2023826-2023-01034 | 2023-03-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.0/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY & RACE UNKNOWN. CLAIM # (B)(4). | -12.0/1.5/092 | OD | Physician | 03/13/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16673867&pc=QCB | 2023826-2023-01036 | 2023-01-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -8.5/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT - UNK. PATIENT ETHNICITY - UNK. PATIENT RACE- UNK. DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM#(B)(4). | -8.5/2.5/092 | OD | Physician | 03/12/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16673884&pc=QCB | 2023826-2023-01046 | 2023-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6M VTICM5 12.6, -11.0/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/1.5/092 | OD | Physician | 03/15/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16673895&pc=QCB | 2023826-2023-01045 | 2023-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -6.0/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, UNDER 21 YRS OF AGE AT DATE OF IMPLANT. (B)(4). | -6.0/1.0/085 | OS | Physician | 03/14/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16673898&pc=QCB | 2023826-2023-01043 | 2022-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.00/1.5/174 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -18.00/1.5/174 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE SAME MODEL, SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. CLAIM #: (B)(4). | -18.00/1.5/174 | OS | Physician | 03/14/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16673903&pc=QCB | 2023826-2023-01055 | 2022-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.0/2.0/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM# (B)(4). | -12.0/2.0/105 | OS | Physician | 03/16/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16673905&pc=QCB | 2023826-2023-01047 | 2023-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM512.6, -10.0/3.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). | -10.0/3.0/088 | OS | Physician | 03/15/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16673926&pc=QCB | 2023826-2023-01033 | 2023-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.0/2.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.0/2.0/079 | OD | Physician | 03/13/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16673937&pc=QCB | 2023826-2023-01056 | 2022-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.5/1.0/069 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, DOCTOR DECIDED TO ROTATE LENS 41 DEGREES CLOCKWISE. Manufacturer Narrative: (B)(4). | -12.5/1.0/069 | OD | Physician | 03/16/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16678114&pc=QCB | 2023826-2023-01244 | 2023-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -14.5/3.5/085 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON REPORTS LENS ROTATION. ON (B)(6) 2022 THE LENS WAS REPOSITIONED BUT THIS ONLY TEMPORARILY RESOLVED THE PROBLEM. THEN ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IT WAS ADDITIONALLY NOTED "FIRST LENS GOT ROTATED VERTICALLY TWICE. EXCHANGED WITH A HIGHER SIZE LENS AND OBSERVED FOR A MONTH AND ITS STABLE NOW". Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # :(B)(4). | -14.5/3.5/085 | OD | Physician | 03/28/2023 | 2023 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16678115&pc=QCB | 2023826-2023-01213 | 2023-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.50/+2.0/112 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THERE WAS LENS DISLOCATION/SUBLUXATION. ON (B)(6) 2023 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED "THE LENS (AND SO AQUAPORT) WAS DISLOCATED TO SUPERIOR (NOT CENTRAL) SO THE PATIENT SAW LIGHTPOINTS TWICE AT THE DARK". Manufacturer Narrative: ASVERSE EVENT PROBLEM - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.50/+2.0/112 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THERE WAS LENS DISLOCATION/SUBLUXATION. ON (B)(6) 2023 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH DIFFERENT MODEL LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED "THE LENS (AND SO AQUAPORT) WAS DISLOCATED TO SUPERIOR (NOT CENTRAL) SO THE PATIENT SAW LIGHTPOINTS TWICE AT THE DARK". ADDITIONAL INFORMATION REPORTED, THE SURGEON HAS DECIDED TO PLACE THE ICL VERTICAL INTO THE EYE. THAT´S WHY THE AXIS FROM THE REPLACEMENT LENS IS DIFFERENT TO THE AXIS FROM THE INITIAL LENS. PATIENT IS VERY HAPPY! D6B: EXPLANTATION DATE: (B)(6) 2023. CLAIM# (B)(4). | -11.50/+2.0/112 | OD | Physician | 03/21/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16678116&pc=QCB | 2023826-2023-01161 |  | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM TMICL12.1 IMPLANTABLE COLLAMER LENS, -14.50/+4.0/108 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO THE LENS WAS TOO SMALL. THE LENS WAS REPLACED WITH A LONGER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. CORRECTED DATA: H6-HEALTH EFFECT- IMPACT CODE: "4629" SHOULD BE REMOVED AND CODE "4627" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -14.50/+4.0/108 | OS | Other Health Care Professional | 03/06/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16678781&pc=QCB | 2023826-2023-01179 | 2023-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -7.0/1.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE CAUSE OF THE EVENT IS OTHER AND THE DEVICE WAS REPORTED AS FAILED TO PERFORM AS INTENDED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PATIENT DEMOGRAPHICS:UNK. ADVERSE EVENT PROBLEM: OFF-LABEL - ACD<3.0 AGE<21. CLAIM# (B)(4). | -7.0/1.0/092 | OS | Physician | 03/06/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16686635&pc=QCB | 2023826-2023-01204 | 2023-02-28 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 13.2MM, VTICMO13.2, -11.0/1.5/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023.THE LENS WAS IMPLANTED, REMOVED AND REPLACED INTRAOPERATIVELY WITH NO PATINET INJURY AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -11.0/1.5/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.0/1.5/084 | OD | Physician | 03/08/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16686636&pc=QCB | 2023826-2023-01265 | 2023-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+3.0/105 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LOW VAULT AND A LONGER LENGTH LENS WAS IMPLANTED, ALTHOUGH IT IS UNKNOWN IF THIS WAS DONE WITHIN THE SAME INITIAL SURGERY. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+3.0/105 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: PER EMAIL RECEIVED ON 11-APR-2023, THE REPLACEMENT LENS WAS IMPLANTED WITHIN THE SAME SURGERY THAT THE INITIAL LENS WAS REMOVED. H6- HEALTH EFFECT - IMPACT CODE: UPDATED FROM 4627 TO 4629. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -13.5/+3.0/105 | OD | Physician | 03/07/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16686638&pc=QCB | 2023826-2023-01268 | 2023-02-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.5/+1.0/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -4.5/+1.0/091 | OS | Physician | 03/08/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16686642&pc=QCB | 2023826-2023-01266 | 2023-01-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LOW VAULT AND LENS ROTATION. A LONGER LENGTH LENS WAS IMPLANTED ON THE SAME DAY, ALTHOUGH IT IS UNKNOWN IF THIS WAS DONE WITHIN THE SAME INITIAL SURGERY. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- PER EMAIL RECEIVED ON 05-APR-2023, THE REPLACEMENT LENS WAS IMPLANTED WITHIN THE SAME SURGERY THAT THE INITIAL LENS WAS REMOVED. H6- HEALTH EFFECT - IMPACT CODE: UPDATED FROM 4627 TO 4629 CLAIM# (B)(4). | -10.0/+2.0/093 | OD | Physician | 03/07/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16686645&pc=QCB | 2023826-2023-01200 | 2022-09-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 12.6MM, VTICM5_12.6, -10.0/2.5/76 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING LOADING, THERE WAS NO PATIENT CONTACT. A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.0/2.5/76 | NA | Physician | 03/08/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16686689&pc=QCB | 2023826-2023-01267 | 2023-02-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.0/+1.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -4.0/+1.0/091 | OD | Physician | 03/08/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16686813&pc=QCB | 2023826-2023-01188 | 2023-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.5/0.5/017 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE LOW VAULT WITHOUT ROTATION. THE PROBLEM IS RESOLVED. REPORTEDLY, "THE SURGEON CONSIDERED THAT THE HEIGHT OF THE LARGEST LENS ARCH WOULD NOT CHANGE, SO HE DECIDED NOT TO REPLACE IT TO AVOID ADVERSE EVENTS." Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC BROKEN. DUE TO SIGNIFICANT LENS HAPTIC DAMAGE, UNABLE TO PERFORM DIMENSIONAL INSPECTION. CLAIM# (B)(4) Manufacturer Narrative: OFF-LABEL ACD <3.0. CLAIM# (B)(4). | -11.5/0.5/017 | OS | Physician | 03/09/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16688612&pc=QCB | 2023826-2023-01203 | 2023-01-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -8.5/2.5/114 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS IMPLANTED, REMOVED AND REPLACED INTRAOPERATIVELY WITH NO PATIENT INJURY. IT WAS REPORTED THAT THE PROBLEM RESOLVED . CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -8.5/2.5/114 | OD | Physician | 03/08/2023 | 2023 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16688761&pc=QCB | 2023826-2023-01201 | 2022-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -8.0/1.5/58 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED AND THE PROBLEM RESOLVED . THE REPORTER STATES THAT THE CAUSE OF THIS EVENT IS DUE TO PATIENT REALTED FACTORS AND USER ERROR. THE REPORTER ALSO STATES THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -8.0/1.5/58 | OD | Physician | 03/08/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16690586&pc=QCB | 2023826-2023-01208 | 2023-03-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS -9.5/1.0/111 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/1.0/111 | OD | Physician | 03/09/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16691033&pc=QCB | 2023826-2023-01171 | 2023-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+0.50/109 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS HAD LOW VAULTING AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.00/+0.50/109 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED A LOW VAULT AND REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. D9: RETURN DATE: 05-SEP-2023. CLAIM (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6: TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). | -18.00/+0.50/109 | OS | Physician | 03/08/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16691040&pc=QCB | 2023826-2023-01176 | 2023-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/+3.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2022 BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY:UNK. RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). Manufacturer Narrative: CORRECTED DATA: B3. DATE OF EVENT: "07/26/2023" SHOULD BE REMOVED. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -11.00/+3.0/090 | OD | Physician | 03/07/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16691065&pc=QCB | 2023826-2023-01175 | 2022-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+2.5/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM # (B)(4). | -10.50/+2.5/086 | OS | Physician | 03/07/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16691110&pc=QCB | 2023826-2023-01172 | 2022-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+2.5/109 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022, THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE LENS IS STABLE AND IN POSITION. Manufacturer Narrative: ADVERSE EVENT PROBLEM - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -11.50/+2.5/109 | OD | Physician | 03/09/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16691151&pc=QCB | 2023826-2023-01170 | 2023-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+0.50/054 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS HAD LOW VAULTING AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTION (FUNCTIONAL) VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2023 FOR A LONGER LENS BUT THE PROBLEM WAS NOT RESOLVED. THE REPLACEMENT LENS STILL HAD A LOW VAULT. SEE MFR. REPORT # 2023826-2023-02274 FOR REPLACEMENT LENS. CLAIM # (B)(4). | -18.00/+0.50/054 | OD | Physician | 03/08/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16691789&pc=QCB | 2023826-2023-01206 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -9.50/4.0/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED LOW VAULT . LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). | -9.50/4.0/086 | OS | Other Health Care Professional | 03/08/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16692725&pc=QCB | 2023826-2023-01209 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 -4.00/2.0/082 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE DURING LOADING, NO PATIENT CONTACT. THE REPORTER STATED " THE LENS GOT STUCK ON THE LIOLI (DELIVERY SYSTEM) WHEN PUSHING IT INTO THE PRELOADING POSITION AND IT EITHER CRINKLED OR RIPPED. THE BACKUP LENS WAS IMPLANTED SUCCESSFULLY. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -4.00/2.0/082 | NA | Other Health Care Professional | 03/09/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16694160&pc=QCB | 2023826-2023-01190 | 2022-12-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM12.6 IMPLANTABLE COLLAMER LENS OF -7.0/1.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS ROTATION WAS OBSERVED. ON (B)(6) 2022, LENS REPOSITIONING WAS PERFORMED, AND WAS REPORTED THAT IT WAS RESOLVED FOR SOMETIME. ON (B)(6) 2022, LENS ROTATION WAS ONCE AGAIN OBSERVED. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. REPORTEDLY, "INITIAL LENS WAS UN STABLE AND GOT ROTATED TWICE VERTICALLY. WITH THE EXCHANGE OF BIGGER SIZE LENS, PATIENT IS DOING WELL." CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -7.0/1.0/091 | OD | Physician | 03/10/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16694161&pc=QCB | 2023826-2023-01191 | 2023-03-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.5/2.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS INTRAOPERATIVELY IMPLANTED AND REMOVED DUE TO A TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT, BUT NO PATIENT INJURY. ON (B)(6) 2023, THE REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "OBSERVED SMALL TEAR ON THE LENS AFTER IMPLANTATION AND REMOVED THE IMMEDIATELY. IMPLANTED A NEW LENS AND IT WENT SMOOTH AND PATIENT IS DOING GOOD." Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -10.5/2.0/094 | OD | Physician | 03/10/2023 | 2023 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16694222&pc=QCB | 2023826-2023-01298 | 2022-12-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.5/+1.5/005 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION. ON(B)(6) 2022 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS BUT THIS ALSO DID NOT RESOLVE THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED "HIPERLAXITY" AND THE DEVICE DID FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4) Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -12.5/+1.5/005 | OS | Physician | 03/31/2023 | 2023 | 50 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16694224&pc=QCB | 2023826-2023-01294 | 2023-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -8.0/+3.0/1 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED THE DEVICE DID FAIL TO PERFORM AS INTENDED; "FAIL IN THE MEASURE OF PREDICTABLE VAULT". Manufacturer Narrative: (B)(4). | -8.0/+3.0/1 | OD | Physician | 03/31/2023 | 2023 | 41 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16695794&pc=QCB | 2023826-2023-01296 | 2023-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+1.0/125 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.0/+1.0/125 | OS | Physician | 03/07/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16695796&pc=QCB | 2023826-2023-01297 | 2023-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+1.0/072 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.0/+1.0/072 | OD | Physician | 03/07/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16696083&pc=QCB | 2023826-2023-01295 | 2023-02-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+3.0/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+3.0/091 | OS | Physician | 03/07/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16705021&pc=QCB | 2023826-2023-01273 | 2023-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+2.0/066 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.5/+2.0/066 | OS | Physician | 03/12/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16705022&pc=QCB | 2023826-2023-01105 | 2022-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.0/118 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -7.0/+1.0/118 | OD | Physician | 03/02/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16710536&pc=QCB | 2023826-2023-01301 | 2023-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+4.0/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.0/+4.0/090 | OS | Physician | 03/14/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16710538&pc=QCB | 2023826-2023-01303 | 2023-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/+5.5/077 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -5.5/+5.5/077 | OS | Physician | 03/14/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16710643&pc=QCB | 2023826-2023-01331 | 2022-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -12.5/3/2 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS IMPLANTED HORIZONTALLY DUE TO LOW VAULT WITH ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. REPORTEDLY, "INITIAL LENS: IMPLANTED VERTICALLY, REPLACEMENT LENS: IMPLANTED HORIZONTALLY." Manufacturer Narrative: D9-RETURN DATE: 04/06/2023. H3:DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND STRETCHED OUT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: H6- 2923 ADDED FOR LENS ROTATION. CLAIM#(B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -12.5/3/2 | OD | Physician | 03/13/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16710644&pc=QCB | 2023826-2023-01330 | 2022-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -10.50/2.0/001 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS IMPLANTED HORIZONTALLY DUE TO LOW VAULT WITH ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. REPORTEDLY, "INITIAL LENS: IMPLANTED VERTICALLY, REPLACEMENT LENS: IMPLANTED HORIZONTALLY. PERFORMED LASEK FOR ASTIGMATISM." Manufacturer Narrative: ADDITIONAL INFORMATION: D9-RETURN DATE: 04/06/2023. H3:DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: H6- 2923 ADDED FOR LENS ROTATION. H6-4625-ADDED FOR CORNEAL REFRACTIVE SURGERY. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -10.50/2.0/001 | OD | Physician | 03/13/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16713981&pc=QCB | 2023826-2023-01220 | 2023-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -5.5/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM, UNDER 21 YRS OF AGE AT DATE OF IMPLANT. CLAIM# (B)(4). | -5.5/1.0/091 | OS | Physician | 03/22/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16713983&pc=QCB | 2023826-2023-01218 | 2023-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -4.0/3.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM # (B)(4). | -4.0/3.5/085 | OD | Physician | 03/19/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16713987&pc=QCB | 2023826-2023-01219 | 2023-03-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -5.5/1.0/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2023 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN ETHNICITY: UNKNOWN RACE: UNKNOWN DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, UNDER 21 YRS OF AGE AT DATE OF IMPLANT CLAIM #(B)(4). | -5.5/1.0/109 | OD | Physician | 03/20/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16713988&pc=QCB | 2023826-2023-01223 | 2023-03-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -5.5/2.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2023 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WEIGHT- UNKNOWN. ETHNICITY - UNKNOWN. RACE - UNKNOWN. CLAIM# (B)(4). | -5.5/2.0/082 | OD | Physician | 03/23/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16713989&pc=QCB | 2023826-2023-01215 | 2023-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.0/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN ETHNICITY: UNKNOWN RACE: UNKNOWN CLAIM # (B)(4). | -9.0/1.0/093 | OS | Physician | 03/19/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16713990&pc=QCB | 2023826-2023-01221 | 2023-03-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.0/3.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN. ETHNICITY:UNKNOWN. RACE: UNKNOWN. CLAIM# (B(4). | -13.0/3.0/094 | OS | Physician | 03/22/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16713996&pc=QCB | 2023826-2023-01222 | 2022-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -10.5/1.5/178 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, VAULT IS ADEQUATE AND PATIENT IS DOING WELL. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY - UNK. RACE - UNK. CLAIM# (B)(4). | -10.5/1.5/178 | OD | Physician | 03/23/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16714003&pc=QCB | 2023826-2023-01224 | 2023-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -13.5/2.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE - UNK. CLAIM# (B)(4). | -13.5/2.5/082 | OS | Physician | 03/23/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16720052&pc=QCB | 2023826-2023-01355 | 2023-03-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -18.0/3.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A BACK-UP LENS OF THE SAME LENGTH. IT WAS ADDITIONALLY NOTED "IN THE PROCESS OF IMPLANTATION THE TICL CRYSTAL IN PATIENT'S LEFT EYE; ONE OF THE LOOPS OF THE CRYSTAL WAS KNOCKED OFF BY THE SURGICAL INSTRUMENT WHILE ADJUSTING THE POSITION OF THE CRYSTAL. LATER THE SPARE CRYSTAL WAS USED FOR EMERGENCY AND THE OPERATION WAS COMPLETED. POSTOPERATIVE ARCH HEIGHT WAS 300UM AND INTRAOCULAR PRESSURE WAS 25MMHG". THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -18.0/3.0/090 | OS | Physician | 04/03/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16720056&pc=QCB | 2023826-2023-01356 | 2023-01-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.0/1.5/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED LENS ROTATION. ON (B)(6) 2021 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. THEN ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT SPHERICAL. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -12.0/1.5/094 | OD | Physician | 04/09/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16725717&pc=QCB | 2023826-2023-01246 | 2023-02-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -04.50/+2.5/087 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. PATIENT IS DOING GREAT. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM #:(B)(4). | -04.50/+2.5/087 | OD | Other Health Care Professional | 03/15/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16725736&pc=QCB | 2023826-2023-01251 | 2022-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.00/+4.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT POWER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT-RELATED FACTOR AND THE DEVICE DID FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTION (FUNCTIONAL) VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.00/+4.5/090 | OS | Physician | 03/17/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16725827&pc=QCB | 2023826-2023-01248 | 2023-03-01 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.00/+2.0/076 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -14.00/+2.0/076 | OD | Physician | 03/14/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16726378&pc=QCB | 2023826-2023-01307 | 2023-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+0.5/040 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.0/+0.5/040 | OD | Physician | 03/14/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16726545&pc=QCB | 2023826-2023-01360 | 2023-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+2.0/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.5/+2.0/092 | OD | Physician | 03/13/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16727246&pc=QCB | 2023826-2023-01347 | 2023-03-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -5.5/0.5/86 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS THAT WAS IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADVERSE EVENT PROBLEM - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: THREE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H6- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID WITH RESIDUE/DEBRIS ON PRODUCT, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS . DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | -5.5/0.5/86 | OD | Physician | 03/14/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16727869&pc=QCB | 2023826-2023-01247 | 2023-02-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -04.50/+1.5/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. PATIENT IS DOING GREAT. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY & RACE UNKNOWN. CLAIM # (B)(4). | -04.50/+1.5/089 | OS | Other Health Care Professional | 03/15/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16727873&pc=QCB | 2023826-2023-01255 |  | Injury | STAAR SURGICAL COMPANY | 2023-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.00/+1.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY & RACE UNKNOWN. EVENT DATE: UNKNOWN. CLAIM # (B)(4). | -13.00/+1.0/094 | OS | Physician | 03/14/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16727880&pc=QCB | 2023826-2023-01254 | 2023-02-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO THE CRYSTAL HAD NO ASTIGMATISM MARKER AND THE POSTOPERATIVE DIOPTER AND ARCH HEIGHT ARE ABNORMAL. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY & RACE UNKNOWN. ADVERSE EVENT PROBLEM - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL WIDTH VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. DIMENSIONAL LENGTH AND REFRACTIVE (FUNCTIONAL) VERIFICATION WAS PERFORMED AND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #(B)(4). | -10.00/+1.0/087 | OD | Physician | 03/14/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16728762&pc=QCB | 2023826-2023-01310 | 2023-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0/+4.5/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -15.0/+4.5/089 | OS | Physician | 03/14/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16728763&pc=QCB | 2023826-2023-01308 | 2023-02-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+1.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.5/+1.0/093 | OD | Physician | 03/14/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16728765&pc=QCB | 2023826-2023-01313 | 2023-02-23 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -2.0/+0.5/099 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -2.0/+0.5/099 | OS | Physician | 03/15/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16728766&pc=QCB | 2023826-2023-01311 | 2022-10-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+1.5/062 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LENS TWO SIZES LONGER IN LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -8.0/+1.5/062 | OS | Physician | 03/14/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16728768&pc=QCB | 2023826-2023-01315 | 2023-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+2.5/080 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -12.5/+2.5/080 | OD | Physician | 03/15/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16735813&pc=QCB | 2023826-2023-01322 | 2023-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+3.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: ADVERSE EVENT PROBLEM DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -7.0/+3.0/091 | OD | Physician | 03/16/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16735842&pc=QCB | 2023826-2023-01320 | 2023-01-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.0/095 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -9.0/+1.0/095 | OS | Physician | 03/16/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16735846&pc=QCB | 2023826-2023-01325 | 2023-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+3.5/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE. (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -9.0/+3.5/094 | OS | Physician | 03/17/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16735922&pc=QCB | 2023826-2023-01321 | 2023-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+2.0/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT. THAT SAME DAY, THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS ALTHOUGH IT IS UNCLEAR WHETHER THIS WAS DONE WITHIN THE SAME PROCEDURE. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) AND (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -9.0/+2.0/092 | OS | Physician | 03/16/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16735931&pc=QCB | 2023826-2023-01324 | 2023-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+1.0/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -14.0/+1.0/092 | OS | Physician | 03/17/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16738041&pc=QCB | 2023826-2023-01338 | 2023-02-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -8/1.5/073 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS A REPLACEMENT LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REPORTEDLY, "THE PATIENT WAS EXCESSIVE VAULTED AND UNDER CLINCICAL OBSERVATION." Manufacturer Narrative: A4-A6: UNK. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -8/1.5/073 | OS | Physician | 03/15/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16738042&pc=QCB | 2023826-2023-01342 | 2023-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -11.5/2.5/077 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2023, THE LENS WAS REMOVED AND LONGER LENGTH LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- IMPACT CODE: "4629" SHOULD BE REMOVED AND CODE "4627" SHOULD BE ADDED. H6- INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -11.5/2.5/077 | OS | Physician | 03/15/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16738043&pc=QCB | 2023826-2023-01339 | 2023-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -11.5/1.0/44 SPHERE/CYLINDER/AXIS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6), 2023. REPORTEDLY, SURGEON USED WRONG LENS POWER FOR CALCULATION +CYLINDER. CAUSE OF THE EVENT WAS USER ERROR. A LENS EXCHANGE IS PLANNED. Manufacturer Narrative: B5 - LENS WAS EXCHANGED ON (B)(6) 2023 WITH A SAME SIZE, DIFFERENT POWER LENS AND PROBLEM WAS RESOLVED. PATIENT IS HAPPY. CLAIM# (B)(4). Manufacturer Narrative: A4-A6: UNK. H6: USER ERROR. CLAIM#: (B)(4). | -11.5/1.0/44 | OD | Physician | 03/15/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16738047&pc=QCB | 2023826-2023-01341 | 2023-03-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.5/2.5/112 (SPHERE/CYLINDER/AXIS) TORE DURING LOADING AND WAS UNABLE TO BE IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- IMPACT CODE: "2199" SHOULD BE ADDED. H6 -MEDICAL DEVICE PROBLEM CODE: "2993" SHOULD BE REMOVED AND "3199" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM#: (B)(4). | -12.5/2.5/112 | OS | Physician | 03/15/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16738079&pc=QCB | 2023826-2023-01340 | 2023-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -11.0/2.5/84 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO LENS DISLOCATION/SUBLUXATION. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: B5 - REPORTEDLY, PROBLEM WAS RESOLVED. PATIENT IS HAPPY. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM#: (B)(4). | -11.0/2.5/84 | OD | Physician | 03/15/2023 | 2023 | 43 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16739726&pc=QCB | 2023826-2023-01371 |  | Injury | STAAR SURGICAL COMPANY | 2023-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -11.0/1.0/167 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2023THE LENS WAS EXPLANTED; AND ON THIS SAME DATE A REPLACEMENT LENS OF A LONGER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.0/1.0/167 | OD | Physician | 04/04/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16744674&pc=QCB | 2023826-2023-01344 | 2023-03-09 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.0/1.0/080 (SPHERE/CYLINDER/AXIS), WHILE LOADING THE LENS, A PIECE OF THE ICLBROKE. THERE WAS NOT PATIENT CONTACT. CAUSE OF THE EVENT WAS USER ERROR. REPORTEDLY, "THE LENS WAS NOT IMPLANTED. PATIENT SENT HOME UNTOUCHED." Manufacturer Narrative: A4-A6:UNK. USER ERROR. CLAIM# (B)(4). | -10.0/1.0/080 | NA | Physician | 03/17/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16744723&pc=QCB | 2023826-2023-01348 |  | Injury | STAAR SURGICAL COMPANY | 2023-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). A PLANNED LENS REMOVAL AND REPLACEMENT FOR AN UNKNOWN REASON WAS REPORTED. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -6.50/1.0/084 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO LOW VAULT. ON (B)(6) 2023 A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED. CLAIM # (B)(6). Manufacturer Narrative: D6A:UNK CLAIM#: (B)(4). | -6.50/1.0/084 | OD | Physician | 03/15/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16744726&pc=QCB | 2023826-2023-01345 | 2023-03-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -11.0/1.0/78 (SPHERE/CYLINDER/AXIS) LENS TORE DURING/BEFORE LOADING. REPORTEDLY, "ON DRAWING LENS INTO THE INJECTOR, THERE WAS RESISTANCE AND SUBSEQUENT SMALL TEAR. BACKUP LENS WAS LOADED AND PROCEDURE COMPLETED WITHOUT COMPLICATIONS. CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. "WHILE GENTLY AND SLOWLY PULLING THE ICL THROUGH THE CARTRIDGE, GRASPING (AS PROTOCOLISED) FROM THE HAPTICS AS CLOSE TO THE OPTIC AS POSSIBLE, THE RESISTANCE MET WITHIN THE CARTRIDGE IS SUCH THAT THE GRASP OF THE HAPTIC SLIPS AND THERE IS A RESULTING TEAR OR CHIP TO THE ICL. I HAVE BEEN USING THE SAME FORCEPS FOR AROUND 350 CASES AND HAVE FOUND THIS TO OCCUR ON OCCASION." THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4-A6:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -11.0/1.0/78 | NA | Physician | 03/17/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16748051&pc=QCB | 2023826-2023-01387 | 2022-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.00/3.5/101 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE SURGEON ROTATED/REPOSITIONED THE LENS DUE TO LENS ROTATION. THE LENS ENDED UP ROTATING AGAIN AND ON 14-MAR-2023 THE LENS WAS REPLACED WITH A SAME LENGTH LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.00/3.5/101 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6)2022. ON (B)(6) 2022 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE PROBLEM DID NOT RESOLVE. ON (B)(6) 2023 THE LENS WAS EXPLANTED. A REPLACEMENT LENS OF THE SAME LENGTH DIFFERENT AXIS WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#:(B)(4). | -11.00/3.5/101 | OS | Other Health Care Professional | 03/15/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16748112&pc=QCB | 2023826-2023-01389 |  | Injury | STAAR SURGICAL COMPANY | 2023-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.50/1.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS EXPLANTED ON (B)(6) 2023 FOR AN UNKNOWN REASON AND THE BACKUP LENS WAS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: D4 - MODEL # VTICM5_13.2 CORRECTED TO VTICM512.6. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -9.50/1.0/091 | NA | Other Health Care Professional | 12/27/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16748570&pc=QCB | 2023826-2023-01352 | 2023-03-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 13.2MM, VTICM5_13.2, -13.0/1.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. IN THE REPORTER'S OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: B5- PER UPDATED INFORMATION A REPLACEMENT LENS WAS IMPLANTED SUCCESSFULLY ON (B)(6) 2023 AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.0/1.0/094 | NA | Physician | 03/16/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16748681&pc=QCB | 2023826-2023-01390 | 2022-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -11.5/1.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE OF EVENT WAS THE DEVICE. Manufacturer Narrative: CLAIM# (B)(4). | -11.5/1.0/094 | OD | Physician | 03/17/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16748683&pc=QCB | 2023826-2023-01353 | 2022-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -11.5/2.0/078 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE OF EVENT WAS THE DEVICE. Manufacturer Narrative: (B)(4). | -11.5/2.0/078 | OS | Physician | 03/17/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16750637&pc=QCB | 2023826-2023-01228 | 2023-03-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -5.5/1.5/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -5.5/1.5/072 | OS | Physician | 03/26/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16750638&pc=QCB | 2023826-2023-01229 | 2023-03-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.0/2.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -12.0/2.5/094 | OD | Physician | 03/26/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16750646&pc=QCB | 2023826-2023-01235 | 2022-12-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.5/5.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4: UNK, A5: UNK, A6: UNK. CLAIM # (B)(4). | -12.5/5.0/096 | OD | Physician | 03/29/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16750649&pc=QCB | 2023826-2023-01233 | 2023-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON VERTICALLY IMPLANTED A 13.2MM VTICMO13.2, -11.5/3.0/178 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. CLAIM # (B)(4). | -11.5/3.0/178 | OS | Physician | 03/27/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16750651&pc=QCB | 2023826-2023-01234 | 2022-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.5/3.0/170 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM # (B)(4). | -10.5/3.0/170 | OD | Physician | 03/27/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16750652&pc=QCB | 2023826-2023-01240 | 2023-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM512.6, -10.0/1.5/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.0/1.5/084 | OD | Physician | 03/31/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16750653&pc=QCB | 2023826-2023-01239 | 2023-03-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.0/1.0/128 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6)2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM# (B)(4). | -9.0/1.0/128 | OD | Physician | 03/31/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16750655&pc=QCB | 2023826-2023-01238 | 2023-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM513.2, -6.5/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: H6 - MEDICAL DEVICE PROBLEM CODE 1494: OFF-LABEL USE (AGE<21YRS). CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -6.5/1.0/090 | OD | Physician | 03/30/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16750659&pc=QCB | 2023826-2023-01241 | 2023-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.0/1.50/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFO: UNK. CLAIM# (B)(4). | -12.0/1.50/102 | OS | Physician | 03/31/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16750661&pc=QCB | 2023826-2023-01236 | 2022-12-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/5.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/5.0/092 | OS | Physician | 03/29/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16757608&pc=QCB | 2023826-2023-01400 | 2023-03-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -7.5/+10/75 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON(B0)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING. ON 20-MAR-2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4) . | -7.5/+10/75 | OD | Physician | 03/29/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16757610&pc=QCB | 2023826-2023-01401 | 2023-03-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -9.0/+1.0/94 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE PATIENT EXPERIENCED LOW VAULTING. ON 23-MAR-2023 THE LENS WAS EXPLANTED AND ON THIS SAME DATE REPLACED WITH A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/+1.0/94 | OS | Physician | 03/29/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16759691&pc=QCB | 2023826-2023-01403 | 2023-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+2.0/063 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 . THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -8.5/+2.0/063 | OS | Physician | 03/20/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16759693&pc=QCB | 2023826-2023-01404 | 2023-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.0/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+2.0/095 | OD | Physician | 03/20/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16762411&pc=QCB | 2023826-2023-01382 | 2022-12-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -09.00/+3.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PANTIENT INFO: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AND HAPTIC BROKEN. DIMENSIONAL WIDTH VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. UNABLE TO COMPLETE LENGTH DIMENSIONAL INSPECTION DUE TO HAPTIC DAMAGE. CLAIM # (B)(4). | -09.00/+3.0/087 | OS | Physician | 03/19/2023 | 2023 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16765978&pc=QCB | 2023826-2023-01436 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -5.00/4.0/149 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. A LENGTH ISSUE THAT WAS LATER RESOLVED WAS REPORTED. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: D6A: UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: H4 - DEVICE MANUFACTURE DATE: 27-MAY-2022 . CLAIM #(B)(4). | -5.00/4.0/149 | NA | Other Health Care Professional | 03/20/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16765979&pc=QCB | 2023826-2023-01452 | 2023-03-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -9.5/0.5/105 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS EXCESSIVE VAULTING AND PRODUCT NON-CONFORMITY. ON (B)(6) 2023THE LENS WAS EXPLANTED AND ON THE SAME DATE REPLACED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: (B)(4). | -9.5/0.5/105 | OD | Physician | 03/31/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16765994&pc=QCB | 2023826-2023-01454 | 2023-01-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -9.5/+2.5/044 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6)2023 THE LENS WAS EXPLANTED; ON THIS SAME DATE A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM#: (B)(4). | -9.5/+2.5/044 | OD | Physician | 04/11/2023 | 2023 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16766107&pc=QCB | 2023826-2023-01192 | 2022-10-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.50/1.5/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS REMOVED AND ON (B)(6)2022, A SHORTER LENGTH LENS WAS IMPLANTED DUE TO EXCESSIVE VAULT THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -10.50/1.5/93 | OD | Physician | 03/21/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16766159&pc=QCB | 2023826-2023-01419 | 2023-01-18 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5/1.5/113 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED, AND REPLACED DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. PROBLEM WAS RESOLVED. CAUSEOF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -9.5/1.5/113 | OS | Physician | 03/21/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16766194&pc=QCB | 2023826-2023-01437 | 2022-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM012.6, -15.50/1.5/068 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLATED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. A REPLACEMENT LENS OF LONGER LENGTH WAS LATER IMPLANTED BUT THE PROBLEM DID NOT RESOLVE. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -15.50/1.5/068 | OS | Physician | 03/21/2023 | 2023 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16766304&pc=QCB | 2023826-2023-01420 | 2023-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -5.0/3.5/85 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LOW VAULT WITH ROTATION. CAUSE OF THE EVENT WAS REPORTED AS OTHER. REPORTEDLY, "THE DIAMETER IS NOT CORRECT." ON (B)(6) 2023, A LONGER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE IS REPORTED AS, "CORRECT." Manufacturer Narrative: A4-A6:UNK. CLAIM# (B)(4). | -5.0/3.5/85 | OS | Physician | 03/21/2023 | 2023 | 42 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16769976&pc=QCB | 2023826-2023-01384 | 2023-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+3.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3:DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -18.00/+3.5/094 | OD | Physician | 03/21/2023 | 2023 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16770054&pc=QCB | 2023826-2023-01380 | 2023-03-06 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.00/+1.5/092 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN AND UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO THE SIGNIFICANT LENS HAPTIC DAMAGE. CLAIM # (B)(4). | -07.00/+1.5/092 | NA | Physician | 03/20/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16770529&pc=QCB | 2023826-2023-01358 | 2023-03-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -6.0/1.0/090 | OS | Physician | 04/10/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16770532&pc=QCB | 2023826-2023-01359 | 2023-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.5/1.5/010 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -16.5/1.5/010 | OD | Physician | 04/12/2023 | 2023 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16773436&pc=QCB | 2023826-2023-01465 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -6/+2.5/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE PATIENT EXPERIENCED LOW VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM#: (B)(4). | -6/+2.5/088 | OS | Physician | 04/11/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16773624&pc=QCB | 2023826-2023-01412 | 2022-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.0/084 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -8.5/+1.0/084 | OS | Physician | 03/23/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16773626&pc=QCB | 2023826-2023-01414 | 2022-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+3.5/115 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -12.0/+3.5/115 | OD | Physician | 03/24/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16773628&pc=QCB | 2023826-2023-01413 | 2022-11-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+3.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS EXCHANGED ON(B)(6)2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: (B)(4). | -12.5/+3.0/091 | OD | Physician | 03/23/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16773629&pc=QCB | 2023826-2023-01416 | 2023-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.0/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -10.5/+2.0/093 | OS | Physician | 03/24/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16773861&pc=QCB | 2023826-2023-01425 | 2022-10-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.0/1.5/074 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NO ASSOCIATED TO LOW VAULT. REPORTEDLY, "A WEEK AFTER SUGERY, IT WAS FOUND THAT THE LENS RETURNED TO ITS ORIGINAL POSITION." THE PROBLEM WAS NOT RESOLVED. ON (B)(6) 2022, THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "ON OCTOBER 4, EXTERNAL FORCE CAUSED THE LEFT EYE LENS TO ROTATE." Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -10.0/1.5/074 | OS | Physician | 03/24/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16777691&pc=QCB | 2023826-2023-01427 | 2023-03-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030017 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023, A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -8.0/1.5/68 (SPHERE/CYLINDER/AXIS) LENS BECAME STUCK IN THE CARTRIDGE OR INJECTION SYSTEM. THERE WAS PATIENT CONTACT WITH THE PATIENT'S RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS INTRAOPERATIVELY REMOVED. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -8.0/1.5/68 | OD | Physician | 03/24/2023 | 2023 | 44 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16777692&pc=QCB | 2023826-2023-01426 | 2022-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.5/3.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS REMOVED AND THE A LONGER LENGTH LENS WAS IMPLANTED DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: WEIGHT:UNKNOWN; ETHNICITY:UNKNOWN; RACE: UNKNOWN. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -9.5/3.0/093 | OS | Physician | 03/24/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16787896&pc=QCB | 2023826-2023-01377 | 2023-01-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+3.5/087 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4: UNKNOWN. A5: UNKNOWN. A6: UNKNOWN. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -16.50/+3.5/087 | OD | Physician | 03/24/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16787907&pc=QCB | 2023826-2023-01378 | 2023-01-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.50/+1.0/145 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS HAD EXCESSIVE VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: AFTER FURTHER RE-MEASUREMENT, THE LENS WAS FOUND TO HAVE A LOWER VAULT. THE LENS REMAINED IMPLANTED AND THE PATIENT WAS HAPPY. THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.50/+1.0/145 | OS | Physician | 03/24/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16787971&pc=QCB | 2023826-2023-01379 | 2023-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -12.50/+0.5/176 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS HAD EXCESSIVE VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNKNOWN. A5: UNKNOWN. A6: UNKNOWN. D6B:UNKNOWN. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: THE LENS WAS REMOVED ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENS (DIFFERENT MODEL) AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. CLAIM # (B)(4). Manufacturer Narrative: H6- HEALTH EFFECT- IMPACT CODE: "4629" SHOULD BE REMOVED AND "4627" SHOULD BE ADDED. CLAIM# (B)(4). | -12.50/+0.5/176 | OD | Physician | 03/24/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16788656&pc=QCB | 2023826-2023-01451 | 2023-03-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 12.1MM, VTICM5_12.1, -7.0/1.5/065 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS STUCK IN INJECTOR. BACKUP LENS IMPLANTED. PATIENT CONTACT IS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.0/1.5/065 | NA | Physician | 03/22/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16792142&pc=QCB | 2023826-2023-01448 | 2022-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -10.5/2.0/013 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM RESOLVED. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A1:(B)(6). (B)(4) | -10.5/2.0/013 | OD | Physician | 03/24/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16792290&pc=QCB | 2023826-2023-01449 | 2023-03-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS -2.00/0.5/017 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5- PER UPDATED INFORMATION THE PROBLEM RESOLVED, ALL IS FINE AND PATIENT IS HAPPY. CLAIM # (B)(4). | -2.00/0.5/017 | OD | Physician | 03/24/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16792315&pc=QCB | 2023826-2023-01450 | 2023-03-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS -1.50/0.5/160 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION THE PROBLEM RESOLVED, ALL IS FINE AND PATIENT IS HAPPY. CLAIM # (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -1.50/0.5/160 | OS | Physician | 03/24/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16792765&pc=QCB | 2023826-2023-01475 | 2023-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.0/2.0/072 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6)2023, THE LENS WAS REMOVED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE CAUSE OF THE EVENT WAS UNKNOWN. REPORTEDLY, "THE PATIENT HAD NO DESIRE TO HAVE ANOTHER SURGERY." THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 17-MAY-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON THE LENS SPECIFICALLY FIBRE. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -10.0/2.0/072 | OS | Physician | 03/27/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16792804&pc=QCB | 2023826-2023-01432 | 2023-03-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.0/1.0/147 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED AND THE REPLACED DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: PT INFO:UNK. OFF-LABEL - ACD<3.0. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATE: B5: LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -12.0/1.0/147 | OS | Physician | 03/26/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16794615&pc=QCB | 2023826-2023-01274 | 2022-12-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.5/1.0/167 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #: (B)(4). | -11.5/1.0/167 | OD | Physician | 04/03/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16794619&pc=QCB | 2023826-2023-01280 | 2023-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.0/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. CLAIM# (B)(4). | -13.0/1.5/095 | OD | Physician | 04/03/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16794621&pc=QCB | 2023826-2023-01279 | 2023-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.5/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -12.5/1.0/089 | OS | Physician | 04/03/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16794622&pc=QCB | 2023826-2023-01290 | 2023-01-11 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1, -10.0/1.5/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REPLACED INTRAOPERATIVELY WITH A LONGER LENGTH LENS ON (B)(6) 2023 FROM PATIENT'S RIGHT EYE (OD) DUE TO LOW VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4: UNK, A5: UNK, A6: UNK. CLAIM # (B)(4). | -10.0/1.5/104 | OD | Physician | 04/06/2023 | 2023 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16794626&pc=QCB | 2023826-2023-01282 | 2023-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/1.0/081 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK, A5 - UNK, A6 - UNK. CLAIM# (B)(4). | -10.5/1.0/081 | OD | Physician | 04/03/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16794631&pc=QCB | 2023826-2023-01288 | 2022-10-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6, -8.0/1.5/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A LONGER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S LEFT EYE (OS) DUE TO LOW VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: CLAIM# (B)(4). | -8.0/1.5/079 | OS | Physician | 04/05/2023 | 2023 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16794636&pc=QCB | 2023826-2023-01285 | 2022-10-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2, -5.0/1.0/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -5.0/1.0/097 | OS | Physician | 04/05/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16794637&pc=QCB | 2023826-2023-01286 | 2022-10-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6, -13.0/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6 - MEDICAL DEVICE PROBLEM CODE 1494: OFF-LABEL USE (AGE>45YRS). CLAIM# (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | -13.0/2.0/088 | OD | Physician | 04/05/2023 | 2023 | 47 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16794640&pc=QCB | 2023826-2023-01289 | 2022-10-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6, -8.0/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A LONGER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S RIGHT EYE (OD) DUE TO LOW VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: CLAIM# (B)(4). | -8.0/1.0/089 | OD | Physician | 04/05/2023 | 2023 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16794668&pc=QCB | 2023826-2023-01284 | 1995-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.0/2.0/133 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/2.0/133 | OD | Physician | 04/05/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16794670&pc=QCB | 2023826-2023-01287 | 2022-10-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6, -12.5/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6- MEDICAL DEVICE CODE: 1494- OFF LABEL USE (AGE>45) . CLAIM# (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | -12.5/2.0/090 | OS | Physician | 04/05/2023 | 2023 | 47 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16802401&pc=QCB | 2023826-2023-01484 | 2023-03-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER, INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF A -11.0/1.0/112 (SPHERE/CYLINDER/AXIS) LENS, TORE BEFORE/DURING LOADING ON (B)(6) 2023. THERE WAS NO PATIENT CONTACT OR INJURY. A REPLACEMENT LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6-MEDICAL DEVICE PROBLEM CODE: "2993" SHOULD BE REMOVED AND CODE "3189" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -11.0/1.0/112 | OD | Physician | 03/31/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16802429&pc=QCB | 2023826-2023-01485 | 2022-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -12.50/3.0/74 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS ROTATION NOT ASSOCIATED TO LOW VAULT WAS OBSERVED. ON (B)(6) 2022, LENS ROTATION WAS PERFORMED. LENS REPOSITION DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023. THE LENS WAS EXCHANGED WITH A REPLACEMENT LENS, AND THE PROBLEM WAS RESOLVED. REPORTEDLY, "THE EXCHANGE LENS WAS IMPLANTED VERTICALLY." Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4-A6: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -12.50/3.0/74 | OS | Physician | 03/31/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16804609&pc=QCB | 2023826-2023-01480 | 2022-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -11.0/3.5/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. LENS ROTATION NOT ASSOCIATED TO LOW VAULT WAS OBSERVED. ON (B)(6)2022, LENS ROTATION WAS PERFORMED. LENS REPOSITION DID NOT RESOLVE THE PROBLEM. ON (B)(6)2023, THE LENS WAS EXCHANGED WITH A REPLACEMENT LENS, AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: H6 - MEDICAL DEVICE PROBLEM CODE 1494: OFF-LABEL USE (UNDER 21 YRS OF AGE AT DATE OF IMPLANTATION) CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -11.0/3.5/091 | OD | Physician | 03/30/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16811557&pc=QCB | 2023826-2023-01535 | 2023-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.0/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) CLAIM# (B)(4). | -10.0/+1.0/089 | OS | Physician | 03/26/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16811559&pc=QCB | 2023826-2023-01536 | 2023-03-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+3.0/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: (B)(4). | -10.5/+3.0/092 | OD | Physician | 03/27/2023 | 2023 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16811640&pc=QCB | 2023826-2023-01492 | 2023-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50/+2.5/071 (SPHERE/CYLINDER/AXIS) (IMPLANTED VERTICALLY), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED HORIZONTALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -08.50/+2.5/071 | OD | Physician | 03/27/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16811652&pc=QCB | 2023826-2023-01494 | 2023-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.50/+1.5/109 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS HAD LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4). Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2023 FOR A LONGER LENS. THE PROBLEM WAS RESOLVED AND THE PATIENT IS HAPPY. CLAIM # (B)(4). | -15.50/+1.5/109 | OD | Physician | 03/28/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16811659&pc=QCB | 2023826-2023-01518 | 2023-03-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 13.2MM, VTICMO13.2, -14.5/1.5/006 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE RIGHT EYE (OD). THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY, WITH NO PATIENT INJURY. A REPLACEMENT LENS WAS IMPLANTED AT A LATER DATE AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -14.5/1.5/006 | OD | Physician | 03/27/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16811913&pc=QCB | 2023826-2023-01519 | 2023-03-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 13.2MM, VTICMO13.2, -14.0/1.5/019 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE RIGHT EYE (OD). THE LENS WAS IMPLANTED, REMOVED AND REPLACED INTRAOPERATIVELY WITH NO PATIENT INJURY AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -14.0/1.5/019 | OD | Physician | 03/27/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16812067&pc=QCB | 2023826-2023-01520 | 2023-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -8.50/2.5/106 WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6), 2023 THE LENS WAS EXPLANTED AND LATER REPLACED WITH A LONGER LENGTH LENS DUE TO "LENS TOO SHORT" AND THE PROBLEM RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITONAL DATA: B5- THE REPORTER INDICATED THE CAUSE OF THE EVENT IS OTHER-"WRONG WTW MEASUREMENTS. CLAIM# (B)(4). | -8.50/2.5/106 | OD | Physician | 03/27/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16812301&pc=QCB | 2023826-2023-01486 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -11.50/1.5/090 SPHERE/CYLINDER/AXIS WAS TORN DURING LOADING. THERE WAS NO PATIENT CONTACT. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS. CLAIM# (B)(4). | -11.50/1.5/090 | NA | Physician | 03/31/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16818064&pc=QCB | 2023826-2023-01458 | 2023-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/2.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. (B)(4). | -14.0/2.5/090 | OD | Physician | 04/17/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16818076&pc=QCB | 2023826-2023-01456 | 2023-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/3.5/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -9.0/3.5/073 | OS | Physician | 04/17/2023 | 2023 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16818078&pc=QCB | 2023826-2023-01457 | 2023-03-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.50/2.0/173 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -10.50/2.0/173 | OD | Physician | 04/17/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16818079&pc=QCB | 2023826-2023-01474 | 2023-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/1.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. (B)(4). | -12.0/1.0/089 | OD | Physician | 04/20/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16818080&pc=QCB | 2023826-2023-01459 | 2022-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.50/1.0/111 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -10.50/1.0/111 | OD | Physician | 04/17/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16818081&pc=QCB | 2023826-2023-01467 | 2023-04-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/1.5/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. (B)(4). | -12.0/1.5/072 | OS | Physician | 04/19/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16818085&pc=QCB | 2023826-2023-01472 | 2023-04-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.50/1.5/112 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -9.50/1.5/112 | OD | Physician | 04/20/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16818086&pc=QCB | 2023826-2023-01473 | 2023-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. (B)(4). | -10.0/1.0/088 | OD | Physician | 04/20/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16818095&pc=QCB | 2023826-2023-01460 | 2023-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.50/1.0/128 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "THE DOCTOR CHOOSES TO PLACE THE LENS IMPLANTED VERTICALLY DURING THE SECOND SURGERY." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. (B)(4). | -7.50/1.0/128 | OS | Physician | 04/17/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16818096&pc=QCB | 2023826-2023-01510 | 2022-12-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.0/5.0/048 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) CLAIM #(B)(4). | -5.0/5.0/048 | OS | Physician | 04/24/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16818102&pc=QCB | 2023826-2023-01508 | 2023-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.5/0.5/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -3.5/0.5/085 | OD | Physician | 04/24/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16818103&pc=QCB | 2023826-2023-01507 | 2023-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. (B)(4) | -9.0/1.0/091 | OS | Physician | 04/24/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16818104&pc=QCB | 2023826-2023-01512 | 2023-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.50/2.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -9.50/2.5/090 | OS | Physician | 04/24/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16818105&pc=QCB | 2023826-2023-01513 | 2022-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. (B)(4). | -10.0/1.5/090 | OD | Physician | 04/23/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16818109&pc=QCB | 2023826-2023-01509 | 2023-01-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.50/1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNKNOWN. CLAIM #(B)(4). | -10.50/1.5/091 | OD | Physician | 04/24/2023 | 2023 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16818115&pc=QCB | 2023826-2023-01511 | 2023-04-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -18.0/2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNKNOWN. CLAIM #(B)(4). | -18.0/2.0/092 | OD | Physician | 04/24/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16818127&pc=QCB | 2023826-2023-01514 | 2022-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.50/3.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -11.50/3.0/091 | OS | Physician | 04/23/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16820284&pc=QCB | 2023826-2023-01523 | 2023-03-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 13.2MM, VTICH13.2, -8.50/3.0/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE RIGHT EYE (OD). THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY, WITH NO PATIENT INJURY. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICH 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER +8.5/+3.0/070 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. REPORTEDLY THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. ON (B)(6) 2023 A REPLACEMENT LENS OF THE SAME MODEL/ LENGTH, SIMILAR POWER WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION PROVIDED: "PATIENT IS HAPPY". THE REPORTER INDICATED THE CAUSE OF THE EVENT IS UNKNOWN. (B)(4). Manufacturer Narrative: B5- PER UPDATED INFORMATION A REPLACEMENT LENS WAS IMPLANTED SUCCESSFULLY ON (B)(6) 2023 AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -8.50/3.0/070 | OD | Physician | 03/28/2023 | 2023 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16820456&pc=QCB | 2023826-2023-01499 | 2023-02-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.00/+4.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS REMOVED ON 11-FEB-2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A DIFFERENT MODEL BUT SAME LENGTH LENS BUT THE PROBLEM WAS NOT RESOLVED. SEE MFR. REPORT # 2023826-2023-01500 FOR REPLACEMENT LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE REPORTER INDICATED THE PATIENT HAS NOT RETURNED TO THE CLINIC SINCE LAST FOLLOW-UP VISIT. ACCORDING TO THE SURGEON, THE PATIENT DOES NOT SEEM TO REQUIRE FURTHER TREATMENT FOR NOW. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -10.00/+4.5/092 | OS | Physician | 03/30/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16820467&pc=QCB | 2023826-2023-01497 | 2022-12-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00/+3.5/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS REPLACED WITH A LONGER LENS BUT THE PROBLEM WAS NOT RESOLVED. THE SURGEON IS CONSIDERING FURTHER TREATMENT. SEE MFR. REPORT # 2023826-2023-01498 FOR REPLACEMENT LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE REPORTER INDICATED THE PATIENT HAS NOT RETURNED TO THE CLINIC SINCE LAST FOLLOW-UP VISIT. ACCORDING TO THE SURGEON, THE PATIENT DOES NOT SEEM TO REQUIRE FURTHER TREATMENT FOR NOW. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC SPLIT AND HAPTIC BROKEN. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.00/+3.5/082 | OS | Physician | 03/30/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16820475&pc=QCB | 2023826-2023-01495 | 2023-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/105 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS HAD LOW VAULTING AND LENS ROTATION. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2023 FOR A LONGER LENS. THE PROBLEM WAS RESOLVED AND THE PATIENT IS HAPPY. CLAIM # (B)(4). Manufacturer Narrative: B5 - THE ISSUE IS NO LONGER CONSIDERED AS RESOLVED. REFRACTIVE SURPRISE WAS REPORTED FOR THE EXCHANGED LENS (CLAIM #(B)(4). CLAIM #(B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.00/+2.0/105 | OS | Physician | 03/28/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16820697&pc=QCB | 2023826-2023-01559 |  | Injury | STAAR SURGICAL COMPANY | 2023-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -3.5/+1.0/055 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). | -3.5/+1.0/055 | OS | Physician | 04/11/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16822025&pc=QCB | 2023826-2023-01540 | 2022-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+3.0/084 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -10.0/+3.0/084 | OS | Physician | 03/29/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16822026&pc=QCB | 2023826-2023-01539 | 2022-12-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+4.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/+4.0/093 | OD | Physician | 03/29/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16822027&pc=QCB | 2023826-2023-01538 | 2022-12-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+2.5/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -7.5/+2.5/093 | OD | Physician | 03/29/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16822028&pc=QCB | 2023826-2023-01541 | 2022-11-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+4.0/083 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -8.0/+4.0/083 | OD | Physician | 03/29/2023 | 2023 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16822029&pc=QCB | 2023826-2023-01543 | 2022-09-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.50/+2.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -16.50/+2.0/091 | OD | Physician | 03/29/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16822031&pc=QCB | 2023826-2023-01545 | 2023-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+0.5/086 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -13.5/+0.5/086 | OS | Physician | 03/31/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16822032&pc=QCB | 2023826-2023-01544 | 2023-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.5/+1.0/080 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#(B)(4). | -15.5/+1.0/080 | OS | Physician | 03/30/2023 | 2023 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16822039&pc=QCB | 2023826-2023-01546 | 2023-03-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.0/+0.5/104 INTO THE PATIENT'S RIGHT EYE (OD) ON 19-MAR-2023. THE LENS WAS EXCHANGED ON 21-MAR-2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) CLAIM#(B)(4). | -5.0/+0.5/104 | OD | Physician | 03/31/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16827518&pc=QCB | 2023826-2023-01564 | 2023-04-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE A 12.6MM VTICMO12.6; -3.5/1.0/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER INTENDED FOR THE PATIENT'S LEFT EYE (OS); HAD TORE. IT WAS NOTED LENS TEAR/BREAK DURING/BEFORE LOADING; AFTER OPENING THE VIAL AND "UNKNOWN WHEN THE DAMAGED OCCURRED". THIS OCCURRED ON (B)(6) 2023. ON (B)(6) 2023 A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -3.5/1.0/094 | OS | Physician | 04/26/2023 | 2023 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16827519&pc=QCB | 2023826-2023-01562 | 2023-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -10.5/+1/177 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2023 THE LENS WAS EXPLANTED; LATER A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). | -10.5/+1/177 | OS | Physician | 04/11/2023 | 2023 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16827761&pc=QCB | 2023826-2023-01565 | 2023-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.00/2.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -15.00/2.5/074 | OD | Physician | 04/26/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16827951&pc=QCB | 2023826-2023-01526 | 2023-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICM5_12.6 IMPLANTABLE COLLAMER LENS OF -2.50/1.50/155 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6)2023 THE LENS WAS EXPLANTED DUE TO THE SURGEON IMPLANTING THE LEFT LENS INTO THE RIGHT EYE. ON (B)(6) 2023 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS USER ERROR ( MIXING UP OD AND OS,ONLY DIFFERENCE WAS AXIS). Manufacturer Narrative: (B)(4). | -2.50/1.50/155 | OD | Physician | 03/30/2023 | 2023 | 41 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16828042&pc=QCB | 2023826-2023-01530 | 2023-03-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.0/2.5/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TORN, BROKEN AND DEFORMED LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN/DEFORMED WITH FIBERS ON THE LENS SURFACE. REHYDRATION OF THE LENS REVEALED SIGNIFICANT LENS DAMAGE; THE SECOND PICTURE HAS BEEN RECORDED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -11.0/2.5/091 | OS | Physician | 03/31/2023 | 2023 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16833323&pc=QCB | 2023826-2023-01501 | 2023-02-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -07.50/+3.5/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2023 BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: H6- HEALTH EFFECT - IMPACT CODE: "4629" SHOULD BE REMOVED AND "4627" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: A4: UNKNOWN. A5: UNKNOWN. A6: UNKNOWN. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -07.50/+3.5/087 | OD | Physician | 03/31/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16833366&pc=QCB | 2023826-2023-01503 | 2023-03-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-04-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -07.00/+1.0/088 (SPHERE/CYLINDER/AXIS), TORE DURING LOADING. THERE WAS NO PATIENT CONTACT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -07.00/+1.0/088 | NA | Physician | 03/31/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16833407&pc=QCB | 2023826-2023-01498 | 2023-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-04-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -12.00/+3.5/079 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-01497, BUT THE PROBLEM WAS NOT RESOLVED. THE LENS REMAINED IMPLANTED. THE SURGEON IS CONSIDERING FURTHER TREATMENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. D6B: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THE PATIENT HAS NOT RETURNED TO THE CLINIC SINCE THE LAST FOLLOW-UP. ACCORDING TO THE SURGEON, THE PATIENT DOES NOT SEEM TO REQUIRE FURTHER TREATMENT FOR NOW. CLAIM # (B)(4). Manufacturer Narrative: NEW INFORMATION: B5 - THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -12.00/+3.5/079 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 WITH A SAME LENGTH BUT DIFFERENT POSITION AXIS LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). | -12.00/+3.5/079 | OS | Physician | 03/30/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16835557&pc=QCB | 2023826-2023-01528 |  | Injury | STAAR SURGICAL COMPANY | 2023-04-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -8.50/3.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON AN UNKNOWN DATE. THE LENS WAS EXPLANTED FOR AN UNKNOWN REASON AND A REPLACEMENT LENS OF LONGER LENGTH AND DIFFERENT POWER WAS LATER IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: D6A:UNK. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#:(B)(4). | -8.50/3.0/090 | OD | Other Health Care Professional | 03/30/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16836033&pc=QCB | 2023826-2023-01529 |  | Injury | STAAR SURGICAL COMPANY | 2023-04-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -9.50/2.5/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON AN UNKNOWN DATE. THE LENS WAS EXPLANTED FOR AN UNKNOWN REASON AND A REPLACEMENT LENS OF LONGER LENGTH, DIFFERENT MODEL AND POWER WAS LATER IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -9.50/2.5/088 | OS | Other Health Care Professional | 03/30/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16850027&pc=QCB | 2023826-2023-01572 | 2023-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.00/3.5/100 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. (B)(4). | -8.00/3.5/100 | OD | Physician | 04/27/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16850031&pc=QCB | 2023826-2023-01574 | 2023-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_ 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.50/2.5/070 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO LOW VAULT WITH ROTATION. A REPLACEMENT LENS OF LONGER LENGTH WAS LATER IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #: (B)(4). | -10.50/2.5/070 | OS | Physician | 04/28/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16850040&pc=QCB | 2023826-2023-01576 | 2023-03-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNKNOWN. CLAIM #(B)(4). | -12.0/2.0/092 | OS | Physician | 04/27/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16850042&pc=QCB | 2023826-2023-01573 | 2023-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.50/3.5/100 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNKNOWN. CLAIM #(B)(4). | -8.50/3.5/100 | OS | Physician | 04/27/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16850053&pc=QCB | 2023826-2023-01579 | 2023-02-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/3.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM (B)(4). | -12.5/3.5/092 | OD | Physician | 04/27/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16850059&pc=QCB | 2023826-2023-01578 | 2023-02-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/1.5/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNKNOWN. CLAIM #(B)(4). | -6.5/1.5/082 | OS | Physician | 04/27/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16850117&pc=QCB | 2023826-2023-01584 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.00/+3.0/109 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS EYE. THE LENS WAS REMOVED DUE TO THE LEADING RIGHT FOOTPLATE HAD A SMALL TEAR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2: UNK. A3: UNK. A4: UNK. A5: UNK. A6: UNK. B3: UNK. D6A: UNK. D6B: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4). | -16.00/+3.0/109 | NA | Physician | 04/04/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16850171&pc=QCB | 2023826-2023-01612 | 2023-03-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.0/1.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT WAS NOTED AND ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL; SAME LENGTH; DIFFERENT AXIS LENS WHICH WAS IMPLANTED VERTICALLY, AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H6: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TORN OPTIC AND THE HAPTICS BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -9.0/1.0/088 | OS | Physician | 04/03/2023 | 2023 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16850728&pc=QCB | 2023826-2023-01618 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -11.00/2.0/035 (SPHERE/CYLINDER/AXIS) WAS REMOVED (EXPLANTED) DUE TO THE LENS TURNING UPSIDE DOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -11.00/2.0/035 | NA | Physician | 04/04/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16852125&pc=QCB | 2023826-2023-01394 | 2023-03-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.0/2.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -11.0/2.5/089 | OS | Physician | 04/12/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16852126&pc=QCB | 2023826-2023-01395 | 2023-03-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.5/3.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #: (B)(4). | -12.5/3.0/091 | OD | Physician | 04/12/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16852127&pc=QCB | 2023826-2023-01396 | 2022-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.5/1.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED VERTICALLY WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM#: (B)(4). | -10.5/1.5/088 | OS | Physician | 04/12/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16852128&pc=QCB | 2023826-2023-01397 | 2022-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.5/1.0/093 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED VERTICALLY WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -9.5/1.0/093 | OD | Physician | 04/12/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16852129&pc=QCB | 2023826-2023-01399 | 2023-03-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.0/1.0/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -10.0/1.0/112 | OD | Physician | 04/13/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16852130&pc=QCB | 2023826-2023-01398 | 2023-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.5/3.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM#: (B)(4). | -9.5/3.0/096 | OD | Physician | 04/13/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16854694&pc=QCB | 2023826-2023-01599 | 2021-04-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.5/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS EXCHANGED ON (B)(6)2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -10.5/+1.5/093 | OD | Physician | 04/02/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16855224&pc=QCB | 2023826-2023-01621 | 2023-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -18.00/3.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -18.00/3.5/090 | OD | Physician | 05/01/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16855225&pc=QCB | 2023826-2023-01623 | 2023-01-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 _13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.50/3.0/096 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -13.50/3.0/096 | OD | Physician | 05/01/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16855227&pc=QCB | 2023826-2023-01625 | 2023-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_ 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.50/0.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -7.50/0.5/088 | OS | Physician | 05/01/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16855873&pc=QCB | 2023826-2023-01661 | 2023-01-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM13.2; -10.00/+0.50/99 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE LEFT EYE (OS) ON (B)(6) 2012. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND LATER ON THIS SAME DATE IN A SEPARATE SURGERY A REPLACEMENT LENS OF THE SAME LENGTH BUT DIFFERENT SPECIFICATIONS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -10.00/+0.50/99 | OS | Physician | 04/11/2023 | 2023 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16855874&pc=QCB | 2023826-2023-01660 | 2023-01-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM13.2; -10.50/+1.50/58 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2012. THE PATIENT EXPERIENCED LOW VAULTING. THE LENS WAS EXPLANTED AND REPLACED WITH A SAME LENGTH BUT DIFFERENT MODEL/SPECS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM#: (B)(4). | -10.50/+1.50/58 | OD | Physician | 04/11/2023 | 2023 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16856812&pc=QCB | 2023826-2023-01591 | 2022-12-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50/+3.5/110 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH ANOTHER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.50/+3.5/110 | OS | Physician | 04/06/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16856824&pc=QCB | 2023826-2023-01602 | 2023-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+4.0/164 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -14.0/+4.0/164 | OS | Physician | 04/03/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16856827&pc=QCB | 2023826-2023-01603 | 2023-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -18.0/+2.0/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -18.0/+2.0/090 | OD | Physician | 04/03/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16856842&pc=QCB | 2023826-2023-01588 | 2022-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.00/+2.5/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: (B)(6) 2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK; A5: UNK; A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | -07.00/+2.5/096 | OS | Physician | 04/03/2023 | 2023 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16857357&pc=QCB | 2023826-2023-01589 | 2023-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50/+2.0/102 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT POWER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 14-JUN-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND RESIDUE/DEBRIS ON THE LENS AND FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -14.50/+2.0/102 | OD | Physician | 04/04/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16857554&pc=QCB | 2023826-2023-01624 | 2022-12-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.00/6.0/118 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNKNOWN. CLAIM #(B)(4). | -7.00/6.0/118 | OD | Physician | 05/01/2023 | 2023 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16857703&pc=QCB | 2023826-2023-01590 | 2022-11-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -06.50/+0.5/042 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT POWER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK; A5: UNK; A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 17-JUL-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -06.50/+0.5/042 | OD | Physician | 04/06/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16857773&pc=QCB | 2023826-2023-01566 | 2023-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/2.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023 . ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.0/2.5/089 | OD | Physician | 04/26/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16864764&pc=QCB | 2023826-2023-01593 | 2023-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.50/+4.5/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED SEVERAL TIMES AND THE PROBLEM WAS RESOLVED EACH TIME BUT THE LENS WAS EVENTUALLY REMOVED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. NO SHAFT MISALIGNMENT AFTER LENS REPLACEMENT. THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 23-MAY-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -15.50/+4.5/095 | OD | Physician | 04/03/2023 | 2023 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16864769&pc=QCB | 2023826-2023-01592 | 2023-01-18 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON WAS INSERTING A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.00/+1.0/166 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) AND THE LENS TORE/BROKE WHILE DISCHARGING. THERE WAS NO PATIENT CONTACT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.00/+1.0/166 | OS | Physician | 04/03/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16865374&pc=QCB | 2023826-2023-01609 | 2023-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.0/062 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023 . THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.0/+2.0/062 | OD | Physician | 04/04/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16865375&pc=QCB | 2023826-2023-01608 | 2023-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+2.0/069 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.5/+2.0/069 | OD | Physician | 04/04/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16865376&pc=QCB | 2023826-2023-01607 | 2023-02-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+4.0/085 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+4.0/085 | OS | Physician | 04/04/2023 | 2023 | 43 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16871600&pc=QCB | 2023826-2023-01594 | 2023-01-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON, BEFORE LOADING A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -07.50/+2.5/180 (SPHERE/CYLINDER/AXIS), FOUND A FOREIGN BODY ADHESION ON THE LENS. THE SURGEON REMOVED THE FOREIGN BODY ADHESION BUT DECIDED NOT TO USE THE LENS. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS EVIDENCE OF RESIDUE AND DEBRIS THOUGHOUT THE LENS. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: D9- "04/26/2023" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON NOTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.50/2.5/180 (SPHERE/CYLINDER/AXIS) HAS A FOREIGN BODY ADHESION ON (B)(6) 2023. REPORTEDLY "BEFORE LOADING THE SURGEON FOUND THE FB ADHESION AND TRIED TO REMOVE IT USING BSS. REMOVED BUT NOT USED". THE LENS WAS NOT IMPLANTED. ON THE SAME DAY A REPLACEMENT LENS WAS IMPLANTED INTO THE LEFT EYE (OS) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK; A5: UNK; A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -07.50/+2.5/180 | OS | Physician | 04/05/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16874802&pc=QCB | 2023826-2023-01672 | 2023-03-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.0/074 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+2.0/074 | OS | Physician | 04/05/2023 | 2023 | 41 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16882518&pc=QCB | 2023826-2023-01669 | 2023-04-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/1.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -12.0/1.5/090 | OS | Physician | 05/04/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16882520&pc=QCB | 2023826-2023-01667 | 2023-04-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/0.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -13.5/0.5/090 | OS | Physician | 05/04/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16882522&pc=QCB | 2023826-2023-01668 | 2023-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/1.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM (B)(4). | -7.0/1.5/092 | OS | Physician | 05/04/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16882607&pc=QCB | 2023826-2023-01674 | 2022-12-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.0/080 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.0/+2.0/080 | OD | Physician | 04/06/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16882609&pc=QCB | 2023826-2023-01675 | 2023-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+1.5/149 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -6.5/+1.5/149 | OD | Physician | 04/06/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16882610&pc=QCB | 2023826-2023-01678 | 2023-01-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.0/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+2.0/094 | OD | Physician | 04/06/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16882612&pc=QCB | 2023826-2023-01677 | 2023-01-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/+3.0/130 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -6.0/+3.0/130 | OD | Physician | 04/06/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16888262&pc=QCB | 2023826-2023-01637 | 2023-04-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF A -15.0/3.0/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OD) ON (B)(6) 2023. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED, AND REPLACED DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT, BUT NO PATIENT INJURY. THE CAUSE OF THE EVENT IS UNKNOWN. THE PROBLEM WAS RESOLVED. REPORTEDLY, "MAIN LENS FLIPPED IN AC AND TORE DURING SURGERY. BACK UP LENS IMPLANTED IN SULCUS." BLURRED VISON AND MILD CORNEAL EDEMA IS OBSERVED. THE STATUS OF THE EYE REPORTED, "MILD CORNEA EDEMA.". Manufacturer Narrative: CORRECTED DATA: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -15/+3/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SAME MODEL/LENGTH LENS BUT THE PROBLEM WAS NOT RESOLVED. IT WAS ADDITIONALLY NOTED "MAIN LENS FLIPPED IN AC AND TEARED DURING SURGERY. BACK UP LENS IMPLANTED IN SULCUS". THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -15.0/3.0/90 | OS | Physician | 04/10/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16888764&pc=QCB | 2023826-2023-01639 | 2022-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -13.0/1.0/86 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT WAS OBSERVED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -13.0/1.0/86 | OD | Physician | 04/10/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16890499&pc=QCB | 2023826-2023-01640 | 2023-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -5.5/1.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6 OFF-LABEL ACD<3.0. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -5.5/1.0/089 | OD | Physician | 04/10/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16891774&pc=QCB | 2023826-2023-01763 | 2023-03-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -13.50/3.0/114 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS PRODUCT NON CONFORMITY AND NOTED "LENS CAME OUT WITH VERY LOW VAULT", PATIENT HAD LOW VAULTING AND ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON THIS SAME DATE IN A SEPARATE SURGERY A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED. IT WAS ALSO INDICATED THERE WAS AN INCISION ENLARGEMENT AND AN ADDITIONAL SUTURE WAS REQUIRED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND THE DEVICE DID FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CLAIM# (B)(4). | -13.50/3.0/114 | OD | Physician | 04/18/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16891775&pc=QCB | 2023826-2023-01762 | 2022-12-25 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT WHILE ATTEMPTING TO LOAD A 12.6MM VTICM5_12.6; -13.50/3.0/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS ON (B)(6) 2022, THE LENS HAD TORE. IT WAS INDICATED THE LENS TORE WHILE PULLING THE LENS WITH THE FORCEPS. THE LENS WAS NOT IMPLANTED. ON (B)(6) 2023 A REPLACEMENT LENS OF THE SAME MODEL/LENGTH WAS IMPLANTED BUT THIS DID NOT RESOLVE THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE; AND THE DEVICE DID FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.50/3.0/095 | NA | Physician | 04/18/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16892051&pc=QCB | 2023826-2023-01706 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-05-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -6.5/2.0/104 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS INTRAOPERATIVELY EXCHANGED FOR A LONGER LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "ISSUE RESOLVED." Manufacturer Narrative: B5: LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. REPORTEDLY, "PATIENT HAVE NO COMPLAINT." CALIM# (B)(4). Manufacturer Narrative: A4-A6:UNKNOWN. CLAIM# (B)(4). | -6.5/2.0/104 | OD | Physician | 04/11/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16892052&pc=QCB | 2023826-2023-01705 | 2022-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM513.7 IMPLANTABLE COLLAMER LENS OF -11.5/2.0/081 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM# (B)(4). | -11.5/2.0/081 | OD | Physician | 04/11/2023 | 2023 | 47 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16892244&pc=QCB | 2023826-2023-01708 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -11.5/4.0/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. REPORTEDLY, "ISSUE RESOLVED." Manufacturer Narrative: (B)(4). | -11.5/4.0/93 | OD | Physician | 04/11/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16892245&pc=QCB | 2023826-2023-01707 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -10.0/2.5/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. REPORTEDLY, "ISSUE RESOLVED." Manufacturer Narrative: CORRECTED DATA: B5: SHOULD BE CORRECTED TO STATE "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO_12.1 IMPLANTABLE COLLAMER LENS OF -10.0/2.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2022 DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED." D4: SHOULD BE CORRECTED TO VTICMO_12.1. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -10.0/2.5/089 | OS | Physician | 04/11/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16892557&pc=QCB | 2023826-2023-01709 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -6.5/4.0/078 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. REPORTEDLY, "ISSUE RESOLVED." Manufacturer Narrative: CLAIM# (B)(4). | -6.5/4.0/078 | OS | Physician | 04/11/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16894686&pc=QCB | 2023826-2023-01547 | 2023-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.5/3.0/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED VERTICALLY WITH A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 14-JUN-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -14.5/3.0/078 | OS | Physician | 04/16/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16894688&pc=QCB | 2023826-2023-01549 | 2023-01-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED UPSIDE DOWN A 12.6MM VTICMO12.6, -14.5/3.5/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE ON (B)(6) 2023. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. BACK UP LENS IMPLANTED INTRAOPERATIVELY . THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK A5 - UNK A6 - UNK H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#(B)(4). | -14.5/3.5/098 | OS | Physician | 04/17/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16894695&pc=QCB | 2023826-2023-01552 | 2022-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -15.5/3.0/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. LENS WAS REPOSITIONED ON (B)(6)2022 BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS, THIS RESOLVED THE PROBLEM. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -15.5/3.0/103 | OD | Physician | 04/18/2023 | 2023 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16894696&pc=QCB | 2023826-2023-01553 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.5/3.0/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. LENS WAS REPOSITIONED ON (B)(6) 2022 BUT DID NOT RESOLVE THE PROBLEM. ON 31/MAR/2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS, THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT.. VISUAL INSPECTION FOUND HAPTIC BROKEN AND BLOOD STAINS ON SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#(B)(4). | -12.5/3.0/90 | OS | Physician | 04/18/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16896828&pc=QCB | 2023826-2023-01713 | 2023-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -9.5/5.5/82 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6)2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. REPORTEDLY, "PATIENT LIVES OUTSIDE OF JORDAN, THAT IS WHY WE HAD DELAYS IN SURGERY DATES." CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR AND UNKNOWN. THE REPORTED STATED, "LOW VAULT CAUSING LENS ROTATION." THE PROBLEM WAS RESOLVED. STATUS OF THE IS REPORTED AS, "EYE IS WELL." Manufacturer Narrative: A4-A6:UNK. H6: PATIENT RELATED FACTOR. CLAIM# (B)(4). | -9.5/5.5/82 | OS | Physician | 04/11/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16896829&pc=QCB | 2023826-2023-01712 | 2023-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -10.00/5.0/78 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. REPORTEDLY, "PATIENT LIVES OUTSIDE OF JORDAN, THAT IS WHY WE HAD DELAYS IN SURGERY DATES." CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR AND UNKNOWN. REPORTER STATED, "LOW VAULT" AS THE CAUSE. THE PROBLEM WAS RESOLVED. STATUS OF THE IS REPORTED AS, "EYE IS WELL." Manufacturer Narrative: H6: PATIENT RELATED FACTOR. CLAIM# (B)(4). | -10.00/5.0/78 | OD | Physician | 04/11/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16897859&pc=QCB | 2023826-2023-01714 | 2023-04-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM513.7 IMPLANTABLE COLLAMER LENS OF A -8.0/4.0/25 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE. REPORTEDLY, "AN OFF-LABEL PSEUDOPHAKIC PATIENT WHO HAD MANY PRIOR EYE SURGERIES... HE TRIED THE 13.7 ICL IMPLANT, DIDN'T SIT WELL DUE TO PATIENT ANATOMY, REMOVED THE LENS, AND TRIED TO IMPLANT THE 13.2 AND ALSO DIDN'T SIT RIGHT SO HE REMOVED THE ICL." Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL - PRE-EXISTING PSUDOPHAKIC PATIENT WITH MANY PRIOR SURGERIES. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6-HEALTH EFFECT- IMPACT CODE: "4627" SHOULD BE REMOVED AND CODE "2199" SHOULD BE ADDED. H6- MEDICAL DEVICE CODE: "2993" SHOULD BE REMOVED AND "3189" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON INTRAOPERATIVELY IMPLANTED AND REMOVED A 13.7MM VTICM5_ 13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.00/4.0/025 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT. REPORTEDLY "WRONG SIZE, UNSTABLE IN SULCUS ON TOP OF IOL." THE LENS WAS REPLACED WITH THE SAME MODEL, SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "QUIET". THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR AND ITS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. CAUSE DETAILS PROVIDED: "BECAME APPARENT INTRAOPERATIVELY THAT THERE WAS COMPROMISE TO THE ZONULE AND THE CAPSULAR BAG WAS UNSTABLE. IT WAS NOT POSSIBLE TO FIT THE ICL IN ITS PROPER ORIENTATION DUE TO SIZING AS WELL AS INSTABILITY OF THE CAPSULE". CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- IMPACT CODE: "4629" SHOULD BE REMOVED AND CODE "4627" SHOULD BE ADDED. H6-MEDICAL DEVICE PROBLEM CODE: "3189" SHOULD BE REMOVED AND CODE "2993" SHOULD BE ADDED. H6- TYPE OF INVESTIGATION CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. G2- "FOREIGN" AND "DISTRIBUTOR/IMPORTER" SHOULD BE REMOVED. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D6B. "(B)(6) 2023" SHOULD BE ADDED. CLAIM# (B)(4). | -8.0/4.0/25 | OD | Physician | 04/12/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16897860&pc=QCB | 2023826-2023-01715 | 2023-04-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF A -9.0/4.0/41 (SPHERE/CYLINDER/AXIS) AS A REPLACEMENT INTO THE PATIENT'S EYE. REPORTEDLY, "AN OFF-LABEL PSEUDOPHAKIC PATIENT WHO HAD MANY PRIOR EYE SURGERIES... HE TRIED THE 13.7 ICL IMPLANT, DIDN'T SIT WELL DUE TO PATIENT ANATOMY, REMOVED THE LENS, AND TRIED TO IMPLANT THE 13.2 AND ALSO DIDN'T SIT RIGHT SO HE REMOVED THE ICL.". Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL - PRE-EXISTING PSUDOPHAKIC PATIENT WITH MANY PRIOR SURGERIES. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-MEDICAL DEVICE PROBLEM CODE: "3189" SHOULD BE REMOVED AND CODE "2993" SHOULD BE ADDED. H6- TYPE OF INVESTIGATION CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. G2- "FOREIGN" AND "DISTRIBUTOR/IMPORTER" SHOULD BE REMOVED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6-HEALTH EFFECT- IMPACT CODE: "4627" SHOULD BE REMOVED AND CODE "2199" SHOULD BE ADDED. H6- MEDICAL DEVICE CODE: "2993" SHOULD BE REMOVED AND "3189" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON A 13.2MM VTICM5_ 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.00/+4.00/041 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE RIGHT (OD) AS A REPLACEMENT LENS ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT. REPORTEDLY "WRONG SIZE, UNSTABLE IN SULCUS ON TOP OF IOL." THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "QUIET". THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR AND ITS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. CAUSE DETAILS PROVIDED: "BECAME APPARENT INTRAOPERATIVELY THAT THERE WAS COMPROMISE TO THE ZONULE AND THE CAPSULAR BAG WAS UNSTABLE. IT WAS NOT POSSIBLE TO FIT THE ICL IN ITS PROPER ORIENTATION DUE TO SIZING AS WELL AS INSTABILITY OF THE CAPSULE". CLAIM# (B)(4). | -9.0/4.0/41 | OD | Physician | 04/12/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16899457&pc=QCB | 2023826-2023-01782 | 2023-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -10.0/1.5/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THERE WAS EXCESSIVE VAULTING. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN BUT NOTED THE DEVICE DID FAIL TO PERFORM AS INTENDED AND "THE SIZE OF ICL CRYSTAL 13.2 IS TOO LARGE AND THE ARCH HEIGHT AFTER SURGERY IS TOO HIGH. FROM A SAFETY PERSPECTIVE, IT IS CONSIDERED TO REPLACE THE SIZE OF CRYSTAL 12.6". Manufacturer Narrative: CLAIM#: (B)(4). | -10.0/1.5/093 | OD | Physician | 05/08/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16904585&pc=QCB | 2023826-2023-01684 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -05.50/+1.0/074 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS FLIPPED UPSIDE DOWN IN THE PATIENTS LEFT EYE (OS). THE LENS TORE UPON REMOVAL FROM THE EYE. THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2: UNK. A3: UNK. A4: UNK. A5: UNK. A6: UNK. B3: UNK. D6A: UNK. D6B: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -05.50/+1.0/074 | OS | Physician | 04/10/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16904590&pc=QCB | 2023826-2023-01698 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -07.00/+2.0/081 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. B3: UNK. (B)(4). | -07.00/+2.0/081 | OD | Physician | 04/11/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16904596&pc=QCB | 2023826-2023-01697 | 2023-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -06.50/+1.5/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. (B)(4). | -06.50/+1.5/086 | OS | Physician | 04/11/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16904617&pc=QCB | 2023826-2023-01692 | 2023-03-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50/+2.0/174 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS WAS IMPLANTED UPSIDE DOWN. THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -08.50/+2.0/174 | NA | Physician | 04/10/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16904640&pc=QCB | 2023826-2023-01694 | 2022-12-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -06.50/+1.5/071 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS REMOVED LATER THAT SAME DAY DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM # (B)(4). | -06.50/+1.5/071 | OD | Physician | 04/11/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16905660&pc=QCB | 2023826-2023-01719 | 2023-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -10.0/3.0/001 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME LENGTH, SPHERICAL LENS, IMPLANTED VERTICALLY DUE TO LOW VAULT WITH ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. REPORTEDLY, "1) INITIAL LENS: IMPLANTED VERTICALLY, 2) EXCHANGE LENS: IMPLANTED HORIZONTALLY, 3) LRI PERFORMED FOR ASTIMATISM." Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. H6- MEDICAL DEVICE PROBLEM CODE: "2923" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. CLAIM# (B)(4). | -10.0/3.0/001 | OS | Physician | 04/12/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16905661&pc=QCB | 2023826-2023-01721 | 2022-11-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -11.0/4.5/109 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON 03-NOV-2022. LENS ROTATION NOT ASSOCIATED TO LOW VAULT WAS OBSERVED. LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2022. THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM DID NOT RESOLVE. REPORTEDLY, "THE SURGEON IS PLANNING TO EXCHANGE THE LENS WITH SPHERICAL ICL AND PERFORM THE LRI FOR ASTIGMATISM. Manufacturer Narrative: DEVICE EVALUATION- H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -11.0/4.5/109 | OS | Physician | 04/12/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16905679&pc=QCB | 2023826-2023-01720 | 2022-05-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -11.0/3.0/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. LENS ROTATION NOT ASSOCIATED TO LOW VAULT WAS OBSERVED. LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2022. THIS DID NOT RESOLVED THE PROBLEM. ON (B)(6) 2022. THE LENS WAS EXCHANGED FOR A SAME LENGTH, SPHERICAL LENS AND THE POBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION- H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -11.0/3.0/088 | OS | Physician | 04/12/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16906150&pc=QCB | 2023826-2023-01796 | 2023-04-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -13.5/+4.0/93 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THERE IS AN OILY SURFACE ON THE LENS, THERE IS NO PATIENT ISSUE WITH NO PLANS TO EXPLANT THE LENS. THE PATIENT IS HAPPY WITH VISION. LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 TYPE OF INVESTIGATION 3331 DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. DEVICE HISTORY RECORD REVIEW FOUND ASSOCIATED SOURCE LOTS WERE MANUFACTURED WITHIN ESTABLISHED PARAMETERS AND LIMITS. (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.5/+4.0/93 | OD | Physician | 04/26/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16906320&pc=QCB | 2023826-2023-01797 | 2023-04-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -11.00/+2.5/88 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS THERE IS AN OILY SURFACE ON THE LENS, THERE IS NO PATIENT ISSUE WITH NO PLANS TO EXPLANT THE LENS. THE PATIENT IS HAPPY WITH VISION. LENS REMAINS IMPLANTED. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION 3331 DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | -11.00/+2.5/88 | OS | Physician | 04/26/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16907193&pc=QCB | 2023826-2023-01786 | 2022-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.00/+1.50/054 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -12.00/+1.50/054 | OD | Physician | 04/10/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16907194&pc=QCB | 2023826-2023-01783 | 2023-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.50/+1.50/097 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -7.50/+1.50/097 | OD | Physician | 04/10/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16907196&pc=QCB | 2023826-2023-01784 | 2023-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+2.0/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: (B)(4). | -13.5/+2.0/094 | OD | Physician | 04/10/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16907197&pc=QCB | 2023826-2023-01788 | 2023-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.50/+2.5/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: (B)(4). | -4.50/+2.5/089 | OS | Physician | 04/10/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16907198&pc=QCB | 2023826-2023-01785 | 2023-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+2.50/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -13.5/+2.50/089 | OS | Physician | 04/10/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16907199&pc=QCB | 2023826-2023-01789 | 2023-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.5/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -11.5/+1.5/091 | OS | Physician | 04/10/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16907200&pc=QCB | 2023826-2023-01787 | 2022-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.0/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -10.5/+1.0/094 | OS | Physician | 04/10/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16907513&pc=QCB | 2023826-2023-01733 | 2023-01-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0/3.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -15.0/3.0/090 | OS | Physician | 05/08/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16907514&pc=QCB | 2023826-2023-01732 | 2023-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/1.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -13.0/1.5/093 | OS | Physician | 05/08/2023 | 2023 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16907515&pc=QCB | 2023826-2023-01735 | 2023-04-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.0/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: UNK CLAIM #(B)(4). | -10.0/1.0/073 | OS | Physician | 05/08/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16907516&pc=QCB | 2023826-2023-01730 | 2023-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.50/3.0/079 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -10.50/3.0/079 | OD | Physician | 05/05/2023 | 2023 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16907520&pc=QCB | 2023826-2023-01734 | 2023-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/2.0/171 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) -2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) CLAIM #(B)(4). | -11.0/2.0/171 | OD | Physician | 05/08/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16907522&pc=QCB | 2023826-2023-01737 | 2023-05-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/2.5/109 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -11.5/2.5/109 | OD | Physician | 05/08/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16907523&pc=QCB | 2023826-2023-01793 | 2023-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 _12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.00/2.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ON (B)(6) 2023 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -12.00/2.0/093 | OD | Physician | 05/09/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16907525&pc=QCB | 2023826-2023-01792 | 2020-03-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.50/4.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -14.50/4.5/091 | OD | Physician | 05/09/2023 | 2023 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16907529&pc=QCB | 2023826-2023-01738 | 2023-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/3.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -12.5/3.0/090 | OD | Physician | 05/08/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16907534&pc=QCB | 2023826-2023-01794 | 2023-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_ 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -12.0/1.5/091 | OS | Physician | 05/09/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16907543&pc=QCB | 2023826-2023-01795 | 2023-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_ 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.00/3.0/074(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION REPORTEDLY, "IT'S QUITE AND PATIENT IS DOING WELL. VAULT IS ADEQUATE AND THE LENS IS STABLE AFTER EXCHANGING WITH THE BIGGER SIZE LENS." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -12.00/3.0/074 | OS | Physician | 05/08/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16908168&pc=QCB | 2023826-2023-01775 | 2022-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.00/+2.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LOW VAULTING AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | -08.00/+2.5/088 | OS | Physician | 04/12/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16908170&pc=QCB | 2023826-2023-01772 | 2023-04-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+2.5/074 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.00/+2.5/074 | OD | Physician | 04/11/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16908181&pc=QCB | 2023826-2023-01700 | 2023-04-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.50/+2.5/089 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE HAPTIC TORE DURING INJECTION INTO THE PATIENTS LEFT EYE (OS). THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.50/+2.5/089 | OS | Physician | 04/11/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16908183&pc=QCB | 2023826-2023-01699 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.50/+1.0/062 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM # (B)(4). | -08.50/+1.0/062 | OS | Physician | 04/11/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16908309&pc=QCB | 2023826-2023-01722 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.MM VTICM513.7 IMPLANTABLE COLLAMER LENS OF -11.0/5.0/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS A REPLACEMENT FOR A PATIENT'S LEFT EYE (OS) ON (B)(6) 2023, DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THE EXCHANGED DID NOT RESOLVE THE PROBLEM. REPORTEDLY, "THE SURGEON IS PLANNING TO EXCHANGE THE LENS WITH SPHERICAL ICL AND PERFORM THE LRI FOR ASTIGMATISM. Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -11.0/5.0/088 | OS | Physician | 04/12/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16908534&pc=QCB | 2023826-2023-01723 | 2023-04-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -11.5/4.5/89(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LOW VAULT WITH ROTATION WAS OBSERVED. CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. THE REPORTED STATED, "SIZE IS TOO SMALL AND LENS STARTED TO ROTATE." THE PROBLEM WAS NOT RESOLVED. REPORTEDLY, "LENS GOT ROTATED OF AROUND 30 DEGREES. MEASURED REFRACTION WAS S+1.25 C-2.50. LARGER LENS IS ALREADY ORDERED. LENS WILL BE REPLACED ONCE THE LENS ARRIVES THE CLINIC." Manufacturer Narrative: B5: LOW VAULT WITH ROTATION AND REFRACTIVE SURPRISE WAS REPORTED. LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS. PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -11.5/4.5/89 | OS | Physician | 04/13/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16912735&pc=QCB | 2023826-2023-01753 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -7.50/1.5/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGED RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4). | -7.50/1.5/094 | OD | Physician | 04/11/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16912742&pc=QCB | 2023826-2023-01752 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -8.00/1.0/124 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGED RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4). | -8.00/1.0/124 | OS | Physician | 04/11/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16913469&pc=QCB | 2023826-2023-01742 | 2023-04-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/1.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -11.5/1.5/092 | OD | Physician | 05/07/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16915068&pc=QCB | 2023826-2023-01757 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -4.50/1/097 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGED RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4). | -4.50/1/097 | OS | Physician | 04/11/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16915805&pc=QCB | 2023826-2023-01754 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -12.50/1.5/089(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGED RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4). | -12.50/1.5/089 | OS | Physician | 04/11/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16916103&pc=QCB | 2023826-2023-01758 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -8.00/2.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGED RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4). | -8.00/2.0/094 | OD | Physician | 04/11/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16916403&pc=QCB | 2023826-2023-01759 | 2023-04-10 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -14.00/1.5/088 (SPHERE/CYLINDER/AXIS) DID NOT OPEN WHEN THE SURGEON WENT TO USE IT. IT WAS REPORTED THAT THERE WAS NO PATIENT CONTACT AND A BACKUP LENS WAS IMPLANTED SUCCESSFULLY WITHIN THE SAME SURGERY. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -14.00/1.5/088 | NA | Other Health Care Professional | 04/11/2023 | 2023 | 68
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16916986&pc=QCB | 2023826-2023-01729 | 2023-04-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.0/3.0/096 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LENS ROTATION NOT ASSOCIATED TO LOW VAULT AND LENS DISLOCATION/SUBLUXATION WERE OBSERVED. LENS REPOSITIONING WAS PERFORMED THAT SAME DAY. THE PROBLEM DID NOT RESOLVE. ON (B)(6)2023, THE LENS WAS EXCHANGED AND THE PROBLEM RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN, HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #: (B)(4). Manufacturer Narrative: H6: 4627 SHOULD BE IN PLACE OF CODE 4629. CLAIM # (B)(4). Manufacturer Narrative: A4-A6:UNK. H6:OFF-LABEL AGE<21. CLAIM# (B)(4). | -12.0/3.0/096 | OD | Physician | 04/14/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16920148&pc=QCB | 2023826-2023-01791 | 2022-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.0/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -10.0/+2.0/094 | OD | Physician | 04/14/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16922077&pc=QCB | 2023826-2023-01691 | 2023-04-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.0/016 (SPHERE/CYLINDER/AXIS), TORE DURING LOADING AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS WAS DISCARDED. THE PATIENT IS DOING WELL AND HAPPY WITH VISION. Manufacturer Narrative: A2: UNK. A3: UNK. A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.00/+1.0/016 | NA | Physician | 04/26/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16922083&pc=QCB | 2023826-2023-01687 | 2023-04-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50/+4.0/070 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS WAS INSERTED UPSIDE DOWN IN THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. POST-OP, THE PATIENT HAD BLURRED VISION AND 1 + CELL. SEE MFR. REPORT # 2023826-2023-01781 FOR REPLACEMENT LENS. Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -17.50/+4.0/070 | OS | Other Health Care Professional | 04/13/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16922307&pc=QCB | 2023826-2023-01780 | 2023-04-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN EVO ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENTS EYE. THE SURGEON REPORTED AN OVERVAULT OF THE LENS. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2: UNK. A3: UNK. A4: UNK. A5: UNK. A6: UNK. B3: UNK. D4: EXPIRATION DATE UNK. D6A: UNK. D6B: UNK. H4: UNK. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.50/+3.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS TOO LONG. SURGEON PREFERENCE TO CHANGE. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS GOOD. D2: COMMON DEVICE NAME: PHAKIC TORIC INTRAOCULAR LENS: PRODUCT CODE: QCB. CLAIM # (B)(4). | -12.50/+3.0/094 | OD | Physician | 04/14/2023 | 2023 | 44 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16922619&pc=QCB | 2023826-2023-01807 | 2023-02-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.00/0.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. PER UPDATED INFORMATION SURGEON DECIDED NOT TO PERFORM SECONDARY INTERVENTION AND TO MONITOR VERY CLOSELY. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.00/0.5/093 | OD | Physician | 04/13/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16922717&pc=QCB | 2023826-2023-01777 | 2023-01-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -14.00/+4.5/110 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND SHALLOWING OF THE ANTERIOR CHAMBER. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS WELL. Manufacturer Narrative: CORRECTED DATA: H6: 4627 SHOULD BE ADDED FOR CORRECTION. H6: 2682 SHOULD BE ADDED FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. CLAIM # (B)(4). | -14.00/+4.5/110 | OD | Physician | 04/13/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16923330&pc=QCB | 2023826-2023-01808 | 2023-02-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.00/0.5/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. PER UPDATED INFORMATION SURGEON DECIDED NOT TO PERFORM SECONDARY INTERVENTION AND TO MONITOR VERY CLOSELY. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.00/0.5/089 | OS | Physician | 04/13/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16924284&pc=QCB | 2023826-2023-01805 | 2023-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -12.50/1.0/106 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND LATER REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6- INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -12.50/1.0/106 | OD | Physician | 04/13/2023 | 2023 | 48 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16924448&pc=QCB | 2023826-2023-01823 | 2023-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -7.0/2.0/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT DIFFERENT AXIS (VERTICALLY). THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: D9: 14-JUN-2023 DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT AND RESIDUE/DEBRIS ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -7.0/2.0/088 | OD | Physician | 05/11/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16924449&pc=QCB | 2023826-2023-01825 | 2023-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.0/1.5/072 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT DIFFERENT AXIS (VERTICALLY). THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 15-JUN-2023. G3: 14/AUG/2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS, SPECIFICALLY FIBRE. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEFORMED WITH FIBRE RESIDUE/DEBRIS ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 TYPE OF INVESTIGATION: 4110 LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -9.0/1.5/072 | OS | Physician | 05/11/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16924451&pc=QCB | 2023826-2023-01821 | 2023-04-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -14.0/1.0/074 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED "ICL SHOWED DAMAGE IN THE HAPTICS". Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4). | -14.0/1.0/074 | OS | Physician | 04/24/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16924690&pc=QCB | 2023826-2023-01804 | 2023-04-10 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6)2023 A 12.6MM, VITICM5_12.6, -8.50/3.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THE LENS WAS IMPLANTED,REMOVED AND REPLACED INTRAOPERATIVELY WITH NO PATIENT INJURY AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS UNKNOWN Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -8.50/3.5/091 | OD | Other Health Care Professional | 04/12/2023 | 2023 | 50 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16924891&pc=QCB | 2023826-2023-01819 | 2023-03-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+1.5/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -12.5/+1.5/093 | OS | Physician | 04/17/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16924892&pc=QCB | 2023826-2023-01817 | 2022-11-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+2.5/089 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -9.0/+2.5/089 | OD | Physician | 04/17/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16924895&pc=QCB | 2023826-2023-01813 | 2023-03-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.5/071 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.5/+1.5/071 | OS | Physician | 04/17/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16924896&pc=QCB | 2023826-2023-01816 | 2023-04-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+2.5/086 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -9.0/+2.5/086 | OS | Physician | 04/17/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16924897&pc=QCB | 2023826-2023-01812 | 2023-03-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.0/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.0/+2.0/090 | OD | Physician | 04/16/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16924898&pc=QCB | 2023826-2023-01820 | 2023-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+2.0/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21 YEARS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -12.0/+2.0/093 | OS | Physician | 04/17/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16924899&pc=QCB | 2023826-2023-01818 | 2023-04-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: (B)(4). | -10.0/+1.5/092 | OD | Physician | 04/17/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16939006&pc=QCB | 2023826-2023-01641 | 2023-04-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Ejection Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A VTICMO13.2, -15.0/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. NO PATIENT CONTACT. A BACKUP LENS WAS IMPLANTED ON (B)(6) 2023 AND THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -15.0/2.0/093 | NA | Physician | 04/24/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16939008&pc=QCB | 2023826-2023-01642 | 2023-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -11.5/3.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. CLAIM # (B)(4). | -11.5/3.5/093 | OS | Physician | 04/24/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16939011&pc=QCB | 2023826-2023-01645 | 2022-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.0/1.0/129 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. REPORTEDLY, THE DOCTOR PUT THE LENS IN PATIENT'S EYE AS INSTRUCTED BY OCOS AND IT WENT TO HIGH VAULT. THEREFORE, DOCTOR THINKS THERE SHOULD BE A PROBLEM WITH THE PRODUCT OR WITH OCOS. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. CLAIM# (B)(4). | -12.0/1.0/129 | OD | Physician | 04/25/2023 | 2023 | 44 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16939013&pc=QCB | 2023826-2023-01646 | 2022-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.0/1.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. REPORTEDLY, THE DOCTOR PUT THE LENS IN PATIENT'S EYE AS INSTRUCTED BY OCOS AND IT WENT TO HIGH VAULT. THEREFORE, DOCTOR THINKS THERE SHOULD BE A PROBLEM WITH THE PRODUCT OR WITH OCOS. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 23-MAY-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS BUT THE FOREIGN MATERIAL ON THE LENS SURFACE WAS SPECIFICALLY FIBRE. DIMENSIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | -12.0/1.0/079 | OS | Physician | 04/25/2023 | 2023 | 44 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16939017&pc=QCB | 2023826-2023-01653 | 2023-04-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND HIGHT VAULT WAS OBSERVED IN PATIENT'S RIGHT EYE (OD). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.2MM VTICM5__13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/3.2/076 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. PER UPDATED INFORMATION REPORTER, "AS SOON AS I HAVE THE FINAL DATA I'LL GET BACK TO YOU". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: A2 - UNK, A3 - UNK, A4 - UNK, A5 - UNK, A6 - UNK, B3 - UNK, D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY), D4 - UNK, D6A - UNK, D6B - UNK, H4 - UNK. CLAIM# (B)(4). | -11.5/3.2/076 | OD | Physician | 04/27/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16939018&pc=QCB | 2023826-2023-01651 | 2023-02-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -6.0/1.5/076 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY ON (B)(6) 2023 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, REPLACEMENT OPERATION DATE TO BE DECIDED. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. (B)(4). | -6.0/1.5/076 | OD | Physician | 04/27/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16939021&pc=QCB | 2023826-2023-01647 | 2023-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.5/3.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED VERTICALLY FOR A LONGER LENGTH LENS, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 14-JUN-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.5/3.0/089 | OD | Physician Assistant | 04/25/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16939025&pc=QCB | 2023826-2023-01648 | 2023-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -9.0/3.0/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. LENS WAS REPOSITIONED ON (B)(6) 2023, THIS DID NOT RESOLVE THE PROBLE. ON 09/MAR/2023 THE LENS WAS EXCHANGED VERTICALLY FOR A LONGER LENGTH LENS, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 14-JUN-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -9.0/3.0/078 | OD | Physician Assistant | 04/25/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16939030&pc=QCB | 2023826-2023-01652 | 2023-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -6.0/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY ON(B)(6) 2023 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, REPLACEMENT OPERATION DATE TO BE DECIDED. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. (B)(4). | -6.0/1.5/089 | OS | Physician | 04/27/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16942467&pc=QCB | 2023826-2023-01858 | 2023-04-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+3.0/089 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS REPORTED AS TEARING DURING LOADING. THE LENS WAS IMPLANTED AND REMAINED IMPLANTED. THERE WAS NO PATIENT INJURY. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. D6B: NA. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.00/+3.0/089 | OS | Physician | 04/28/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16942854&pc=QCB | 2023826-2023-01856 | 2023-04-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.50/+4.0/071 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6)2023. THERE WAS IRRITATION-FREE PHAKIC DUOPHAKIA WITH GOOD VAULTING. THERE WAS SOME REFRACTIVE SURPRISE. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-01855. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | -06.50/+4.0/071 | OS | Physician | 04/16/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16943135&pc=QCB | 2023826-2023-01827 | 2022-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/0.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -8.0/0.5/090 | OD | Physician | 05/11/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16943137&pc=QCB | 2023826-2023-01826 | 2023-03-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/1.5/087(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. REPORTEDLY, "PATIENT UNDERWENT TICL IMPLANTATION IN OUR HOSPITAL ON (B)(6) 2022. ON OCTOBER 23, SHE FOUND THAT THE RIGHT EYE LENS HAD ROTATED 90 DEGREES, SO SHE PLACED AN ORDER ON THE SAME DAY. DURING THE SUBSEQUENT FOLLOW-UP VISIT, SHE FOUND THAT THE ARCH HEIGHT WAS LOWER THAN EXPECTED, AND THEN SHE ADJUSTED THE ORDER AND ORDERED A 12.6 LENS. ON MARCH 19 THIS YEAR, SHE FOUND THAT THE LEFT EYE HAD ROTATED AGAIN, AND THEN SHE ORDERED AN ORDER FOR THE LEFT EYE." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: B4: CORRECTED TO 07-MAY-2023 IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM #(B)(4). | -9.5/1.5/087 | OS | Physician | 05/07/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16943139&pc=QCB | 2023826-2023-01831 | 2022-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.0/173 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -10.0/1.0/173 | OD | Physician | 05/11/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16943140&pc=QCB | 2023826-2023-01830 | 2023-04-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/1.5/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK CLAIM # (B)(4). | -11.0/1.5/072 | OD | Physician | 05/11/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16943146&pc=QCB | 2023826-2023-01835 | 2023-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/1.0/108 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #: (B)(4). | -12.5/1.0/108 | OD | Physician | 05/11/2023 | 2023 | 44 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16943151&pc=QCB | 2023826-2023-01828 | 2022-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8..50/1.0/089(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -8..50/1.0/089 | OD | Physician | 05/11/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16950114&pc=QCB | 2023826-2023-01842 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -8.0/1.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS IMPLANTED AND REMOVED FOR A SHORTER LENGTH LENS. Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -8.0/1.5/094 | OS | Physician | 04/18/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16950116&pc=QCB | 2023826-2023-01841 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF A -7.50/1.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD). THE LENS WAS IMPLANTED AND REMOVED FOR A SHORTER LENGTH LENS. Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -7.50/1.5/089 | OD | Physician | 04/18/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16950307&pc=QCB | 2023826-2023-01896 | 2023-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/2.0/029 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -7.0/2.0/029 | OS | Physician | 05/13/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16950309&pc=QCB | 2023826-2023-01895 | 2023-05-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/1.0/075(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -6.5/1.0/075 | OS | Physician | 05/15/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16950312&pc=QCB | 2023826-2023-01893 | 2023-05-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -10.0/1.0/091 | OD | Physician | 05/15/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16950849&pc=QCB | 2023826-2023-01840 | 2023-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -6.5/1.5/98 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED. Manufacturer Narrative: H6: WORK ORDER FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5- ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL, SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "PATIENT IS HAPPY". CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: G2: GERMANY, CLAIM# (B)(4). | -6.5/1.5/98 | OD | Physician | 04/19/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16950909&pc=QCB | 2023826-2023-01899 | 2023-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0/+2.0/088 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN AND THE REPORTER STATED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CLAIM# (B)(4). | -15.0/+2.0/088 | OD | Physician | 04/19/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16950911&pc=QCB | 2023826-2023-01903 | 2023-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.0/074 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -10.0/+1.0/074 | OS | Physician | 04/20/2023 | 2023 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16950912&pc=QCB | 2023826-2023-01904 | 2022-11-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+3.0/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -9.0/+3.0/092 | OD | Physician | 04/20/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16950913&pc=QCB | 2023826-2023-01901 | 2022-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.0/054 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -8.5/+1.0/054 | OS | Physician | 04/19/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16950946&pc=QCB | 2023826-2023-01844 | 2023-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -12.0/1.0/113 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS LATER REMOVED DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT, BUT NO PATIENT INJURY. ON (B)(6) 2023, A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. STATUS OF THE EYE IS REPORTED AS, "NORMAL." Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -12.0/1.0/113 | OS | Physician | 04/21/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16954888&pc=QCB | 2023826-2023-01917 | 2023-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -8.5/6.0/100 WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(6). Manufacturer Narrative: ADDITIONAL DATA: B5- ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL, SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "PATIENT IS HAPPY". CLAIM# (B)(4). | -8.5/6.0/100 | OS | Physician | 04/18/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16955147&pc=QCB | 2023826-2023-01902 | 2023-03-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.5/+2.0/113 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -17.5/+2.0/113 | OD | Physician | 04/19/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16956397&pc=QCB | 2023826-2023-01920 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-05-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00/1.0/014 (SPHERE/CYLINDER/AXIS), TORE DURING LOADING AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT IS DOING WELL AND HAPPY WITH VISION. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -12.00/1.0/014 | NA | Physician | 04/18/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16956639&pc=QCB | 2023826-2023-01918 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-05-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/1.5/109 (SPHERE/CYLINDER/AXIS), TORE DURING LOADING, THERE WAS NO PATIENT CONTACT. IT WAS REPORTED THAT THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -10.00/1.5/109 | NA | Physician | 04/18/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16957512&pc=QCB | 2023826-2023-01919 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-05-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.50/1.0/035 (SPHERE/CYLINDER/AXIS), TORE DURING LOADING AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT IS DOING WELL AND HAPPY WITH VISION. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -12.50/1.0/035 | NA | Physician | 04/18/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16959162&pc=QCB | 2023826-2023-01905 | 2022-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.0/+3.0/087 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. LATER ON (B)(6) 2023 A SHORTER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -5.0/+3.0/087 | OS | Physician | 04/20/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16959165&pc=QCB | 2023826-2023-01908 | 2022-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.50/+2.5/080 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -12.50/+2.5/080 | OD | Physician | 04/20/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16959166&pc=QCB | 2023826-2023-01907 | 2023-01-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.50/+2.5/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -13.50/+2.5/093 | OS | Physician | 04/20/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16960315&pc=QCB | 2023826-2023-01868 | 2023-03-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.00/+3.5/010 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.00/+3.5/010 | OD | Physician | 04/21/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16963853&pc=QCB | 2023826-2023-01926 | 2023-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/2.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -12.5/2.0/091 | OD | Physician | 05/15/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16963856&pc=QCB | 2023826-2023-01930 | 2022-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.0/066 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -10.5/1.0/066 | OD | Physician | 05/16/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16963862&pc=QCB | 2023826-2023-01929 | 2022-09-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/2.0/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. ON (B)(6)2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -11.0/2.0/085 | OD | Physician | 05/16/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16963863&pc=QCB | 2023826-2023-01927 | 2023-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/1.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -14.0/1.5/090 | OS | Physician | 05/15/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16963865&pc=QCB | 2023826-2023-01931 | 2022-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -10.5/1.5/091 | OS | Physician | 05/16/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16967949&pc=QCB | 2023826-2023-01913 | 2023-04-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.0/045 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM (B)(4). | -9.0/+1.0/045 | OD | Physician | 04/21/2023 | 2023 | 42 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16967950&pc=QCB | 2023826-2023-01912 | 2023-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/+2.0/025 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -5.5/+2.0/025 | OD | Physician | 04/21/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16967961&pc=QCB | 2023826-2023-01911 | 2023-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.0/+2.0/159 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -5.0/+2.0/159 | OS | Physician | 04/21/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16969387&pc=QCB | 2023826-2023-01928 | 2023-05-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/2.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -8.0/2.0/090 | OS | Physician | 05/16/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16969923&pc=QCB | 2023826-2023-01848 | 2023-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -6.5/2.0/70 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6)2023, THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS, IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT. REPORTEDLY, "THE SECOND ICL IS A VERTICAL IMPLANT." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM #(B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL DRY AND WITH RESIDUE/DEBRI ON THE LENS. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: UNK. CLAIM# (B)(4). | -6.5/2.0/70 | OS | Physician | 04/24/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16970382&pc=QCB | 2023826-2023-01849 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Improper or Incorrect Procedure or Method | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -12.0/1.0/074 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS A REPLACEMENT INTO THE PATIENT'S RIGHT EYE (OD) (B)(6) 2023, DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. CAUSE OF THE EVENT WAS USER ERROR. THE REPORTER SAID, "I PUT IN A 13.2MM AND SHOULD HAVE BEEN 12.6 IN THE SIGHT." THE PROBLEM WAS RESOLVED. REPORTEDLY, "THE CALCULATION CALLED FOR 13.2MM BUT IT WAS WAY TOO BIG. 12.6MM STILL HAD A 500+M VAULT." Manufacturer Narrative: ADDITONAL DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.00/+1.00/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023 AS A REPLACEMENT LENS. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT- "12.6MM STILL HAD A 500+M VAULT". REPORTEDLY "THIS HAS BEEN RESOLVED. PER DR. SHARPE: THE 13.2 MM LENS WAS WAY TOO BIG. THE 12.6 MM LENS WAS SLIGHTLY BIG BUT JUST FINE. IT'S ALL GOOD NOW. THE PATIENT IS SEEING WELL POSTOPERATIVELY." THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS USER ERROR. REPORTEDLY " I THINK IT WAS TOTALLY MY FAULT IN THE MISMEASURE OF HER WHITE TO WHITE. NOT A DEVICE PROBLEM". H6- INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. G2- "FOREIGN" AND "DISTRUBUTOR/IMPORTER" SHOULD BE REMOVED. CLAIM# (B)(4). Manufacturer Narrative: H6: USER ERROR. (B)(4). | -12.0/1.0/074 | OD | Physician | 04/24/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16970485&pc=QCB | 2023826-2023-01850 | 2023-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -9.0/1.5/107 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. OUTPUT AXIS TRANSMISSION ERROR WAS NOTED. REPORTEDLY, "THE OUTPUT AXIS WAS 111 NOT 11. TRANSMISSION ERROR. LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITONAL DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.5/107 (SPHERE/CYLINDER/AXIS) INTO THE RIGHT EYE (OD) ON (B)(6) 2023. REPORTEDLY "OUTPUT AXIS TRANSMISSION ERROR". ON (B)(6) 2023 THE LENS WAS EXPLANTED AND LATER REPLACED WITH THE SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION PROVIDED: "PATIENT IS HAPPY". THE REPORTER INDICATED THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. CAUSE DETAILS: "THE OUTPUT AXIS WAS 111° NOT 11°. TRANSMISSION ERROR!". H6-HEALTH EFFECT - IMPACT CODE: "2199" SHOULD BE REMOVED AND "4627" SHOULD BE ADDED. H6- MEDCIAL DEVICE PROBLEM CODE: "2993" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT:UNK, ETHNICITY:UNK, RACE:UNK H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM#: (B)(4). | -9.0/1.5/107 | OD | Physician | 04/25/2023 | 2023 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16976347&pc=QCB | 2023826-2023-01933 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.00/+4.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD). THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: B5: THE LENS WAS IMPLANTED ON (B)(6). 2022. CLAIM # (B)(4). | -17.00/+4.0/089 | OD | Physician | 04/25/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16976426&pc=QCB | 2023826-2023-01864 | 2023-04-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.50/+3.0/104 (SPHERE/CYLINDER/AXIS), TORE DURING INJECTION. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4). Manufacturer Narrative: B5: THE LENS WAS INSERTED AND REMOVED DURING THE SAME SURGERY. CLAIM # (B)(4). | -13.50/+3.0/104 | NA | Physician | 04/25/2023 | 2023 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16976437&pc=QCB | 2023826-2023-01861 | 2023-03-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.00/+3.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LOW VAULTING AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4: UNK A5: UNK A6: UNK H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WERE FIBERS ON THE LENS SURFACE. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). | -13.00/+3.0/094 | OD | Physician | 04/27/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16976438&pc=QCB | 2023826-2023-01866 | 2023-04-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.50/+4.0/091 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. (B)(4). | -13.50/+4.0/091 | NA | Physician | 04/26/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16976473&pc=QCB | 2023826-2023-01862 | 2023-03-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.00/+2.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LOW VAULTING AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). | -13.00/+2.0/091 | OS | Physician | 04/27/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16976485&pc=QCB | 2023826-2023-01865 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/070 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). | -10.50/+1.0/070 | OD | Physician | 04/25/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16984879&pc=QCB | 2023826-2023-01851 | 2023-02-28 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -17.5/4.0/81 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON 28-FEB-2023. THE LENS WAS INTRAOPERATIVELY IMPLANTED AND REMOVED DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. ON 09-MAR-2023, THE REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSEOF THE EVENT IS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: HEALTH EFFECT - IMPACT CODE: "4629" SHOULD BE REMOVED AND "2199" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "2993" SHOULD BE REMOVED AND "3189" SHOULD BE ADDED. (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS. CLAIM# (B)(4). | -17.5/4.0/81 | OD | Physician | 04/25/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16986524&pc=QCB | 2023826-2023-01852 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.0/2.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6)2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. REPORTEDLY, "HIGH VAULT. EXCHANGED FROM 13.2 TO 12.6. ISSUE RESOLVED." Manufacturer Narrative: CLAIM#: (B)(4). | -10.0/2.0/092 | OD | Physician | 04/25/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16986610&pc=QCB | 2023826-2023-01853 | 2023-03-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5/3.0/078 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. ON (B)(6)2023, THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS, IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT. REPORTEDLY, "THE SECOND IENS WAS IMPLANTED VERTICALLY." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL USE - ACD <3.0. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC/HAPTIC DEFORMED. REHYDRATION OF THE LENS HAS CONFORMED SIGNIFICANTLY + LENS DAMAGE, THUS UNABLE TO PERFORM DIMENSION INSPECTION FOR LENS HEIGHT. DIMENSIONAL INSPECTION FOUND THE OVERALL WIDTH WITHIN SPECIFICATIONS BUT OVERALL LENGTH COULD NOT BE PERFORMED DUE TO THE DAMAGED STATE OF THE LENS. CLAIM# (B)(4). | -9.5/3.0/078 | OD | Physician | 04/26/2023 | 2023 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16987460&pc=QCB | 2023826-2023-01854 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -5.0/4.5/80 (SPHERE/CYLINDER/AXIS) WAS TO BE INSERTED INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS NOT IMPLANTED. REPORTEDLY, "WAS TORN LENS INTRA-OP, SCHEDULED TO DO REPLACEMENT LATER AND THEN PATIENT GOT PREGNANT." Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: AFTER CLARIFICATON FROM REPORTER, THIS EVENT WAS FOUND TO BE NOT REPORTABLE DUE TO LENS NOT BEING USED. HOWEVER, THE CLAIM CANNOT BE VOIDED AS AN MDR HAS ALREADY BEEN SUBMITTED. CLAIM# (B)(4). | -5.0/4.5/80 | OS | Physician | 04/26/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16988092&pc=QCB | 2023826-2023-01991 | 2023-03-31 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM513.2, -11.0/5.0/114 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT THERE WAS NO PATIENT INJURY. THE REPORTER STATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B1- CORRECTED REPORT TYPE FROM ADVERSE EVENT TO PRODUCT PROBLEM. B5-THE REPORTER INDICATED THAT A 13.2MM, VTICM513.2, -11.0/5.0/114 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY WITH NO PATIENT INJURY. ON (B)(6) 2023 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. THE REPORTER STATED THAT THE CAUSE OF THE EVENT WAS UNKNOWN. H1- CORRECTED TYPE OF REPORT EVENT FROM SERIOUS INJURY TO MALFUNCTION. H6- HEALTH EFFECT-IMPACT CODE: 4629 TO 2199, PREVIOUS CODE NOT APPLICABLE. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -11.0/5.0/114 | OD | Physician | 04/24/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16988707&pc=QCB | 2023826-2023-01873 | 2023-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -17.00/+3.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6)2023. THE LENS WAS EXPLANTED ON (B)(6)2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -17.00/+3.0/093 | OS | Physician | 04/26/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16988817&pc=QCB | 2023826-2023-01867 | 2023-04-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.50/+2.0/153 (SPHERE/CYLINDER/AXIS) DURING THE SAME SURGERY DUE TO THE LENS WAS STUCK IN THE CARTRIDGE OR INJECTION SYSTEM AND TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS ON (B)(6) 2023 AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING FINE. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.50/+2.0/153 | OD | Physician | 04/26/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16989731&pc=QCB | 2023826-2023-01940 | 2023-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/1.5/015 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -12.5/1.5/015 | OS | Physician | 05/17/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16989732&pc=QCB | 2023826-2023-01936 | 2023-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ON (B)(6) 2023 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -10.0/1.0/091 | OD | Physician | 05/16/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16989733&pc=QCB | 2023826-2023-01935 | 2023-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/2.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -7.5/2.5/093 | OD | Physician | 05/16/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16989734&pc=QCB | 2023826-2023-01937 | 2023-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.0/086(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -10.5/1.0/086 | OS | Physician | 05/16/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16989735&pc=QCB | 2023826-2023-01939 | 2023-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/2.0/002 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -13.5/2.0/002 | OD | Physician | 05/17/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16989741&pc=QCB | 2023826-2023-01961 | 2023-03-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/2.0/070 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -10.5/2.0/070 | OD | Physician | 05/19/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16989742&pc=QCB | 2023826-2023-01958 | 2022-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/2.0/013 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO LOW VAULT WITHOUT. ON (B)(6) 2023 A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED. REPORTEDLY, "THE INITIAL LENS WAS VERTICALLY IMPLANTED." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -10.0/2.0/013 | OS | Physician | 05/18/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16989745&pc=QCB | 2023826-2023-01963 | 2023-03-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/1.5/168 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. REPORTEDLY, "THE FIRST TICL WAS IMPLANTED VERTICALLY, AND THE SECOND TICL WAS CHANGED TO BE IMPLANTED HORIZONTALLY." Manufacturer Narrative: A4-A6: UNK. H6: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #: (B)(4). | -6.0/1.5/168 | OD | Physician | 05/19/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16989905&pc=QCB | 2023826-2023-02013 | 2023-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+4.0/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. PER THE MEDICAL REVIEW THIS COMPLAINT IS DEEMED SERIOUS AND DEVICE CAUSALITY IS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) AND (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -8.5/+4.0/094 | OD | Physician | 04/24/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16989906&pc=QCB | 2023826-2023-02014 | 2022-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.5/075 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -10.5/+1.5/075 | OS | Physician | 04/24/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16991206&pc=QCB | 2023826-2023-01768 | 2023-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.0/1.0/114 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THERE ARE NO PLANS TO IMPLANT ANOTHER LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 14-JUN-2023. G3: 14/AUG/2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT AND RESIDUE/DEBRIS ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. CLAIM# (B)(4). | -10.0/1.0/114 | OS | Physician | 05/04/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16991208&pc=QCB | 2023826-2023-01767 | 2023-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.0/1.5/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE LENS WAS REPLACED VERTICALLY WITH A SAME LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: (B)(6) 2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. CLAIM#(B)(4). | -9.0/1.5/107 | OD | Physician | 05/04/2023 | 2023 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16991213&pc=QCB | 2023826-2023-01770 | 2023-05-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.0/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED VERTICALLY WITH A SAME LENGTH LENS DUE TO EXCESSIVE VAULT IN A SECOND SURGERY ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM# (B)(4). | -7.0/2.0/093 | OD | Physician | 05/04/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16997233&pc=QCB | 2023826-2023-02020 | 2023-03-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.5/113 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE LENS WAS EXCHANGED ON (B)(6)2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+2.5/113 | OD | Physician | 04/24/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16997235&pc=QCB | 2023826-2023-02017 | 2023-03-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0/+4.5/077 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE LENS WAS EXCHANGED ON (B)(6)2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -15.0/+4.5/077 | OS | Physician | 04/24/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16997236&pc=QCB | 2023826-2023-02021 | 2023-03-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+2.0/112 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LOW VAULT. IN A SEPARATE SURGERY THAT DAY A LONGER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.5/+2.0/112 | OD | Physician | 04/24/2023 | 2023 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16997239&pc=QCB | 2023826-2023-02016 | 2023-03-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+5.5/114 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.0/+5.5/114 | OD | Physician | 04/24/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16997253&pc=QCB | 2023826-2023-02019 | 2023-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.0/145 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -10.0/+1.0/145 | OS | Physician | 04/24/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16998311&pc=QCB | 2023826-2023-01943 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+1.5/008 (SPHERE/CYLINDER/AXIS), WAS REPORTED AS "APPEARS TO BE FAULTY". THE LENS WAS NOT USED AND THE BACKUP LENS WAS USED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6:SHOULD BE CORRECTED TO STATE "THE DHR REVIEW INDICATES THAT THE PRODUCT HAS NOT BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS AND THAT THERE IS INDICATION THAT THE MANUFACTURING AND/OR PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. A DEFINITIVE ROOT CAUSE FOR THIS EVENT WAS UNABLE TO BE DETERMINED. AS A PREVENTATIVE ACTION, WI-0716 AND WI-0801 HAS BEEN UPDATED TO REDUCE THE LIKELIHOOD OF DAMAGE CAUSED BY INADVERTENT CONTACT BETWEEN THE LENS AND THE DRY GLASS SURFACE OF THE 2ML VIAL". CLAIM# (B)(4). Manufacturer Narrative: A2: UNK. A3: UNK. A4: UNK. A5: UNK. A6: UNK. B3: UNK. H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -16.50/+1.5/008 | NA | Physician | 05/03/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16998322&pc=QCB | 2023826-2023-01874 | 2023-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.00/+4.0/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | -18.00/+4.0/080 | OD | Physician | 04/26/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16998336&pc=QCB | 2023826-2023-01945 | 2023-03-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.00/+1.0/087 (SPHERE/CYLINDER/AXIS), TORE DURING LOADING. THE LENS RUPTURE WAS DUE TO SNAGGING WHILE LOADING INJECTOR. THE LENS WAS NOT IMPLANTED. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -08.00/+1.0/087 | NA | Physician | 04/30/2023 | 2023 | 44 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16998837&pc=QCB | 2023826-2023-01964 | 2023-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.0/035 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.0/035 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. REPORTEDLY, ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SAME LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION STATED, "FIRST VERTICAL POSITION". CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.0/035 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REPORTEDLY, ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: D9: RETURN DATE: 17-JUL-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN, BROKEN, AND BENT WITH RESIDUE ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -10.5/1.0/035 | OS | Physician | 05/19/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16998838&pc=QCB | 2023826-2023-01971 | 2023-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/1.5/068 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "THE SECOND PIECE IS VERTICAL." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.0/1.5/068 | OS | Physician | 05/19/2023 | 2023 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16998840&pc=QCB | 2023826-2023-01968 | 2023-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/1.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE (B)(4). (B)(4). | -11.0/1.0/086 | OS | Physician | 05/19/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16998841&pc=QCB | 2023826-2023-01967 | 2023-03-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/2.5/095(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED A LOW VAULT WITHOUT ROTATION. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM (B)(4). | -13.0/2.5/095 | OD | Physician | 05/19/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16998842&pc=QCB | 2023826-2023-01973 | 2023-03-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.5/1.0/091 | OD | Physician | 05/22/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16998843&pc=QCB | 2023826-2023-01969 | 2023-01-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/5.0/055 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "THE FIRST TICL WAS IMPLANTED HORIZONTALLY, AND THE CRYSTAL WAS LABELED AT 45 CCW." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -8.0/5.0/055 | OS | Physician | 05/19/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16998844&pc=QCB | 2023826-2023-01972 | 2023-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/2.5/068 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -5.5/2.5/068 | OD | Physician | 05/19/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16998845&pc=QCB | 2023826-2023-01974 | 2023-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/2.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL/LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. REPORTEDLY, "BECAUSE OF THE PATIENT'S CONCERNS ABOUT THE ROTATION OF THE ASTIGMATIC CRYSTALS, THE NON-DISPERSIVE CRYSTALS WERE REPLACED AFTER COMMUNICATION WITH THE DOCTOR." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -7.5/2.0/094 | OD | Physician | 05/22/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16998846&pc=QCB | 2023826-2023-01970 | 2023-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.5/111 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "THE SECOND PIECE IS VERTICAL." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -10.0/1.5/111 | OD | Physician | 05/19/2023 | 2023 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16999553&pc=QCB | 2023826-2023-01995 | 2023-04-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 12.6MM, VTICMO12.6, -8.5/1.5/067 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS LEFT EYE (OS). THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY, WITH NO PATIENT INJURY. ON (B)(6) 2023 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: D4: MODEL-VTICMO_12.6. D9: RETURN DATE: 17-JUL-2023. CLAIM# (B)(4). Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -8.5/1.5/067 | OS | Physician | 04/28/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16999842&pc=QCB | 2023826-2023-01996 | 2023-04-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 13.2MM, VTICM513.2 -13.0/2.0/092 (SPHERE/CYLINDER/AXIS) IMPLATABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE RIGHT (OD) EYE. THERE WAS NO PATIENT CONTACT OR INJURY. ON (B)(6) 2023 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.0/2.0/092 | OD | Physician | 04/28/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17007064&pc=QCB | 2023826-2023-02046 | 2023-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT AN TORIC IMPLANTABLE COLLAMER LENS HAD ROTATED ON (B)(6) 2023. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 12.6 VTICM5_12.6 -11.50/2.0/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED ON (B)(6) 2023 INTO THE PATIENT'S LEFT EYE (OS). LOW VAULT WITH ROTATION WAS OBSERVED. LENS WAS EXCHANGED ON (B)(6) 2023 WITH A LONGER LENGTH LENS AND PROBLEM WAS RESOLVED. QUIET LENS IN POSITION, PATIENT HAPPY. CLAIM# (B)(4). Manufacturer Narrative: CLAIM (B)(4). | -11.50/2.0/91 | OS | Physician | 05/01/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17007355&pc=QCB | 2023826-2023-02044 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 -5.00/2.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO HIGH VAULT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4) . Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- IMPACT CODE: "4629" SHOULD BE REMOVED AND "4627" SHOULD BE ADDED. DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -5.00/2.5/090 | OD | Other Health Care Professional | 04/25/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17008097&pc=QCB | 2023826-2023-01975 | 2023-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/2.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL/LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. REPORTEDLY, "DUE TO CONCERNS ABOUT THE ROTATION OF THE ASTIGMATIC CRYSTALS, THE NON-ASTIGMATIC CRYSTALS WERE REPLACED AFTER COMMUNICATION WITH THE DOCTOR." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -8.0/2.0/090 | OS | Physician | 05/22/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17008100&pc=QCB | 2023826-2023-01976 | 2023-05-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/1.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -6.5/1.0/088 | OS | Physician | 05/22/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17008112&pc=QCB | 2023826-2023-01984 | 2023-03-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/0.5/075 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -11.0/0.5/075 | OS | Physician | 05/21/2023 | 2023 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17008116&pc=QCB | 2023826-2023-01987 | 2023-03-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -10.0/1.5/094 | OS | Physician | 05/21/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17008118&pc=QCB | 2023826-2023-01986 | 2023-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED, THAT THE SURGEON IMPLANTED A 13.2MM VTICMO, 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/1.0/127 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS, DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. (B)(4). | -8.5/1.0/127 | OS | Physician | 05/21/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17008129&pc=QCB | 2023826-2023-02001 | 2023-01-29 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023, A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF A -6.5/1.0/093 (SPHERE/CYLINDER/AXIS) TORE DURING LOADING. REPORTEDLY, "LENS RUPTURE DUE TO SNAGGING WHILE LOADING INJECTOR." A REPLACEMENT LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE(OS) AND THE PROBLEM WAS RESOLVED. REPORTEDLY, STATUS OF THE EYE IS, "GOOD." Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM (B)(4). | -6.5/1.0/093 | OS | Physician | 04/30/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17008130&pc=QCB | 2023826-2023-02003 | 2023-02-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -5.00/2.00/176 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED, AND REPLACED FOR A SAME SIZE LENS IMPLANTED VERTICALLY, DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. STATUS OF THE EYE IS, "GOOD." THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: H6- INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. HEALTH EFFECT - IMPACT CODE: "4629" SHOULD BE REMOVED AND "2199" SHOULD BE ADDED. MEDICAL DEVICE PROBLEM CODE: "3189" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: CLAIM (B)(4). | -5.00/2.00/176 | OS | Physician | 04/30/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17008131&pc=QCB | 2023826-2023-02002 | 2023-02-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -6.0/3.0176 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED, AND REPLACED FOR A SAME SIZE LENS IMPLANTED VERTICALLY, DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. STATUS OF THE EYE IS, "GOOD." THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIA IN LIQUID WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VIISBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. CLAIM# (B)(4). | -6.0/3.0176 | OD | Physician | 04/30/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17008162&pc=QCB | 2023826-2023-02004 | 2023-02-23 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023, A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF A -5.50/1.50/087 (SPHERE/CYLINDER/AXIS) TORE DURING LOADING. REPORTEDLY, "LENS RUPTURE DUE TO SNAGGING WHILE LOADING INJECTOR." A REPLACEMENT LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE(OD) AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS OTHER. REPORTEDLY, STATUS OF THE EYE IS, "GOOD." Manufacturer Narrative: ADDITIONAL DATA: H6- HEALTH EFFECT - IMPACT CODE: "4629" SHOULD BE REMOVED AND "2199" SHOULD BE ADDED. D6A.- "(B)(6) 2023" SHOULD BE REMOVED. D6B.- "(B)(6) 2023" SHOULD BE REMOVED. B5- THE REPORTER INDICATED THE SURGEON NOTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/+1.5/087 (SPHERE/CYLINDER/AXIS) TORE/BROKE DURING LOADING ON (B)(6) 2023. THE DAMAGE OCCURRED DURING LOADING AND WAS NOTED AS "LENS RUPTURE DUE TO SNAGGING." THE LENS WAS NOT IMPLANTED. ON THE SAME DAY A REPLACEMENT LENS OF THE SAME MODEL AND LENGTH, WITH A DIFFERENT POWER WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS NOTED TO BE "GOOD". THE CAUSE OF THE EVENT IS "OTHER" AND ITS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. CAUSE DETAILS: "LENS RUPTURE DUE TO SNAGGING WHILE LOADING INJECTOR". CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -5.50/1.50/087 | OD | Physician | 04/30/2023 | 2023 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17008282&pc=QCB | 2023826-2023-02023 | 2023-03-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/073 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -9.5/+1.0/073 | OS | Physician | 04/26/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17011910&pc=QCB | 2023826-2023-02007 | 2023-02-23 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -1.50/5.5/121 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS TORE DURING LOADING AND INJECTION/DELIVERY INTO THE EYE. NO PATIENT INJURY WAS NOTED. REPORTELY, "CLEAR CORNEA, AC Q&D, ICL ROTATED ALMOST 80 DEGREE REFRACTION WITH ICL ROTATED (+2.75/-8.00@10 FOR OD)," AND "VA WAS 20/25 FOR 10 DAYS POST SURGERY, THEN ON 23RD OF FEB. 2023 PATIENT COMPLAINED OF SUDDEN DECREASE OF VA THEN PATIENT CAME TO CLINIC, AND ICL ROTATION WAS DETECTED, THEN DECISION FOR ICL REPLACEMENT WAS DONE, AFTER RECIEVING THE NEW ICL TO BE IMPLANTED, AND JUST 2 DAYS BEFORE PLANNED SURGERY FOR ICL REPLACEMENT, PATIENT GOT PREGNANT BEFORE SCHEDULED SURGERY, SO THE SURGERY WAS POSTPONED AFTER THE CHILD BIRTH BY 3-6 MONTHS AT LEAST." Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -1.50/5.5/121 | OD | Physician | 05/02/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17012390&pc=QCB | 2023826-2023-02050 | 2023-01-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2. -12.0/2.0/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY WITH NO INJURY DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. ON (B)(6) 2023 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H4 - CORRECTED TO (B)(6) 2021 IN INITIAL MDR. CLAIM #(B)(4). | -12.0/2.0/073 | OS | Physician | 04/27/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17013767&pc=QCB | 2023826-2023-01946 | 2023-03-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50/+1.5/094 (SPHERE/CYLINDER/AXIS), TORE DURING LOADING, THE LENS RUPTURE WAS DUE TO SNAGGING WHILE LOADING THE INJECTOR. THE LENS WAS NOT IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.50/+1.5/094 | NA | Physician | 04/30/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17013966&pc=QCB | 2023826-2023-02064 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: CUSTOMER RETURNED A 12.1MM VTICM5_12.1 -11.00/1.0/126 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WITH NO INFORMATION, OPEN BOX. NO FURTHER INFORMATION HAS BEEN PROVIDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. LENS VIAL AND CASE RETURNED UNOPENED. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -11.00/1.0/126 | NA | Other Health Care Professional | 04/26/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17015026&pc=QCB | 2023826-2023-02008 | 2023-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.50/3.5/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. EXCESSIVE VAULT WITHOUT ROTATION WAS OBSERVED. REPORTEDLY, "THE PATIENT COMPLAINED OF DIZZINESS AT NIGHT, NEVER ADAPTED, AND WAS RELUCTANT TO RE-IMPLANT. ON (B)(6)2023, THE LENS WAS REMOVED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH PERFORMED AND 1 SIMILAR COMPLAINT FOUND. INVESTIGATION FOR COMPLAINT #(B)(4) FOUND NO INDICATION THAT MANUFACTURING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 18-JUL-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -10.50/3.5/090 | OD | Physician | 05/03/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17015027&pc=QCB | 2023826-2023-02009 | 2023-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5/3.5/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED. REPORTEDLY, "THE PATIENT COMPLAINED OF DIZZINESS AT NIGHT, NEVER ADAPTED, AND WAS RELUCTANT TO RE-IMPLANT. ON (B)(6) 2023, THE LENS WAS REMOVED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM#: (B)(4). | -11.5/3.5/086 | OS | Physician | 05/03/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17015192&pc=QCB | 2023826-2023-02025 | 2023-03-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.5/+3.5/175 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -16.5/+3.5/175 | OD | Physician | 04/27/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17015194&pc=QCB | 2023826-2023-02024 | 2022-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.0/109 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS TWO SIZES LONGER IN LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -8.5/+1.0/109 | OD | Physician | 04/27/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17015207&pc=QCB | 2023826-2023-02026 | 2023-03-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.5/175 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.5/+1.5/175 | OS | Physician | 04/27/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17015225&pc=QCB | 2023826-2023-02052 | 2022-12-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.5/1.5/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -15.5/1.5/103 | OD | Physician | 05/23/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17015226&pc=QCB | 2023826-2023-02053 | 2023-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.0/2.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -5.0/2.5/092 | OD | Physician | 05/23/2023 | 2023 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17015229&pc=QCB | 2023826-2023-02056 | 2023-03-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_ 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8..5/6.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -8..5/6.0/093 | OD | Physician | 05/24/2023 | 2023 | 43 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17015230&pc=QCB | 2023826-2023-02051 | 2023-04-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/2.0/070 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -9.5/2.0/070 | OS | Physician | 05/23/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17015232&pc=QCB | 2023826-2023-02057 | 2023-04-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/1.5/166 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXPLANTED DUE TO LOW VAULT WITHOUT ROTATION. ON (B)(6)2023 A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED. REPORTEDLY, "THE POSTERIOR ARCH HEIGHT OF THE PATIENT WAS IMPLANTED WITH 12.1 SIZE LENS. AFTER SELECTING THE LENS ACCORDING TO THE PREOPERATIVE OPTOMETRY DEGREE, THE PATIENT HAD POOR SELF-ADJUSTMENT ABILITY AND INTOLERANCE. FINALLY, THE LENS WAS SELECTED ACCORDING TO THE PATIENT'S PREVIOUS EYE DEGREE, AND THE POSTOPERATIVE BCVA WAS 20/20 WITHOUT DISCOMFORT." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -11.5/1.5/166 | OD | Physician | 05/24/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17015233&pc=QCB | 2023826-2023-02058 | 2023-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/2.0/179 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -11.5/2.0/179 | OS | Physician | 05/24/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17015321&pc=QCB | 2023826-2023-02072 | 2022-12-24 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS WITH DIOPTER -9.50/+2.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED DUE TO LOW VAULT. IN A SEPARATE SURGERY THAT SAME DAY THE LENS WAS REPLACED WITH A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS WITH DIOPTER -9.50/+2.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: PER THE UPDATED COMPLAINT QUESTIONNAIRE THE REPORTER CONFIRMED THE INITIAL LENS WAS IMPLANTED, REMOVED AND REPLACED WITHIN THE INITIAL SURGERY. CLAIM# (B)(4). | -9.50/+2.0/091 | OD | Physician | 04/30/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17015324&pc=QCB | 2023826-2023-02071 | 2023-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/089 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -10.0/+1.5/089 | OD | Physician | 04/28/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17015327&pc=QCB | 2023826-2023-02074 | 2023-02-03 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/+1.0/085 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -5.5/+1.0/085 | OS | Physician | 04/30/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17015338&pc=QCB | 2023826-2023-02076 | 2023-03-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.00/+2.00/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.00/+2.00/093 | OS | Physician | 04/30/2023 | 2023 | 45 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17015474&pc=QCB | 2023826-2023-02012 | 2022-12-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -6.5/0.5/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS, IMPLANTED IN A OBLIQUE POSITION DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BROKEN/BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: H6: OFF-LABEL - ACD<3.0. CLAIM# (B)(4). | -6.5/0.5/93 | OD | Physician | 05/04/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17019245&pc=QCB | 2023826-2023-01952 | 2022-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.50/+1.0/109 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.50/+1.0/109 | OD | Physician | 05/04/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17019268&pc=QCB | 2023826-2023-01951 | 2023-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT POWER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FIBERS WERE FOUND ON THE LENS SURFACE. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. REFRACTIVE (FUNCTIONAL) VERIFICATION WAS PERFORMED AND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6- TYPE OF INVESTIGATION CODE:3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.00/+2.0/080 | OS | Physician | 05/03/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17019302&pc=QCB | 2023826-2023-01949 | 2022-01-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -07.50/+2.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2022 BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH A SAME LENGTH BUT DIFFERENT MODEL AND POWER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). | -07.50/+2.5/090 | OD | Physician | 05/03/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17019341&pc=QCB | 2023826-2023-01950 | 2021-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -06.00/+2.0/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6)2021. THE LENS WAS REMOVED ON (B)(6)2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6)2021 BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH A SAME LENGTH BUT DIFFERENT MODEL AND POWER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6: WORK ORDER SEARCH: ONE ADDITIONAL SIMILAR COMPLAINT TYPE EVENT WITHIN ASSOCIATED LOTS WAS FOUND. CLAIM #(B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -06.00/+2.0/096 | OS | Physician | 05/03/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17019360&pc=QCB | 2023826-2023-01953 | 2022-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/078 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/078 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE LENS WAS REPOSITIONED ON (B)(6) 2022 BUT THE PROBLEM WAS NOT RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. CLAIM # (B)(4). | -11.00/+1.5/078 | OS | Physician | 05/04/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17022936&pc=QCB | 2023826-2023-02137 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT 13.2MM VTICM_13.2 IMPLANTABLE COLLAMER LENS OF -5.00/4.0/159 WAS IMPLANTED INTO THE PATIENTS' RIGHT EYE (OD) AND THE PATIENT EXPERIENCED EXCESSIVE VAULT. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -5.00/4.0/159 WAS IMPLANTED INTO THE PATIENTS' RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULT. LENS EXPLANTED ON AN UNKNOWN DATE WITH PLANNED REPLACEMENT. CLAIM#(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON (B)(6) 2023, THE LENS WAS EXPLANTED; IN A SEPARATE SURGERY THAT SAME DAY THE LENS WAS REPLACED WITH A SMALLER LENGTH LENS. THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -5.00/4.0/159 | OD | Physician | 05/02/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17024814&pc=QCB | 2023826-2023-02139 | 2023-03-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -17.50/2.5/114 (SPHERE/CYLINDER/AXIS) TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT WITH NO PATIENT INJURY. ON (B)(6) 2023 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -17.50/2.5/114 | NA | Physician | 05/02/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17025657&pc=QCB | 2023826-2023-02138 |  | Injury | STAAR SURGICAL COMPANY | 2023-05-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -13.5/2.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -13.5/2.5/094 | OD | Physician | 05/02/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17027841&pc=QCB | 2023826-2023-02152 | 2022-11-15 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6; -15/-2.50/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS HAD TORE DURING INJECTIVE/DELIVERY INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THERE WAS PATIENT CONTACT WITH NO INJURY, BUT THE LENS WAS NOT IMPLANTED. ON (B)(6) 2022 A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED "LENS DID NOT SEPARATE FROM THE FOAM TIP DURING INJECTION". Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -15/-2.50/89 | OS | Physician | 05/26/2023 | 2023 | 43 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17027842&pc=QCB | 2023826-2023-02151 | 2023-05-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -9.0/1.0/169 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON THIS SAME DATE A REPLACEMENT LENS OF THE SAME LENGTH BUT IMPLANTED HORIZONTALLY. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 17-JUL-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -9.0/1.0/169 | OS | Physician | 05/30/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17027843&pc=QCB | 2023826-2023-02150 | 2023-05-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.5/1.5/013 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON THIS SAME DATE A REPLACEMENT LENS OF THE SAME LENGTH BUT IMPLANTED HORIZONTALLY. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 17-JUL-2023 . H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). | -11.5/1.5/013 | OD | Physician | 05/30/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17027845&pc=QCB | 2023826-2023-02149 | 2023-04-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -14.0/3.0/001 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS TORE DURING INJECTION/DELIVERY INTO THE EYE. INTRAOPERATIVELY THE LENS WAS REMOVED WITH NO PATIENT INJURY. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND INDICATED THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -14.0/3.0/001 | OS | Physician | 05/24/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17027928&pc=QCB | 2023826-2023-02114 | 2023-04-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-05-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -11.0/1.0/106 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED UPSIDE DOWN INTO THE PATIENT'S RIGHT EYE (OD). ON (B)(6) 2023, THE LENS WAS INTRAOPERATIVELY REMOVED AND REPLACED FOR A SAME SIZE LENS. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "FLIPPED AND NEEDED TO REPLACED. HAPTIC DAMAGED DURING REMOVAL." THE PROBLEM WAS RESOLVED. STATUS OF THE EYE IS, "SUCCESSFUL SUGERY, LENS FLIPPED REPLACED." Manufacturer Narrative: CORRECTED DATA: H6: HEALTH EFFECT - IMPACT CODE: "4629" SHOULD BE REMOVED AND "2199" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "3189" SHOULD BE ADDED. B5: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/1.0/106 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED UPSIDE DOWN INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. REPORTEDLY "FLIPPED LENS DURING INSERTION, REPLACED. HAPTICS DAMAGED DURING REMOVAL". THE LEN WAS IMPLANTED, REMOVED AND REPLACED INTRAOPERATIVELY WITH THE SAME MODEL, LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "SUCCESSFUL SURGERY, LENS FLIPPED AND REPLACED". THE CAUSE OF THE EVENT IS UNKNOWN AND THE DEVICE FAILED TO PERFORM AS INTENDED. CAUSE DETAILS: "FLIPPED LENS DURING INSERTION, REPLACED. HAPTIC DAMAGED DURING REMOVAL." CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -11.0/1.0/106 | OD | Physician | 05/05/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17027931&pc=QCB | 2023826-2023-02092 | 2023-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.50/+1.0/083 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT. IN A SEPARATE SURGERY THAT DAY A SHORTER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -16.50/+1.0/083 | OS | Physician | 05/01/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17027932&pc=QCB | 2023826-2023-02093 | 2023-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+3.0/084 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. TTHE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT. THAT SAME DAY, THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS ALTHOUGH IT IS UNCLEAR WHETHER THIS WAS WITHIN THE SAME PROCEDURE. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+3.0/084 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2023 DUE TO EXCESSIVE VAULT. THIS PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- UPDATED TO REFLECT AN EXCHANGE. H6 HEALTH EFFECT - IMPACT CODE: CHANGED TO 4629. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -7.5/+3.0/084 | OD | Physician | 05/01/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17027934&pc=QCB | 2023826-2023-02096 | 2023-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+2.0/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -7.5/+2.0/095 | OD | Physician | 05/03/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17028495&pc=QCB | 2023826-2023-02145 | 2023-05-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -14.00/1.0/077 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: H6- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -14.00/1.0/077 | OD | Physician | 05/05/2023 | 2023 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17028508&pc=QCB | 2023826-2023-02146 | 2023-05-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-05-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -16.00/1.0/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TORN HAPTIC. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -16.00/1.0/095 | OS | Physician | 05/05/2023 | 2023 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17039572&pc=QCB | 2023826-2023-01886 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.00/1.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -17.00/1.5/087 | OS | Physician | 05/12/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17051837&pc=QCB | 2023826-2023-02125 | 2022-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -9.5/4.5/84 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME LENGTH, IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. REPORTEDLY, "THE REPLACEMENT LENS WAS IMPLANTED VERTICALLY.". Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: H6- INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. (B)(4). | -9.5/4.5/84 | OS | Physician | 05/11/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17051993&pc=QCB | 2023826-2023-02084 | 2023-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50/+1.5/175 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC BENT/DEFORMED WITH FIBRE ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | -07.50/+1.5/175 | OS | Physician | 05/08/2023 | 2023 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17052007&pc=QCB | 2023826-2023-02082 | 2022-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6)2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION: D9: RETURN DATE: 15-JUN-2023. H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE/DEBRIS ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.50/+1.0/088 | OD | Physician | 05/07/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17052043&pc=QCB | 2023826-2023-02088 | 2023-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+2.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6)2023. THE LENS WAS REMOVED ON (B)(6)2023 DUE TO THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE LENS IS STABLE WITH GOOD VAULT AND DEEP AC. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -18.00/+2.0/092 | OD | Physician | 05/09/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17052045&pc=QCB | 2023826-2023-02085 | 2023-01-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.50/+2.0/129 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -12.50/+2.0/129 | OS | Physician | 05/10/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17055437&pc=QCB | 2023826-2023-02091 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -04.50/+0.5/154 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS EYE, ON (B)(6) 2023. THE LENS WAS REMOVED DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -04.50/+0.5/154 | NA | Physician | 05/12/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17055529&pc=QCB | 2023826-2023-02090 | 2023-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50/+2.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2023 BUT THE PROBLEM WAS NOT RESOLVED, AND THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH A SAME LENGTH BUT DIFFERENT MODEL AND DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE LRI PERFORMED FOR ASTIGMATISM. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: \WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.50/+2.0/091 | OD | Physician | 05/11/2023 | 2023 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17056603&pc=QCB | 2023826-2023-02109 | 2023-01-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+2.5/109 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -12.5/+2.5/109 | OD | Physician | 05/08/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17056632&pc=QCB | 2023826-2023-02131 | 2023-01-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -9.50/2.5/099(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED AND IN A SUBSEQUENT SURGERY REPLACED DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: H6-HEALTH EFFECT - IMPACT CODE: "4628" SHOULD BE ADDED. H6-HEALTH EFFECT - IMPACT CODE: "4629" SHOULD BE REMOVED AND "4627" SHOULD BE ADDED. D 4.MODEL#: "VTICM5_13.2" SHOULD BE REMOVED AND "VTICM5_12.6" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+2.5/099 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED WITH LOW VAULT AND LENS REPOSITIONING. REPORTEDLY "THE LENS WAS OFF AXIS, AFTER 1ST REPOSITIONING IT RETURNED BACK TO THE FIRST AXIS (CCW 60°). THE SURGEON DECIDED TO CHANGE FROM TORIC LENS TO SPHERICAL LENS". ON (B)(6) 2023 THE LENS WAS EXPLANTED AND REPLACED WITH A SPHERICAL MODEL AND POWER, SAME LENGTH AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -9.50/2.5/099 | OS | Physician | 05/15/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17056647&pc=QCB | 2023826-2023-02110 | 2023-03-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+2.5/087 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.5/+2.5/087 | OS | Physician | 05/08/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17056962&pc=QCB | 2023826-2023-02136 | 2023-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Improper or Incorrect Procedure or Method; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -8.50/3.5/109 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6)2023, THE LENS WAS REMOVED AND IN A SUBSEQUENT SURGERY LATER THAT DAY, REPLACED DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS A USER ERROR. REPORTEDLY, "A ROTATION OF MORE THAN 70 DEGREES COUNTERCLOCKWISE DUE TO THE SMALL SIZE OF THE LENS DISPITE HAVING CALCULATED THE WTW TO ORBSCAN." STATUS OF THE EYE IS STATED, "ASINTOMATIC. BIOMICROCOPY LE: TRANSPARENT CORNEA WITHOUT EDEMA. GRADE III ANTERIOR CHAMBER. CENTRAL AND ROUND BLACK PUPIL. TYNDALL -/+-. ICL VAULT 0/100 IN 180 DEGREES LE. TRANSPARENT CRYSTALINE," AND "GOOD EVOLUTION LE." Manufacturer Narrative: H6: USER ERROR. CLAIM#: (B)(4). | -8.50/3.5/109 | OS | Physician | 05/16/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17057007&pc=QCB | 2023826-2023-02179 | 2023-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS -13.0/4.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS THAT WAS IMPLANTED IN OBLIQUE POSITION DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6: 1494- IT SHOULD BE NOTED THAT THE PATIENT'S AGE FELL OUTSIDE THE INDICATED AGE RANGE AT THE TIME OF IMPLANTATION. ADDITIONAL INFORMATION: D9: RETURN DATE: 15-SEP-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.0/4.5/093 | OD | Physician | 05/08/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17057045&pc=QCB | 2023826-2023-02134 | 2023-05-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -11.0/2.0/134 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LOW VAULT WITHOUT ROTATON WAS OBSERVED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -11.0/2.0/134 | OD | Physician | 05/16/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17057153&pc=QCB | 2023826-2023-02135 | 2023-05-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.0/1.5/81 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LOW VAULT WITHOUT ROTATON WAS OBSERVED. Manufacturer Narrative: A4-A6:UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -10.0/1.5/81 | OS | Physician | 05/16/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17057420&pc=QCB | 2023826-2023-02180 | 2023-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -11.0/1.5/167 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO LOW VAULT AND THE PROBLEM RESOLVED. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC TORN AND HAPTIC BROKEN. DIMENSIONAL INSPECTION WAS UNABLE TO BE PERFORMED DUE TO SIGNIFICANT LENS DAMAGE. CLAIM#(B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -11.0/1.5/167 | OS | Physician | 05/08/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17060987&pc=QCB | 2023826-2023-02033 | 2023-03-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -7.5/0.5/028 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, NO PLANS TO IMPLANT NEW LENS. Manufacturer Narrative: H3 - H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: H6 - DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -7.5/0.5/028 | OD | Physician | 05/16/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17061401&pc=QCB | 2023826-2023-02034 | 2023-01-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.0/1.00/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED VERTICALLY WITH A SAME LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6:WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC STRETCHED OUT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. CLAIM# (B)(4). | -10.0/1.00/086 | OD | Physician | 05/16/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17065129&pc=QCB | 2023826-2023-02036 | 2023-01-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED OBLIQUED-AXIS A 12.1MM VTICMO12.1, -8.5/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SAME LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM# (B)(4). | -8.5/1.5/090 | OS | Physician | 05/17/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17065338&pc=QCB | 2023826-2023-02217 | 2023-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS -15.00/1.5/100 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6)2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS THAT WAS IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: H6: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -15.00/1.5/100 | OD | Physician | 05/09/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17066110&pc=QCB | 2023826-2023-02220 | 2022-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS -11.00/3.5/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS THAT WAS IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -11.00/3.5/089 | OD | Physician | 05/10/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17066332&pc=QCB | 2023826-2023-02060 | 2023-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/3.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -9.5/3.5/093 | OD | Physician | 05/25/2023 | 2023 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17066377&pc=QCB | 2023826-2023-02219 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 12.6MM, VTICM5_12.6, -9.0/3.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE RIGHT EYE (OD). THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. THERE WAS NO PATIENT INJURY. ON (B)(6) 2023 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6: TYPE OF INVESTIGATION LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.0/3.0/094 | OD | Physician | 05/08/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17066560&pc=QCB | 2023826-2023-02035 | 2023-05-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED UPSIDE DOWN A 12.6MM VTICM5_12.6, -11.5/2.0/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE(OD) ON (B)(6) 2023. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED AND TORE OUTSIDE THE PATIENTS EYE AND BACK UP LENS WAS IMPLANTED INTRAOPERATIVLY ON (B)(6) 2023. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS DEVICE, UNFOLDED UPSIDE DOWN AND TORE OUTSIDE EYE. REPORTEDLY PATIENT REMAINS IN STUDY. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.5/2.0/110 | OD | Physician | 05/17/2023 | 2023 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17066581&pc=QCB | 2023826-2023-02193 | 2023-03-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/2.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -11.0/2.0/090 | OS | Physician | 05/29/2023 | 2023 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17066766&pc=QCB | 2023826-2023-02037 | 2022-12-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED VERTICALLY A 12.6MM VTICM5_12.6, -10.5/2.5/180 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT ON (B)(6) 2023. THIS DID NOT RESOLVE THE PROBLEM. THE LENS WAS REPLACED HORIZONTALLY WITH A SAME LENGTH SPHERICAL LENS ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -10.5/2.5/180 | OD | Physician | 05/19/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17066779&pc=QCB | 2023826-2023-02061 | 2023-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/3.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -12.5/3.0/093 | OD | Physician | 05/25/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17067070&pc=QCB | 2023826-2023-02218 | 2023-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS -16.00/2.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS THAT WAS IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIA DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -16.00/2.0/091 | OS | Physician | 05/09/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17067148&pc=QCB | 2023826-2023-02078 | 2023-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.0/2.5/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #: (B)(4). | -17.0/2.5/083 | OS | Physician | 05/26/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17076061&pc=QCB | 2023826-2023-02197 | 2023-05-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/1.0/112 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4) | -7.5/1.0/112 | OS | Physician | 05/30/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17076062&pc=QCB | 2023826-2023-02198 | 2023-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5 /6.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | -6.5 /6.0/092 | OS | Physician | 05/30/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17076063&pc=QCB | 2023826-2023-02200 | 2023-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/1.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -9.0/1.5/090 | OD | Physician | 05/30/2023 | 2023 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17076064&pc=QCB | 2023826-2023-02201 | 2023-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/0.5/105 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -7.5/0.5/105 | OS | Physician | 05/30/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17076065&pc=QCB | 2023826-2023-02196 | 2023-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/2.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -11.5/2.0/088 | OD | Physician | 05/27/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17076066&pc=QCB | 2023826-2023-02199 | 2023-05-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.5/5.0/077 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -15.5/5.0/077 | OD | Physician | 05/30/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17076069&pc=QCB | 2023826-2023-02202 | 2023-05-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_ 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.50/1.0/107 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -6.50/1.0/107 | OD | Physician | 05/31/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17076191&pc=QCB | 2023826-2023-02227 | 2023-02-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.0/1.5/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. ON (B)(6)2023, THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS, IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT. REPORTEDLY, "THE SECOND TICL IS PLACED VERTICALLY." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -9.0/1.5/091 | OD | Physician | 05/19/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17076402&pc=QCB | 2023826-2023-02261 | 2023-05-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.50/0.5/014, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.50/0.5/014 | OD | Physician | 05/11/2023 | 2023 | 45 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17076404&pc=QCB | 2023826-2023-02262 | 2022-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -6.50/1.0/062 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)96) 2022. LOW VAULT WITH ROTATION WAS OBSERVED. LENS REPOSITIONING WAS PERFORMED (ON (B)(6) 2022). ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -6.50/1.0/062 | OS | Physician | 05/11/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17083051&pc=QCB | 2023826-2023-02264 | 2023-03-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -13.5/-2.5/103 WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. IT WAS REPORTED THAT THE PROBLEM DID NOT RESOLVE. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-MEDICAL DEVICE PROBLEM CODE: "2923" SHOULD BE ADDED. B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS -13.5/+2.5/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED A LOW VAULT WITH ROTATION. ON (B)(6) 2023 THE LENS EXCHANGED WITH THE SAME MODEL AND SIMILAR POWER BUT LONGER LENGTH, BUT THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE: "CORRECT". THE CAUSE OF THE EVENT IS UNKNOWN. DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -13.5/+2.5/103 | OS | Physician | 05/12/2023 | 2023 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17083052&pc=QCB | 2023826-2023-02265 | 2023-03-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -12.0/1.50/114 WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO LOW VAULT. IT WAS REPORTED THAT THIS EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- MEDICAL DEVICE PROBLEM CODE: "2923" SHOULD BE ADDED. ADDITONAL DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -12.0/+1.5/114 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED A LOW VAULT WITH ROTATION. ON (B)(6) 2023 THE LENS EXPLANTED, AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "CORRECT". THE CAUSE OF THE EVENT IS UNKNOWN. CAUSE DETAILS PROVIDED: "WE CHOSE THE BIGGER DIAMETER, BUT THE VAULT CONTINUED LOW." D6B. (B)(6) 2023 SHOULD BE REMOVED AND (B)(6) 2023 SHOULD BE ADDED. DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC AND HAPTIC DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -12.0/1.50/114 | OD | Physician | 05/12/2023 | 2023 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17083202&pc=QCB | 2023826-2023-02266 | 2023-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -13.5/2.0/092 WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 AS A REPLACEMENT LENS TO ADDRESS LOW VAULT BUT THE VAULT DID NOT RESOLVE. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. SEE MFR # 2023826-2023-02264 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: CORRECTED DATA: H6-MEDICAL DEVICE PROBLEM CODE: "2923" SHOULD BE ADDED. B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -13.5/+2.0/92 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 AS A REPLACEMENT LENS. THE PATIENT EXPERIENCED A LOW VAULT WITH ROTATION. ON (B)(6) 2023 THE LENS EXPLANTED AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "CORRECT". THE CAUSE OF THE EVENT IS UNKNOWN. DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC/HAPTIC TORN WITH FOREIGN MATERIAL ON LENS (UNABLE TO IDENTIFY). DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.5/2.0/092 | OS | Physician | 05/12/2023 | 2023 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17083433&pc=QCB | 2023826-2023-02275 | 2023-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.0/1.5/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON REPORTS LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH LENS BUT DIFFERENT AXIS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE:18-JUL-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -11.0/1.5/095 | OD | Physician | 06/03/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17083874&pc=QCB | 2023826-2023-02183 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON WAS LOADING A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.50/+1.0/079 (SPHERE/CYLINDER/AXIS), AND THE LENS TORE. UNKNOWN IF THERE WAS ANY PATIENT CONTACT. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON NOTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.50/1.0/079 (SPHERE/CYLINDER/AXIS) TORE DURING LOADING. IT IS UNKNOWN IF THERE WAS PATIENT CONTACT. A BACK UP LENS WAS ATTEMPTED TO BE IMPLANTED INTO THE LEFT EYE (OS) BUT THAT ALSO TORE. REPORTEDLY "ON THE SECOND EYE FROM THE SAME LOT I USED 2 AND ALL TORN". THE PROBLEM WAS NOT RESOLVED. CAUSE OF EVENT IS UNKNOWN: "I AM NOT SURE. I NEVER HAD THIS HAPPEN BEFORE AND I HAVE USED THE LIOLI INJECTOR IN THE PAST. I AM WORKING WITH (B)(6) TO TRY AND FIGURE THIS OUT. TRYING TO SEE IF I NEED TO CHANGE TECHNIQUE." H6-HEALTH EFFECT- IMPACT CODE: "4614" SHOULD BE ADDED. H6- MEDICAL DEVICE CODE: "3189" SHOULD BE REMOVED AND "2993" SHOULD BE ADDED. CLAIM# (B)(4). | -06.50/+1.0/079 | OS | Physician | 05/18/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17084049&pc=QCB | 2023826-2023-02280 | 2023-03-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.00/+2.0/071 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -11.00/+2.0/071 | OS | Physician | 05/10/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17084053&pc=QCB | 2023826-2023-02284 | 2023-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.5/111 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+2.5/111 | OS | Physician | 05/11/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17084054&pc=QCB | 2023826-2023-02285 | 2023-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.5/180 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/+1.5/180 | OD | Physician | 05/11/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17084055&pc=QCB | 2023826-2023-02286 | 2023-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+2.5/088 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.5/+2.5/088 | OD | Physician | 05/12/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17084056&pc=QCB | 2023826-2023-02287 | 2023-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+1.5/082 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.0/+1.5/082 | OD | Physician | 05/12/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17084176&pc=QCB | 2023826-2023-02231 | 2022-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -10.0/1.5/110 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED DUE TO LOW VAULT. REPORTEDLY, "AFTER THE PATIENT'S LENS WAS REMOVED, THERE WAS NO FURTHER IMPLANTATION OF THE LENS. WAIT FOR THE PATIENT'S CONDITION TO STABILIZE BEFORE PERFORMING THE LENS IMPLANTATION." THE PROBLEM WAS NOT RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL: ACD>3.0. CLAIM# (B)(4). | -10.0/1.5/110 | OS | Physician | 05/22/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17084377&pc=QCB | 2023826-2023-02267 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. LENS ROTATION WAS REPORTED. THE LENS WAS REPOSITIONED BUT ROTATED AGAIN. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: D4 MODEL#: VTCIM5_12.6. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.5/2.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED LENS ROTATION AND LENS REPOSITIONING. REPORTEDLY "THE PATIENT THAT WE SENT YOU INFORMATION PREVIOUSLY UNDERWENT ROTATION OF THE ICL. DAY ONE AFTER ROTATION, UNCORRECTED ACUITY WAS 20/25 OS AND MANIFEST REVEALED REFRACTION OF -0.50+0.75X85. ONE WEEK FOLLOWING THE PROCEDURE, IT ROTATED AGAIN HIS REFRACTION IS -15.+1.75X110 TO YIELD 20/20- VISION. THE TORIC ICL APPEARS TO BE AT THE 85 DEGREE AXIS. CYCLOPLEGIC REFRACTION REVIEWED -1.13+2.75X95." THE LENS REMAINS IMPLANTED. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -4.5/2.0/094 | OS | Physician | 05/11/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17089798&pc=QCB | 2023826-2023-02291 | 2023-04-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -8/1.0/113 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR, AND NOTED THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -8/1.0/113 | OS | Physician | 05/12/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17089802&pc=QCB | 2023826-2023-02292 | 2023-04-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.0/2.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS HAD TORE DURING IMPLANTATION. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH/MODEL LENS AND THE PROBLEM WAS RESOLVED. THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -13.0/2.5/090 | OS | Physician | 06/05/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17089814&pc=QCB | 2023826-2023-02293 | 2023-01-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.0/1.5/115 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. INTRAOPERATIVELY THE LENS WAS REMOVED DUE TO THE LENS HAD TORN DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. ON (B)(6) 2023 A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6. TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -13.0/1.5/115 | OS | Physician | 06/06/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17091340&pc=QCB | 2023826-2023-02294 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6; -8.50/2.0/080 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS EXPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6; -8.50/2.0/080 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR. H6: 1494: OFF-LABEL USAGE: ACD WAS LESS THAN 3.0MM AT THE TIME OF IMPLANTATION. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -8.50/2.0/080 | NA | Physician | 05/12/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17093605&pc=QCB | 2023826-2023-02237 | 2023-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.5/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SAME SIZE LENS, IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT. REPORTEDLY, "ADJUST HORIZONTAL POSITION TO VERTICAL POSITION." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN WITH FOREIGN MATERIAL ON THE LENS SURFACE, SPECIFICALLY FIBERS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -11.5/2.0/089 | OS | Physician | 05/24/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17093867&pc=QCB | 2023826-2023-02211 | 2023-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/3.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -13.0/3.5/090 | OD | Physician | 06/02/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17093868&pc=QCB | 2023826-2023-02212 | 2023-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/2.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -9.0/2.0/091 | OS | Physician | 06/02/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17100629&pc=QCB | 2023826-2023-02190 | 2022-11-25 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+0.5/090 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS WAS IMPLANTED UPSIDE DOWN IN THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK . H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.00/+0.5/090 | OD | Physician | 05/16/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17101309&pc=QCB | 2023826-2023-02184 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON WAS LOADING A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.50/+1.0/090 (SPHERE/CYLINDER/AXIS), AND THE LENS TORE. UNKNOWN IF THERE WAS ANY PATIENT CONTACT. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5-THE REPORTER INDICATED THAT THE SURGEON NOTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.50/1.0/090(SPHERE/CYLINDER/AXIS) TORE DURING LOADING. IT IS UNKNOWN IF THERE WAS PATIENT CONTACT. A BACK UP LENS WAS ATTEMPTED TO BE IMPLANTED INTO THE LEFT EYE (OS) BUT THAT ALSO TORE. REPORTEDLY "ON THE SECOND EYE FROM THE SAME LOT I USED 2 AND ALL TORN". THE PROBLEM WAS NOT RESOLVED. CAUSE OF EVENT IS UNKNOWN: "I AM NOT SURE. I NEVER HAD THIS HAPPEN BEFORE AND I HAVE USED THE LIOLI INJECTOR IN THE PAST. I AM WORKING WITH JOSH GOLDING TO TRY AND FIGURE THIS OUT. TRYING TO SEE IF I NEED TO CHANGE TECHNIQUE." H6-HEALTH EFFECT- IMPACT CODE: "4614" SHOULD BE ADDED. H6- MEDICAL DEVICE CODE: "3189" SHOULD BE REMOVED AND "2993" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: A2: UNK. A3: UNK. A4: UNK. A5: UNK. A6: UNK. B3: UNK. D6A: UNK. D6B: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # 433009 | -06.50/+1.0/090 | OS | Physician | 05/18/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17102193&pc=QCB | 2023826-2023-02214 | 2023-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/1.0/173 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -5.5/1.0/173 | OD | Physician | 06/02/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17102195&pc=QCB | 2023826-2023-02213 | 2023-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/1.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -11.5/1.5/089 | OS | Physician | 06/02/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17102196&pc=QCB | 2023826-2023-02215 | 2023-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/1.0/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -6.0/1.0/085 | OS | Physician | 06/02/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17102197&pc=QCB | 2023826-2023-02216 | 2023-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/1.0/035 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -6.0/1.0/035 | OD | Physician | 06/02/2023 | 2023 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17102198&pc=QCB | 2023826-2023-02240 | 2023-05-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/2.0/075 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -14.5/2.0/075 | OS | Physician | 06/05/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17102364&pc=QCB | 2023826-2023-02185 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-06-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON WAS LOADING A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -07.50/+1.0/087 (SPHERE/CYLINDER/AXIS), AND THE LENS TORE. THERE WAS PATIENT CONTACT AND THE BACKUP LENS WAS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -07.50/+1.0/087 | NA | Physician | 05/18/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17105577&pc=QCB | 2023826-2023-02271 | 2023-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -6.5/5.5/104 WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4). Manufacturer Narrative: B5: PER UPDATED INFORMATION THE LENS WAS EXCHANGED. H6- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: ON (B)(6) 2023, THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL, SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. D6B: "ON (B)(6) 2023" SHOULD BE ADDED. H6: MEDICAL DEVICE PROBLEM CODE:"2993" SHOULD BE ADDED. H6: HEALTH EFFECT- IMPACT CODE: "2199" SHOULD BE REMOVED. CLAIM#: (B)(4). | -6.5/5.5/104 | OD | Physician | 05/15/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17106440&pc=QCB | 2023826-2023-01957 | 2022-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/2.0/136 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -9.5/2.0/136 | OD | Physician | 05/18/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17106591&pc=QCB | 2023826-2023-02059 | 2023-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/3.0/178 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -10.0/3.0/178 | OD | Physician | 05/24/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17108660&pc=QCB | 2023826-2023-02274 | 2023-05-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00/+2.00/112 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6)2023. THE LENS HAD LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-01170. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: H1: SHOULD BE CORRECTED TO SERIOUS INJURY. CLAIM#(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00/+2.00/112 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS HAD LOW VAULTING. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS UNKNOWN. D9: RETURN DATE: 05-SEP-2023. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -18.00/+2.00/112 | OD | Physician | 05/31/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17108699&pc=QCB | 2023826-2023-02191 | 2023-01-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.00/+1.0/113 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT POWER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.00/+1.0/113 | OS | Physician | 05/16/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17108708&pc=QCB | 2023826-2023-02189 | 2022-01-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.50/+3.5/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2022 BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -15.50/+3.5/087 | OS | Physician | 05/16/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17112245&pc=QCB | 2023826-2023-02248 | 2023-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.5/099 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -10.0/1.5/099 | OD | Physician | 06/05/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17112249&pc=QCB | 2023826-2023-02249 | 2023-04-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.0/2.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -17.0/2.0/089 | OD | Physician | 06/05/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17112252&pc=QCB | 2023826-2023-02251 | 2022-12-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/3.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -12.5/3.0/089 | OS | Physician | 06/06/2023 | 2023 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17112258&pc=QCB | 2023826-2023-02247 | 2023-03-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/2.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -11.0/2.0/089 | OD | Physician | 06/05/2023 | 2023 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17112272&pc=QCB | 2023826-2023-02243 | 2023-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/2.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -13.0/2.5/092 | OD | Physician | 06/05/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17112273&pc=QCB | 2023826-2023-02245 | 2023-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.0/4.0/177 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ON (B)(6) 2023 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6- DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -4.0/4.0/177 | OS | Physician | 06/05/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17112275&pc=QCB | 2023826-2023-02250 | 2022-08-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -6.0/2.5/091 | OD | Physician | 06/06/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17114589&pc=QCB | 2023826-2023-02332 | 2022-12-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.0/113 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.0/+2.0/113 | OD | Physician | 05/16/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17114592&pc=QCB | 2023826-2023-02335 | 2023-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+1.5/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -12.5/+1.5/092 | OS | Physician | 05/16/2023 | 2023 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17114593&pc=QCB | 2023826-2023-02334 | 2023-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+1.5/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.5/+1.5/090 | OD | Physician | 05/16/2023 | 2023 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17119686&pc=QCB | 2023826-2023-02158 | 2022-12-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -14.0/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. LENS REPOSITIONED ON (B)(6) 2023, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -14.0/1.0/090 | OS | Physician | 05/23/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17119688&pc=QCB | 2023826-2023-02157 | 2022-12-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.00/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. LENS REPOSITIONED ON (B)(6) 2023, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.00/1.0/091 | OD | Physician | 05/23/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17119693&pc=QCB | 2023826-2023-02160 | 2020-11-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.5/1.0/006 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS DEVICE. LENS TOO LONG. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.5/1.0/006 | OS | Physician | 05/23/2023 | 2023 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17119694&pc=QCB | 2023826-2023-02159 | 2020-11-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.5/0.5/024 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON 16/NOV/2020. EXCESSIVE VAULT WAS OBSERVED. PROBLEM NOT RESOLVED. CAUSE OF THE EVENT IS REPORTED AS DEVICE. LENS TOO LONG. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.5/0.5/024 | OD | Physician | 05/23/2023 | 2023 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17122143&pc=QCB | 2023826-2023-02352 | 2022-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/2.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -13.0/2.0/090 | OD | Physician | 06/09/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17122145&pc=QCB | 2023826-2023-02347 | 2023-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/3.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -9.5/3.0/095 | OD | Physician | 06/09/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17122146&pc=QCB | 2023826-2023-02351 | 2023-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/0.5/148 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "THE FIRST IMPLANT IN THIS PATIENT WAS A HORIZONTAL IMPLANT." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -7.5/0.5/148 | OS | Physician | 06/09/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17122147&pc=QCB | 2023826-2023-02350 | 2023-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/1.0/111 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "THE FIRST IMPLANT IN THIS PATIENT WAS A 90-DEGREE IMPLANT." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -7.5/1.0/111 | OD | Physician | 06/09/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17122148&pc=QCB | 2023826-2023-02344 | 2023-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/3.0/108 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -13.5/3.0/108 | OD | Physician | 06/08/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17122149&pc=QCB | 2023826-2023-02349 | 2023-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/0.5/170 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH/DIFFERENT MODEL LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "ASTIGMATIC CRYSTALS WERE IMPLANTED FOR THE FIRST TIME. THE PATIENT'S SUPPLY WAS TOO HIGH, AND THE DOCTOR FINALLY ADJUSTED 12.6 IMPLANTATION OF NON-ASTIGMATIC CRYSTALS BECAUSE SUITABLE ASTIGMATIC CRYSTALS COULD NOT BE MATCHED." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -6.0/0.5/170 | OS | Physician | 06/09/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17122152&pc=QCB | 2023826-2023-02346 | 2022-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/1.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.0/1.5/092 | OS | Physician | 06/08/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17122154&pc=QCB | 2023826-2023-02343 | 2023-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/3.5/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -11.0/3.5/085 | OS | Physician | 06/08/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17122155&pc=QCB | 2023826-2023-02255 | 2023-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/1.5/114 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -8.0/1.5/114 | OS | Physician | 06/06/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17131362&pc=QCB | 2023826-2023-02395 | 2023-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -13.0/2.0/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE AND NOTED IT HAD ROTATED DAYS AFTER. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS AND THIS RESOLVED THE PROBLEM. THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: : (B)(4). | -13.0/2.0/094 | OD | Physician | 06/09/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17131635&pc=QCB | 2023826-2023-02361 | 2023-05-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/1.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # : (B)(4). | -6.0/1.5/095 | OD | Physician | 06/12/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17131636&pc=QCB | 2023826-2023-02358 | 2023-04-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/2.0/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #: (B)(4). | -9.5/2.0/092 | OD | Physician | 06/09/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17131637&pc=QCB | 2023826-2023-02359 | 2023-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/1.5/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # : (B)(4). | -9.5/1.5/073 | OS | Physician | 06/09/2023 | 2023 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17131639&pc=QCB | 2023826-2023-02360 | 2023-05-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/1.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -11.5/1.5/092 | OD | Physician | 06/12/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17131643&pc=QCB | 2023826-2023-02354 | 2023-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.5/2.5/100 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # : (B)(4). | -17.5/2.5/100 | OD | Physician | 06/09/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17131645&pc=QCB | 2023826-2023-02355 | 2023-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.0/2.5/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # : (B)(4). | -17.0/2.5/073 | OS | Physician | 06/09/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17131867&pc=QCB | 2023826-2023-02368 | 2023-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+1.0/107 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# : (B)(4). | -12.5/+1.0/107 | OD | Physician | 05/18/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17131959&pc=QCB | 2023826-2023-02370 | 2023-01-06 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.50/+2.5/079 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED DUE TO LOW VAULT. THAT SAME DAY A LONGER LENGTH LENS WAS IMPLANTED BUT IT IS UNKNOWN IF THIS WAS DONE WITHIN THE INITIAL SURGERY. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+2.5/079 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: B5- PER AN EMAIL IT WAS CONFIRMED THAT THE INITAL LENS WAS IMPLANTED, REMOVED, AND REPLACED WITHIN THE INITIAL SURGERY. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -7.50/+2.5/079 | OS | Physician | 05/18/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17131960&pc=QCB | 2023826-2023-02369 | 2023-01-06 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.5/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED DUE TO LOW VAULT. THAT SAME DAY A LONGER LENGTH LENS WAS IMPLANTED BUT IT IS UNKNOWN IF THIS WAS DONE WITHIN THE INITIAL SURGERY. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.5/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) JAN 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: B5- PER AN EMAIL IT WAS CONFIRMED THAT THE INITAL LENS WAS IMPLANTED, REMOVED, AND REPLACED WITHIN THE INITIAL SURGERY. CLAIM# (B)(4). | -11.0/+1.5/093 | OD | Physician | 05/18/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17132057&pc=QCB | 2023826-2023-02390 | 2023-01-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS -11.5/1.5/166 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN VERTICAL POSITION INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS LATER EXCHANGED WITH A SAME LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6- DEVICE EVALUATION: LENS WAS RETURNED DRY IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -11.5/1.5/166 | OS | Physician | 05/17/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17132127&pc=QCB | 2023826-2023-02391 | 2022-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -9.0/4.5/007 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN VERTICAL POSITION INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS AND THE PROBLEM RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# : (B)(4). | -9.0/4.5/007 | OS | Physician | 05/18/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17138047&pc=QCB | 2023826-2023-02311 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/024 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS EYE ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023. THE LENS WAS REPLACED WITH A LONGER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK A5: UNK. A6: UNK. B3: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.00/+1.5/024 | NA | Other Health Care Professional | 05/24/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17138308&pc=QCB | 2023826-2023-02314 | 2022-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50/+6.0/175 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTION (FUNCTIONAL) VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: A4: UNK; A5: UNK; A6: UNK; H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.50/+6.0/175 | OS | Physician | 05/24/2023 | 2023 | 47 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17140142&pc=QCB | 2023826-2023-02372 | 2023-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+5.0/104 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE LENS WAS IMPLANTED VERTICALLY AND THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+5.0/104 | OD | Physician | 05/19/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17140143&pc=QCB | 2023826-2023-02371 | 2023-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+1.0/007 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE LENS WAS IMPLANTED VERTICALLY AND THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -7.5/+1.0/007 | OD | Physician | 05/19/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17140144&pc=QCB | 2023826-2023-02373 | 2022-12-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.5/+0.5/130 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -4.5/+0.5/130 | OD | Physician | 05/19/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17140146&pc=QCB | 2023826-2023-02374 | 2022-12-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.0/+0.5/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH SPHERICAL LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -5.0/+0.5/089 | OS | Physician | 05/19/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17140149&pc=QCB | 2023826-2023-02378 | 2023-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/+3.5/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | -14.5/+3.5/092 | OD | Physician | 05/19/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17140150&pc=QCB | 2023826-2023-02376 | 2023-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+2.5/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -12.0/+2.5/093 | OS | Physician | 05/19/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17140152&pc=QCB | 2023826-2023-02379 | 2023-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/+0.5/175 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -5.5/+0.5/175 | OS | Physician | 05/19/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17140262&pc=QCB | 2023826-2023-02404 | 2022-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6; -9.50/+1.50/049 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2022 THE LENS WAS REPOSITIONED BUT IT DID NOT RESOLVE THE PROBLEM AS THE LENS HAD ROTATED VERTICALLY. LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE: KERATOCONUS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6; -9.50/+1.50/049 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2022 THE LENS WAS REPOSITIONED BUT IT DID NOT RESOLVE THE PROBLEM AS THE LENS HAD ROTATED VERTICALLY. REPORTEDLY, THE LENS WAS EXPLANTED ON AN UNKNOWN DATE. PATIENT WAS DOING WELL BUT COMPLAINED ABOUT DECREASED VA. HOWEVER, "THE DOCTOR THINKS THAT THIS IS A SURFACE ISSUE." CLAIM# (B)(4). | -9.50/+1.50/049 | OS | Physician | 05/22/2023 | 2023 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17140469&pc=QCB | 2023826-2023-02383 | 2022-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.0/+1.5/091 | OD | Physician | 05/19/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17140470&pc=QCB | 2023826-2023-02382 | 2022-12-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.0/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -11.0/+2.0/089 | OS | Physician | 05/19/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17140471&pc=QCB | 2023826-2023-02381 | 2023-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+1.5/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.0/+1.5/092 | OD | Physician | 05/19/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17140691&pc=QCB | 2023826-2023-02375 | 2023-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.5/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.5/+1.5/092 | OD | Physician | 05/19/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17148655&pc=QCB | 2023826-2023-02407 | 2023-04-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO_12.1; -10.5/2.0/068 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND LATER ON THIS SAME DATE REPLACED WITH A SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THERE WAS NO PATIENT INJURY. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -10.5/2.0/068 | OS | Physician | 06/12/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17148767&pc=QCB | 2023826-2023-02402 | 2023-05-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/2.5/109 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -14.0/2.5/109 | OD | Physician | 06/14/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17148771&pc=QCB | 2023826-2023-02363 | 2023-05-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.0/0.5/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -5.0/0.5/103 | OD | Physician | 06/12/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17148772&pc=QCB | 2023826-2023-02398 | 2023-03-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/1.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -12.5/1.5/094 | OS | Physician | 06/13/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17148773&pc=QCB | 2023826-2023-02397 | 2023-03-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/3.0/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -13.5/3.0/082 | OS | Physician | 06/13/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17148774&pc=QCB | 2023826-2023-02399 | 2023-03-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/1.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "THE DOCTOR CHOOSES VERTICAL IMPLANTATION FOR THE SECOND IMPLANTED LENS." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #: (B)(4). | -9.5/1.5/093 | OS | Physician | 06/12/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17148775&pc=QCB | 2023826-2023-02396 | 2023-03-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/3.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -13.5/3.0/094 | OD | Physician | 06/13/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17151989&pc=QCB | 2023826-2023-02324 | 2023-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3: TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -09.50/+1.0/094 | OD | Physician | 05/21/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17152016&pc=QCB | 2023826-2023-02322 | 2023-03-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+1.0/128 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+1.0/128 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH A SHORTER LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS THROUGHOUT THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH RESIDUE/DEBRIS THROUGHOUT THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -09.00/+1.0/128 | OS | Physician | 05/21/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17159825&pc=QCB | 2023826-2023-02409 | 2023-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+2.5/096 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -7.0/+2.5/096 | OS | Physician | 05/21/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17159826&pc=QCB | 2023826-2023-02408 | 2023-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+2.0/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.0/+2.0/094 | OS | Physician | 05/21/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17159827&pc=QCB | 2023826-2023-02440 | 2023-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.0/1.5/091 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. IT WAS REPORTED THAT THIS EXCHANGE DID NOT RESOLVE THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). Manufacturer Narrative: H6: TYPE OF INVESTIGATION CODE 4110 SHOULD BE REMOVED FROM ADDITIONAL MDR AS IT IS NOT APPLICABLE H6: INVESTIGATION FINDINGS CODE 213 TO 114 PREVIOUS CODE NOT APPLICABLE. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN/BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -11.0/1.5/091 | OD | Physician | 05/23/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17159934&pc=QCB | 2023826-2023-02413 | 2023-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+2.5/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -12.5/+2.5/092 | OD | Physician | 05/21/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17159991&pc=QCB | 2023826-2023-02417 | 2023-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/+1.0/088 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -5.5/+1.0/088 | OS | Physician | 05/21/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17160071&pc=QCB | 2023826-2023-02447 | 2023-01-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+2.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | -8.0/+2.0/091 | OD | Physician | 05/23/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17160072&pc=QCB | 2023826-2023-02423 | 2022-09-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.5/074 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.5/+1.5/074 | OS | Physician | 05/22/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17160076&pc=QCB | 2023826-2023-02420 | 2023-04-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.0/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# : (B)(4). | -7.0/+1.0/092 | OS | Physician | 05/22/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17160357&pc=QCB | 2023826-2023-02449 | 2023-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+3.0/075 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: (B)(4). | -13.0/+3.0/075 | OS | Physician | 05/24/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17160370&pc=QCB | 2023826-2023-02459 | 2023-03-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.0/110 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.5/+1.0/110 | OD | Physician | 05/29/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17160837&pc=QCB | 2023826-2023-02450 | 2023-03-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.50/+0.5/071 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD 3.0MM). CLAIM# : (B)(4). | -12.50/+0.5/071 | OS | Physician | 05/25/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17160838&pc=QCB | 2023826-2023-02453 | 2023-04-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+2.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D4 MODEL #: CORRECTED TO VTICMO12.1 CLAIM# (B)(4). | -12.5/+2.0/091 | OD | Physician | 05/25/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17160841&pc=QCB | 2023826-2023-02454 | 2023-04-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+4.0/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -10.5/+4.0/091 | OS | Physician | 05/25/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17163477&pc=QCB | 2023826-2023-02299 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM512.1, -6.0/5.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS OTHER, DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CLAIM# (B)(4). | -6.0/5.0/089 | OS | Physician | 05/30/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17163481&pc=QCB | 2023826-2023-02298 | 2023-05-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 13.2MM VTICM5 13.2, -6.0/5.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE BEFORE/DURING LOADING. DAMAGE WAS NOTED DURING LOADING. NO PATIENT CONTACT. BACK UP LENS IMPLANTED ON (B)(6) 2023. PROBLEM RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -6.0/5.0/089 | NA | Physician | 05/30/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17163482&pc=QCB | 2023826-2023-02300 | 2001-01-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -14.5/3.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LOW VAULT WITH ROTATION WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2023, BUT DID NOT RESOLVE THE PROBLEM. LENS REMAIN IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS OTHER, DEVICE DID NOT FAIL TO PERFORM AS INTENDED, SMALL SIZE LENS WITH ROTATION AND CHANGING ASTIGMATISM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -14.5/3.0/086 | OD | Physician | 05/30/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17163484&pc=QCB | 2023826-2023-02306 | 2023-05-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED VERTICALLY A 13.2MM VTICMO13.2, -10.5/3.0/003 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED HORIZONTALLY WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | -10.5/3.0/003 | OS | Physician | 06/02/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17163485&pc=QCB | 2023826-2023-02303 | 2023-06-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -13.0/2.5/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023.THE LENS TORE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY ON (B)(6) 2020. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, DEVICE FAILED TO PERFORM AS INTENDED. REPORTEDLY, TEAR SPLITTING THE CENTRAL OPTIC WHICH PRGINATED FROM THE LEADING EDGE. PATIENT DOING EXCELLENT WITH NO COMPLICATIONS. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -13.0/2.5/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023.THE LENS TORE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY ON (B)(6) 2023. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, DEVICE FAILED TO PERFORM AS INTENDED. REPORTEDLY, TEAR SPLITTING THE CENTRAL OPTIC WHICH ORIGINATED FROM THE LEADING EDGE. PATIENT DOING EXCELLENT WITH NO COMPLICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.0/2.5/110 | OD | Physician | 06/01/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17163490&pc=QCB | 2023826-2023-02301 | 2001-01-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5 12.6, -14.5/3.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LOW VAULT WITH ROTATION WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2023, BUT DID NOT RESOLVE THE PROBLEM. LENS REMAIN IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS OTHER, DEVICE DID NOT FAIL TO PERFORM AS INTENDED, SMALL SIZE LENS WITH ROTATION AND CHANGING ASTIGMATISM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -14.5/3.5/088 | OS | Physician | 05/30/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17163496&pc=QCB | 2023826-2023-02307 | 2023-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED VERTICALLY A 13.7MM VTICMO13.7, -10.5/2.5/038 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED HORIZONTALLY WITH A SHORTER LENGTH LENS DUE TO LOW VAULT WITH ROTATION ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED VERTICALLY A 13.7MM VTICMO13.7, -12.5/2.5/038 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED HORIZONTALLY WITH A SAME SIZE LENS DUE TO LOW VAULT WITH ROTATION ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -10.5/2.5/038 | OD | Physician | 06/02/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17163499&pc=QCB | 2023826-2023-02308 | 2023-05-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO 13.2, -10.0/3.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023.THE LENS TORE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS EXPLANTED ON (B)(6) 2023. A REPLACEMENT LENS WAS IMPLANTED IN A SECOND SURGERY ON (B)(6) 2023. PROBLEM RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6- MEDICAL DEVICE CODE: 1494- OFF LABEL USE (AGE<21). H6-HEALTH EFFECT- IMPACT CODE: "4629" SHOULD BE REMOVED. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.0/3.0/094 | OS | Physician | 06/02/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17164497&pc=QCB | 2023826-2023-02412 | 2023-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.0/073 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -10.5/+2.0/073 | OD | Physician | 05/21/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17164593&pc=QCB | 2023826-2023-02441 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.0/1.5/106 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023 AS A REPLACEMENT LENS TO ADDRESS LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. THE REPORTER STATES THAT THE ARCH HEIGHT IS STILL VERY LOW. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -11.0/1.5/106 | OD | Physician | 05/23/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17167159&pc=QCB | 2023826-2023-02442 | 2023-04-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -9.5/2.0/176 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN A VERTICAL POSITION INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS THAT WAS IMPLANTED IN A HORIZONTAL POSITION DUE TO LOW VAULT AND THE PROBLEM RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -9.5/2.0/176 | OD | Physician | 05/23/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17167505&pc=QCB | 2023826-2023-02461 | 2023-05-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -12.0/4.0/071 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. IT WAS REPORTED THAT THIS LENS WAS INCORRECTLY IMPLANTED INTO THE PATIENT'S RIGHT EYE, AS IT WAS INTENDED FOR THE PATIENT'S LEFT EYE. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND ON THIS SAME DATE REPLACED WITH A SAME MODEL/LENGTH LENS BUT THIS DID NOT RESOLVE THE PROBLEM. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.0/4.0/071 | OD | Physician | 06/08/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17174869&pc=QCB | 2023826-2023-02520 | 2023-06-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -11.00/1.5/106 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS HAD TORE DURING INSERTION. IT WAS INTRAOPERATIVELY REMOVED AND REPLACED WITH A BACK UP LENS. THIS RESOLVED THE PROBLEM. THERE WAS NO PATIENT INJURY. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.00/1.5/106 | OD | Physician | 06/06/2023 | 2023 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17176048&pc=QCB | 2023826-2023-02478 | 2023-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_ 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/4.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -13.0/4.5/095 | OS | Physician | 06/15/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17176049&pc=QCB | 2023826-2023-02484 | 2023-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/5.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -7.0/5.0/089 | OS | Physician | 06/16/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17176050&pc=QCB | 2023826-2023-02483 | 2023-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/2.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -10.5/2.5/095 | OS | Physician | 06/16/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17176051&pc=QCB | 2023826-2023-02487 | 2023-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/1.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "THE INITIAL LENS WAS IMPLANTED IN HORIZONTAL POSITION, REPLACEMENT LENS WAS IMPLANTED IN VERTICAL POSITION." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -7.5/1.5/086 | OS | Physician | 06/19/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17176052&pc=QCB | 2023826-2023-02486 | 2023-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.5/2.0/073 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXPLANTED DUE TO LOW VAULT WITH ROTATION. ON (B)(6)2023 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -15.5/2.0/073 | OS | Physician | 06/19/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17176053&pc=QCB | 2023826-2023-02482 | 2023-04-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/2.0/087(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -11.5/2.0/087 | OD | Physician | 06/16/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17176057&pc=QCB | 2023826-2023-02480 | 2023-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/1.5/179 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -5.5/1.5/179 | OS | Physician | 06/16/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17176058&pc=QCB | 2023826-2023-02475 | 2023-02-15 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/2.0/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -10.0/2.0/072 | OD | Physician | 06/15/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17176060&pc=QCB | 2023826-2023-02479 | 2023-01-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_ 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/2.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -6.5/2.0/089 | OS | Physician | 06/15/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17176061&pc=QCB | 2023826-2023-02481 | 2023-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -10.0/1.0/093 | OS | Physician | 06/16/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17176067&pc=QCB | 2023826-2023-02473 | 2023-03-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/1.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -12.0/1.5/074 | OD | Physician | 06/15/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17177099&pc=QCB | 2023826-2023-02443 | 2023-02-21 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023, A 13.2MM, VTICMO13.2, -12.5/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE LEFT EYE (OS). THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY WITH NO PATIENT INJURY. ON (B)(6) 2023, A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: A1:(B)(6). H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.5/2.5/091 | OS | Physician | 05/23/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17185672&pc=QCB | 2023826-2023-02523 | 2023-05-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -15+1.0/13 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND IT WAS INDICATED THAT A REPLACEMENT IMPLANT WILL OCCUR IN (B)(6) 2023. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - A SHORTER REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2023 AND PROBLEM WAS RESOLVED. REPORTEDLY "ALL GOOD. PATIENT IS HAPPY." CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -15+1.0/13 | OS | Physician | 06/21/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17185698&pc=QCB | 2023826-2023-02522 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -15+1.0/13 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND IT WAS INDICATED THAT A REPLACEMENT IMPLANT WILL OCCUR IN (B)(6) 2023. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - A SHORTER REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2023 AND PROBLEM WAS RESOLVED. REPORTEDLY "ALL GOOD. PATIENT IS HAPPY." CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -15+1.0/13 | OD | Physician | 06/21/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17187714&pc=QCB | 2023826-2023-02499 | 2023-05-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/1.5/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #: (B)(4). | -14.0/1.5/072 | OS | Physician | 06/20/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17187715&pc=QCB | 2023826-2023-02501 | 2023-05-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -10.0/1.0/093 | OD | Physician | 06/20/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17187716&pc=QCB | 2023826-2023-02498 | 2023-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/2.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #: (B)(4). | -12.0/2.0/094 | OD | Physician | 06/20/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17187720&pc=QCB | 2023826-2023-02492 | 2023-05-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/1.0/052 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -11.5/1.0/052 | OS | Physician | 06/19/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17187721&pc=QCB | 2023826-2023-02491 | 2023-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/1.0/106 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -11.5/1.0/106 | OD | Physician | 06/19/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17187723&pc=QCB | 2023826-2023-02494 | 2023-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/2.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -7.5/2.0/088 | OS | Physician | 06/19/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17187725&pc=QCB | 2023826-2023-02490 | 2023-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/2.5/138(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -14.0/2.5/138 | OS | Physician | 06/19/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17187726&pc=QCB | 2023826-2023-02488 | 2023-06-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.5/1.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -3.5/1.0/094 | OD | Physician | 06/19/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17187728&pc=QCB | 2023826-2023-02489 | 2023-05-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/1.0/170 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -8.5/1.0/170 | OD | Physician | 06/19/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17188718&pc=QCB | 2023826-2023-02444 | 2023-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -10.00/0.5/073 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.00/0.5/073 | OD | Physician | 05/26/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17188719&pc=QCB | 2023826-2023-02445 | 2023-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -11.50/1.0/069 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.50/1.0/069 | OS | Physician | 05/26/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17191830&pc=QCB | 2023826-2023-02502 | 2023-05-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -15.0/2.0/055 (SPHERE/CYLINDER/AXIS) LENS TORE. THERE WAS NO PATIENT CONTACT AND INJURY REPORTED. Manufacturer Narrative: B5: CAUSE OF THE EVENT WAS USER ERROR. REPORTEDLY, "LENS TORN WHILST LOADING. DID NOT TOUCH PATIENT." THE PROBLEM WAS RESOLVED. THE REPORTER ALSO STATED, "LENS WAS TORN DURING LOADING. DID NOT TOUCH PATIENT'S EYE. THE BACKUP LENS WAS USED INSTEAD. ALL'S WELL. H6: WORK ORDER SEARCH WAS PERFORMED BUT WAS NOT REQUIRED. CLAIM#(B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS. CLAIM# (B)(4). | -15.0/2.0/055 | NA | Physician | 05/26/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17191940&pc=QCB | 2023826-2023-02525 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2; -17.00/3.5/096 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS HAD TORN WHILE IN A LIOLI-24 DEVICE. THERE WAS NO PATIENT CONTACT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -17.00/3.5/096 | NA | Physician | 05/30/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17194726&pc=QCB | 2023826-2023-02530 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2; -12.00/2.0/071 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS HAD TORE DURING LOADING. THERE WAS NO PATIENT CONTACT. A BACK UP LENS WAS USED. THIS OCCURRED ON (B)(6) 2023. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.00/2.0/071 | NA | Physician | 06/02/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17195739&pc=QCB | 2023826-2023-02533 | 2023-02-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+1.5/009 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND THE REPORTER STATED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -6.5/+1.5/009 | OS | Physician | 05/29/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17195740&pc=QCB | 2023826-2023-02548 | 2023-02-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.5/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.5/+1.5/095 | OD | Physician | 05/29/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17200403&pc=QCB | 2023826-2023-02467 | 2023-03-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -04.00/+4.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND LENS ROTATION. THE LENS WAS REPOSITIONED ON (B)(6) 2023 BUT THE PROBLEM WAS NOT RESOLVED. THE LENS REMAINED IMPLANTED. THE INITIAL CILIARY SHUTDOWN THEN SYSTEMIC ALLERGY TO "PREDFORTE". THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. D6B: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -04.00/+4.0/093 | OD | Physician | 05/31/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17205954&pc=QCB | 2023826-2023-02505 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -10.5/4.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE LEFT EYE (OS) OF A PATIENT WITH PRE-EXISTING PROGRESSIVE MYOPIA ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE/DEBRIS ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL - PRE-EXISTING PROGRESSIVE MYOPIA. CLAIM# (B)(4). | -10.5/4.0/089 | OS | Physician | 05/30/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17205955&pc=QCB | 2023826-2023-02507 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-06-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -9.0/0.5/098 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO PATIENT PREFERENCE. Manufacturer Narrative: B5: LENS WAS NOT USED. REPORTEDLY, "IT HAS NOT BEEN USED, IS IN ITS ORIGINAL PACKAGING, HAS NOT BEEN IN CONTACT WITH THE PATIENT'S EYE BECAUSE THIS IS WHAT 'EXCHANGE FOR PREFERENCE' MEANS." CLAIM# (B)(4). Manufacturer Narrative: H6: PATIENT RELATED FACTOR. CLAIM # (B)(4). | -9.0/0.5/098 | NA | Physician | 05/30/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17205956&pc=QCB | 2023826-2023-02506 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -10.0/2.5/102 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE RIGHT EYE (OD) OF A PATIENT WITH PRE-EXISTING PROGRESSIVE MYOPIA ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL - PRE-EXISTING PROGRESSIVE MYOPIA. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND RESIDUE/DEBRIS ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -10.0/2.5/102 | OD | Physician | 05/30/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17206118&pc=QCB | 2023826-2023-02551 |  | Injury | STAAR SURGICAL COMPANY | 2023-06-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -14.00/1.0/113 FLIPPED INSIDE THE EYE AFTER LENS INSERTION. THE LENS WAS REMOVED AND BACKUP LENS WAS IMPLANTED. THE SURGEON REPORTS THAT THE LENS FLIPPED DUE TO INCISION SIZE. Manufacturer Narrative: D6A:UNK. D6B:UNK. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -14.00/1.0/113 | NA | Other Health Care Professional | 05/29/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17206273&pc=QCB | 2023826-2023-02552 | 2023-05-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 12.6MM, VTICMO12.6, -18.00/2.0/092, IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE RIGHT EYE (OD). THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY WITH NO PATIENT INJURY. ON (B)(6) 2023 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS THE DEVICE. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -18.00/2.0/092 | OD | Physician | 05/30/2023 | 2023 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17210600&pc=QCB | 2023826-2023-02571 | 2023-06-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -6.5/1/101 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE LENS HAD TORE BEFORE/DURING LOADING, AFTER OPENING THE VIAL, THE LENS WAS STUCK IN CARTRIDGE OR INJECTION SYSTEM AND TORE DURING INJECTION/DELIVERY INTO THE EYE. INTRAOPERATIVELY THE LENS WAS REMOVED WITH NO PATIENT INJURY. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#:(B)(4). | -6.5/1/101 | OD | Physician | 06/22/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17210779&pc=QCB | 2023826-2023-02540 | 2023-04-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.5/114 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE LENS WAS EXCHANGED ON (B)(6)2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+1.5/114 | OS | Physician | 05/30/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17210780&pc=QCB | 2023826-2023-02537 | 2023-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+1.0/074 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.0/+1.0/074 | OD | Physician | 05/30/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17210781&pc=QCB | 2023826-2023-02538 | 2023-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+1.5/095 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.0/+1.5/095 | OS | Physician | 05/30/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17210786&pc=QCB | 2023826-2023-02541 | 2023-02-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+2.0/105 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -11.5/+2.0/105 | OD | Physician | 05/31/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17210787&pc=QCB | 2023826-2023-02539 | 2023-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.0/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+2.0/090 | OS | Physician | 05/30/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17213408&pc=QCB | 2023826-2023-02553 | 2023-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 OF -9.5/1.0/008 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN VERTICAL POSITION INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS OF THE SAME LENGTH AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT; ACD<3.0MM). CLAIM# (B)(4). | -9.5/1.0/008 | OD | Physician | 05/30/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17213630&pc=QCB | 2023826-2023-02470 | 2023-04-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.50/+2.5/083 (SPHERE/CYLINDER / AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO THE LENS WAS TOO LONG. IT WAS THE SURGEONS PREFERENCE TO CHANGE LENS. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS GOOD. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -13.50/+2.5/083 | OS | Physician | 06/01/2023 | 2023 | 44 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17213814&pc=QCB | 2023826-2023-02514 | 2022-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -9.0/4.0/169 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023, THE LENS WAS REMOVED, AND A REPLACEMENT LENS WAS IMPLANTED. REPORTEDLY, "LENS REPLACED FROM VERTICAL TO HORIZONTAL." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -9.0/4.0/169 | OD | Physician | 06/01/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17213816&pc=QCB | 2023826-2023-02512 | 2023-04-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -11.0/2.0/105 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS INTRAOPERATIVELY IMPLANTED AND REMOVED DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. ON (B)(6) 2023, THE REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "THE LENS LOADED SUCCESSFULLY BUT DURING THE INJECTION/IMPLANTATION IT WAS CAUGHT BY THE FOAM TIP PLUNGER AND TORE IT APART, THE LENS WAS TORN COMPLETELY AND WE REMOVED IT IMMEDIATELY." STATUS OF THE EYE IS, "AS EXPECTED EYE IS WELL." Manufacturer Narrative: CORRECTED DATA: H6- MEDICAL DEVICE CODE: 1494- OFF LABEL USE (AGE<21). H6- MEDICAL DEVICE PROBLEM CODE: "2993" SHOULD BE REMOVED AND "3189" SHOULD BE ADDED. H6-HEALTH EFFECT- IMPACT CODE: "4629" SHOULD BE REMOVED AND CODE "2199" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS. CLAIM#(B)(4). | -11.0/2.0/105 | OD | Physician | 05/31/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17216362&pc=QCB | 2023826-2023-02426 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. LENS ROTATION WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2 - UNKNOWN. A4 - UNKNOWN. A5 - UNKNOWN. A6 - UNKNOWN. B3 - UNKNOWN. D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). D4 - UNKNOWN. D6A - UNKNOWN. D6B - UNKNOWN. H4 - UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.5/074 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2023 THE LENS WAS REMOVED AND IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. CLAIM# (B)(4). | -8.5/+1.5/074 | OS | Physician | 06/05/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17216371&pc=QCB | 2023826-2023-02430 | 2023-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.5/3.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK,A5 - UNK.,A6 - UNK. CLAIM# (B)(4). | -9.5/3.0/096 | OD | Physician | 06/06/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17216377&pc=QCB | 2023826-2023-02433 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT WHILE THE SURGEON WAS INJECTING A 13.2MM VTICMO13.2, -17.00/3.0/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE LENS PINCHED BETWEEN THE CARTRIDGE AND PUSHER. THE BACK UP LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON NOTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.00/3.0/074 (SPHERE/CYLINDER/AXIS) TORE WHILE INJECTING. REPORTEDLY "LENS TEAR TODAY IN SURGERY WHILE INJECTING IT. IT GOT PINCHED BETWEEN THE CARTRIDGE AND THE GREEN PUSHER OF A LIOLI INJECTOR, LOT #FCK1703. HE INJECTED THE BACKUP LENS WITH A LIOLI INJECTOR FROM THE SAME LOT WITHOUT ANY ISSUE." THE LENS WAS IMPLANTED, REMOVED AND REPLACED INTRAOPERATIVELY. CLAIM# (B)(4). Manufacturer Narrative: A2 - UNK. A3 - UNK. A4 - UNK. A5 - UNK. A6 - UNK. B3 - UNK. D6A - UNK. D6B - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -17.00/3.0/074 | NA | Physician | 06/08/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17216379&pc=QCB | 2023826-2023-02436 | 2023-01-18 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED, THAT THE SURGEON IMPLANTED. A 13.2MM VTICMO13.2, -14.50/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS OTHER: HIGHER THAN EXPECTED VAULT, DUE TO LARGER SIZE LENS BEING SELECTED. REPORTEDLY, EXCHANGE POSSIBLE. Manufacturer Narrative: A4, A5, A6. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -14.50/1.0/085 | OD | Physician | 06/09/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17216387&pc=QCB | 2023826-2023-02435 | 2023-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -17.0/1.5/003 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS OPACITY(CATARACT /ASC)OBSERVED ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, NO PLANS TO IMPLANT A NEW LENS. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM, OVER 45 YRS OF AGE AT DATE OF IMPLANT. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -17.0/1.5/003 | OD | Physician | 06/08/2023 | 2023 | 50 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17221564&pc=QCB | 2023826-2023-02557 | 2023-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -10.5/1.5/110 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS THAT WAS IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: A1: (B)(6). H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -10.5/1.5/110 | OS | Physician | 05/30/2023 | 2023 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17221690&pc=QCB | 2023826-2023-02517 | 2023-05-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | UNK | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. REPORTEDLY, "AN ICL PATIENT FROM (B)(6) 2023 SURGERY COMPLAINING OF PAIN AND POOR VISION IN ONE EYE," AND "USED MIOSTAT ON THIS PATIENT AND FEELS THIS MAY BE THE ISSUE." Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS WITH A -10.5/1.0/165 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED AND LATER REPLACED FOR A SPHERICAL LENS. REPORTEDLY, "OD HORIZONTAL JUXTAPOSED METAMORPHIC EYE," AND "MIOCHOL WAS ON BACK ORDER MIOSTATE USED OU FOR INITIAL SURGERY. PT HAS RETINAL EVAL PENDING ALTHOUGH MOCULAR OCT SHOWS NO ABNORMALITIES WHEN COMPARED TO PREOP MOC OCT. APPT (B)(6) 2023 RETINA" THE PROBLEM WAS NOT RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK B3: UNK D4 (EXPERIATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM# (B)(4). | -10.5/1.0/165 | OD | Physician | 06/02/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17221767&pc=QCB | 2023826-2023-02558 | 2023-05-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 12.6MM, VTICMO12.6, -13.5/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE LEFT EYE (OS). THE LENS WAS IMPLANTED, REMOVED AND REPLACED INTRAOPERATIVELY WITH NO PATIENT INJURY AND THE PROBLEM RESOLVED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- MEDICAL DEVICE CODE: 1494- OFF LABEL USE (UNDER 21 YRS OF AGE AT DATE OF IMPLANTATION), CLAIM# (B)(4). | -13.5/2.0/091 | OS | Physician | 05/30/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17221877&pc=QCB | 2023826-2023-02400 | 2023-02-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/2.0/071 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "THE SECOND CRYSTAL IS PLACED AT AN OBLIQUE POSITION, 30 DEGREES." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -9.5/2.0/071 | OS | Physician | 06/12/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17221974&pc=QCB | 2023826-2023-02564 | 2023-05-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 12.6MM, VTICMO12.6, -5.5/1.0/011 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE LEFT EYE (OS). THE LENS WAS IMPLANTED, REMOVED AND REPLACED INTRAOPERATIVELY WITH NO PATIENT INJURY AND THE PROBLEM RESOLVED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | -5.5/1.0/011 | OS | Physician | 06/02/2023 | 2023 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17222199&pc=QCB | 2023826-2023-02602 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -9.50/3.0/140 TORE DURING DELIVERY INTO THE EYE. REPORTEDLY, "WAS INJECTING THE ICL IN THE EYE, THE ICL GOT STUCK IN THE INJECTOR AS HE WAS PULLING OUT THE PLUNGER. HE THEN PULLED OUT THE INJECTOR, PLUNGER AND ICL," AND "THERE WAS A BRIEF PATIENT CONTACT. ICL WAS PULLED OUT WITH NO ISSUE." Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -9.50/3.0/140 TORE DURING DELIVERY INTO THE EYE. ON (B)(6) 2023 THE LENS WAS REMOVED INTRAOPERATIVELY. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SIMILAR LENS MODEL. THIS RESOLVED THE PROBLEM. REPORTEDLY, "WAS INJECTING THE ICL IN THE EYE, THE ICL GOT STUCK IN THE INJECTOR AS HE WAS PULLING OUT THE PLUNGER. HE THEN PULLED OUT THE INJECTOR, PLUNGER AND ICL," AND "THERE WAS A BRIEF PATIENT CONTACT. ICL WAS PULLED OUT WITH NO ISSUE." CAUSE OF THE EVENT WAS USER ERROR. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS. CLAIM# (B)(6). | -9.50/3.0/140 | NA | Physician | 06/05/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17223223&pc=QCB | 2023826-2023-02565 | 2023-05-31 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-06-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2 -3.00/1.5/067 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE DURING LOADING. IT WAS REPORTED THAT THE DAMAGED WAS NOTED WHILE PULLING LENS THROUGH THE INSERTER. THERE WAS NO PATIENT CONTACT. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS THE DEVICE, FOR CAUSE DETAILS THE REPORTER STATES " WAS STUCK UPON PULLING THROUGH". Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -3.00/1.5/067 | NA | Other Health Care Professional | 06/02/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17225350&pc=QCB | 2023826-2023-02468 | 2023-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/112 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT POWER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 03-JUL-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.50/+1.0/112 | OS | Physician | 06/02/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17232295&pc=QCB | 2023826-2023-02569 | 2023-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-06-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -11.0/0.5/092 (SPHERE / CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS OF LONGER LENGTH AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD 3.0MM). CLAIM#: (B)(4). | -11.0/0.5/092 | OD | Physician | 06/05/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17244376&pc=QCB | 2023826-2023-02590 | 2023-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+2.0/174 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE LENS WAS EXCHANGED ON (B)(6)2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/+2.0/174 | OD | Physician | 06/02/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17244377&pc=QCB | 2023826-2023-02594 | 2022-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.0/+0.5/100 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -4.0/+0.5/100 | OS | Physician | 06/02/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17244378&pc=QCB | 2023826-2023-02591 | 2023-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.0/075 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -11.0/+1.0/075 | OD | Physician | 06/02/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17244383&pc=QCB | 2023826-2023-02597 | 2023-03-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.5/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+1.5/091 | OD | Physician | 06/02/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17248285&pc=QCB | 2023826-2023-02606 | 2022-12-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 OF A -10.0/2.0/095 (SPHERE/CYLINDER/AXIS) MAY HAVE A POSSIBLE ADVERSE EVENT OR MALFUNCTION. NO FURTHER INFORMATION AT THIS TIME HAS BEEN PROVIDED. Manufacturer Narrative: A4-A6:UNK CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 -10.0/2.0/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS WAS EXPLANTED ON (B)(6) 2023. REPORTEDLY, "THE PATIENT HAD MUSCAE VOLITANTES AND OFTEN COMPLAINED THE PAIN AROUND THE EYE SO SHE REQUIRED THE LENS EXTRACTED." CLAIM# (B)(4). Manufacturer Narrative: H6 - INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -10.0/2.0/095 | OS | Physician | 06/06/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17249526&pc=QCB | 2023826-2023-02656 | 2023-05-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -18.00/2.0/174(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED AND A SHORTER LENGTH LENS WAS IMPLANTED DUE TO EXCESSIVE VAULT. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: HEALTH EFFECT - IMPACT CODE: "4628" SHOULD BE REMOVED AND "4627" SHOULD BE ADDED/. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -18.00/2.0/174 | OS | Physician | 06/07/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253009&pc=QCB | 2023826-2023-02578 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/129 (SPHERE / CYLINDER/AXIS), WAS NOTED TO BE TORN WHEN REMOVED FROM THE VIAL. THERE WAS NO PATIENT CONTACT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE WITH FOREIGN MATERIAL ON LENS SURFACE, SPECIFICALLY, DEBRIS ON THE LENS. CLAIM # (B)(4). | -10.00/+1.5/129 | NA | Physician | 12/20/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253086&pc=QCB | 2023826-2023-02699 |  | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -9.5/+1.5/88 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON THIS SAME DATE A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. IT WAS ADDITIONALLY NOTED "AFTER INCREASING THE SIZE OF THE ICL, THE LENS NO LONGER ROTATES". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -9.5/+1.5/88 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2023 THE LENS WAS REPOSITIONED AND REPORTEDLY RESOLVED THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON THIS SAME DATE A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. IT WAS ADDITIONALLY NOTED "AFTER INCREASING THE SIZE OF THE ICL, THE LENS NO LONGER ROTATES". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. D9: RETURN DATE: (B)(6) 2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND SCRATCHES ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6: 4628 LENS REPOSITIONED. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.5/+1.5/88 | OS | Physician | 06/30/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253089&pc=QCB | 2023826-2023-02691 | 2023-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -6.5/1.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON (B)(6) 2023 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER". Manufacturer Narrative: CLAIM#: (B)(4). | -6.5/1.0/092 | OS | Physician | 06/30/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253103&pc=QCB | 2023826-2023-02528 | 2023-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.5/178 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -10.5/1.5/178 | OD | Physician | 06/21/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253105&pc=QCB | 2023826-2023-02625 | 2023-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/1.0/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -6.5/1.0/074 | OD | Physician | 06/24/2023 | 2023 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253106&pc=QCB | 2023826-2023-02624 | 2023-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/0.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -6.0/0.5/089 | OS | Physician | 06/24/2023 | 2023 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253108&pc=QCB | 2023826-2023-02627 | 2023-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/2.0/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -11.5/2.0/074 | OS | Physician | 06/25/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253109&pc=QCB | 2023826-2023-02630 | 2023-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.5/1.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -16.5/1.5/074 | OS | Physician | 06/27/2023 | 2023 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253110&pc=QCB | 2023826-2023-02626 | 2023-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/1.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -7.5/1.5/093 | OD | Physician | 06/25/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253111&pc=QCB | 2023826-2023-02628 | 2023-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/3.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -12.5/3.0/090 | OD | Physician | 06/27/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253112&pc=QCB | 2023826-2023-02631 | 2023-04-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/1.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -7.0/1.0/093 | OD | Physician | 06/27/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253113&pc=QCB | 2023826-2023-02629 | 2023-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.5/1.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -15.5/1.0/086 | OD | Physician | 06/27/2023 | 2023 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253116&pc=QCB | 2023826-2023-02634 | 2023-04-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/3.0/095(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -11.0/3.0/095 | OD | Physician | 06/27/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253118&pc=QCB | 2023826-2023-02635 | 2023-04-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/3.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -10.0/3.0/094 | OS | Physician | 06/27/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253209&pc=QCB | 2023826-2023-02692 | 2023-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -9.0/1.0/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE WITH FOREIGN MATERIAL ON THE LENS SURFACE, SPECIFICALLY FIBERS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM #(B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.0/1.0/095 | OS | Physician | 06/06/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253277&pc=QCB | 2023826-2023-02546 | 2023-04-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+2.5/074 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#(B)(4). | -12.5/+2.5/074 | OS | Physician | 06/05/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253278&pc=QCB | 2023826-2023-02677 | 2023-04-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.0/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+1.0/093 | OS | Physician | 06/05/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253279&pc=QCB | 2023826-2023-02547 | 2023-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+1.0/007 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) AND (OVER 45YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -6.5/+1.0/007 | OS | Physician | 06/05/2023 | 2023 | 49 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253281&pc=QCB | 2023826-2023-02681 | 2023-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+1.0/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -7.5/+1.0/093 | OS | Physician | 06/06/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253284&pc=QCB | 2023826-2023-02679 | 2023-02-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.0/+2.5/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -16.0/+2.5/093 | OD | Physician | 06/06/2023 | 2023 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17253285&pc=QCB | 2023826-2023-02680 | 2023-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/+1.5/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -14.5/+1.5/094 | OD | Physician | 06/06/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17259072&pc=QCB | 2023826-2023-02695 | 2022-08-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -7.0/2.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 AS A REPLACEMENT LENS TO ADDRESS EXCESSIVE VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -7.0/2.0/093 | OS | Physician | 06/06/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17259706&pc=QCB | 2023826-2023-02694 | 2022-08-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -7.0/2.0/081 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE DID NOT RESOLVE THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -7.0/2.0/081 | OS | Physician | 06/06/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17262092&pc=QCB | 2023826-2023-02582 | 2023-05-31 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.50/+0.5/170 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK A5: UNK A6: UNK H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | -10.50/+0.5/170 | NA | Physician | 06/12/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17262104&pc=QCB | 2023826-2023-02583 | 2023-03-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+3.5/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: A4: UNK, A5: UNK, A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.50/+3.5/074 | OS | Physician | 06/09/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17262123&pc=QCB | 2023826-2023-02686 | 2023-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+3.0/115 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.0/+3.0/115 | OD | Physician | 06/06/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17262173&pc=QCB | 2023826-2023-02738 |  | Injury | STAAR SURGICAL COMPANY | 2023-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2; -14.50/1.5/131 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE, IT HAD FOLDED ON INSERTION. WHEN THE LENS WAS REMOVED THE LENS TORE. IT WAS REPLACED WITH A SAME LENGTH LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -14.50/1.5/131 | NA | Physician | 06/12/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17262272&pc=QCB | 2023826-2023-02581 | 2023-05-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+1.0/092 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO LENS TORE/BROKE AFTER OPENING VIAL. THERE WAS PATIENT CONTACT BUT NO INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: B5 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN MATERIAL ON THE LENS SURFACE, SPECIFICALLY DEBRIS THOUGHOUT THE LENS. CLAIM # (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.50/+1.0/092 | NA | Physician | 06/09/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17262468&pc=QCB | 2023826-2023-02683 | 2022-12-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.5/109 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -7.0/+1.5/109 | OS | Physician | 06/06/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17262713&pc=QCB | 2023826-2023-02684 | 2023-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+2.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.0/+2.0/091 | OD | Physician | 06/06/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17264621&pc=QCB | 2023826-2023-02666 | 2023-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -7.0/1.0/161 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, STATUS OF THE EYE IS SAID TO BE, "GOOD." ALSO, "SCHEDULED TO BE REPLACED AT A LATER DATE AFTER LENS SELECTION." Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE RETURNED LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -7.0/1.0/161 | OS | Physician | 06/20/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17264622&pc=QCB | 2023826-2023-02667 | 2023-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -6.5/1.0/161 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, STATUS OF THE EYE IS SAID TO BE, "GOOD." ALSO, "SCHEDULED TO BE REPLACED AT A LATER DATE AFTER LENS SELECTION." Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE RETURNED LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -6.5/1.0/161 | OD | Physician | 06/20/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17274413&pc=QCB | 2023826-2023-02670 | 2023-02-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -11.0/3.5/089 (SPHERE / CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2023, THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SPHERICAL LENS AND THE PROBLEM WAS RESOLVED. REPORTEDLY, "LASEK WAS PERFORMED FOR ASTIGMATISM." CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: OFF-LABEL AGE 21. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE/DEBRIS ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). | -11.0/3.5/089 | OS | Physician | 06/21/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17274414&pc=QCB | 2023826-2023-02668 | 2022-12-27 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022, A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -12.00/1.0/073 (SPHERE/CYLINDER/AXIS) TORE DURING LOADING. THE LENS WAS INTRAOPERATIVELY IMPLANTED AND REMOVED FOR A REPLACEMENT LENS INTO THE PATIENT'S RIGHT EYE(OD). THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT ON (B)(6) 2022, A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -12.00/1.0/073 (SPHERE/CYLINDER/AXIS) TORE DURING LOADING. THERE WAS NO PATIENT CONTACT. A BACKUP REPLACEMENT LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CLAIM#(B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -12.00/1.0/073 | OD | Physician | 06/21/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17274417&pc=QCB | 2023826-2023-02669 | 2023-04-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023, A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -6.50/3.5/087 (SPHERE/CYLINDER/AXIS) TORE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS INTRAOPERATIVELY IMPLANTED AND REMOVED FOR A LARGER LENGTH LENS INTO THE PATIENT'S LEFT EYE (OS). NO INJURY WAS REPORTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -6.50/3.5/087 | OS | Physician | 06/21/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17276357&pc=QCB | 2023826-2023-02742 | 2023-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.0/+3.0/085 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM#: (B)(4). | -16.0/+3.0/085 | OS | Physician | 06/08/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17276359&pc=QCB | 2023826-2023-02743 | 2023-05-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.50/+2.0/083 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -12.50/+2.0/083 | OS | Physician | 06/08/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17276361&pc=QCB | 2023826-2023-02744 | 2023-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+1.5/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -13.0/+1.5/091 | OD | Physician | 06/08/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17276362&pc=QCB | 2023826-2023-02745 | 2023-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+3.5/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -13.0/+3.5/094 | OD | Physician | 06/08/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17276369&pc=QCB | 2023826-2023-02747 | 2023-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+3.0/096 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.5/+3.0/096 | OS | Physician | 06/09/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17279002&pc=QCB | 2023826-2023-02615 | 2023-06-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -14.50/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE LENS TORE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY ON (B)(6)2023. PROBLEM RESOLVED. CAUSE OF THE EVENT IS REPORTED AS DEVICE. REPORTEDLY, PATIENT IS HAPPY. Manufacturer Narrative: A4 - UNK A5 - UNK A6 - UNK H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -14.50/2.0/089 | OS | Physician | 06/23/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17279004&pc=QCB | 2023826-2023-02612 | 2023-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED VERTICALLY A 12.6MM VTICMO12.6, -3.5/1.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED HORIZONTALLY WITH A SAME LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3 - LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTTIC DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -3.5/1.0/094 | OS | Physician | 06/14/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17279007&pc=QCB | 2023826-2023-02616 | 2023-05-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -13.00/4.50/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING LOADING LENS FOR DELIVERY INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REPLACEMENT LENS IMPLANTED INTRAOPERATIVELY ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: D4 - EXPIRATION DATE: 28-FEB-2026. Manufacturer Narrative: A4 - UNK, A5 - UNK, A6 - UNK. (B)(4) | -13.00/4.50/093 | OS | Physician | 06/15/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17279030&pc=QCB | 2023826-2023-02620 | 2023-04-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.5/2.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS STUCK IN CARTRIDGE/INJECTION SYSTEM DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED AND REMOVED REPLACED INTRAOPERATIVELY ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, 12.6 LENS TORE AND BACK UP LENS 13.2 USED. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. | -10.5/2.5/095 | OS | Physician | 06/15/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17280830&pc=QCB | 2023826-2023-02585 | 2019-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.50/+2.0/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2019 BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS REPLACED WITH A LONGER LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. B3: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A HAPTIC BROKEN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -13.50/+2.0/100 | OD | Physician | 06/12/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17280838&pc=QCB | 2023826-2023-02584 | 2023-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -06.50/+1.0/005 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5: THE REPLACEMENT LENS WAS IMPLANTED VERTICALLY. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -06.50/+1.0/005 | OS | Physician | 06/12/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283697&pc=QCB | 2023826-2023-02637 | 2023-01-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.0/069 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -10.5/1.0/069 | OS | Physician | 06/28/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283698&pc=QCB | 2023826-2023-02636 | 2023-04-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.0/071 (SPHERE / CYLINDER / AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM#: (B)(4). | -10.5/1.0/071 | OS | Physician | 06/27/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283700&pc=QCB | 2023826-2023-02640 | 2023-04-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.5/1.5/086 (SPHERE / CYLINDER / AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "STATUS OF THE EYE: CORNEA CLEAR, ACD DEEP, ICL CENTRAL, VAULTY 0.8 CT, LENS CLEAR, RETINA NORMAL. REPLACEMENT DEGREE BECAUSE PATIENT'S MYOPIA DEGREE INCREASED A LITTLE MORE THAN THREE MONTHS AGO." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM#: (B)(4). | -3.5/1.5/086 | OD | Physician | 06/28/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283703&pc=QCB | 2023826-2023-02642 | 2023-06-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/2.0/109 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -11.5/2.0/109 | OS | Physician | 06/28/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283704&pc=QCB | 2023826-2023-02700 | 2023-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/3.0/ 174 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "THE FIRST PIECE IS A VERTICAL IMPLANT." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM (B)(4). | -12.5/3.0/ 174 | OD | Physician | 06/29/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283705&pc=QCB | 2023826-2023-02703 | 2023-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/2.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -10.0/2.5/093 | OD | Physician | 06/30/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283706&pc=QCB | 2023826-2023-02704 | 2023-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/0.5/152 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM (B)(4). | -5.5/0.5/152 | OS | Physician | 07/01/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283707&pc=QCB | 2023826-2023-02701 | 2023-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/1.5/010 (SPHERE / CYLINDER / AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD 3.0MM). CLAIM#: (B)(4). | -9.0/1.5/010 | OD | Physician | 06/29/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283708&pc=QCB | 2023826-2023-02702 | 2023-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/2.0/160(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -9.0/2.0/160 | OS | Physician | 06/29/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283709&pc=QCB | 2023826-2023-02705 | 2023-04-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/0.5/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - CORRECTED TO : "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/0.5/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN." IN THE INITIAL MDR. D6A - CORRECTED TO: (B)(6) 2023 IN THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -6.0/0.5/085 | OS | Physician | 07/01/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283710&pc=QCB | 2023826-2023-02707 | 2023-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/2.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -12.0/2.0/095 | OS | Physician | 07/02/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283711&pc=QCB | 2023826-2023-02706 | 2023-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/0.5/105 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM (B)(4). | -11.0/0.5/105 | OD | Physician | 07/02/2023 | 2023 | 17 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283713&pc=QCB | 2023826-2023-02708 | 2023-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/2.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -12.0/2.0/095 | OD | Physician | 07/02/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283714&pc=QCB | 2023826-2023-02709 | 2023-05-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/3.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -7.5/3.0/091 | OD | Physician | 07/02/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283715&pc=QCB | 2023826-2023-02711 | 2023-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/2.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ON (B)(6) 2023 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM (B)(4). | -13.5/2.5/089 | OS | Physician | 07/03/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283718&pc=QCB | 2023826-2023-02712 | 2023-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/1.5/115 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN Manufacturer Narrative: A4-A6: UNK. CLAIM (B)(4). | -9.0/1.5/115 | OD | Physician | 07/03/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283719&pc=QCB | 2023826-2023-02715 | 2023-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/1.5/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -11.5/1.5/073 | OS | Physician | 07/04/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283721&pc=QCB | 2023826-2023-02713 | 2023-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/2.0/115 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNKNOWN. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -10.5/2.0/115 | OD | Physician | 07/03/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283722&pc=QCB | 2023826-2023-02714 | 2023-04-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/2.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) . CLAIM # (B)(4). | -14.5/2.5/092 | OS | Physician | 07/04/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283822&pc=QCB | 2023826-2023-02750 | 2023-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: A 12.6MM VTICM5_12.6 OF -12.0/4.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLER LENS WAS NOTED TO HAVE AN OVAL- SHAPED MISSING SECTION DURING UNFOLDING/INJECTION INTO THE PATIENTS EYE. THE PIECE WAS LATER DISCOVERED TO BE INSIDE THE LIOLI INJECTOR. THE LENS WAS EXPLANTED AND REPLACED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -12.0/4.0/089 | NA | Non-Healthcare Professional | 06/09/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283880&pc=QCB | 2023826-2023-02674 | 2023-06-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -3.5/2.5/27 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT AND REFRACTIVE SUPRISE WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL - ACD<3.0. (B)(4). | -3.5/2.5/27 | OS | Physician | 06/22/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17283883&pc=QCB | 2023826-2023-02673 | 2023-06-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -5.0/3.5/164 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT AND REFRACTIVE SUPRISE WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -5.0/3.5/164 | OD | Physician | 06/22/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17284960&pc=QCB | 2023826-2023-02751 | 2023-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.5/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.5/+1.5/092 | OS | Physician | 06/09/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17292414&pc=QCB | 2023826-2023-02727 | 2023-06-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50/+3.5/068 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO THERE WAS A RIP IN THE MIDDLE OF THE LENS AS IT WAS GOING THROUGH THE PLUNGER. THE BACKUP LENS WAS USED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN VIAL IN LIQUID. VISUAL INSPECTION FOUND OPTIC TORN. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | -14.50/+3.5/068 | NA | Other Health Care Professional | 06/26/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17292415&pc=QCB | 2023826-2023-02718 | 2023-06-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50/+2.0/088 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS RIGHT EYE (OD). A MARK WAS NOTED ON THE LENS OPTIC AFTER INJECTION. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. THE LENS WAS VIEWED UNDER 20X MAGNIFICATION; NO MARKS WERE NOTED ON THE OPTIC. CLAIM # (B)(4). | -10.50/+2.0/088 | OD | Other Health Care Professional | 06/16/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17292449&pc=QCB | 2023826-2023-02654 | 2023-06-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -05.00/+0.5/074 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO SURGEON IMPLANTED THE LENS FOR THE RIGHT EYE (OD) INTO THE LEFT EYE (OS). THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT POWER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR; THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. . Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. CLAIM # (B)(4). | -05.00/+0.5/074 | OS | Physician | 06/16/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17292474&pc=QCB | 2023826-2023-02651 | 2023-05-31 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00/+1.0/113 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6)2023. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. D6B: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.00/+1.0/113 | OS | Physician | 06/15/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17292616&pc=QCB | 2023826-2023-02652 | 2023-06-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.00/+1.5/002 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6)2023. THE LENS HAD LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK A5: UNK A6: UNK D6B:UNK H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2023 FOR A LONGER LENS. THE PROBLEM WAS RESOLVED AND THE PATIENT WAS HAPPY. CLAIM # (B)(4). | -09.00/+1.5/002 | OD | Physician | 06/16/2023 | 2023 | 43 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17292646&pc=QCB | 2023826-2023-02653 | 2023-06-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -07.00/+1.0/108 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6)2023. THE LENS WAS REMOVED ON (B)(6)2023 DUE TO SURGEON IMPLANTED THE LENS FOR THE LEFT EYE (OS) INTO THE RIGHT EYE (OD). THE LENS WAS REPLACED WITH A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR; THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. . Manufacturer Narrative: CLAIM # (B)(4). | -07.00/+1.0/108 | OD | Physician | 06/16/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17295353&pc=QCB | 2023826-2023-02831 | 2022-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -15.0/+1.5/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND ON THIS SAME DATE, A LONGER LENGTH LENS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: B5 - LENS WAS REPOSITIONED ON (B)(6) 2022 BUT THAT DID NOT RESOLVED THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND ON THIS SAME DATE, A LONGER LENGTH LENS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -15.0/+1.5/89 | OS | Physician | 07/06/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17295354&pc=QCB | 2023826-2023-02820 | 2022-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -13.5/+20/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND ON THIS SAME DATE, A LONGER LENGTH LENS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: B5 - LENS WAS REPOSITIONED ON (B)(6) 2022 BUT THAT DID NOT RESOLVED THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND ON THIS SAME DATE, A LONGER LENGTH LENS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: H6: TYPE OF INVESTIGATION: 4110: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.5/+20/92 | OD | Physician | 07/06/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17295355&pc=QCB | 2023826-2023-02832 | 2023-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.5/1.0/013 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING WITHOUT ROTATION. ON (B)(6) 2023 THE LENS WAS EXPLANTED. ON (B)(6) 2023 A REPLACEMENT LENS OF THE SAME LENGTH BUT IMPLANTED HORIZONTALLY (INITIAL WAS IMPLANTED VERTICALLY) AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC BENT AND DEFORMED.DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). | -9.5/1.0/013 | OD | Physician | 07/06/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17295508&pc=QCB | 2023826-2023-02788 | 2023-03-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.5/+1.0/178 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -4.5/+1.0/178 | OD | Physician | 06/12/2023 | 2023 | 45 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17295509&pc=QCB | 2023826-2023-02790 | 2023-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+1.5/088 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -12.5/+1.5/088 | OS | Physician | 06/12/2023 | 2023 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17295512&pc=QCB | 2023826-2023-02789 | 2023-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+5.0/097 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -10.5/+5.0/097 | OS | Physician | 06/12/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17295514&pc=QCB | 2023826-2023-02794 | 2023-04-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.0/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -10.0/+1.0/089 | OS | Physician | 06/12/2023 | 2023 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17295516&pc=QCB | 2023826-2023-02795 | 2023-04-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+4.0/095 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -8.0/+4.0/095 | OS | Physician | 06/12/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17295991&pc=QCB | 2023826-2023-02834 | 2023-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.0/2.0/073 (SPHERE/CYLINDER/AXIS) TORE/BROKE DURING INJECTION/DELIVERY INTO THE LEFT EYE (OS). ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH NO PATIENT INJURY AND THE PROBLEM RESOLVED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -10.0/2.0/073 | OS | Physician | 06/12/2023 | 2023 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17297494&pc=QCB | 2023826-2023-02780 | 2023-06-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED UPSIDE DOWN A 13.2MM VTICM5_13.2, -630/1.0/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE ON (B)(6) 2023. THE LENS TEAR/BREAK DURING INJECTION/DELIVER INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY ON (B)(6) 2023 WITH A SAME SIZE, SIMILAR (SPHERE/CYLINDER/AXIS) LENS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK, A5 - UNK, A6 - UNK, H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -6.30/1.0/101 | OD | Physician | 06/20/2023 | 2023 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17297510&pc=QCB | 2023826-2023-02783 | 2023-05-31 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.50/3.0/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023.THE LENS TORE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY ON (B)(6) 2023. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, EYE IS DOING WELL, NO CONCERNS. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.50/3.0/110 | OS | Physician | 06/21/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17299686&pc=QCB | 2023826-2023-02796 | 2023-02-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.5/+2.0/008 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -4.5/+2.0/008 | OD | Physician | 06/12/2023 | 2023 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17299687&pc=QCB | 2023826-2023-02798 | 2023-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: HE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.5/+2.0/008 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -4.5/+2.0/008 | OD | Physician | 06/12/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17301939&pc=QCB | 2023826-2023-02838 | 2023-06-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: A 13.2MM VTICM13.2 OF -16.0/1.0/124 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS PARTIALLY IMPLANTED UPSIDE DOWN INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6)2023. THE LENS WAS PARTIALLY IMPLANTED, REMOVED AND REPLACED INTRAOPERATIVELY WITH NO PATIENT INJURY AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS USER ERROR, LOADING ERROR WAS REPORTED FOR CAUSE DETAILS. THE REPORTER ALSO STATED THAT IT WAS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -16.0/1.0/124 | OD | Physician | 06/12/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17302268&pc=QCB | 2023826-2023-02760 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -5.5/3.0/118 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS A REPLACEMENT LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE EXCHANGED WAS DUE TO LOW VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY, "SURGEON DECIDED TO MONITOR THE PATIENT CLOSELY AND CHECK IF THE RX IS STABLE." Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS FOUND. CLAIM# (B)(4). | -5.5/3.0/118 | OD | Physician | 06/23/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17302544&pc=QCB | 2023826-2023-02725 | 2023-03-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.5/168 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "THE FIRST PIECE IS A VERTICAL IMPLANT." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -10.0/1.5/168 | OD | Physician | 07/06/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17302545&pc=QCB | 2023826-2023-02485 | 2023-05-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_ 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -2.5/0.5/070 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2023 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -2.5/0.5/070 | OS | Physician | 06/19/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17302546&pc=QCB | 2023826-2023-02753 | 2023-03-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/2.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -10.5/2.5/089 | OS | Physician | 07/06/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17302547&pc=QCB | 2023826-2023-02755 | 2023-03-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/3.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "SECOND VERTICAL POSITION." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -11.5/3.0/094 | OD | Physician | 07/06/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17302548&pc=QCB | 2023826-2023-02754 | 2023-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/2.0/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "SECOND VERTICAL POSITION." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -12.0/2.0/072 | OS | Physician | 07/06/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17302552&pc=QCB | 2023826-2023-02758 | 2023-06-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/2.0/070 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING WITHOUT ROTATION. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM (B)(4). | -14.5/2.0/070 | OS | Physician | 07/06/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17302553&pc=QCB | 2023826-2023-02756 | 2023-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.5/1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -15.5/1.5/091 | OD | Physician | 07/06/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17302732&pc=QCB | 2023826-2023-02823 | 2023-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, 12.00/+4.0/109 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. THE SURGEON USED THE WRONG AXIS: 17° INSTEAD OF 170°. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2023 FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+4.0/109 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL AND POWER, LONGER LENGTH, DIFFERENT AXIS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS USER ERROR. CAUSE DETAILS PROVIDED: "SURGEON USED THE WRONG AXIS: 17° INSTEAD 170°". CLAIM# (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. D6B: UNK. H6. WORK ORDER SEARCH: NO. SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4). | -12.0/+4.0/109 | OS | Physician | 06/21/2023 | 2023 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17302934&pc=QCB | 2023826-2023-02877 | 2022-12-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.0/2.0/110 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT DIFFERENT AXIS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.0/2.0/110 | OD | Physician | 07/09/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17303119&pc=QCB | 2023826-2023-02801 | 2023-04-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+1.0/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -7.5/+1.0/094 | OD | Physician | 06/13/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17303120&pc=QCB | 2023826-2023-02800 | 2023-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.0/072 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) AND (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -10.0/+2.0/072 | OS | Physician | 06/13/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17303124&pc=QCB | 2023826-2023-02804 | 2023-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+2.5/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE LENS WAS EXCHANGED ON (B)(6)2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -11.5/+2.5/095 | OD | Physician | 06/14/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17303125&pc=QCB | 2023826-2023-02802 | 2023-05-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+2.0/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -8.0/+2.0/089 | OS | Physician | 06/13/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17303176&pc=QCB | 2023826-2023-02805 | 2023-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.0/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE (B)(4) : OFF-LABEL USE (ACD < 3.0MM) CLAIM# (B)(4). | -7.0/+1.0/094 | OS | Physician | 06/14/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17303177&pc=QCB | 2023826-2023-02806 | 2023-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.5/+1.5/146 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -4.5/+1.5/146 | OS | Physician | 06/14/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17307587&pc=QCB | 2023826-2023-02843 | 2023-04-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 12.6MM VTICMO12.6 OF -10.0/1.5/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE LEFT EYE (OS). THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY WITH NO PATIENT INJURY. ON (B)(6) 2023 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. THE REPORTER ALSO STATES THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.0/1.5/086 | OS | Physician | 06/15/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17308922&pc=QCB | 2023826-2023-02765 | 2023-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -11.0/5.5/101 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LOW VAULT WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: ON 18-SEP-2023 THE LENS WAS EXCHANGED FOR A LONGER LENS; THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -11.0/5.5/101 | OD | Physician | 06/26/2023 | 2023 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17308924&pc=QCB | 2023826-2023-02763 | 2023-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -2.5/4/083 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT AND REFRACTIVE SUPRISE WERE OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE LENS WAS REPOSITIONED. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# 730208. | -2.5/4/083 | OD | Physician | 06/26/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17308942&pc=QCB | 2023826-2023-02764 | 2023-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -2.00/3.0/076 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SUPRISE WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE LENS HAS BEEN REPOSITIONED. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# : (B)(4). | -2.00/3.0/076 | OS | Physician | 06/26/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17313014&pc=QCB | 2023826-2023-02863 | 2023-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/2.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -7.0/2.0/093 | OS | Physician | 07/07/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17313015&pc=QCB | 2023826-2023-02864 | 2023-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/2.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -13.0/2.0/090 | OD | Physician | 07/07/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17313016&pc=QCB | 2023826-2023-02865 | 2023-03-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2IMPLANTABLE COLLAMER LENS OF DIOPTER -8.50/3.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -8.50/3.0/093 | OD | Physician | 07/09/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17313018&pc=QCB | 2023826-2023-02871 | 2023-05-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -10.0/2.0/092 | OD | Physician | 07/11/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17313027&pc=QCB | 2023826-2023-02870 | 2023-04-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/0.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -6.0/0.5/074 | OD | Physician | 07/10/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17313028&pc=QCB | 2023826-2023-02869 | 2023-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/4.0/112 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -10.0/4.0/112 | OS | Physician | 06/21/2023 | 2023 | 16 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17313029&pc=QCB | 2023826-2023-02872 | 2023-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/0.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -7.5/0.5/090 | OS | Physician | 07/11/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17313133&pc=QCB | 2023826-2023-02813 | 2023-05-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+2.5/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.5/+2.5/090 | OS | Physician | 06/15/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17313135&pc=QCB | 2023826-2023-02808 | 2022-12-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.5/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.5/+1.5/093 | OD | Physician | 06/15/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17313136&pc=QCB | 2023826-2023-02812 | 2022-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+0.5/070 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -7.0/+0.5/070 | OS | Physician | 06/15/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17313137&pc=QCB | 2023826-2023-02809 | 2023-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.5/109 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -9.5/+1.5/109 | OS | Physician | 06/15/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17321072&pc=QCB | 2023826-2023-02844 | 2023-05-31 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 12.1MM VTICM5_12.1 OF -14.0/2.5/002 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE/BROKE DURING LOADING AND WAS STUCK IN THE CARTRIDGE OR INJECTION SYSTEM. THERE WAS NO PATIENT CONTACT. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -14.0/2.5/002 | NA | Physician | 06/15/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17321925&pc=QCB | 2023826-2023-02900 | 2023-04-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2; -14.0/+2.0/81 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS THAT WAS INTENDED FOR A PATIENT'S LEFT EYE (OS) HAD TORN WHILE LOADING. THIS OCCURRED ON (B)(6) 2023. ON THE SAME DATE A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) (B)(4). | -14.0/+2.0/81 | OS | Physician | 07/12/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17322071&pc=QCB | 2023826-2023-02845 |  | Injury | STAAR SURGICAL COMPANY | 2023-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -7.00/1.0/017 (SPHERE/CYLINDER/AXIS) TORE AND A BACK UP LENS WAS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -7.00/1.0/017 | NA | Other Health Care Professional | 06/15/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17322212&pc=QCB | 2023826-2023-02769 |  | Injury | STAAR SURGICAL COMPANY | 2023-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -14.5/3.0/072 (SPHERE / CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE PROBLEM WAS RESOLVED. REPORTEDLY, "GOOD VAULT." THE CAUSE OF THE EVENT WAS OTHER AND STATING, "CALCULATOR ERROR." Manufacturer Narrative: H6 - HEALTH EFFECT CLINICAL CODE: 4581 SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES - SHOULD BE INCLUDED IN THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -14.5/3.0/072 | OS | Physician | 06/26/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17322214&pc=QCB | 2023826-2023-02771 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -7.00/2.0/090 TORE DURING LOADING. REPORTEDLY, "ICL GOT CAUGHT IN AN INJECTOR WHICH CAUSED THE LENS TO RIP. TURNS OUT THE LOT#FCK1701 WAS USED. I WAS JUST INFORMED THAT THESE ARE ON RECALL." THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS. (B)(4). | -7.00/2.0/090 | NA | Physician | 06/27/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17322216&pc=QCB | 2023826-2023-02768 | 2023-06-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -6.5/1.5/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. REPORTEDLY, "THE DOCTOR FOUND A FLOATING LINEAR STICKING TO THE SURFACE OF THE LENS DURING THE SURGERY. THE FOREIGN PARTICULATE WAS REMOVED AND SURGERY WAS DONE." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL - AGE<21 WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | -6.5/1.5/091 | OS | Physician | 06/26/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17322330&pc=QCB | 2023826-2023-02815 | 2023-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED, THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.0/+1.0/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS, DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21 YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -4.0/+1.0/094 | OS | Physician | 06/16/2023 | 2023 | 17 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17322331&pc=QCB | 2023826-2023-02816 | 2023-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+3.0/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -9.0/+3.0/094 | OD | Physician | 06/16/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17329934&pc=QCB | 2023826-2023-02775 | 2023-04-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF A -8.5/1.0/094 (SPHERE/AXIS/CYLINDER) WAS IMPLANTED INTOT HE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. IN A SUBSEQUENT SURGERY LATER THAT DAY, THE LENS WAS REMOVED AND LATER REPLACED FOR A REPLACEMENT LENS DUE TO PRODUCT NON-CONFORMITY. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B5: REPORTEDLY. "FOUND FOOT LOOP DEFECT WHEN REMOVING FROM THE BOTTLE." CLAIM# (B)(4). Manufacturer Narrative: B5 - LENS WAS NOT IMPLANTED. CLAIM# (B)(4). | -8.5/1.0/094 | OD | Physician | 06/28/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17335289&pc=QCB | 2023826-2023-02918 | 2023-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.0/097 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: (B)(4). | -10.0/+1.0/097 | OD | Physician | 06/21/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17335290&pc=QCB | 2023826-2023-02920 | 2023-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/+5.5/071 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -5.5/+5.5/071 | OD | Physician | 06/22/2023 | 2023 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17335291&pc=QCB | 2023826-2023-02917 | 2023-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+2.0/093 | OD | Physician | 06/20/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17335292&pc=QCB | 2023826-2023-02915 | 2023-02-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/080 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) (B)(4). | -9.5/+1.0/080 | OD | Physician | 06/19/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17335293&pc=QCB | 2023826-2023-02910 | 2022-11-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0/+1.0/044 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -15.0/+1.0/044 | OD | Physician | 06/19/2023 | 2023 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17335296&pc=QCB | 2023826-2023-02907 | 2023-04-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/+0.5/086 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -6.0/+0.5/086 | OS | Physician | 06/19/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17335298&pc=QCB | 2023826-2023-02909 | 2023-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.0/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.0/+1.0/090 | OS | Physician | 06/19/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17335303&pc=QCB | 2023826-2023-02914 | 2023-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/089 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.0/+1.5/089 | OD | Physician | 06/19/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17335305&pc=QCB | 2023826-2023-02903 | 2023-02-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.0/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -7.0/+1.0/092 | OD | Physician | 06/18/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17335307&pc=QCB | 2023826-2023-02902 | 2023-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.0/068 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -8.5/+1.0/068 | OS | Physician | 06/18/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17335311&pc=QCB | 2023826-2023-02906 | 2023-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: BC THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+2.0/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT. ON THAT THE SAME DAY A SHORTER LENGTH LENS WAS IMPLANTED ALTHOUGH IT IS UNKNOWN IF THIS WAS PERFORMED WITHIN THE SAME SURGERY THE LENS WAS REMOVED. THE PROBLEM WAS RESOLVED. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM #: (B)(4). | -14.0/+2.0/090 | OS | Physician | 06/19/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17335316&pc=QCB | 2023826-2023-02911 | 2023-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+3.5/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.0/+3.5/089 | OS | Physician | 06/19/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17335318&pc=QCB | 2023826-2023-02913 | 2023-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+2.5/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS TWO SIZES SHORTER IN LENGTH DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -14.0/+2.5/094 | OD | Physician | 06/19/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17335323&pc=QCB | 2023826-2023-02919 | 2023-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+1.5/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS TWO SIZES SHORTER IN LENGTH DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THE EVENT. Manufacturer Narrative: (B)(4). | -7.5/+1.5/090 | OD | Physician | 06/21/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17337292&pc=QCB | 2023826-2023-02933 | 2023-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.0/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE LENS WAS EXCHANGED ON (B)(6)2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#(B)(4). | -10.5/+1.0/094 | OD | Physician | 06/25/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17337299&pc=QCB | 2023826-2023-02940 | 2023-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -9.0/+1.0/091 | OD | Physician | 06/27/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17337308&pc=QCB | 2023826-2023-02951 | 2023-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+3.0/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -9.0/+3.0/092 | OD | Physician | 06/29/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17338888&pc=QCB | 2023826-2023-02981 | 2023-05-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -16.0/+3.0/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS ROTATION. ON (B)(6) 2023 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT MODEL/POWER LENS AND THE PROBLEM WAS RESOLVED. IT WAS ADDITIONALLY NOTED THERE IS A PLANNED LRI FOR ASTIGMATISM. Manufacturer Narrative: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -16.0/+3.0/90 | OD | Physician | 07/12/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17339602&pc=QCB | 2023826-2023-02964 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A REPLACEMENT 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -11.0/1.0/114 (SPHERE/CYLINDER/AXIS) TORE DURING INJECTION. Manufacturer Narrative: A4-A6: UNK. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS. CLAIM# (B)(4). | -11.0/1.0/114 | NA | Physician | 06/29/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17339604&pc=QCB | 2023826-2023-02963 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -11.50/1.0/094 (SPHERE/CYCLINDER/AXIS) TORE DURING INJECTION. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS. (B)(4). | -11.50/1.0/094 | NA | Physician | 06/29/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17340553&pc=QCB | 2023826-2023-02934 | 2023-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.5/156 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/+1.5/156 | OS | Physician | 06/26/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17340554&pc=QCB | 2023826-2023-02932 | 2023-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+4.0/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE LENS WAS EXCHANGED ON (B)(6)2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -7.0/+4.0/090 | OS | Physician | 06/25/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17340555&pc=QCB | 2023826-2023-02936 | 2023-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+1.0/089 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -6.5/+1.0/089 | OD | Physician | 06/27/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17340559&pc=QCB | 2023826-2023-02941 | 2023-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.0/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE LENS WAS EXCHANGED ON (B)(6)2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -7.0/+1.0/090 | OD | Physician | 06/27/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17340561&pc=QCB | 2023826-2023-02943 | 2023-01-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+1.0/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE LENS WAS EXCHANGED ON (B)(6)2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -8.0/+1.0/093 | OS | Physician | 06/27/2023 | 2023 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17340568&pc=QCB | 2023826-2023-02935 | 2023-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/+1.0/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -6.0/+1.0/090 | OS | Physician | 06/27/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17340569&pc=QCB | 2023826-2023-02948 | 2023-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+3.0/081 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -7.5/+3.0/081 | OD | Physician | 06/27/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17342517&pc=QCB | 2023826-2023-02879 | 2023-06-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.00/+2.0/065 (SPHERE/CYLINDER/AXIS). AFTER THE LENS WAS INSERTED, THE SURGEON NOTED THERE WAS A HOLE IN THE ICL LENS. THE LENS WAS REMOVED AND THE BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #: (B)(4) | -17.00/+2.0/065 | NA | Physician | 06/27/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17342626&pc=QCB | 2023826-2023-02880 | 2023-04-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+5.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2023 BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H6 DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AND HAPTIC BROKEN. DIMENSIONAL INSPECTION UNABLE TO BE PERFORMED DUE TO SIGNIFICANT LENS PART MISSING. CLAIM # (B)(4). | -09.00/+5.5/089 | OS | Physician | 06/27/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17342664&pc=QCB | 2023826-2023-02883 | 2023-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50/+2.5/115 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS NOT RESOLVED. SEE MFR. REPORT # 2023826-2023-02884 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -07.50/+2.5/115 | OS | Physician | 06/28/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17342673&pc=QCB | 2023826-2023-02882 | 2023-03-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+0.5/116 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.5/+0.5/116 (SPHERE/CYLINDER/AXIS) INTO THE RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED A LOW VAULT AND REFRACTIVE SURPRISE. THE LENS REMAINS IMPLANTED. REPORTEDLY "THE PATIENT HAS DECIDED NOT TO PROCEED WITH ANY FURTHER INTERVENTION". CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -16.50/+0.5/116 | OD | Physician | 06/27/2023 | 2023 | 45 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17342691&pc=QCB | 2023826-2023-02888 | 2023-01-18 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -04.50/+0.5/109 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -04.50/+0.5/109 | OD | Physician | 06/29/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17342721&pc=QCB | 2023826-2023-02886 | 2023-05-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.00/+2.0/090 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS CRUMBLED UP. (B)(4). Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -07.00/+2.0/090 | NA | Physician | 06/29/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17344440&pc=QCB | 2023826-2023-02851 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.00/4.0/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS FLIPPED ON INSERTION INTO THE PATIENT'S LEFT EYE (OS). LENS WAS REMOVED AND BACK UP LENS IMPLANTED WITH NO PATIENT INJURY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2 - UNK. A3 - UNK. A4 - UNK. A5 - UNK. A6 - UNK. B3 - UNK. D6A - UNK. D6B - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTCIM5_ 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.00/4.0/101 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE. REPORTEDLY "LENS FLIPPED ON INSERTION, REMOVED AND REPLACED WITH NO PATIENT INJURY". THE LENS WAS IMPLANTED, REMOVED AND REPLACED INTRAOPERATIVELY WITH A BACK UP LENS. THERE WAS NO PATIENT INJURY. CLAIM# (B)(4). | -11.00/4.0/101 | OS | Physician | 06/26/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17344442&pc=QCB | 2023826-2023-02857 | 2023-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.50/3.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT AND SURGEON HAPPY. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. CLAIM# (B)(4). | -16.50/3.5/090 | OS | Physician | 06/28/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17344443&pc=QCB | 2023826-2023-02856 | 2023-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -13.0/3.0/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT AND SURGEON ARE HAPPY. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -13.0/3.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT AND SURGEON ARE HAPPY. CLAIM# (B)(4). | -13.0/3.0/105 | OD | Physician | 06/28/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17344444&pc=QCB | 2023826-2023-02854 | 2023-06-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICH12.1, -2.50/3.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK, A5 - UNK, A6 - UNK, H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | -2.50/3.5/085 | OD | Physician | 06/27/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17344445&pc=QCB | 2023826-2023-02855 | 2023-06-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICH12.1, -1.0/5.0/115 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK, A5 - UNK, A6 - UNK, H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -1.0/5.0/115 | OS | Physician | 06/27/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17344613&pc=QCB | 2023826-2023-02853 | 2023-06-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.50/3.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). THE LENS DID NOT UNFOLD CORRECTLY. LENS WAS IMPLANTED AND REMOVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIPPER -10.50/3.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. REPORTEDLY "THE LENS DID NOT UNFOLD CORRECTLY". THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. D6A. "(B)(6) 2023" SHOULD BE ADDED. D6B. "(B)(6) 2023" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.50/3.0/090 | OD | Physician | 06/26/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17344620&pc=QCB | 2023826-2023-02862 | 2023-05-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A VTICM5 12.6, -9.0/1.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE BEFORE/DURING LOADING. DAMAGE WAS NOTED BEFORE LOADING. A BACKUP LENS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: : B5- THE REPORTER INDICATED THE SURGEON NOTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.00/1.5/082 (SPHERE/CYLINDER/AXIS) TORE/ BROKE BEFORE LOADING. THE DAMAGE OCCURRED AFTER OPENING THE VIAL. A BACKUP LENS WAS IMPLANTED INTO THE LEFT EYE (OS) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS THE DEVICE. CAUSE DETAILS: "LENS WAS DAMAGED". H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.0/1.5/082 | OS | Physician | 06/30/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17347235&pc=QCB | 2023826-2023-02985 | 2023-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -10.5/1.5/009 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATED THE INITIAL LENS WAS IMPLANTED IN VERTICAL POSITION AND THE REPLACEMENT LENS IN HORIZONTAL POSITION. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICORCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | -10.5/1.5/009 | OS | Physician | 06/19/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17347360&pc=QCB | 2023826-2023-03024 | 2023-06-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -6.50/1.5/073 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE LENS HAD TORN DURING INJECTION/DELIVERY. INTRAOPERATIVELY THE LENS WAS REMOVED WITHOUT PATIENT INJURY. ON (B)(6) 2023 A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -6.50/1.5/073 | OS | Physician | 06/28/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17347361&pc=QCB | 2023826-2023-03026 | 2023-05-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -13.0/4.5/159 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS ROTATION. ON (B)(6) 2023 THE SURGEON REPOSITIONED THE LENS BUT THIS DID NOT RESOLVE THE PROBLEM. THEN ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT DIFFERENT MODEL (SPHERICAL) AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IT WAS ADDITIONALLY LASIK WAS PERFORMED TO CORRECT THE ASTIGMATISM. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). | -13.0/4.5/159 | OD | Physician | 07/16/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17347765&pc=QCB | 2023826-2023-02988 | 2023-05-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 12.6MM, VTICM5_12.6, -10.00/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE RIGHT EYE (OD). THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY WITH NO PATIENT INJURY. ON (B)(6) 2023 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS DUE TO THE DEVICE. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.00/2.0/091 | OD | Physician | 06/21/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17347799&pc=QCB | 2023826-2023-02989 | 2023-04-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.5/1.5/074 (SPHERE/CYLINDER/AXIS) TORE/BROKE DURING INJECTION/DELIVERY INTO THE LEFT EYE (OS). THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY WITH NO PATIENT INJURY. ON (B)(6) 2023 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -9.5/1.5/074 | OS | Physician | 06/21/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17347841&pc=QCB | 2023826-2023-02991 | 2023-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -5.50/2.0/176 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. LENS ROTATION 2X WAS OBSERVED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -5.50/2.0/176 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023.ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE EXCHANGED RESOLVED THE PROBLEM. THE REPORTER STATES THE CAUSE OF THIS EVENT IS DUE TO PATIENT RELATED FACTORS. THE REPORTER ALSO STATES THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. CLAIM # (B)(4). | -5.50/2.0/176 | OS | Physician | 06/21/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17351613&pc=QCB | 2023826-2023-02993 |  | Injury | STAAR SURGICAL COMPANY | 2023-07-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -13.00/1.5/066 (SPHERE/CYLINDER/AXIS) WAS EXPLANTED FROM THE RIGHT EYE (OD) FOR AN UNKNOWN REASON. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.00/1.5/066 | OD | Physician | 06/22/2023 | 2023 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17351925&pc=QCB | 2023826-2023-02994 |  | Injury | STAAR SURGICAL COMPANY | 2023-07-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.00/1.5/113 (SPHERE/CYLINDER/AXIS) WAS EXPLANTED FROM THE LEFT EYE (OS) FOR AN UNKNOWN REASON. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: D6A: UNK. D6B: UNK. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.00/1.5/113 | OS | Physician | 06/22/2023 | 2023 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17358239&pc=QCB | 2023826-2023-02874 | 2022-12-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP: 23 MMHG WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -10.0/1.5/094 | OD | Physician | 07/11/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17358249&pc=QCB | 2023826-2023-02998 | 2023-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/2.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -6.0/2.0/093 | OD | Physician | 07/14/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17358253&pc=QCB | 2023826-2023-02873 | 2023-07-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/1.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6), 2023. ON (B)(6), 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #730476. | -6.0/1.0/092 | OD | Physician | 07/11/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17358254&pc=QCB | 2023826-2023-02921 | 2023-06-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/4.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -6.5/4.0/086 | OS | Physician | 07/13/2023 | 2023 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17358280&pc=QCB | 2023826-2023-02996 | 2023-01-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.5/1.0/015 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -15.5/1.0/015 | OD | Physician | 07/14/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17358297&pc=QCB | 2023826-2023-02997 | 2023-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/4.0/174 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -13.5/4.0/174 | OD | Physician | 07/14/2023 | 2023 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17358306&pc=QCB | 2023826-2023-02891 | 2023-04-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/2.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -7.0/2.5/090 | OS | Physician | 07/12/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17358307&pc=QCB | 2023826-2023-02923 | 2023-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/2.0/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -13.5/2.0/085 | OD | Physician | 07/13/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17358308&pc=QCB | 2023826-2023-02967 | 2023-05-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -7.0/3.0/089 (SPHERE/CYLINDER/AXIS) TORE DURING LOADING. REPORTEDLY, "LENS TEAR DUE TO CATCHING ON INJECTOR" ON (B)(6) 2023, A REPLACEMENT LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) AND THE PROBLEM WAS RESOLVED. REPORTEDLY, STATUS OF THE EYE IS, "NO PROBLEM." Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | -7.0/3.0/089 | OS | Physician | 06/30/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17358320&pc=QCB | 2023826-2023-02924 | 2023-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/1.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -13.0/1.5/088 | OS | Physician | 07/13/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17358330&pc=QCB | 2023826-2023-02922 | 2023-06-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/1.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -8.0/1.0/095 | OS | Physician | 07/13/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17358366&pc=QCB | 2023826-2023-02966 | 2023-05-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -10.50/1.0/093 (SPHERE/CYLINDER/AXIS) TORE DURING LOADING. REPORTEDLY, "LENS TEAR DUE TO CATCHING ON INJECTOR" ON (B)(6) 2023, A REPLACEMENT LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE(OS) AND THE PROBLEM WAS RESOLVED. REPORTEDLY, STATUS OF THE EYE IS, "NO PROBLEM." Manufacturer Narrative: A4-A6:UNKNOWN. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | -10.50/1.0/093 | OS | Physician | 06/30/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17358495&pc=QCB | 2023826-2023-02971 | 2023-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -14.0/1.5/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "REPLACE THE LENS AND PLACE IT IN A VERTICAL POSITION." Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -14.0/1.5/089 | OD | Physician | 07/03/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17358525&pc=QCB | 2023826-2023-02970 | 2023-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.0/1.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6)2023, THE LENS WAS EXCHANGED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #: (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | -10.0/1.5/092 | OS | Physician | 07/03/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17367766&pc=QCB | 2023826-2023-03001 | 2023-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "SECOND VERTICAL POSITION." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -12.5/2.5/091 | OS | Physician | 07/14/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17367767&pc=QCB | 2023826-2023-03000 | 2023-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "SECOND VERTICAL POSITION." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -9.0/2.0/092 | OD | Physician | 07/14/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17367768&pc=QCB | 2023826-2023-02999 | 2023-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/1.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -9.5/1.5/074 | OD | Physician | 07/14/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17367769&pc=QCB | 2023826-2023-03002 | 2023-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/3.0/094(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -9.5/3.0/094 | OD | Physician | 07/14/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17367771&pc=QCB | 2023826-2023-03004 | 2023-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/2.0/019(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -7.5/2.0/019 | OD | Physician | 07/14/2023 | 2023 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17367775&pc=QCB | 2023826-2023-03005 | 2023-05-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/3.5/089 (SPHERE / CYLINDER / AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -10.5/3.5/089 | OS | Physician | 07/14/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17367776&pc=QCB | 2023826-2023-03006 | 2023-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/1.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -14.0/1.0/094 | OS | Physician | 07/17/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17372785&pc=QCB | 2023826-2023-02954 | 2023-07-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00/+1.0/173 (SPHERE/CYLINDER/AXIS), WAS TORN WHILE BEING LOADED. THERE WAS NO PATIENT CONTACT. THE CAUSE WAS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -15.00/+1.0/173 | NA | Physician | 07/05/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17372896&pc=QCB | 2023826-2023-02884 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50/+2.5/085 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. LENS ROTATION WAS NOTED AND THE LENS REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-02883. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THE ISSUE OF LENS ROTATION WAS RESOLVED. CLAIM # (B)(4). | -09.50/+2.5/085 | OS | Physician | 06/28/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17373013&pc=QCB | 2023826-2023-03020 | 2023-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -01.50/+3.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2023. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT POWER LENS AND THE PROBLEM WAS RESOLVED. PATIENT IS HAPPY. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. D6B:UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -01.50/+3.0/090 | OD | Physician | 12/04/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17373033&pc=QCB | 2023826-2023-02956 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -17.50/+2.0/064 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON NOTED THERE WERE NO TORIC MARKERS ON THE LENS AND THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER, INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. B3: UNK. D6B: NA. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -17.50/+2.0/064 | OS | Physician | 07/03/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17373161&pc=QCB | 2023826-2023-02959 | 2023-06-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO IMPLANT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -05.50/+1.0/074, INTO THE PATIENTS LEFT EYE (OS) BUT THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT BUT NO INJURY. THE LENS WAS REPLACED WITH A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -05.50/+1.0/074 | OS | Physician | 07/05/2023 | 2023 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17377769&pc=QCB | 2023826-2023-02955 | 2021-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/112 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS REPLACED WITH A DIFFERENT MODEL AND LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. CLAIM # (B)(4). | -09.50/+1.0/112 | OS | Physician | 07/03/2023 | 2023 | 46 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17378247&pc=QCB | 2023826-2023-02957 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -18.00/+1.5/089 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON NOTED THERE WERE NO TORIC MARKERS ON THE LENS AND THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. B3: UNK. D6B: NA. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER, INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). | -18.00/+1.5/089 | OD | Physician | 07/03/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17378430&pc=QCB | 2023826-2023-02961 | 2021-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.00/+1.5/106 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A DIFFERENT MODEL/DIOPTER POWER BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.00/+1.5/106 | OD | Physician | 07/05/2023 | 2023 | 45 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17378646&pc=QCB | 2023826-2023-03118 | 2023-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+2.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS FIBER ON THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.50/+2.0/089 | OD | Physician | 07/13/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17378794&pc=QCB | 2023826-2023-03117 | 2023-05-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -08.50/+2.5/008 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK . H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -08.50/+2.5/008 | OS | Physician | 07/14/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17378942&pc=QCB | 2023826-2023-02979 | 2022-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/135 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "NO SIGNIFICANTLY FINDINGS." Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL - ACD<3.0. CLAIM# (B)(4). | -9.0/1.0/135 | OD | Physician | 07/06/2023 | 2023 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17379040&pc=QCB | 2023826-2023-03111 | 2023-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+2.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.50/+2.5/090 | OS | Physician | 07/11/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17379041&pc=QCB | 2023826-2023-03021 | 2023-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -02.50/+4.0/102 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REMOVED ON (B)(6) 2023. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT POWER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM #(B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -02.50/+4.0/102 | OS | Physician | 12/21/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17379065&pc=QCB | 2023826-2023-03102 | 2023-05-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -08.50/+1.5/148 (SPHERE / CYLINDER / AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS BUT THE PROBLEM WAS NOT RESOLVED. SEE MFR. REPORT#: 2023826-2023-03103 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND BENT. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). | -08.50/+1.5/148 | OS | Physician | 07/09/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17379100&pc=QCB | 2023826-2023-03104 | 2023-06-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -03.00/+0.5/156 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS BUT THE PROBLEM WAS NOT RESOLVED. SEE MFR. REPORT # 2023826-2023-03105 FOR REPLACEMENT LENS. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK . H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). | -03.00/+0.5/156 | NA | Physician | 07/10/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17379222&pc=QCB | 2023826-2023-02978 | 2023-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.50/1.0/058 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS INTRAOPERATIVELY IMPLANTED AND REMOVED DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.50/1.0/058 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS INTRAOPERATIVELY IMPLANTED AND REMOVED DUE TO THE LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT. NO INJURY REPORTED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. CAUSE WAS REPORTED AS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR LOTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -10.50/1.0/058 | OS | Physician | 10/04/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17379325&pc=QCB | 2023826-2023-03100 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.00/+2.0/089 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO THE LENS FLIPPED INSIDE THE PATIENTS RIGHT EYE (OD). THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2: UNK, A3: UNK, A4: UNK, A5: UNK, A6: UNK, B3: UNK, D6A: UNK, D6B: UNK, H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -08.00/+2.0/089 | OD | Physician | 07/13/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17379791&pc=QCB | 2023826-2023-03098 | 2023-07-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -8.0/1.0/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUDN NO VISBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -8.0/1.0/088 | OD | Physician | 07/17/2023 | 2023 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17381108&pc=QCB | 2023826-2023-03058 | 2023-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/+1.5/103 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LOW VAULT. THAT SAME DAY, A LONGER LENGTH LENS WAS IMPLANTED, ALTHOUGH IT IS UNCLEAR WHETHER THIS WAS PERFORMED WITHIN THE SAME SURGERY THAT THE LENS WAS REMOVED. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/+1.5/103 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. IT WAS CONFIRMED THAT THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- UPDATED TO REFLECT AN EXCHANGE . H6 HEALTH EFFECT - IMPACT CODE: CHANGED TO 4629. CLAIM# (B)(4). | -6.0/+1.5/103 | OS | Physician | 06/30/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17381139&pc=QCB | 2023826-2023-03063 | 2023-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+1.0/106 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.5/+1.0/106 | OD | Physician | 06/30/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17381141&pc=QCB | 2023826-2023-03060 | 2023-04-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS WITH DIOPTER -3.0/+1.0/011 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -3.0/+1.0/011 | OS | Physician | 06/30/2023 | 2023 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17381144&pc=QCB | 2023826-2023-03061 | 2023-04-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS WITH DIOPTER -7.0/+1.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -7.0/+1.0/093 | OD | Physician | 06/30/2023 | 2023 | 47 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17381159&pc=QCB | 2023826-2023-03059 | 2023-04-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS WITH DIOPTER -4.0/+2.0/172 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -4.0/+2.0/172 | OD | Physician | 06/30/2023 | 2023 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17381185&pc=QCB | 2023826-2023-03062 | 2023-04-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS WITH DIOPTER -9.0/+2.0/088 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -9.0/+2.0/088 | OS | Physician | 06/30/2023 | 2023 | 47 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17386087&pc=QCB | 2023826-2023-03068 | 2023-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.0/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) AND (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -10.5/+2.0/092 | OD | Physician | 07/02/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17386130&pc=QCB | 2023826-2023-03065 | 2023-06-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS WITH DIOPTER -9.0/+2.0/095 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.0/+2.0/095 | OS | Physician | 06/30/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17386139&pc=QCB | 2023826-2023-03070 | 2023-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/095 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.0/+1.5/095 | OS | Physician | 07/02/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17386154&pc=QCB | 2023826-2023-03071 | 2023-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.5/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -7.0/+1.5/092 | OD | Physician | 07/02/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17386175&pc=QCB | 2023826-2023-03066 | 2023-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.5/089 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -10.5/+1.5/089 | OD | Physician | 06/30/2023 | 2023 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17386196&pc=QCB | 2023826-2023-03073 | 2023-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.5/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/+1.5/091 | OD | Physician | 07/02/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17386200&pc=QCB | 2023826-2023-03113 | 2023-07-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.00/+1.0/175 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD ON (B)(6)2023. THE SURGEON OBSERVED THREE POSITIONING MARKS ON THE LENS FOOTPLATE. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK D6B: NA H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -07.00/+1.0/175 | OD | Physician | 07/12/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17386207&pc=QCB | 2023826-2023-03069 | 2023-06-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.0/075 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND RESIDUE THROUGHOUT THE LENS. CLAIM# (B)(4). Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -11.5/+1.0/075 | OS | Physician | 07/02/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17386298&pc=QCB | 2023826-2023-03109 | 2023-06-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.5/069 (SPHERE/CYLINDER/AXIS). DURING THE SAME SURGERY DUE TO THE LENS BECAME STUCK IN THE CARTRIDGE AND WAS TORN/BROKEN DURING INJECTION/DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AT A LATER DATE AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK, A5: UNK, A6: UNK, H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -09.50/+1.5/069 | OS | Physician | 07/12/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17386301&pc=QCB | 2023826-2023-03119 | 2023-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+2.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: A4: UNK . A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.50/+2.0/089 | OS | Physician | 07/13/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17386320&pc=QCB | 2023826-2023-03110 | 2023-06-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/091 (SPHERE/CYLINDER/AXIS) DURING THE SAME SURGERY DUE TO THE LENS WAS IMPLANTED UPSIDE DOWN IN THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS DAMAGED WHEN REMOVED. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AT A LATER DATE AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.00/+1.5/091 | OD | Physician | 07/12/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17386350&pc=QCB | 2023826-2023-03103 | 2023-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50/+1.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (THE LENS WAS PLACED HORIZONTALLY TO OBLIQUELY) AND THE PROBLEM WAS RESOLVED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-03102. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -08.50/+1.5/089 | OS | Physician | 07/09/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17386362&pc=QCB | 2023826-2023-03105 | 2023-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -03.00/+0.5/167 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-03104. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. PRESSURE AND ARCH HIGH IS GOOD. Manufacturer Narrative: A4: UNK. A5: UNK . A6: UNK . H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B3: DATE OF EVENT: 23-JUN-2023. B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, 03.00/+0.5/167 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-03104. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. PRESSURE AND ARCH HEIGHT WERE GOOD. THE CAUSE OF THE EVENT WAS UNKNOWN. CLAIM # (B)(4). | -03.00/+0.5/167 | OD | Physician | 07/10/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17386559&pc=QCB | 2023826-2023-03142 |  | Injury | STAAR SURGICAL COMPANY | 2023-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient Device Interaction Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A TORIC EVO IMPLANTABLE COLLAMER LENS WAS EXPLANTED PER PATIENTS REQUEST. PATIENT WAS DISSATISFIED WITH VISION. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2 VTICM5_13.2; -14.00/1.5/116 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS EXPLANTED ON (B)(6) 2023 PER PATIENTS REQUEST. PATIENT WAS DISSATISFIED WITH VISION. D4: SERIAL# (B)(6). CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -14.00/1.5/116 | NA | Physician | 06/27/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17386560&pc=QCB | 2023826-2023-03143 |  | Injury | STAAR SURGICAL COMPANY | 2023-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient Device Interaction Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A TORIC EVO IMPLANTABLE COLLAMER LENS WAS EXPLANTED PER PATIENTS REQUEST. PATIENT WAS DISSATISFIED WITH VISION. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFO: B5: THE REPORTER INDICATED THAT A 13.2 VTICM5_13.2; -14.00/2.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS EXPLANTED ON (B)(6) 2023 PER PATIENTS REQUEST. PATIENT WAS DISSATISFIED WITH VISION. D4: SN#: (B)(6). H6: 4110: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS FOUND. CLAIM#: (B)(6). | -14.00/2.5/090 | NA | Physician | 06/27/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17386571&pc=QCB | 2023826-2023-03141 |  | Injury | STAAR SURGICAL COMPANY | 2023-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -13.50/1.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AND REMOVED FOR AN UNKNOWN REASON. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: D4- "VTICM5_13.2" SHOULD BE REMOVED AND "VTICM5_12.6" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -13.50/1.0/090 | NA | Other Health Care Professional | 06/27/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17386578&pc=QCB | 2023826-2023-03072 | 2023-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.0/+1.5/090 | OS | Physician | 07/02/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17391255&pc=QCB | 2023826-2023-03034 | 2023-05-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO_13.2, -17.50/3.5/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023.THE LENS TORE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED AND REMOVED. CAUSE OF THE EVENT IS REPORTED AS DEVICE, MAY BE DUE TO WRONG LENS LOADING OR INJECTION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.5/+3.5/73 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. REPORTEDLY THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. ON THE SAME DAY SEPARATE SURGERY, A REPLACEMENT LENS OF THE SAME PARAMETERS WAS IMPLANTED. THE CAUSE OF THE EVENT WAS THE DEVICE. CAUSE DETAILS PROVIDED: "THE LENS GOT TEARED AFTER INJECTION. HAD TO EXPLANT AND EXCHANGE WITH A NEW ONE". CLAIM# (B)(4). | -17.50/3.5/073 | OS | Physician | 07/05/2023 | 2023 | 43 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17391256&pc=QCB | 2023826-2023-03029 | 2023-05-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -14.0/2.0/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS FROM THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT . A SHORTER LENGTH LENS WAS IMPLANTED ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, AS EXPECTED EYE IS WELL AND IOP IS STABLE WITH LOWER VAULT. Manufacturer Narrative: PT INFO: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -14.0/2.0/102 | OD | Physician | 07/02/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17391258&pc=QCB | 2023826-2023-03097 | 2022-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.50/4.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS ROTATION WAS OBSERVED. LENS WAS ROTATED TWICE AND THIS DID NOT RESOLVE THE PROBLEM. LENS EXPLANTED ON (B)(6) 2023. A LONGER LENGTH LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2- UNK. A3- UNK. A4 - UNK. A5 - UNK. A6 - UNK. B3- UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - LENS ROTATION PERFORMED ON (B)(6) 2022 AND (B)(6) 2023. LENS EXPLANTED ON (B)(6) 2023. A NEW LONGER LENGTH LENS WAS IMPLANTED IN A SECOND SURGERY ON (B)(6) 2023. CLAIM# (B)(4). | -9.50/4.0/095 | OD | Physician | 08/08/2023 | 2023 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17391259&pc=QCB | 2023826-2023-03030 | 2023-04-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -09.5/3.0/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING INJECTION/DELIVERY INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. ON (B)(6) 2023, A REPLACEMENT LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: D9 - RETURNED TO MANUFACTURER DATE CORRECTED TO 12-JUL-2023 IN INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -09.5/3.0/102 | OS | Physician | 07/03/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17400412&pc=QCB | 2023826-2023-03147 | 2023-06-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.00/1.5/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) CLAIM# (B)(4). | -10.00/1.5/094 | OS | Physician | 06/30/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17401848&pc=QCB | 2023826-2023-03131 | 2023-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT AN 13.7 LENGTH IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). A POSSIBLE LENS EXCHANGE FOR A 13.2 LENGTH LENS MAY BE PLANNED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 OF A -15.0/1.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS. THE PROBLEM WAS RESOLVED. REPORTEDLY, "PT NOTES LESS ABBERRATIONS, SORE, STILL BLURRY BUT BETTER" H6: OFF-LABEL - ACD<3.0. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. B3: UNK. D4 (EXPERIATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM#: (B)(4). | -15.0/1.0/089 | OS | Physician | 07/10/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17401849&pc=QCB | 2023826-2023-03130 | 2023-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT AN 13.7 LENGTH IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). A POSSIBLE LENS EXCHANGE FOR A 13.2 LENGTH LENS MAY BE PLANNED. Manufacturer Narrative: B3: UNK. D4 (EXPIRATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK. H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM(B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 OF A -15.0/2.0/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS. THE PROBLEM WAS RESOLVED. REPORTEDLY, "PT NOTES IMPROVED VISION AND DECREASE IN ABBERATIONS." H6: OFF-LABEL - ACD<3.0. CLAIM# (B)(4). | -15.0/2.0/92 | OD | Physician | 07/10/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410534&pc=QCB | 2023826-2023-03009 | 2023-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/2.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -6.5/2.5/092 | OD | Physician | 07/17/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410536&pc=QCB | 2023826-2023-03008 | 2023-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/2.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) CLAIM #(B)(4). | -11.5/2.0/093 | OD | Physician | 07/17/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410537&pc=QCB | 2023826-2023-03011 | 2023-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/5.0/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4).. | -14.0/5.0/080 | OD | Physician | 07/17/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410538&pc=QCB | 2023826-2023-03012 | 2023-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/3.0/145(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. (B)(4). | -13.5/3.0/145 | OS | Physician | 07/16/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410539&pc=QCB | 2023826-2023-03010 | 2023-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/1.5/092 (SPHERE / CYLINDER / AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -9.5/1.5/092 | OS | Physician | 07/17/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410540&pc=QCB | 2023826-2023-03049 | 2023-05-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #: (B)(4). | -10.0/1.0/087 | OD | Physician | 07/18/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410541&pc=QCB | 2023826-2023-03015 | 2023-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/3.0/169 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. (B)(4). | -12.0/3.0/169 | OS | Physician | 07/16/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410542&pc=QCB | 2023826-2023-03013 | 2023-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/3.5/170 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #:(B)(4). | -14.5/3.5/170 | OD | Physician | 07/16/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410543&pc=QCB | 2023826-2023-03014 | 2023-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/2.5/158 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. (B)(4). | -10.0/2.5/158 | OD | Physician | 07/16/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410548&pc=QCB | 2023826-2023-03055 | 2023-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/2.0/110(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -13.0/2.0/110 | OD | Physician | 07/17/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410549&pc=QCB | 2023826-2023-03051 | 2023-07-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.5/148 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "TWO PIECES OF ICL ULTRA AXIAL, WITHOUT IOD DIAGRAM." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -10.0/1.5/148 | OS | Physician | 07/17/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410550&pc=QCB | 2023826-2023-03092 | 2022-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/3.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -13.0/3.0/095 | OS | Physician | 07/19/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410553&pc=QCB | 2023826-2023-03095 | 2023-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -10.0/1.5/091 | OS | Physician | 07/19/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410555&pc=QCB | 2023826-2023-03054 | 2023-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #:(B)(4). | -9.0/1.0/091 | OD | Physician | 07/17/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410556&pc=QCB | 2023826-2023-03056 | 2023-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/3.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "THE ARCH IS LOW." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -12.5/3.5/091 | OS | Physician | 07/17/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410558&pc=QCB | 2023826-2023-03159 | 2023-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.0/1.0/091 (SPHERE / CYLINDER / AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | -5.0/1.0/091 | OD | Physician | 07/21/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410560&pc=QCB | 2023826-2023-03164 | 2023-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/2.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #: (B)(4). | -12.0/2.0/094 | OS | Physician | 07/21/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410562&pc=QCB | 2023826-2023-03158 | 2023-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/2.0/072(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -9.0/2.0/072 | OD | Physician | 07/20/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410563&pc=QCB | 2023826-2023-03096 | 2023-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/1.5/106 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -11.5/1.5/106 | OD | Physician | 07/19/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410565&pc=QCB | 2023826-2023-03162 | 2023-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/3.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM (B)(4). | -9.5/3.0/094 | OD | Physician | 07/21/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410566&pc=QCB | 2023826-2023-03168 | 2023-01-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/0.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -10.0/0.5/091 | OS | Physician | 07/21/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410568&pc=QCB | 2023826-2023-03169 | 2023-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/3.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -14.0/3.0/088 | OS | Physician | 07/24/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410570&pc=QCB | 2023826-2023-03163 | 2023-06-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/2.0/096 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #:(B)(4). | -10.0/2.0/096 | OS | Physician | 07/20/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410571&pc=QCB | 2023826-2023-03165 | 2023-07-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/1.0/100 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -8.0/1.0/100 | OS | Physician | 07/21/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410574&pc=QCB | 2023826-2023-03170 | 2023-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.0/3.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #: (B)(4). | -16.0/3.0/089 | OD | Physician | 07/24/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410578&pc=QCB | 2023826-2023-03167 | 2023-05-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/1.0/075 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM (B)(4). | -7.5/1.0/075 | OS | Physician | 07/21/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410579&pc=QCB | 2023826-2023-03172 | 2023-07-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/1.0/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING WITHOUT ROTATION. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -9.5/1.0/072 | OD | Physician | 07/22/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410583&pc=QCB | 2023826-2023-03174 | 2023-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_ 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/0.5/087 (SPHERE / CYLINDER / AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -6.5/0.5/087 | OS | Physician | 07/24/2023 | 2023 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410589&pc=QCB | 2023826-2023-03173 | 2023-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_ 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/0.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH/ DIFFERENT MODEL LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -8.0/0.5/095 | OD | Physician | 07/24/2023 | 2023 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17410610&pc=QCB | 2023826-2023-03182 | 2023-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.50/1.0/110 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -14.50/1.0/110 | OD | Physician | 07/25/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17417147&pc=QCB | 2023826-2023-03134 |  | Injury | STAAR SURGICAL COMPANY | 2023-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -16.0/3.0/172 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS A REPLACEMENT LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023 DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. REPORTEDLY, "THE PATIENT WENT TO THE HOSPITAL FOR PREOPERATIVE EXAMINATION OF ICL SURGERY ON (B)(6) 2023. THE EXAMINATION FOUND THAT THERE WAS A SIGNIFICANT DIFFERENCE BETWEEN COMPUTER OPTOMETRY AND COMPREHENSIVE OPTOMETRY, AND REPEATED COMPREHENSIVE OPTOMEOTRY RESULTS WERE CONSISTENT. ACCORDING TO THE COMPREHENSIVE OPTOMETRY RESULTS BEFORE SURGERY, THE VISUAL ACUITY COULD REACH NORMAL CORRECTED VISION. POSTOPERATIVE VISUAL ACUITY: 1.0 IN THE RIGHT EYE AND 0.6 IN THE LEFT EYE. COMPUTER OPTOMETRY SHOWED THAT BOTH EYES WERE UNDER CORRECTION FOR ASTIGMATISM. THREE MONTHS AFTER SURGERY, THE PATIENT REPORTED A DECREASE IN BINOCULAR VISION, AND REPEATED COMPUTER OPTOMETRY RESULTED IN UNDERCORRECTED MYOPIA AND ASTIGMASTISM IN BOTH EYES. THE SUFFICIENT DIOPTER WAS C.ORRECTED TO 1.0. AFTER COMMUNICATING WITH THE PATIENT, THE PATIENT UNDERWENT LENS REPLACEMENT SURGERY. BILATERAL TICL LENS REPLACEMENT SURGERY WAS PERFORMED ON (B)(6) 2023. POSTOPERATIVE VISION IMPROVED, WITH SLIGHT CORNEAL EDEMA. THE PATIENT WAS INSTRUCTED TO TAKE MEDICATION ON TIME AND UNDERGO REGULAR FOLLOW UP." Manufacturer Narrative: CORRECTED DATA: B1- CORRECTED REPORT TYPE FROM ADVERSE EVENT TO PRODUCT PROBLEM. B2-OMIT "REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE" FOR CORRECTION. B5-CORRECTED DIOPTER TO -16.50/3.0/172 (SPHERE/CYLINDER/AXIS). H1- CORRECTED TYPE OF REPORT EVENT FROM SERIOUS INJURY TO MALFUNCTION. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -16.0/3.0/172 | OD | Physician | 07/10/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17417149&pc=QCB | 2023826-2023-03135 |  | Injury | STAAR SURGICAL COMPANY | 2023-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -18.0/4.0/112 (SPHERE / CYLINDER / AXIS) WAS IMPLANTED AS A REPLACEMENT LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS NOT RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. REPORTEDLY, "THE PATIENT WENT TO THE HOSPITAL FOR PREOPERATIVE EXAMINATION OF ICL SURGERY ON (B)(6) 2023. THE EXAMINATION FOUND THAT THERE WAS A SIGNIFICANT DIFFERENCE BETWEEN COMPUTER OPTOMETRY AND COMPREHENSIVE OPTOMETRY, AND REPEATED COMPREHENSIVE OPTOMEOTRY RESULTS WERE CONSISTENT. ACCORDING TO THE COMPREHENSIVE OPTOMETRY RESULTS BEFORE SURGERY, THE VISUAL ACUITY COULD REACH NORMAL CORRECTED VISION. POSTOPERATIVE VISUAL ACUITY: 1.0 IN THE RIGHT EYE AND 0.6 IN THE LEFT EYE. COMPUTER OPTOMETRY SHOWED THAT BOTH EYES WERE UNDER CORRECTION FOR ASTIGMATISM. THREE MONTHS AFTER SURGERY, THE PATIENT REPORTED A DECREASE IN BINOCULAR VISION, AND REPEATED COMPUTER OPTOMETRY RESULTED IN UNDERCORRECTED MYOPIA AND ASTIGMASTISM IN BOTH EYES. THE SUFFICIENT DIOPTER WAS CORRECTED TO 1.0. AFTER COMMUNICATING WITH THE PATIENT, THE PATIENT UNDERWENT LENS REPLACEMENT SURGERY. BILATERAL TICL LENS REPLACEMENT SURGERY WAS PERFORMED ON (B)(6) 2023. POSTOPERATIVE VISION IMPROVED, WITH SLIGHT CORNEAL EDEMA. THE PATIENT WAS INSTRUCTED TO TAKE MEDICATION ON TIME AND UNDERGO REGULAR FOLLOW UP." Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B1- CORRECTED REPORT TYPE FROM ADVERSE EVENT TO PRODUCT PROBLEM. B2- OMIT "REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE" FOR CORRECTION. H1- CORRECTED TYPE OF REPORT EVENT FROM SERIOUS INJURY TO MALFUNCTION. CLAIM#(B)(4). | -18.0/4.0/112 | OS | Physician | 07/10/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17420572&pc=QCB | 2023826-2023-03050 | 2023-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/1.0/070 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -9.0/1.0/070 | OS | Physician | 07/18/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17421038&pc=QCB | 2023826-2023-03197 | 2023-07-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -18.00/4.0/085 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. IT WAS REPORTED THAT THE LENS FLIPPED DURING SURGERY WITH NO PATIENT INJURY. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -18.00/4.0/085 | NA | Physician | 07/10/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17432102&pc=QCB | 2023826-2023-03202 | 2023-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/1.5/011 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ADN SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -11.5/1.5/011 | OD | Physician | 07/28/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17434420&pc=QCB | 2023826-2023-03190 | 2023-02-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.5/1.5/069 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #: (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -11.5/1.5/069 | OD | Physician | 07/13/2023 | 2023 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17434424&pc=QCB | 2023826-2023-03191 | 2023-02-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.5/2.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2022, THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H6 - 4628 SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #: (B)(4). | -12.5/2.0/092 | OS | Physician | 07/13/2023 | 2023 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17435777&pc=QCB | 2023826-2023-03076 | 2023-03-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+3.0/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.5/+3.0/090 | OS | Physician | 07/06/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17438529&pc=QCB | 2023826-2023-03084 | 2023-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+2.5/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -14.0/+2.5/094 | OD | Physician | 07/06/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17438531&pc=QCB | 2023826-2023-03079 | 2023-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.5/+0.5/158 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -4.5/+0.5/158 | OD | Physician | 07/06/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17438532&pc=QCB | 2023826-2023-03082 | 2023-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.5/112 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LOW VAULT. THAT SAME DAY, THE LENS WAS REPLACED WITH A LONGER LENGTH LENS ALTHOUGH IT IS UNCLEAR WHETHER THIS WAS WITHIN THE SAME PROCEDURE. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - LENS EXCHANGED ON (B)(6) 2023 DUE TO LOW VAULT AND PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+2.5/112 | OD | Physician | 07/06/2023 | 2023 | 51 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17438533&pc=QCB | 2023826-2023-03078 | 2023-03-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/103 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+1.0/103 | OS | Physician | 07/06/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17438534&pc=QCB | 2023826-2023-03085 | 2023-04-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.50/+4.0/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B3- 20-APR-2023. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -12.50/+4.0/092 | OS | Physician | 07/06/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17438537&pc=QCB | 2023826-2023-03086 | 2023-04-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+4.5/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B3- 20-APR-2023. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -12.0/+4.5/092 | OD | Physician | 07/06/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17438538&pc=QCB | 2023826-2023-03083 | 2023-04-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+3.0/006 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+3.0/006 | OS | Physician | 07/06/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17438540&pc=QCB | 2023826-2023-03246 | 2023-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.0/075 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -7.0/+1.0/075 | OS | Physician | 07/07/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17438905&pc=QCB | 2023826-2023-03276 | 2023-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.5/1.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -4.5/1.0/092 | OS | Physician | 07/31/2023 | 2023 | 17 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17438906&pc=QCB | 2023826-2023-03277 | 2023-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/2.0/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "T985638 IMPLANTATION IN VERTICAL POSITION." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -11.5/2.0/074 | OD | Physician | 07/31/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17438907&pc=QCB | 2023826-2023-03275 | 2023-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/1.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -12.0/1.5/074 | OS | Physician | 07/30/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17438908&pc=QCB | 2023826-2023-03278 | 2023-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -6.5/1.0/090 | OD | Physician | 07/31/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17439057&pc=QCB | 2023826-2023-03208 | 2023-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -14.0/4.0/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS OTHER, AND THE DEVICE FAILING TO PERFORM AS INTENDED. REPORTEDLY, "UNDERSIZED ICL." STATUS OF THE EYE IS REPORTED AS, "REDUCED VISION AND HALO VISION (COMA)." Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -14.0/4.0/90 | OD | Physician | 07/12/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17439059&pc=QCB | 2023826-2023-03210 | 2023-07-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -11.5/3.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LENS ROTATION NOT ASSOCIATED TO LOW VAULT WAS OBSERVED. REPOSITIONING IS PLANNED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: ADDITONAL INFORMATION: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+3.0/91 (SPHERE/CYLINDER/AXIS) INTO THE RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED A LOW VAULT WITHOUT ROTATION. THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION PROVIDED: "PATIENT IS HAPPY, NO FURTHER TREATMENTS PLANNED". THE REPORTER INDICATED THE CAUSE OF THE EVENT IS UNKNOWN. H6- MEDICAL DEVICE PROBLEM CODE: "2923" SHOULD BE REMOVED AND "1583" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -11.5/3.0/091 | OD | Physician | 07/12/2023 | 2023 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17439060&pc=QCB | 2023826-2023-03209 | 2023-04-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.0/1.5/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS, IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT. REPORTEDLY, "SECOND VERTICAL POSITION." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED WITHIN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: ONE SIMILAR COMPLAINT WAS FOUND WITHIN ASSOCIATED LOTS. CLAIM# (B)(4). CLAIM# (B)(4). | -11.0/1.5/089 | OD | Physician | 07/11/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17439480&pc=QCB | 2023826-2023-03189 | 2021-04-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO13.2, -11.0/2.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2022. THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS, THIS RESOLVED THE PROBLEM. Manufacturer Narrative: B5 - VTICMO13.2 SHOULD BE CORRECTED TO VTICMO12.1 . H6 - HEALTH EFFECT - IMPACT CODE: 4628 SHOULD HAVE BEEN INCLUDED. H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.0/2.5/089 | OS | Physician | 07/13/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17440084&pc=QCB | 2023826-2023-03077 | 2023-03-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.5/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+1.5/090 | OD | Physician | 07/06/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17440423&pc=QCB | 2023826-2023-03245 | 2023-05-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.5/095 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+1.5/095 | OS | Physician | 07/07/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17444299&pc=QCB | 2023826-2023-03232 | 2023-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -9.0/2.0/087(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS THAT WAS IMPLANTED IN VERTICAL POSITION DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT UNKNOWN. Manufacturer Narrative: H6: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.0/2.0/087 | OD | Physician | 07/06/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17445926&pc=QCB | 2023826-2023-03234 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS -9.0/4.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT UNKNOWN. Manufacturer Narrative: (B)(4). | -9.0/4.0/092 | OD | Physician | 07/07/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17446076&pc=QCB | 2023826-2023-03214 | 2023-06-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -8.0/1.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED, AND REPLACED DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS. CLAIM#: (B)(4). | -8.0/1.5/092 | OS | Physician | 07/13/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17446529&pc=QCB | 2023826-2023-03283 | 2023-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/1.5/032 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -11.5/1.5/032 | OD | Physician | 07/31/2023 | 2023 | 47 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17446530&pc=QCB | 2023826-2023-03282 | 2023-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/1.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -11.5/1.5/095 | OD | Physician | 07/31/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17446534&pc=QCB | 2023826-2023-03290 | 2022-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_ 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. (B)(4). | -11.5/2.5/091 | OS | Physician | 08/02/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17446536&pc=QCB | 2023826-2023-03289 | 2023-05-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_ 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/2.0/137 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. (B)(4). | -11.0/2.0/137 | OS | Physician | 08/02/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17446541&pc=QCB | 2023826-2023-03288 | 2022-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.5/2.0/055 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -15.5/2.0/055 | OD | Physician | 08/02/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17446542&pc=QCB | 2023826-2023-03284 | 2023-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/1.0/150 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -11.0/1.0/150 | OS | Physician | 07/31/2023 | 2023 | 47 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17446632&pc=QCB | 2023826-2023-03279 | 2023-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/1.5/107 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON(B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -6.5/1.5/107 | OS | Physician | 07/31/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17446690&pc=QCB | 2023826-2023-03235 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -9.5/2/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS INTRAOPERATIVELY REMOVED DUE TO LOW VAULT. A REPLACEMENT LENS OF LONGER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/2/091 | OD | Physician | 07/07/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17447291&pc=QCB | 2023826-2023-03216 | 2023-07-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -9.5/1.5/084 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. REPORTEDLY, "THE ICL SIZE MAYBE OVER-ESTIMATE PARTICULARLY WITH BORDERLINE WTW MEASUREMENT." Manufacturer Narrative: CLAIM# (B)(4). | -9.5/1.5/084 | OS | Physician | 07/14/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17447293&pc=QCB | 2023826-2023-03215 | 2023-07-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.0/1.5/111 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS OTHER. REPORTEDLY, "THE ICL SIZE MAYBE OVER-ESTIMATE PARTICULARLY WITH BORDERLINE WTW MEASUREMENT." Manufacturer Narrative: CLAIM# (B)(4). | -10.0/1.5/111 | OD | Physician | 07/14/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17453898&pc=QCB | 2023826-2023-03247 | 2023-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+2.5/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS TWO SIZES SHORTER IN LENGTH DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -12.0/+2.5/091 | OD | Physician | 07/10/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17453900&pc=QCB | 2023826-2023-03251 | 2023-05-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+3.0/084 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: (B)(4). | -11.5/+3.0/084 | OD | Physician | 07/11/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17453902&pc=QCB | 2023826-2023-03250 | 2023-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+0.5/119 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -8.5/+0.5/119 | OS | Physician | 07/10/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17453903&pc=QCB | 2023826-2023-03252 | 2023-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+4.0/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: (B)(4). | -9.0/+4.0/091 | OS | Physician | 07/11/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17453907&pc=QCB | 2023826-2023-03248 | 2023-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+3.0/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -8.0/+3.0/090 | OD | Physician | 07/10/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17453908&pc=QCB | 2023826-2023-03254 | 2023-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+2.5/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: (B)(4). | -13.0/+2.5/095 | OD | Physician | 07/11/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17453912&pc=QCB | 2023826-2023-03258 | 2023-04-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+1.0/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -6.5/+1.0/090 | OS | Physician | 07/11/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17453913&pc=QCB | 2023826-2023-03299 | 2023-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/089 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -10.0/+1.5/089 | OD | Physician | 07/13/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17453914&pc=QCB | 2023826-2023-03255 | 2023-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.0/007 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -10.5/+2.0/007 | OD | Physician | 07/11/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17453915&pc=QCB | 2023826-2023-03295 | 2023-02-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.5/+0.5/029 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -4.5/+0.5/029 | OS | Physician | 07/13/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17453924&pc=QCB | 2023826-2023-03253 | 2023-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/089 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) (B)(4). | -10.0/+1.5/089 | OD | Physician | 07/11/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17453936&pc=QCB | 2023826-2023-03303 | 2023-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.5/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -11.5/+1.5/091 | OS | Physician | 07/13/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17453937&pc=QCB | 2023826-2023-03305 | 2023-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+3.0/095 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -12.5/+3.0/095 | OS | Physician | 07/13/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17453950&pc=QCB | 2023826-2023-03298 | 2023-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.0/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -7.0/+1.0/090 | OS | Physician | 07/13/2023 | 2023 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17454908&pc=QCB | 2023826-2023-03256 | 2023-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.50/+3.0/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.50/+3.0/095 | OD | Physician | 07/11/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17454909&pc=QCB | 2023826-2023-03294 | 2023-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+1.0/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. IN A SEPARATE SURGERY THAT DAY THE LENS WAS REMOVED DUE TO LOW VAULT. ON (B)(6) 2023 A LONGER LENGTH LENS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -6.5/+1.0/094 | OS | Physician | 07/13/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17454910&pc=QCB | 2023826-2023-03300 | 2023-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/088 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+1.0/088 | OS | Physician | 07/13/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17454911&pc=QCB | 2023826-2023-03302 | 2023-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+2.0/074 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.0/+2.0/074 | OS | Physician | 07/13/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17454912&pc=QCB | 2023826-2023-03297 | 2023-04-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS WITH DIOPTER -9.0/+3.5/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/+3.5/092 | OS | Physician | 07/13/2023 | 2023 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17454913&pc=QCB | 2023826-2023-03306 | 2023-01-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+1.0/060 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -14.0/+1.0/060 | OD | Physician | 07/13/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17454915&pc=QCB | 2023826-2023-03304 | 2023-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.5/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+2.5/090 | OD | Physician | 07/13/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17455345&pc=QCB | 2023826-2023-03309 | 2023-07-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -6.5/2.5/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: D9: 19-SEP-2023 H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.5/2.5/087 | OS | Physician | 08/01/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17455346&pc=QCB | 2023826-2023-03307 | 2023-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.0/4.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2023 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. THE LENS WAS EXCHANGED ON (B)(6) 2023 WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FIBRE ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.0/4.5/090 | OS | Physician | 07/31/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17455347&pc=QCB | 2023826-2023-03308 | 2023-06-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -11.0/1.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE SURGEON NOTES THE LENS HAD TORE DURING LOADING. INTRAOPERATIVELY THE LENS WAS REMOVED. ON (B)(6) 2023 A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE. EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.5/90 (SPHERE/CYLINDER/AXIS) INTO THE RIGHT EYE (OD) ON (B)(6) 2023. REPORTEDLY THE LENS TORE/BROKE DURING LOADING. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. ON 21/JUNE/2023 A REPLACEMENT LENS OF THE SAME PARAMETERS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). | -11.0/1.5/090 | OD | Physician | 07/31/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17455640&pc=QCB | 2023826-2023-03223 | 2023-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -11.0/1.0/102 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR WITH THE DEVICE NOTED A FAILED AS INTENDED. REPORTEDLY, "DESPITE DETAILED INFORMATION BEFOREHAND, THE PATIENT COULD NOT COPE WITH THE POST-OP PRESBYOPIC LIMITATIONS AND DECIDED TO WITHDRAW FROM SURGERY, WHEREUPON OD WAS EXPLANTED AND OS WAS NOT INSERTED." Manufacturer Narrative: A4-A6: UNK. H6 PATIENT RELATED FACTOR. CLAIM# (B)(4). | -11.0/1.0/102 | OD | Physician | 07/17/2023 | 2023 | 53 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17455642&pc=QCB | 2023826-2023-03221 | 2023-07-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -7.0/3.0/089 (SPHERE/CYLINDER/AXIS) LENS TORE DURING LOADING AFTER OPENING THE VIAL. THE LENS WAS NOT IMPLANTED. ON (B)(6) 2023, A LONGER LENGTH LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "PATIENT UBM PARTICULARITY, CHOOSE TWO MODELS FOR BACKUP, 13.2 CRYSTAL BREAKAGE, IMPLANTATION 1.7, ARCH HEIGHT IDEA." Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -7.0/3.0/089 | OS | Physician | 07/17/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17455679&pc=QCB | 2023826-2023-03222 | 2023-05-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -13.5/2.5/091 (SPHERE / CYLINDER / AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS, IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT. REPORTEDLY, "THE LENS EXCHANGED WAS IMPLANTED VERTICALLY." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. STATUS OF THE EYE IS REPORTED AS, "OK." Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN/BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL - ACD<3.0. CLAIM#: (B)(4). | -13.5/2.5/091 | OD | Physician | 07/16/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17455942&pc=QCB | 2023826-2023-03301 | 2023-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -10.0/+1.0/091 | OD | Physician | 07/13/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17466632&pc=QCB | 2023826-2023-03314 | 2023-07-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/1.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "THE CENTRAL VAULT IS HIGH, ACCORDING TO THE DOCTOR'S JUDGMENT, THE CRYSTAL WAS CHANGED INTO SMALLER SIZE." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -11.5/1.5/092 | OD | Physician | 08/04/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17477530&pc=QCB | 2023826-2023-03349 | 2022-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -12.50/2.0/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS OF THE SAME LENGTH AND THE PROBLEM RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION - WITH RESIDUE/DEBRIS ON THE PRODUCT SHOULD ALSO HAVE ALSO BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: H6- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | -12.50/2.0/90 | OS | Physician | 07/11/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17477732&pc=QCB | 2023826-2023-03353 | 2023-07-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF -12.50/2.5/030 WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED, HOWEVER THE VAULT NORMALIZED AND THE PROBLEM RESOLVED. NO FURTHER INTERVENTION NEEDED. CUASE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.50/2.5/030 | OD | Physician | 07/11/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17477766&pc=QCB | 2023826-2023-03354 | 2023-07-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF -11.00/2.0/088 WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED, HOWEVER THE VAULT NORMALIZED AND THE PROBLEM RESOLVED. NO FURTHER INTERVENTION NEEDED. CUASE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6: TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.00/2.0/088 | OS | Physician | 07/11/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17477841&pc=QCB | 2023826-2023-03357 | 2023-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -7.0/0.5/160 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H6: DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -7.0/0.5/160 | OD | Physician | 07/13/2023 | 2023 | 46 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17477887&pc=QCB | 2023826-2023-03358 | 2023-06-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.0/2.0/073 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. INTRAOPERATIVELY THE LENS WAS REMOVED DUE TO LOW VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTICS BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) CLAIM#: (B)(4). | -13.0/2.0/073 | OS | Physician | 07/13/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17484054&pc=QCB | 2023826-2023-03389 | 2023-05-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/+1.5/073 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS TWO SIZES SHORTER IN LENGTH DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -6.0/+1.5/073 | OD | Physician | 07/14/2023 | 2023 | 44 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17484055&pc=QCB | 2023826-2023-03390 | 2023-03-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.0/+1.0/082 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -5.0/+1.0/082 | OD | Physician | 07/14/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17484056&pc=QCB | 2023826-2023-03391 | 2023-03-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/+1.5/112 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -5.5/+1.5/112 | OS | Physician | 07/14/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17484238&pc=QCB | 2023826-2023-03431 | 2023-06-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS WITH DIOPTER -7.0/+1.0/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -7.0/+1.0/091 | OS | Physician | 07/13/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17484323&pc=QCB | 2023826-2023-03318 | 2023-07-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF A -8.0/2.5/93 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. REPORTEDLY, "THERE WAS A DEPOSIT ON THE LENS" AND PRODUCT NON-CONFORMITY WAS NOTED. CAUSE OF THE EVENT IS UNKNOWN. THE PROBLEM WAS NOT RESOLVED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATE LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). | -8.0/2.5/93 | OS | Physician | 11/29/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17484691&pc=QCB | 2023826-2023-03335 | 2022-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/112 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2023. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5: THE LENS WAS REMOVED DUE TO LOW VAULTING. CLAIM # (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.50/+2.0/112 | OS | Physician | 07/07/2023 | 2023 | 51 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17484695&pc=QCB | 2023826-2023-03410 | 2023-07-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.50/+3.0/090 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO THE LENS FLIPPED. THE BACKUP LENS WAS IMPLANTED WITH NO ISSUES. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -07.50/+3.0/090 | NA | Physician | 07/18/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17484701&pc=QCB | 2023826-2023-03409 | 2023-07-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50/+2.5/076 (SPHERE/CYLINDER/AXIS), WAS DAMAGED WHILE LOADING AND THERE WAS NO PATIENT CONTACT. ANOTHER ICL LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING WELL. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE CAUSE OF THE EVENT WAS DUE TO THE DOCTOR LOADED THE LENS INCORRECTLY. CLAIM # (B)(4). Manufacturer Narrative: A2: UNK; A3: UNK; A4: UNK; A5: UNK; A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | -11.50/+2.5/076 | NA | Physician | 07/18/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17484800&pc=QCB | 2023826-2023-03408 | 2023-07-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2; -16/+1.00/37 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS HAD TORE DURING LOADING, AND NOTED AFTER REMOVING FROM THE VIAL. THE LENS WAS NOT IMPLANTED. THIS OCCURRED ON (B)(6) 2023. THE CAUSE OF THE EVENT WAS REPORTED AS AN ISSUE WITH THE INJECTOR/PLUNGER, LENS TORE. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- ON (B)(6) 2023 A REPLACEMENT LENS WAS IMPLANTED INTO THE LEFT EYE (OS) AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "PATIENT VERY SATISFIED". CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -16/+1.00/37 | OS | Physician | 07/19/2023 | 2023 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17484801&pc=QCB | 2023826-2023-03414 | 2023-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -7.5/2.0/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: CLAIM# (B)(4). | -7.5/2.0/086 | OD | Physician | 06/18/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17484802&pc=QCB | 2023826-2023-03413 | 2023-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -8.0/2.0/073 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: CLAIM# (B)(4). | -8.0/2.0/073 | OS | Physician | 06/18/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17484968&pc=QCB | 2023826-2023-03360 | 2023-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.5/1.5/012 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATED THE INITIAL LENS WAS IMPLANTED IN VERTICAL POSITION. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6- DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE OPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.5/1.5/012 | OD | Physician | 07/14/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17485013&pc=QCB | 2023826-2023-03361 | 2023-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.0/1.5/178 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATED THE INITIAL LENS WAS IMPLANTED IN VERTICAL POSITION. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H6- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A TORN HAPTIC. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | -10.0/1.5/178 | OS | Physician | 07/14/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17485814&pc=QCB | 2023826-2023-03362 | 2023-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -18.0/3.5/110 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM RESOLVED. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: B5 - ON AN UCLEAR DATE THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. CLAIM # (B)(4). | -18.0/3.5/110 | OD | Physician | 11/15/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17485863&pc=QCB | 2023826-2023-03363 | 2023-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.0/1.0/075 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. NO PATIENT INJURY WAS REPORTED. THE REPORTER STATES THAT THE CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.0/1.0/075 | OS | Physician | 07/14/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17486810&pc=QCB | 2023826-2023-03260 | 2023-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.0/1.5/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, NO NEW LENS WILL BE IMPLANTED. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: SHOULD BE CORRECTED TO STATE: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.0/1.5/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, NO NEW LENS WILL BE IMPLANTED. CLAIM#(B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). | -10.0/1.5/108 | OD | Physician | 07/17/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17486811&pc=QCB | 2023826-2023-03365 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). HIGH VAULT HAS BEEN OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION: A2 - 03-MAR-2023 A3 - M B3 - 30-JUN-2023 B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. AFTER PATIENT RE-ASSESSMENT. PATIENT IS HAPPY WITH VISION AND IS ASYMPTOMATIC. PROBLEM IS NOW RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. D4 - MODEL#; VTICMO13.2, SERIAL#; (B)(6), EXPIRATION DATE; 31/MAY/2026. 6A - 29-JUN-2023 H4 - 10-JUN-2023 H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: D4 -UDI NUMBER ADDED. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4) | -10.5/2.0/090 | OS | Physician | 07/15/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17486812&pc=QCB | 2023826-2023-03364 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). HIGH VAULT HAS BEEN OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION: A2 - 03-MAR-2023 A3 - M B3 - 30-JUN-2023 B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -10/1.5/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) /2023. AFTER PATIENT RE-ASSESSMENT. PATIENT IS HAPPY WITH VISION AND IS ASYMPTOMATIC. PROBLEM IS NOW RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. D4 - MODEL#; VTICM5_13.2, SERIAL#; (B)(6), EXPIRATION DATE; 31/MAY/2026, UDI#; (B)(4). 6A - 29-JUN-2023 H4 - 14-JUN-2023 H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: A2 - UNK, A3 - UNK, A4 - UNK, A5 - UNK ,A6 - UNK, B3 - UNK, D4 - UNK, D6A - UNK, H4 - UNK, CLAIM# (B)(4). | -10/1.5/075 | OD | Physician | 07/15/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17486814&pc=QCB | 2023826-2023-03259 | 2023-05-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.5/1.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SAME LENGTH SPHERICAL LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -11.5/1.0/075 | OS | Physician | 07/17/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17486817&pc=QCB | 2023826-2023-03265 | 2023-07-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.63, -10.0/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT HAVE BEEN OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.0/2.0/095 | OD | Physician | 07/17/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17486821&pc=QCB | 2023826-2023-03261 | 2023-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/1.0/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, NO NEW LENS WILL BE IMPLANTED. Manufacturer Narrative: CORRECTED DATA: B5: SHOULD BE CORRECTED TO STATE: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/1.0/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXPLANTED DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, NO NEW LENS WILL BE IMPLANTED. CLAIM#(B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.5/1.0/070 | OS | Physician | 07/17/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17486839&pc=QCB | 2023826-2023-03272 | 2023-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.0/0.5/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. CLAIM# (B)(4). | -8.0/0.5/074 | OD | Physician | 06/18/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17491503&pc=QCB | 2023826-2023-03320 | 2023-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -10.0/3.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE PROBLEM WAS RESOLVED. REPORTEDLY, "AFTER SURGERY, THE LENS OF THE RIGHT EYE ROTATED 10 DEGREES, THE NAKED EYE VISUAL ACUITY WAS 0.8, THE COMPREHENSIVE OPTOMETRY COULD REACH 1.0, AND "THE" ROTATION WAS STILL 10 DEGREES THREE MONTHS AFTER SURGERY. VISUAL ACUITY 0.8, COMBINED OPTOMETRY TO 1.0," AND "THE VISION BECAME BETTER AFTER THE LENS WAS REPLACED. THE CRYSTALS ARE IN THE RIGHT PLACE." Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 19-SEP-2023. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -10.0/3.0/089 | OD | Physician | 07/21/2023 | 2023 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17494054&pc=QCB | 2023826-2023-03438 | 2022-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -4.5/1.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED AND THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS OF THE SAME SIZE AND -3.50 DIOPTER AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS DEVICE. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS RETURNED WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -4.5/1.0/091 | OD | Physician | 08/08/2023 | 2023 | 41 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17495142&pc=QCB | 2023826-2023-03016 | 2023-03-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/6.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ON (B)(6) 2023 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM #(B)(4). | -12.0/6.0/089 | OD | Physician | 07/16/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17495147&pc=QCB | 2023826-2023-03369 | 2023-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/2.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -14.5/2.5/094 | OD | Physician | 08/06/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17495150&pc=QCB | 2023826-2023-03371 | 2023-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/1.0/002 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -6.0/1.0/002 | OS | Physician | 08/07/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17495151&pc=QCB | 2023826-2023-03377 | 2023-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/3.0/070 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -14.5/3.0/070 | OD | Physician | 08/07/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17495152&pc=QCB | 2023826-2023-03372 | 2023-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/2.5/176 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -8.0/2.5/176 | OD | Physician | 08/07/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17495154&pc=QCB | 2023826-2023-03374 | 2023-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/2.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -11.0/2.0/088 | OS | Physician | 08/07/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17495156&pc=QCB | 2023826-2023-03375 | 2023-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -10.5/1.0/088 | OD | Physician | 08/07/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17495491&pc=QCB | 2023826-2023-03322 | 2023-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -12.5/3.50/085 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). LOW VAULT WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -12.5/3.50/085 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LOW VAULT WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENS; THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# : (B)(4). | -12.5/3.50/085 | OS | Physician | 07/21/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17495492&pc=QCB | 2023826-2023-03321 | 2023-06-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -8.5/1.0/072 (SPHERE/CYLINDER/AXIS) WAS INTRAOPERATIVELY IMPLANTED AND REMOVED DUE TO A LENS TEAR DURING INJECTION/DELIVERY INTO THE EYE ON 24-JUN-2023. NO PATIENT INJURY WAS REPORTED. A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "NO DISCOMFORT IN THE SURGICAL EYE," AND "NO SUPPLEMENTATION." Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# : (B)(4). | -8.5/1.0/072 | NA | Physician | 07/21/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17504606&pc=QCB | 2023826-2023-03397 | 2023-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE LENS WAS EXCHANGED ON (B)(6)2023 FOR A LENS TWO SIZES LONGER IN LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+1.0/091 | OD | Physician | 07/16/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17504608&pc=QCB | 2023826-2023-03393 | 2023-05-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+2.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.5/+2.0/091 | OD | Physician | 07/16/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17504609&pc=QCB | 2023826-2023-03396 | 2023-05-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.5/+1.5/111 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS TWO SIZES LONGER IN LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -3.5/+1.5/111 | OS | Physician | 07/16/2023 | 2023 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17504610&pc=QCB | 2023826-2023-03395 | 2023-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.5/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: (UNDER 21YRS OF AGE AT DATE OF IMPLANT) CLAIM#(B)(4). | -10.5/+2.5/094 | OS | Physician | 07/16/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17505259&pc=QCB | 2023826-2023-03381 | 2023-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/2.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. (B)(4). | -8.5/2.0/095 | OS | Physician | 08/07/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17505260&pc=QCB | 2023826-2023-03382 | 2023-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -10.5/1.5/092 | OD | Physician | 08/07/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17505262&pc=QCB | 2023826-2023-03379 | 2023-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -11.0/1.0/090 | OD | Physician | 08/07/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17505263&pc=QCB | 2023826-2023-03378 | 2023-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/4.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -13.0/4.0/094 | OS | Physician | 08/07/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17505814&pc=QCB | 2023826-2023-03325 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.5/2.5/110 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS A REPLACEMENT LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023, DUE TO LENS TEAR. THE PROBLEM WAS RESOLVED. REPORTEDLY, "CORTICAL SPOKING ON LENS. DOING WELL OTHERWISE." Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS. CLAIM# : (B)(4). | -10.5/2.5/110 | OS | Physician | 07/24/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17505815&pc=QCB | 2023826-2023-03328 | 2023-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.0/2.0/175 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SAME SIZE LENS, IMPLANTED VERTICALLY DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "T1190558 WAS IMPLANTED IN A VERTICAL POSITION T1415603 WAS IMPLANTED IN A HORIZONTAL POSITION." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL: AGE<21. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -11.0/2.0/175 | OS | Physician | 07/25/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17510013&pc=QCB | 2023826-2023-03327 | 2023-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.0/1.0/003 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SAME SIZE LENS, IMPLANTED VERTICALLY DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "T1360849 WAS IMPLANTED IN A VERTICAL POSITION T1354973 WAS IMPLANTED IN A HORIZONTAL POSITION." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6: OFF-LABEL: AGE<21 CLAIM# (B)(4). | -11.0/1.0/003 | OD | Physician | 07/25/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17512007&pc=QCB | 2023826-2023-03329 | 2023-07-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -14.5/2.5/112 (SPHERE/CYLINDER/AXIS) LENS TORE DURING LOADING. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#(B)(4). | -14.5/2.5/112 | NA | Physician | 07/25/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17512008&pc=QCB | 2023826-2023-03331 | 2023-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -13.54.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "THE ARCH HEIGHT OF THE CRYSTAL WAS TOO HIGH AFTER IMPLANTATION, AND THE PATIENT WAS WORRIED ABOUT THE RISK OF A SECOND OPERATION AND ASKED TO REMOVE IT. AFTER REPEATED COMMUNICATION WITH THE DOCTOR, THE CRYSTAL WAS FINALLY REMOVED." Manufacturer Narrative: A4-A6: UNK. H6: WORK ORDER SEARCH - NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -13.54.5/093 | OS | Physician | 07/25/2023 | 2023 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17512014&pc=QCB | 2023826-2023-03330 | 2023-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -13.0/4.5/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "THE ARCH HEIGHT OF THE CRYSTAL WAS TOO HIGH AFTER IMPLANTATION, AND THE PATIENT WAS WORRIED ABOUT THE RISK OF A SECOND OPERATION AND ASKED TO REMOVE IT. AFTER REPEATED COMMUNICATION WITH THE DOCTOR, THE CRYSTAL WAS FINALLY REMOVED." Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH - NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -13.0/4.5/086 | OD | Physician | 07/25/2023 | 2023 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17512100&pc=QCB | 2023826-2023-03400 | 2023-07-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+1.0/095 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -13.0/+1.0/095 | OS | Physician | 07/17/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17512102&pc=QCB | 2023826-2023-03399 | 2023-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.0/071 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS TWO SIZES LONGER IN LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4) . | -10.0/+2.0/071 | OS | Physician | 07/17/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17515429&pc=QCB | 2023826-2023-03417 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.50/+1.5/072 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. AN LRI WAS PERFORMED TO CORRECT RESIDUAL ASTIGMATISM ON (B)(6) 2023. THE PATIENTS LATEST EXAM ON (B)(6) 2023 INDICATED VA SC 20/40-2. THE SURGEON PLANS TO ROTATE (REPOSITION) THE LENS. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - LRI SURGERY. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: THE LENS WAS REPOSITIONED ON (B)(6) 2023 AND THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). | -08.50/+1.5/072 | OD | Other Health Care Professional | 08/03/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17515518&pc=QCB | 2023826-2023-03411 | 2023-07-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+1.0/090 (SPHERE/CYLINDER/AXIS), DUE TO THE LENS TORE DURING SURGERY. THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK, A5: UNK, A6: UNK, D6A: UNK, D6B: UNK, H6 - WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.50/+1.0/090 | NA | Other Health Care Professional | 07/18/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17515535&pc=QCB | 2023826-2023-03426 | 2023-06-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+4.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND THE PROBLEM WAS RESOLVED. THE PATIENT WANTED THE LENS REMOVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUTION - LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -15.00/+4.0/095 | OD | Physician | 07/16/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17515540&pc=QCB | 2023826-2023-03423 | 2023-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LOW VAULTING AND THE PROBLEM WAS RESOLVED. THE PATIENT FELT UNCOMFORTABLE AND URGED REMOVAL OF THE ICL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.50/+1.5/088 | OS | Physician | 07/16/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17515541&pc=QCB | 2023826-2023-03425 | 2023-03-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/179 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/1.5/179 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS ON (B)(6) 2023 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | -11.00/+1.5/179 | OS | Physician | 07/16/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17515544&pc=QCB | 2023826-2023-03424 | 2023-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LOW VAULTING AND THE PROBLEM WAS RESOLVED. THE PATIENT FELT UNCOMFORTABLE AND URGED REMOVAL OF THE ICL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -11.50/+1.5/094 | OD | Physician | 07/16/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17515547&pc=QCB | 2023826-2023-03419 | 2023-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+1.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -11.50/+1.0/089 | OD | Physician | 07/17/2023 | 2023 | 17 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17515550&pc=QCB | 2023826-2023-03420 | 2023-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.50/+2.5/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -11.50/+2.5/090 | OD | Physician | 07/17/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17516053&pc=QCB | 2023826-2023-03383 | 2023-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/1.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -9.5/1.0/093 | OD | Physician | 08/07/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17516054&pc=QCB | 2023826-2023-03384 | 2023-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/1.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -11.0/1.5/093 | OS | Physician | 08/07/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17516056&pc=QCB | 2023826-2023-03385 | 2023-06-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #: (B)(4). | -12.5/1.0/090 | OS | Physician | 08/07/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17541715&pc=QCB | 2023826-2023-03421 | 2023-06-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50/+1.0/093 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK, A5: UNK, A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -07.50/+1.0/093 | OS | Physician | 07/21/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17541731&pc=QCB | 2023826-2023-03427 | 2023-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -07.00/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO THE LENS TORE DURING LOADING. THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -07.00/+2.0/090 | OD | Physician | 07/21/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17541734&pc=QCB | 2023826-2023-03428 | 2023-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/109 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -10.00/+1.0/109 | OD | Physician | 07/21/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17543377&pc=QCB | 2023826-2023-03516 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.00/1.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS INTRAOPERATIVELY REMOVED DUE TO LOW VAULT. A REPLACEMENT LENS OF LONGER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED TO: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/1.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS INTRAOPERATIVELY REMOVED DUE TO LOW VAULT. A REPLACEMENT LEN OF A LONGER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM #(B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT - IMPACT CODE: 4629 CORRECTED TO 2199 IN SUPP3. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6 - HEALTH EFFECT - IMPACT CODE: 4627 CORRECTED TO 4629 INITIAL MDR. CLAIM # (B)(4). | -9.00/1.0/092 | OS | Physician | 07/17/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17546204&pc=QCB | 2023826-2023-03518 | 2023-06-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 12.6MM VTICMO12.6 -13.00/2.5/077 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE DURING LOADING. THERE WAS NO PATIENT CONTACT. ON (B)(6) 2023 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | -13.00/2.5/077 | NA | Physician | 07/17/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17546359&pc=QCB | 2023826-2023-03486 | 2023-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.0/2.0/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -4.0/2.0/074 | OD | Physician | 08/09/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17546363&pc=QCB | 2023826-2023-03487 | 2023-06-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/1.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) (B)(4). | -12.0/1.5/094 | OD | Physician | 08/09/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17547222&pc=QCB | 2023826-2023-03405 | 2023-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+1.0/171 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -6.5/+1.0/171 | OS | Physician | 07/17/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17547225&pc=QCB | 2023826-2023-03407 | 2023-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.0/074 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -8.5/+1.0/074 | OD | Physician | 07/17/2023 | 2023 | 17 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17547226&pc=QCB | 2023826-2023-03406 | 2023-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.5/172 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21 YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.5/172 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL AND POWER BUT SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). | -9.0/+1.5/172 | OD | Physician | 07/17/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549361&pc=QCB | 2023826-2023-03445 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -2.0/2.0/169 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. B3 - UNK. (B)(4). | -2.0/2.0/169 | OD | Physician | 07/24/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549363&pc=QCB | 2023826-2023-03444 | 2023-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.0/0.5/134 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK A5 - UNK A6 - UNK H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE PATIENT IS HAPPY ALL IS FINE! CLAIM#: (B)(4). | -7.0/0.5/134 | OS | Physician | 07/24/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549372&pc=QCB | 2023826-2023-03455 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -8.50/1.0/003 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. D6A - UNK. CLAIM (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B1- REPORTED AS PRODUCT PROBLEM - UPDATE TO ADVERSE EVENT. B2 - UPDATE TO REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. B5 - REPORTED AS THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -8.50/1.0/003 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. - UPDATE TO THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -8.50/1.0/003 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON 06/OCT/2023. LENS WAS EXPLANTED ON (B)(6) 2023. D6A - UPDATE TO (B)(6) 2023. D6B - UPDATE TO (B)(6) 2023. H1 - REPORTED AS MALFUNCTION - UPDATE TO SERIOUS INJURY. H6 - REPORTED AS 2199 - UPDATE TO 4627. CLAIM# (B)(4). | -8.50/1.0/003 | OS | Physician | 07/28/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549374&pc=QCB | 2023826-2023-03457 | 2023-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.0/1.5/131 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SAME LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -10.0/1.5/131 | OS | Physician | 07/28/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549379&pc=QCB | 2023826-2023-03456 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -8.50/1.0/003 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION: B1- REPORTED AS PRODUCT PROBLEM - UPDATE TO ADVERSE EVENT B2 - UPDATE TO REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE B5 - REPORTED AS THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -8.50/1.0/003 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. - UPDATE TO THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -8.50/1.0/003. (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LENS WAS EXPLANTED ON (B)(6) 2023. D6A - UPDATE TO (B)(6) 2023. D6B - UPDATE TO (B)(6) 2023. H1 - REPORTED AS MALFUNCTION - UPDATE TO SERIOUS INJURY. H6 - REPORTED AS 2199 - UPDATE TO 4627. CLAIM# (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. D6A - UNK. (B)(4). | -8.50/1.0/003 | OD | Physician | 07/28/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549393&pc=QCB | 2023826-2023-03459 | 2023-07-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE(OS). HIGH VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM513.2, -12.0/1.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION: CLAIM# (B)(4). Manufacturer Narrative: A2 - UNK. A3 - UNK. A4 - UNK. A5 - UNK. A6 - UNK. B3 - UNK. D2 - PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). D4 - UNK. D6A - UNK. H4 - UNK. CLAIM#: (B)(4). | -12.0/1.0/081 | OS | Physician | 07/28/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549693&pc=QCB | 2023826-2023-03500 | 2023-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS REMOVED AND REPLACED DUE TO SEVERE VAULT AND ANGLE CLOSURE. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM513.7 IMPLANTABLE COLLAMER LENS OF A -7.50/1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS THE DEVICE. REPORTEDLY, "13.7MM LENS TOO BIG." H6: OFF-LABEL - ACD<3.0. CLAIM# (B)(4). Manufacturer Narrative: B3: DATE OF EVENT CORRECTED TO 05-JUL-2023 IN SUPP2 MDR. CLAIM#: (B)(4). Manufacturer Narrative: A4-A6:UNK. B3: UNK. D4 (EXPERIATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK. H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM# (B)(4). | -7.50/1.0/090 | OS | Physician | 07/27/2023 | 2023 | 45 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549695&pc=QCB | 2023826-2023-03499 | 2023-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). THE LENS WAS REMOVED AND REPLACED DUE TO SEVERE VAULT AND ANGLE CLOSURE. REPORTEDLY, "THE EXCHANGE WENT WELL FOR RIGHT EYE (ORIGINALLY HAD TRAUMA WITH RETINAL LASER REPAIR) WITH INCREASED IOP." Manufacturer Narrative: (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM513.7 IMPLANTABLE COLLAMER LENS OF A -7.50/1.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ANGLE CLOSURE WITH ELEVATED IOP. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS THE DEVICE. REPORTEDLY, "13.7MM LENS TOO BIG." CLAIM# (B)(4). Manufacturer Narrative: A2: DOB CORRECTED TO ON (B)(6) 1978 IN SUPP1 MDR. B3: DATE OF EVENT CORRECTED TO 05-JUL-2023 IN SUPP1 MDR. H4: MANUFACTURE DATE CORRECTED TO 16-NOV-2022. CLAIM#: (B)(4). | -7.50/1.0/092 | OD | Physician | 07/27/2023 | 2023 | 44 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549723&pc=QCB | 2023826-2023-03464 | 2023-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.5/070 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+1.5/070 | OS | Physician | 07/18/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549726&pc=QCB | 2023826-2023-03469 | 2023-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.00/+1.0/177 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -7.00/+1.0/177 | OS | Physician | 07/19/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549727&pc=QCB | 2023826-2023-03467 | 2023-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+2.5/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: (B)(4). | -13.0/+2.5/094 | OD | Physician | 07/18/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549728&pc=QCB | 2023826-2023-03465 | 2023-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+5.0/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -12.0/+5.0/090 | OS | Physician | 07/18/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549729&pc=QCB | 2023826-2023-03466 | 2023-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/+2.0/088 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -5.5/+2.0/088 | OD | Physician | 07/18/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17549730&pc=QCB | 2023826-2023-03468 | 2023-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/+3.5/078 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -14.5/+3.5/078 | OS | Physician | 07/18/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17552300&pc=QCB | 2023826-2023-03346 | 2023-07-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -9.0/2.0/093 (SPHERE/CYLINDER/AXIS) LENS TEAR/BREAK BEFORE/DURING LOADING AFTER OPENING VIAL. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY ON (B)(6) 2023. A REPLACEMENT LENS WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN AND THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: ADDITIONAL INFORMATION: A4, A5, A6-UNK. H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -9.0/2.0/093 | NA | Physician | 08/07/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17552801&pc=QCB | 2023826-2023-03501 | 2023-06-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -12.0/3.5/093 (SPHERE/CYLINDER/AXIS) TORE IN HALF. THE LENS WAS INTENDED FOR THE PATIENT'S LEFT EYE (OS). REPORTEDLY, "LENS TORN IN HALF." Manufacturer Narrative: H6- MEDICAL DEVICE CODE: "2993" SHOULD BE REMOVED AND "3189" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# : (B)(4). | -12.0/3.5/093 | OS | Physician | 07/27/2023 | 2023 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17553356&pc=QCB | 2023826-2023-03528 | 2023-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICH13.2 IMPLANTABLE COLLAMER LENS OF -2.00/6.00/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 AND (B)(6) 2023 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE REPOSITIONING DID NOT RESOLVE THE PROBLEM. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -2.00/6.00/090 | OS | Physician | 07/19/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17553513&pc=QCB | 2023826-2023-03491 | 2023-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/3.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -12.0/3.0/087 | OD | Physician | 08/10/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17553514&pc=QCB | 2023826-2023-03493 | 2023-05-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/1.5/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -6.5/1.5/072 | OS | Physician | 08/10/2023 | 2023 | 44 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17553520&pc=QCB | 2023826-2023-03495 | 2023-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.0/035 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -10.0/1.0/035 | OS | Physician | 08/11/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17553521&pc=QCB | 2023826-2023-03494 | 2023-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/0.5/167 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -8.5/0.5/167 | OD | Physician | 08/11/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17558283&pc=QCB | 2023826-2023-03472 | 2023-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.5/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT REPORT A CAUSE FOR THIS EVENT. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) CLAIM# (B)(4). | -11.0/+2.5/094 | OS | Physician | 07/19/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17558284&pc=QCB | 2023826-2023-03471 | 2023-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.5/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT REPORT A CAUSE FOR THIS EVENT. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -11.0/+2.5/090 | OD | Physician | 07/19/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17558285&pc=QCB | 2023826-2023-03473 | 2023-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+3.0/086 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.0/+3.0/086 | OS | Physician | 07/19/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17559247&pc=QCB | 2023826-2023-03545 | 2023-07-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 13.2MM VTICMO13.2 -11.5/2.0/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE RIGHT EYE (OD). INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH NO PATIENT INJURY AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -11.5/2.0/095 | OD | Physician | 07/20/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17559776&pc=QCB | 2023826-2023-03546 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 12.6MM VTICMO12.6 -12.5/2.5/87 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE RIGHT EYE (OD). ON (B)(6) 2023 THE LENS WAS EXCHANGED AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT THERE WAS NO PATIENT INJURY. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.5/2.5/87 | OD | Physician | 07/20/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17560526&pc=QCB | 2023826-2023-03480 | 2023-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+2.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.0/+2.0/093 | OD | Physician | 07/21/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17560527&pc=QCB | 2023826-2023-03483 | 2023-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+1.5/070 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -8.0/+1.5/070 | OS | Physician | 07/21/2023 | 2023 | 17 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17560529&pc=QCB | 2023826-2023-03481 | 2023-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+2.5/073 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.0/+2.5/073 | OS | Physician | 07/21/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17560533&pc=QCB | 2023826-2023-03478 | 2023-06-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.5/083 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -9.0/+1.5/083 | OD | Physician | 07/21/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17561390&pc=QCB | 2023826-2023-03551 | 2023-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -10.0/1.5/091 | OD | Physician | 08/14/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17561391&pc=QCB | 2023826-2023-03552 | 2023-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/3.0/110 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -11.5/3.0/110 | OD | Physician | 08/14/2023 | 2023 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17561392&pc=QCB | 2023826-2023-03553 | 2023-07-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/3.5/112 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -11.0/3.5/112 | OS | Physician | 08/14/2023 | 2023 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17561393&pc=QCB | 2023826-2023-03548 | 2023-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/2.0/074 (SPHERE / CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT.. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -12.0/2.0/074 | OS | Physician | 08/14/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17561395&pc=QCB | 2023826-2023-03549 | 2023-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/1.0/095 (SPHERE / CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -9.5/1.0/095 | OD | Physician | 08/14/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17561565&pc=QCB | 2023826-2023-03476 | 2023-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+4.5/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LOW VAULT. THAT SAME DAY, THE LENS WAS REPLACED WITH A LONGER LENGTH LENS ALTHOUGH IT IS UNCLEAR WHETHER THIS WAS WITHIN THE SAME PROCEDURE. THE PROBLEM WAS RESOLVED. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -14.0/+4.5/094 | OS | Physician | 07/21/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17561599&pc=QCB | 2023826-2023-03507 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -8.00/4.0/088 BECAME DAMAGED DURING LOADING. A REPLACEMENT LENS WAS IMPLANTED. REPORTEDLY, "NO PATIENT CONTACT." Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5- THE REPORTER INDICATED THE SURGEON NOTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.00/4.0/088 (SPHERE/CYLINDER/AXIS) REPORTEDLY WAS "DAMAGED DURING LOADING. NO, PATIENT CONTACT"AND "IT JUST GOT DROPPED OR SOMETHING". A BACK UP LENS OF THE SAME PARAMETERS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS USER ERROR. REPORTEDLY "USER ERROR IN LOADING". CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- ORIGINALLY THE CLAIM WAS DETERMINED TO BE REPORTABLE, BUT UPON FURTHER REVIEW WAS DETERMINED TO NOT MEET DEFINITION OF A COMPLAINT. HOWEVER, THE CLAIM CANNOT BE VOIDED AS AN MDR HAS ALREADY BEEN SUBMITTED. REPORTEDLY WAS "DAMAGED DURING LOADING. NO, PATIENT CONTACT" AND "IT JUST GOT DROPPED OR SOMETHING". THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS USER ERROR. REPORTEDLY "USER ERROR IN LOADING". CLAIM# (B)(4). | -8.00/4.0/088 | NA | Physician | 07/28/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17561600&pc=QCB | 2023826-2023-03508 | 2023-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.5/1.0/073 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "THE PATIENT DID NOT IMPLANT A SECOND ICL." Manufacturer Narrative: D4 - MODEL # CORRECTED TO VTICMO13.2 IN THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -9.5/1.0/073 | OD | Physician | 07/28/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17568860&pc=QCB | 2023826-2023-03511 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -8.0/5.5/115 (SPHERE/CYLINDER/AXIS) WAS SUPPRESSED WHILE INJECTING. REPORTEDLY, "TREATMENT DISCONTINUED DUE TO HIGH HYPEROPIA. SUPPRESSES THE LENS WHILE INJECTING." Manufacturer Narrative: A4-A6:UNK H6: PATIENT RELATED FACTOR. CLAIM# (B)(4). | -8.0/5.5/115 | NA | Physician | 07/28/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17569169&pc=QCB | 2023826-2023-03514 | 2023-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -11.0/1.0/117 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -11.0/1.0/117 | OD | Physician | 07/31/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17569310&pc=QCB | 2023826-2023-03567 | 2023-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.0/073 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.5/+1.0/073 | OS | Physician | 07/24/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17569312&pc=QCB | 2023826-2023-03565 | 2023-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+1.0/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | -13.0/+1.0/095 | OD | Physician | 07/24/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17569314&pc=QCB | 2023826-2023-03571 | 2023-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+4.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -7.5/+4.0/093 | OD | Physician | 07/25/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17569362&pc=QCB | 2023826-2023-03555 | 2023-07-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/2.0/067 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.5/2.0/067 | OS | Physician | 08/14/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17569363&pc=QCB | 2023826-2023-03559 | 2023-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/2.5/113 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -11.0/2.5/113 | OD | Physician | 08/14/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17569364&pc=QCB | 2023826-2023-03554 | 2023-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/1.0/070 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -8.5/1.0/070 | OD | Physician | 08/14/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17569367&pc=QCB | 2023826-2023-03386 | 2023-05-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/2.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #: (B)(4). | -11.0/2.0/087 | OD | Physician | 08/07/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17569368&pc=QCB | 2023826-2023-03387 | 2023-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/1.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -8.0/1.0/087 | OD | Physician | 08/07/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17569370&pc=QCB | 2023826-2023-03561 | 2023-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/1.0/110 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -11.0/1.0/110 | OD | Physician | 08/15/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17569371&pc=QCB | 2023826-2023-03564 | 2023-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/1.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -12.0/1.5/090 | OD | Physician | 08/15/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17569372&pc=QCB | 2023826-2023-03563 | 2023-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/2.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -14.5/2.0/088 | OD | Physician | 08/15/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17569373&pc=QCB | 2023826-2023-03558 | 2023-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -13.0/2.5/091 | OS | Physician | 08/14/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17569374&pc=QCB | 2023826-2023-03560 | 2023-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/1.0/106 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -12.0/1.0/106 | OS | Physician | 08/15/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17569375&pc=QCB | 2023826-2023-03562 | 2023-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/6.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -11.0/6.0/092 | OD | Physician | 08/15/2023 | 2023 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17575198&pc=QCB | 2023826-2023-03504 | 2023-06-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF A -12.50/3.5/118 (SPHERE/CYLINDER/AXIS) WAS INTENDED TO BE IMPLANTED INTO THE RIGHT EYE (OD) OF A PATIENT WITH PRE-EXISTING CATARACT IN THE OPERATIVE EYE AND PROGRESSIVE MYOPIA. THE LENS WAS NOT IMPLANTED. Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL USE: PRE-EXISTING CATARACT IN THE OPERATIVE EYE AND PROGRESSIVE MYOPIA. CLAIM# (B)(4). Manufacturer Narrative: H4 - DEVICE MANFUCTURE DATE - 01-FEB-2023. CLAIM#: (B)(4). | -12.50/3.5/118 | OD | Physician | 07/27/2023 | 2023 | 47 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17581259&pc=QCB | 2023826-2023-03590 | 2023-07-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 13.2MM, VTICMO13.2 -10.5/1.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS GOT STUCK IN THE CARTRIDGE OR INJECTION SYSTEM, PATIENT CONTACT WAS REPORTED. PER REPORTER THE LENS WAS INTRAOPERATIVELY REMOVED WITH NO PATIENT INJURY. ON (B)(6) 2023 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.5/1.5/089 | NA | Physician | 07/26/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17581261&pc=QCB | 2023826-2023-03594 | 2023-04-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -7.5/1.0/172 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED UPSIDE DOWN INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023 . ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH NO PATIENT INJURY AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT IS UNKNOW. IT WAS REPORTED THAT THE REPLACEMENT LENS WAS IMPLANTED IN A VERTICAL POSITION. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -7.5/1.0/172 | OS | Physician | 07/26/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17581266&pc=QCB | 2023826-2023-03591 | 2023-07-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 13.2MM, VTICMO13.2 -11.00/1.5/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS INTRAOPERATIVELY REMOVED AND REPLACED WITH NO PATIENT INJURY. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.00/1.5/092 | NA | Physician | 07/26/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17591894&pc=QCB | 2023826-2023-03619 | 2023-04-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -4.5/1.0/83 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT IMPLANTED VERTICALLY AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND THERE WAS NO DEVICE CAUSALITY. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 02-AUG-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN WITH RESIDUE ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -4.5/1.0/83 | OD | Physician | 07/25/2023 | 2023 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17591903&pc=QCB | 2023826-2023-03618 | 2023-03-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -11.00/1.5/100 (SPHERE / CYLINDER / AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND IT WAS NOTED THERE WAS NO DEVICE CAUSALITY. Manufacturer Narrative: CLAIM#: (B)(4). | -11.00/1.5/100 | OD | Physician | 07/25/2023 | 2023 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17592021&pc=QCB | 2023826-2023-03622 | 2023-07-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -10.0/3.5/095 TORE/BROKE DURING INJECTION INTO THE PATIENT'S LEFT EYE (OS). THERE WAS PATIENT CONTACT. NO INJURY WAS REPORTED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -10.0/3.5/095 | OS | Physician | 08/02/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17592489&pc=QCB | 2023826-2023-03579 | 2020-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.5/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.5/+1.5/090 | OS | Physician | 07/25/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17592493&pc=QCB | 2023826-2023-03574 | 2023-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+2.0/073 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -11.5/+2.0/073 | OS | Physician | 07/25/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17592494&pc=QCB | 2023826-2023-03573 | 2023-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.0/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) AND (ACD < 3.0MM). (B)(4). | -10.0/+1.0/094 | OD | Physician | 07/25/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17592497&pc=QCB | 2023826-2023-03583 | 2023-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.5/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.0/+2.5/090 | OD | Physician | 07/27/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17592498&pc=QCB | 2023826-2023-03575 | 2023-06-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+1.0/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -6.5/+1.0/092 | OS | Physician | 07/25/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17592501&pc=QCB | 2023826-2023-03584 | 2023-06-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+2.0/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023, FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+2.0/091 | OS | Physician | 07/27/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17592541&pc=QCB | 2023826-2023-03634 | 2023-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/1.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -8.0/1.0/089 | OS | Physician | 08/17/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17592544&pc=QCB | 2023826-2023-03636 | 2023-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.50/3.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -12.50/3.5/093 | OS | Physician | 08/17/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17592552&pc=QCB | 2023826-2023-03632 | 2023-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/1.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -11.0/1.5/092 | OD | Physician | 08/15/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17592553&pc=QCB | 2023826-2023-03635 | 2023-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/3.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) .(B)(4). | -11.0/3.0/089 | OD | Physician | 08/17/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17592562&pc=QCB | 2023826-2023-03633 | 2023-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/1.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -9.0/1.5/093 | OD | Physician | 08/17/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17597799&pc=QCB | 2023826-2023-03663 | 2023-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.5/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.5/+1.5/092 | OS | Physician | 07/28/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17597800&pc=QCB | 2023826-2023-03664 | 2023-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+1.5/104 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.0/+1.5/104 | OD | Physician | 07/28/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17597801&pc=QCB | 2023826-2023-03665 | 2023-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+2.5/116 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. PER THE MEDICAL REVIEW THIS COMPLAINT IS DEEMED SERIOUS AND DEVICE CAUSALITY IS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -11.5/+2.5/116 | OD | Physician | 07/28/2023 | 2023 | 41 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17597803&pc=QCB | 2023826-2023-03666 | 2023-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+2.0/111 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -11.5/+2.0/111 | OS | Physician | 07/28/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17599554&pc=QCB | 2023826-2023-03677 | 2023-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.0/2.0/075 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -16.0/2.0/075 | OS | Physician | 08/21/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17599555&pc=QCB | 2023826-2023-03640 | 2023-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/2.5/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -12.0/2.5/073 | OS | Physician | 08/17/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17599556&pc=QCB | 2023826-2023-03641 | 2022-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/1.5/109 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -7.0/1.5/109 | OD | Physician | 08/17/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17599752&pc=QCB | 2023826-2023-03627 | 2023-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -5.5/2.0/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6)2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -5.5/2.0/086 | OS | Physician | 08/07/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17600535&pc=QCB | 2023826-2023-03683 | 2023-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -10.0/1.5/091 | OD | Physician | 08/21/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17600762&pc=QCB | 2023826-2023-03660 | 2023-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -12.5/+3.0/93 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING WITH NO LENS ROTATION. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/3.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS ON (B)(6) 2023 DUE TO LOW VAULT WITHOUT ROTATION. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT IMPACT CODE 4629: LENS WAS EXCHANGED SHOULD BE ADDED TO SUPPLEMENTAL MDR #1. CLAIM #(B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND FOREIGN MATERIAL ON THE LENS SURFACE, SPECIFICALLY, DEBRIS ON THE LENS. CLAIM # (B)(4). | -12.5/+3.0/93 | OD | Physician | 12/19/2023 | 2023 | 44 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17600765&pc=QCB | 2023826-2023-03661 | 2023-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -9.5/+3.0/122 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING WITH NO LENS ROTATION. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/3.0/122 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS ON (B)(6) 2023 DUE TO LOW VAULT WITHOUT ROTATION. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT IMPACT CODE 4629: LENS WAS EXCHANGED SHOULD BE ADDED TO SUPPLEMENTAL MDR #1. CLAIM #(B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN AND FOREIGN MATERIAL ON THE LENS SURFACE, SPECIFICALLY, RESIDUE ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: G3 - 20-DEC-2023 CORRECTED TO 28-NOV-2023. CLAIM # (B)(4). | -9.5/+3.0/122 | OS | Physician | 12/20/2023 | 2023 | 44 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17606662&pc=QCB | 2023826-2023-03596 | 2023-02-11 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023, A 13.2MM VTICM5_13.2 -9.5/2.5/083 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE/BROKE DURING LOADING. THE DAMAGED WAS NOTED DURING LENS SETTING. THERE WAS NO PATIENT CONTACT. A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -9.5/2.5/083 | NA | Physician | 07/27/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17607595&pc=QCB | 2023826-2023-03629 | 2023-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -5.0/1.5/029 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SPHERICAL LENS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY, "BLURRED VISION DUE TO RESIDUAL ASTIGMATISM. SUBJECTIVE REFRACTION AT FINAL VISIT. +0-1.00X130." CAUSE OF THE EVENT WAS THE DEVICE. THE REPORTED STATED, "SPONTANEOUS ROTATION IN FIRST WEEK AFTER INITIAL SURGERY." Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNH3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4) | -5.0/1.5/029 | OS | Physician | 08/08/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17612094&pc=QCB | 2023826-2023-03597 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -10.50/2.0/101 WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2023. THE REPORTER STATES THAT THE LENS WAS TOO SMALL AND THE DOCTOR HAD TO GO WITH A LONGER LENGTH LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: 433258. | -10.50/2.0/101 | NA | Other Health Care Professional | 07/27/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17618234&pc=QCB | 2023826-2023-03599 | 2023-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -9.5/1.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | -9.5/1.0/089 | OD | Physician | 12/04/2023 | 2023 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17619190&pc=QCB | 2023826-2023-03669 | 2023-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+1.5/097 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -8.0/+1.5/097 | OD | Physician | 07/30/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17619514&pc=QCB | 2023826-2023-03719 | 2023-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -11.0/+2.0/087 (SPHERE / CYLINDER / AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE PATIENT EXPERIENCED LOW VAULTING WITH ROTATION, ON (B)(6) 2022 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -11.0/+2.0/087 | OS | Physician | 08/03/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17619861&pc=QCB | 2023826-2023-03688 | 2023-05-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -10.0/1.5/090 | OS | Physician | 08/22/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17619862&pc=QCB | 2023826-2023-03686 | 2023-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/3.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -13.0/3.0/091 | OD | Physician | 08/22/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17619864&pc=QCB | 2023826-2023-03709 | 2023-02-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_ 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/2.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -8.5/2.0/091 | OD | Physician | 08/23/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17619867&pc=QCB | 2023826-2023-03708 | 2023-06-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_ 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/1.5/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -7.0/1.5/083 | OS | Physician | 08/23/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17619870&pc=QCB | 2023826-2023-03710 | 2023-02-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_ 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/1.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -9.5/1.5/088 | OS | Physician | 08/23/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17619876&pc=QCB | 2023826-2023-03707 | 2023-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_ 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/2.5/084(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -10.5/2.5/084 | OS | Physician | 08/23/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17624289&pc=QCB | 2023826-2023-03671 | 2023-06-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+0.5/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.5/+0.5/094 | OD | Physician | 07/31/2023 | 2023 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17624290&pc=QCB | 2023826-2023-03670 | 2023-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+2.0/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -14.0/+2.0/089 | OS | Physician | 07/31/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17624291&pc=QCB | 2023826-2023-03674 | 2023-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+2.0/069 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.5/+2.0/069 | OD | Physician | 07/31/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17624292&pc=QCB | 2023826-2023-03672 | 2023-06-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+1.5/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -6.5/+1.5/091 | OS | Physician | 07/31/2023 | 2023 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17624293&pc=QCB | 2023826-2023-03675 | 2023-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+2.0/111 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.5/+2.0/111 | OS | Physician | 07/31/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17624303&pc=QCB | 2023826-2023-03676 | 2022-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+2.0/082 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.5/+2.0/082 | OD | Physician | 07/31/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17627754&pc=QCB | 2023826-2023-03715 | 2023-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/2.0/096(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -9.0/2.0/096 | OD | Physician | 08/22/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17627755&pc=QCB | 2023826-2023-03716 | 2023-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/2.5/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -12.0/2.5/073 | OS | Physician | 08/22/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17627756&pc=QCB | 2023826-2023-03714 | 2023-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/1.5/014 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -9.5/1.5/014 | OD | Physician | 08/22/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17627864&pc=QCB | 2023826-2023-03712 | 2023-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/1.5/111 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -11.5/1.5/111 | OD | Physician | 08/22/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17628087&pc=QCB | 2023826-2023-03722 | 2023-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -11.5/3/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2023, THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS "LENS SIZE CHANGE". Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -11.5/3/91 | OD | Physician | 08/09/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17628423&pc=QCB | 2023826-2023-03732 | 2023-07-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 12.6MM VTICM5_12.6, -9.5/2.0/090, IMPLANTABLE COLLAMER LENS GOT STUCK IN THE CARTRIDGE AND TORE, THERE WAS PATIENT CONTACT. THE LENS WAS INTRAOPERATIVELY REMOVED WITH NO PATIENT INJURY. ON (B)(6) 2023 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6: TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -9.5/2.0/090 | NA | Physician | 07/31/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17628481&pc=QCB | 2023826-2023-03603 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -4.5/2.0/096 (SPHERE/CYLINDER/AXIS) TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THE LENS WAS IMPLANTED, REMOVED AND REPLACED INTRAOPERATIVELY WITH NO PATIENT INJURY AND THE PROBLEM RESOLVED. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: H6: TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -4.5/2.0/096 | OD | Physician | 07/31/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17633293&pc=QCB | 2023826-2023-03534 | 2023-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED VERTICALLY A 12.6MM VTICMO12.6, -10.0/1.5/174 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED HORIZONTALLY WITH A SAME LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #: (B)(4). Manufacturer Narrative: H6 - DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM# (B)(4). | -10.0/1.5/174 | OS | Physician | 08/06/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17633306&pc=QCB | 2023826-2023-03535 | 2023-02-27 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.5/1.0/148 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS TORE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED AND REMOVED ON (B)(6) 2023. ON (B)(6)2023 A REPLACEMENT LENS WAS IMPLANTED. PROBLEM RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, UNDER 21 YRS OF AGE AT DATE OF IMPLANT. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.5/1.0/148 | OD | Physician | 08/06/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17633313&pc=QCB | 2023826-2023-03537 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.50/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LENS ROTATION ON (B)(6) 2023. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2 - UNK, A3 - UNK, A4 - UNK, A5 - UNK, A6 - UNK, B3 - UNK, D6A - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: D6A - UNK THERE WAS NO CONFIRMATION THAT LENS WAS EXPLANTED/EXCHANGED ON (B)(6) 2023. H6 - HEALTH EFFECT - IMPACT CODE: REMOVED FROM INITAL MDR CLAIM # (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.50/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). LENS ROTATION WAS OBSERVED. IT IS UNKNOWN IF THE LENS REMAINS IMPLANTED. THERE WAS NO CONFIRMATION THAT LENS WAS EXPLANTED/EXCHANGED ON (B)(6) 2023. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). | -10.50/1.5/093 | OD | Other | 08/08/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17633323&pc=QCB | 2023826-2023-03539 | 2023-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.5/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED VERTICALLY WITH A A SAME LENGTH LENS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY. VISUAL INSPECTION FOUND RESIDUE/DEBRIS ON PRODUCT. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: H3 - PRODUCT WAS RETURNED IN A MICROCENTRIFUGE VIAL, DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.5/2.0/090 | OS | Physician | 08/09/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17633387&pc=QCB | 2023826-2023-03532 | 2022-11-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.5/1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED VERTICALLY WITH A SAME LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#:(B)(4). | -10.5/1.0/092 | OD | Physician | 08/06/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17635277&pc=QCB | 2023826-2023-03608 | 2022-11-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -10.5/4.0/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6)2022 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.5/4.0/087 | OD | Physician | 08/02/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17636337&pc=QCB | 2023826-2023-03609 | 2023-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -12.00/3.5/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS OF THE SAME LENGTH AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -12.00/3.5/094 | OD | Physician | 08/02/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17637150&pc=QCB | 2023826-2023-03744 | 2023-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -7.0/+1.0/108 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6-HEALTH EFFECT- IMPACT CODE: "2199" SHOULD BE REMOVED AND CODE "4629" SHOULD BE ADDED. H6- MEDICAL DEVICE CODE: "2993" SHOULD BE ADDED. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.00/1.0/108 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED A LOW VAULT. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL, LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "ALL FINE" AND "PATIENT IS HAPPY". CLAIM# (B)(4). | -7.0/+1.0/108 | OS | Physician | 08/09/2023 | 2023 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17637548&pc=QCB | 2023826-2023-03610 | 2023-06-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -16.0/5.0/098 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS LATER EXPLANTED DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -16.0/5.0/098 | OD | Physician | 08/02/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17637557&pc=QCB | 2023826-2023-03611 | 2023-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -10.00/3.00/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.00/3.00/086 | OS | Physician | 08/03/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17638480&pc=QCB | 2023826-2023-03726 | 2023-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/4.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON(B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -13.5/4.5/074 | OS | Physician | 08/11/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17638481&pc=QCB | 2023826-2023-03724 | 2023-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/1.5/032 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ON (B)(6) 2023, A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -11.0/1.5/032 | OD | Physician | 08/24/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17638483&pc=QCB | 2023826-2023-03725 | 2023-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/1.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ON (B)(6) 2023 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -9.0/1.5/092 | OS | Physician | 08/24/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17638484&pc=QCB | 2023826-2023-03729 | 2023-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_ 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/3.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING WITHOUT ROTATION. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -11.5/3.0/086 | OD | Physician | 08/23/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17638722&pc=QCB | 2023826-2023-03345 | 2023-06-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -17.5/2.0/109 (SPHERE/CYLINDER/AXIS) WAS REMOVED FROM THE PATIENTS RIGHT EYE (OD) WITHIN SAME SURGERY INTRA-OPERATIVELY ON (B)(6) 2023 DUE TO LENS TEAR/BREAK BEFORE/DURING LOADING AFTER OPENING THE VIAL. A REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2023 OF THE SAME MODEL AND SIZE. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -17.5/2.0/109 | OD | Physician | 08/04/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17638815&pc=QCB | 2023826-2023-03746 | 2023-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.5/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+1.5/092 | OD | Physician | 08/01/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17638816&pc=QCB | 2023826-2023-03749 | 2023-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.5/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -11.5/+1.5/091 | OS | Physician | 08/02/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17638818&pc=QCB | 2023826-2023-03745 | 2023-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+2.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.5/+2.0/091 | OD | Physician | 08/01/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17638819&pc=QCB | 2023826-2023-03751 | 2023-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.0/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS TWO SIZES LONGER IN LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -10.0/+1.0/094 | OD | Physician | 08/02/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17638820&pc=QCB | 2023826-2023-03748 | 2023-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+2.5/085 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -7.0/+2.5/085 | OD | Physician | 08/02/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17638870&pc=QCB | 2023826-2023-03614 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -12.00/1.5/152 (SPHERE/CYLINDER/AXIS) WAS PART OF A LENS EXCHANGE DUE TO THE LENS BEING TOO SMALL, DOCTOR WENT WITH A 12.6MM LENS INSTEAD. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -12.00/1.5/152 | NA | Other Health Care Professional | 08/02/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17638921&pc=QCB | 2023826-2023-03615 |  | Injury | STAAR SURGICAL COMPANY | 2023-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -13.00/2.0/081 (SPHERE/CYLINDER/AXIS) WAS PART OF A LENS EXCHANGE DUE TO THE LENS BEING TOO SMALL, DOCTOR WENT WITH A 12.6MM LENS INSTEAD. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.00/2.0/081 | NA | Other Health Care Professional | 08/02/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17646186&pc=QCB | 2023826-2023-03796 | 2023-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.5/2.0/174 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT IMPLANTED HORIZONTALLY AS OPPOSED TO THE INITIAL VERTICAL IMPLANT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) CLAIM# (B)(4). | -9.5/2.0/174 | OD | Physician | 08/22/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17646187&pc=QCB | 2023826-2023-03793 | 2023-04-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -9.00/+4.0/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING AND LENS ROTATION. THE SURGEON HAD REPOSITIONED THE LENS 3 TIMES, ON (B)(6) 2023, (B)(6) 2023, AND (B)(6) 2023. BUT THIS DID NOT RESOLVE THE PROBLEM. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONL DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+4.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED A LOW VAULT WITH ROTATION AND REPOSITIONING. THE LENS WAS REPOSITIONED 3 TIMES, ON (B)(6) 2023, 05/JUNE/2023, AND (B)(6) 2023 BUT THAT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH THE SAME MODEL, LONGER LENGTH AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "ALL FINE" AND "PATIENT IS HAPPY". THE CAUSE OF THE EVENT WAS UNKNOWN. H6-HEALTH EFFECT- IMPACT CODE: "4629" SHOULD BE ADDED. H6- MEDICAL DEVICE CODE:"2993" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B3: "04/07/2023" SHOULD BE CORRECTED TO "04/27/2023". CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.00/+4.0/094 | OD | Physician | 08/09/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17647018&pc=QCB | 2023826-2023-03797 | 2023-08-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.5/3.0/171 (SPHERE / CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT IMPLANTED HORIZONTALLY AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21 YRS OF AGE AT DATE OF IMPLANT). H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -13.5/3.0/171 | OS | Physician | 08/22/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17647193&pc=QCB | 2023826-2023-03755 | 2023-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_ 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -7.0/2.0/092 | OD | Physician | 08/25/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17647194&pc=QCB | 2023826-2023-03175 | 2023-07-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.0/2.0/161(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -17.0/2.0/161 | OD | Physician | 07/25/2023 | 2023 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17647195&pc=QCB | 2023826-2023-03756 | 2023-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/0.5/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM (B)(4). | -6.5/0.5/085 | OS | Physician | 08/24/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17647196&pc=QCB | 2023826-2023-03754 | 2023-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_ 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/5.5/136 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ON (B)(6) 2023 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -6.0/5.5/136 | OD | Physician | 08/25/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17653303&pc=QCB | 2023826-2023-03376 | 2023-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/1.5/173 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: D4 - CORRECTED TO VTICMO12.6. CLAIM#: (B)(4). Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). H6-DEVICE CODE 1494: OFF-LABEL USE (OVER 45 YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -11.5/1.5/173 | OS | Physician | 08/07/2023 | 2023 | 46 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17653647&pc=QCB | 2023826-2023-03616 | 2023-07-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023, A 12.6MM, VTICMO12.6, -8.0/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE RIGHT EYE (OD). THE LENS WAS REMOVED INTRAOPERATIVELY WITH NO PATIENT INJURY. ON (B)(6) 2023 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # 730968 | -8.0/2.0/089 | OD | Physician | 08/07/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17655250&pc=QCB | 2023826-2023-03769 | 2023-03-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -15.0/5.5/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE PROBLEM DID NOT RESOLVE. ON (B)(6) 2023 THE LENS WAS EXPLANTED. CAUSE OF EVENT WAS UNKNOWN. PER ADDITIONAL INFORMATION THE PATIENT REQUESTED THAT IT BE REMOVED AND NOT RE-IMPLANTED. Manufacturer Narrative: ADDITIONAL DATA: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -15.0/5.5/089 | OD | Physician | 08/08/2023 | 2023 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17655778&pc=QCB | 2023826-2023-03770 | 2023-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -14.0/4.0/081 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE PROBLEM DID NOT RESOLVE. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS OF DIFFERENT AXIS BUT THE PROBLEM DID NOT RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -14.0/4.0/081 | OS | Physician | 08/08/2023 | 2023 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17656068&pc=QCB | 2023826-2023-03771 | 2023-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -14.0/4.0/176 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 AS A REPLACEMENT LENS TO ADDRESS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. THE REPORTER STATES THAT THE LENS STILL ROTATED. ON (B)(6) 2023 THE LENS WAS EXPLANTED. PER ADDITIONAL INFORMATION THE PATIENT REQUESTED THAT IT BE REMOVED AND NOT RE-IMPLANTED .CAUSE OF EVENT WAS UNKNOWN. SEE MFR # 2023826-2023-03770 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -14.0/4.0/176 | OS | Physician | 08/08/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17656079&pc=QCB | 2023826-2023-03700 | 2023-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 OF -5.50/2.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS NOT RESOLVED. REFERENCE MFR # 2023826-2023-03766 FOR EXCHANGE LENS THAT RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: B4- DATE OF THIS REPORT: "08/01/1991" SHOULD BE REMOVED AND "08/18/2023" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -5.50/2.0/091 | OS | Physician | 08/18/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17656779&pc=QCB | 2023826-2023-03803 | 2023-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/1.5/173 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -8.0/1.5/173 | OS | Physician | 08/28/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17656780&pc=QCB | 2023826-2023-03806 | 2023-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.5/1.0/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -16.5/1.0/072 | OS | Physician | 08/26/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17657367&pc=QCB | 2023826-2023-03779 | 2023-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.0/2.0/113 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6)2023, THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS, IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT. REPORTEDLY, "ICL LENS IN VERTICAL POSITION." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | -9.0/2.0/113 | OD | Physician | 08/09/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17658848&pc=QCB | 2023826-2023-03781 | 2023-04-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/175 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS, IMPLANTED VERTICALLY DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "FIRST VERTICAL POSITION" THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -12.0/1.5/175 | OS | Physician | 08/10/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17669063&pc=QCB | 2023826-2023-03529 | 2023-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED OBLIQUELY A 12.6MM VTICMO12.6, -11.0/1.5/052 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED HORIZONTALLY WITH A SAME LENGTH, SIMILAR POWER (SPHERE/CYLINDER/AXIS) LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: H6 - DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM, UNDER 21 YRS OF AGE AT DATE OF IMPLANT CLAIM# (B)(4). | -11.0/1.5/052 | OS | Physician | 08/14/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17669064&pc=QCB | 2023826-2023-03454 | 2023-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.5/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.5/1.5/091 | OD | Physician | 08/14/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17669073&pc=QCB | 2023826-2023-03741 | 2023-06-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.0/1.5/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - A6: UNK. H6 - DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, UNDER 21 YRS OF AGE AT DATE OF IMPLANT. (B)(4). | -12.0/1.5/070 | OD | Physician | 08/17/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17672680&pc=QCB | 2023826-2023-03655 | 2023-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/081 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL LENGTH VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). | -10.00/+2.0/081 | OS | Physician | 08/08/2023 | 2023 | 43 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17672687&pc=QCB | 2023826-2023-03653 | 2023-06-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.50/+1.5/099 (SPHERE/CYLINDER/AXIS), TORE DURING LOADING AND THERE WAS NO PATIENT CONTACT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.50/+1.5/099 | NA | Physician | 08/08/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17672691&pc=QCB | 2023826-2023-03835 |  | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50/+1.5/151 (SPHERE/CYLINDER/AXIS). THE LENS FLIPPED INSIDE THE EYE, THE SURGEON ATTEMPTED TO REPOSITION THE LENS BUT WAS UNSUCCESSFUL. THE LENS WAS REMOVED AND THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B5-THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/+1.5/151 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS), REPORTEDLY, THE "LENS FLIPPED INSIDE THE EYE, SURGEON TRIED REPOSITION THE LENS, HOWEVER HE COULDN'T SO THE BACKUP LENS WAS USED". THE LENS WAS REPOSITIONED. THE LENS WAS EXPLANTED AND A BACK UP LENS OF THE SAME MODEL, LENGTH WAS IMPLANTED. REPORTEDLY, "I HAVE HAD NO NEWS THAT THERE ARE ANY ISSUES OTHERWISE WITH THE PATIENT/IMPLANTED LENSES, I WOULD SAY THAT NO NEWS IS GOOD NEWS!". H6-HEALTH EFFECT- IMPACT CODE: "4629" SHOULD BE REMOVED AND CODE "4627" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -14.50/+1.5/151 | OS | Physician | 08/03/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17672700&pc=QCB | 2023826-2023-03836 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.50/+1.5/174 (SPHERE/CYLINDER/AXIS). THE LENS FLIPPED INSIDE THE EYE, THE SURGEON ATTEMPTED TO REPOSITION THE LENS BUT WAS UNSUCCESSFUL. THE LENS WAS REMOVED AND THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B5: THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -15.50/+1.5/174 (SPHERE/CYLINDERL/AXIS) INTO THE PATIENTS LEFT EYE (OS). THIS WAS THE REPLACEMENT LENS FOR MFR REPORT # 2023-2023-03835 AND THE LENS REMAINED IMPLANTED. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- ORIGINALLY THE CLAIM WAS DETERMINED TO BE REPORTABLE, BUT UPON FURTHER REVIEW WAS DETERMINED TO NOT MEET DEFINITION OF A COMPLAINT. HOWEVER, THE CLAIM CANNOT BE VOIDED AS AN MDR HAS ALREADY BEEN SUBMITTED. REPLACEMENT/BACK UP LENS FOR MFR REPORT # 2023-2023-03835. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -15.50/+1.5/174 | OS | Physician | 08/03/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17672746&pc=QCB | 2023826-2023-03654 | 2023-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/101 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). | -11.50/+2.0/101 | OD | Physician | 08/08/2023 | 2023 | 43 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17672865&pc=QCB | 2023826-2023-03830 | 2023-05-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.5/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -10.0/+2.5/092 | OD | Physician | 08/06/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17672866&pc=QCB | 2023826-2023-03762 | 2023-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+3.5/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS TWO SIZES SHORTER IN LENGTH DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.0/+3.5/093 | OD | Physician | 08/07/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17672870&pc=QCB | 2023826-2023-03761 | 2023-07-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS WITH DIOPTER -10.0/+3.0/001 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.5/2.0/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. CLAIM#: (B)(4). | -10.0/+3.0/001 | OD | Physician | 08/07/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17672873&pc=QCB | 2023826-2023-03829 | 2023-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+2.0/084 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6 - MEDICAL DEVICE PROBLEM CODE 1494: OFF-LABEL USE (UNDER 21 YRS OF AGE AT DATE OF IMPLANTATION) . CLAIM# (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | -14.0/+2.0/084 | OS | Physician | 08/06/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17672883&pc=QCB | 2023826-2023-03759 | 2023-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.0/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN IT IS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: H6-DEVICE CODE 1494: (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.5/2.5/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. - CORRECTION TO INITIAL MDR. CLAIM#: (B)(4). | -12.5/2.5/92 | OS | Physician | 08/07/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17672892&pc=QCB | 2023826-2023-03764 | 2023-07-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0/+4.0/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) AND (ACD < 3.0MM). CLAIM# (B)(4). | -15.0/+4.0/094 | OD | Physician | 08/07/2023 | 2023 | 17 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17672912&pc=QCB | 2023826-2023-03766 | 2023-07-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/+2.0/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: REFERENCE MFR # 2023826-2023-03700 FOR INITIAL LENS IMPLANT. CLAIM#(B)(4). Manufacturer Narrative: H6: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -5.5/+2.0/091 | OS | Physician | 08/07/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17673543&pc=QCB | 2023826-2023-03657 | 2023-06-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.50/+3.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -15.50/+3.0/095 | OD | Physician | 08/08/2023 | 2023 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17673569&pc=QCB | 2023826-2023-03763 | 2023-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+3.5/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS TWO SIZES SHORTER IN LENGTH DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.0/+3.5/092 | OS | Physician | 08/07/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17674130&pc=QCB | 2023826-2023-03789 | 2023-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.0/4.5/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LENS ROTATION NOT ASSOCIATED TO LOW VAULT WAS OBSERVED. ON (B)(6) 2023, THE LENS WAS REPOSITIONED, AND THIS RESOLVED THE PROBLEM. REPORTEDLY, "EYE IS QUIET VAULT 410 MICRON BUT AGAIN THE LENS HAS ROTATED BY MORE THEN 30 DEGREES" AND "PATIENT NOT HAPPY WITH VISION." THE PROBLEM WAS NOT RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM (B)(4). | -10.0/4.5/089 | OD | Physician | 08/14/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17674133&pc=QCB | 2023826-2023-03790 | 2023-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -6.50/3.5/84 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LENS ROTATION NOT ASSOCIATED TO LOW VAULT WAS OBSERVED. ON (B)(6) 2023, THE LENS WAS REPOSITIONED, AND THIS RESOLVED THE PROBLEM. REPORTEDLY, "EYE IS QUIET VAULT 420 MICRON BUT AGAIN THE LENS HAS ROTATED BY MORE THEN 30 DEGREES" AND "PATIENT NOT HAPPY WITH VISION." THE PROBLEM WAS NOT RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. THE REPORTER STATED "LENS ROTATED VAULT 420 MICRON." Manufacturer Narrative: A4-A6:UNK H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -6.50/3.5/84 | OS | Physician | 08/14/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17674550&pc=QCB | 2023826-2023-03861 | 2023-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -8.0/1.0/170 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2023, THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT IMPLANTED HORIZONTAL AS OPPOSED TO THE INITIAL VERTICAL IMPLANTATION. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -8.0/1.0/170 | OS | Physician | 08/24/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17674551&pc=QCB | 2023826-2023-03860 | 2023-08-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -8.5/1.0/74 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS WITHOUT PATIENT INJURY. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED THE DEVICE DID FAIL TO PERFORM AS INTENDED. IT WAS ADDITIONALLY NOTED "THE PATIENT'S LEFT EYE WAS PLANNED FOR IMPLANTATION OF A 13.7MM IN THE VERTICAL POSITION. HOWEVER, DUE TO CRYSTALLINE LENS RUPTURE THERE WAS NO AVAILABLE ICL WITH THE SAME SIZE AND REFRACTIVE POWER. THEREFORE A 13.MM ICL WAS SELECTED AND PLACED IN THE HORIZONTAL POSITION. THE RESIDUAL ASTIGMATISM CORRECTION WAS 0.5 DIOPTERS AND THE REMAINING REFRACTIVE ERROR WAS DEEMED ACCEPTABLE". Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -8.5/1.0/74 | OS | Physician | 08/24/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17674677&pc=QCB | 2023826-2023-03847 | 2023-07-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS WITH DIOPTER -10.0/+3.0/001 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. THAT SAME DAY, THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS ALTHOUGH IT IS UNCLEAR WHETHER THIS WAS WITHIN THE SAME PROCEDURE. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN AND IT IS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: (B)(4). | -10.0/+3.0/001 | OD | Physician | 08/07/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17674680&pc=QCB | 2023826-2023-03842 | 2023-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+3.0/089 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023, FOR A LENS TWO SIZES LONGER IN LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -7.0/+3.0/089 | OD | Physician | 08/07/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17674685&pc=QCB | 2023826-2023-03850 | 2023-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.5/070 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023, DUE TO LOW VAULT. A LONGER LENGTH LENS WAS IMPLANTED ON (B)(6) 2023, AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -8.5/+1.5/070 | OS | Physician | 08/07/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17675136&pc=QCB | 2023826-2023-03858 | 2023-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/168 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+1.0/168 | OD | Physician | 08/08/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17675137&pc=QCB | 2023826-2023-03856 | 2023-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+2.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN AND IT IS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: CLAIM# (B)(4). | -8.0/+2.0/093 | OD | Physician | 08/08/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17675139&pc=QCB | 2023826-2023-03857 | 2023-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+2.0/070 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN AND IT IS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: CLAIM# (B)(4). | -7.5/+2.0/070 | OS | Physician | 08/08/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17675141&pc=QCB | 2023826-2023-03852 | 2023-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.0/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN AND IT IS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -10.0/+2.0/092 | OD | Physician | 08/07/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17675142&pc=QCB | 2023826-2023-03854 | 2023-05-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/102 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: (B)(4). | -10.0/+1.5/102 | OS | Physician | 08/08/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17675143&pc=QCB | 2023826-2023-03859 | 2023-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+1.0/069 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN AND IT IS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: A2- DOB: CORRECTED TO (B)(6) 2004 . H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) . CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -12.0/+1.0/069 | OD | Physician | 08/09/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17684951&pc=QCB | 2023826-2023-03747 | 2023-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.0/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM (B)(4). | -10.5/+2.0/091 | OS | Physician | 08/02/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17686935&pc=QCB | 2023826-2023-03907 | 2023-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -15.0/1.0/085 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS ROTATION. ON (B)(6) 2023, THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT SPHERICAL LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IT WAS ADDITIONALLY NOTED "AFTER THE OPERATION, THE CRYSTAL ROTATES, AND THE POSITION IS NOT FIXED WITH TIME. SO THE SURGEON DECIDED TO TAKE OUT TICL CRYSTAL AND REPLACE IT WITH ICL CRYSTAL. AFTER REPLACEMENT, THE VISION WAS 20/20, THE INTRAOCULAR PRESSURE WAS 22MMHG, AND THE ARCH HEIGHT OF THE CRYSTAL WAS 0.8CT." Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -15.0/1.0/085 | OS | Physician | 08/31/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17686936&pc=QCB | 2023826-2023-03906 | 2023-01-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -9.5/3.0/081 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION. ON (B)(6) 2023, THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023, THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT IMPLANTED VERTICALLY AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -9.5/3.0/081 | OS | Physician | 08/27/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17686937&pc=QCB | 2023826-2023-03905 | 2023-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -15.0/2.5/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT DIFFERENT AXIS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -15.0/2.5/095 | OS | Physician | 08/28/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17686938&pc=QCB | 2023826-2023-03904 | 2023-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -16.0/1.0/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS ROTATION. ON (B)(6) 2023 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. THEN ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT DIFFERENT AXIS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -16.0/1.0/095 | OD | Physician | 08/25/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17686965&pc=QCB | 2023826-2023-03911 | 2023-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.0/071 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4) | -10.0/+2.0/071 | OS | Physician | 08/10/2023 | 2023 | 45 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17686966&pc=QCB | 2023826-2023-03910 | 2023-02-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+2.0/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN AND IT IS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4) | -8.5/+2.0/093 | OS | Physician | 08/10/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17686967&pc=QCB | 2023826-2023-03908 | 2023-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.5/+1.5/068 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023, FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -16.5/+1.5/068 | OD | Physician | 08/10/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17686968&pc=QCB | 2023826-2023-03912 | 2023-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+3.5/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. DUE TO LOW VAULT, THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS AND IT WAS IMPLANTED VERTICALLY. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+3.5/090 | OD | Physician | 08/10/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17686970&pc=QCB | 2023826-2023-03909 | 2023-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+2.0/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.0/+2.0/095 | OD | Physician | 08/10/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17687289&pc=QCB | 2023826-2023-03913 | 2023-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+3.0/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. DUE TO LOW VAULT, THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.0/+3.0/094 | OS | Physician | 08/10/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17687423&pc=QCB | 2023826-2023-03823 | 2023-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/1.5/134 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -13.5/1.5/134 | OD | Physician | 08/28/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17687424&pc=QCB | 2023826-2023-03821 | 2023-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK, (B)(4). | -10.0/1.0/091 | OD | Physician | 08/28/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17687425&pc=QCB | 2023826-2023-03825 | 2023-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/3.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -8.0/3.5/091 | OD | Physician | 08/29/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17687426&pc=QCB | 2023826-2023-03824 | 2023-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/1.0/075 (SPHERE / CYLINDER / AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -14.0/1.0/075 | OS | Physician | 08/28/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17687427&pc=QCB | 2023826-2023-03826 | 2023-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/3.5/093(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK, (B)(4). | -8.5/3.5/093 | OS | Physician | 08/29/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17687432&pc=QCB | 2023826-2023-03827 | 2023-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/4.0/102(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. (B)(4). | -12.5/4.0/102 | OD | Physician | 08/29/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17687433&pc=QCB | 2023826-2023-03828 | 2023-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 132MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/2.5/087(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. CLAIM #(B)(4). | -12.0/2.5/087 | OD | Physician | 08/30/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17695342&pc=QCB | 2023826-2023-03933 | 2023-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 132MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/1.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -13.5/1.0/086 | OD | Physician | 08/30/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17695343&pc=QCB | 2023826-2023-03935 | 2023-05-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/2.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON 30-MAY-2023. ON 31-MAY-2023 THE LENS WAS EXCHANGED FOR A SAME MODEL/ LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "ACCORDING TO THE ASSESSMENT OF POST-OP ARCH, CHANGE THE ICL POSITION FROM HORIZONTAL TO VERTICAL." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -7.5/2.5/092 | OD | Physician | 08/31/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17695344&pc=QCB | 2023826-2023-03934 | 2023-05-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.0/2.0/079 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -3.0/2.0/079 | OD | Physician | 08/30/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17695353&pc=QCB | 2023826-2023-03937 | 2023-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/1.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -12.5/1.5/094 | OD | Physician | 09/01/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17695355&pc=QCB | 2023826-2023-03936 | 2023-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/1.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL/ LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "ACCORDING TO THE ASSESSMENT OF POST-OP ARCH, CHANGE THE ICL POSITION FROM HORIZONTAL TO VERTICAL." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | -12.5/1.0/093 | OD | Physician | 08/31/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17695773&pc=QCB | 2023826-2023-03888 | 2023-08-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF A -6.5/1.5/090 (SPHERE/CYLINDER/AXIS) WAS REPORTEDLY, "CRACKED." A REPLACEMENT LENS WAS IMPLANTED. Manufacturer Narrative: A4-A6:UNK H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM (B)(4). Manufacturer Narrative: B5: AN IMPLANTABLE COLLAMER LENS WAS INTRAOPERATIVELY IMPLANTED AND REMOVED FROM THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023, DUE TO A LENS TEAR/BREAK DURING INJECTION DELIVERY INTO THE EYE. NO PATIENT INJURY. THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). | -6.5/1.5/090 | OD | Physician | 08/15/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17695774&pc=QCB | 2023826-2023-03889 | 2023-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.0/2.0/101 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL AGE<21, (B)(4). | -11.0/2.0/101 | OD | Physician | 08/15/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17695796&pc=QCB | 2023826-2023-03945 | 2023-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.5/174 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "THE LENS POSITION IS ADJUSTED FROM VERTICAL TO HORIZONTAL." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK, (B)(4). | -10.0/1.5/174 | OS | Physician | 09/03/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17695797&pc=QCB | 2023826-2023-03942 | 2023-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/1.0/105 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -9.0/1.0/105 | OD | Physician | 09/01/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17695798&pc=QCB | 2023826-2023-03940 | 2023-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.50/3.0/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: HEIGHT-WEIGHT: UNK. CLAIM #(B)(6). | -14.50/3.0/072 | OS | Physician | 09/01/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17695800&pc=QCB | 2023826-2023-03941 | 2023-08-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_ 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/3.5/081 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -8.0/3.5/081 | OD | Physician | 09/01/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17695801&pc=QCB | 2023826-2023-03939 | 2023-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/1.5/106 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -11.5/1.5/106 | OD | Physician | 09/01/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17695802&pc=QCB | 2023826-2023-03943 | 2023-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/2.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -13.5/2.5/095 | OD | Physician | 09/03/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17696302&pc=QCB | 2023826-2023-03774 | 2023-05-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -15.5/1.0/102 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - PRODUCT EVALUATION FOUND LENS IN RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -15.5/1.0/102 | OD | Physician | 08/10/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17699141&pc=QCB | 2023826-2023-03949 | 2023-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM512.6 -4.00/1.50/020 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REPLACED WITH A SAME LENGTH BUT DIFFERENT MODEL AND DIOPTER LENS DUE TO THE SURGEON IMPLANTING THE LENS FOR THE RIGHT EYE (OD) INTO THE LEFT EYE (OS). THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CLAIM#(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: PER THE REPORTER: STATUS OF THE EYE IS "ALL FINE. PATIENT IS HAPPY! THE SURGEON WILL CLOSELY MONITOR THE VAULT." CLAIM# (B)(4). | -4.00/1.50/020 | OS | Physician | 08/09/2023 | 2023 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17704398&pc=QCB | 2023826-2023-03877 | 2023-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -09.50/+2.0/095 (SPHERE/CYLINDER/AXIS), (IMPLANTED HORIZONTALLY) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: 08-SEP-2023 H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE WITH RESIDUE/DEBRIS ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) | -09.50/+2.0/095 | OS | Physician | 08/16/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17704431&pc=QCB | 2023826-2023-03882 | 2023-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.50/+1.5/104 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023, DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT WAS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. CLAIM # (B)(4). | -08.50/+1.5/104 | OD | Physician | 08/16/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17704439&pc=QCB | 2023826-2023-03876 | 2023-03-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.50/+2.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THERE WAS LENS DISLOCATION/SUBLUXATION. THE LENS REMAINED IMPLANTED. POST-OP, THE PATIENT HAD DIPLOPIA. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -16.50/+2.5/095 | OD | Physician | 08/15/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17705477&pc=QCB | 2023826-2023-03776 | 2023-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -11.0/3.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXPLANTED DUE TO LOW VAULT WITH ROTATION. A REPLACEMENT LENS OF LONGER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED.THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. THE REPORTER ALSO STATES THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.0/3.0/091 | OD | Physician | 08/11/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17705823&pc=QCB | 2023826-2023-03778 | 2023-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -6.5/5.0/100 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS LATER EXPLANTED DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. THE REPORTER STATES THE CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLIAM# (B)(4). | -6.5/5.0/100 | OS | Physician | 08/15/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17715163&pc=QCB | 2023826-2023-03903 | 2023-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -9.5/1.0/064 (SPHERE/CYLINDER/AXIS) LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENT'S RIGHT EYE (OD). THE HAPTICS ARE DAMAGED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT WAS USER ERROR. Manufacturer Narrative: A4-A6:UNK. H6: USER ERROR. CLAIM#: (B)(4). | -9.5/1.0/064 | OD | Physician | 08/25/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17715568&pc=QCB | 2023826-2023-03869 | 2023-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.00/+1.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2023, DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PROBLEM WAS RESOLVED. THE LENS WAS REPOSITIONED AND ROTATED MULTIPLE TIMES BEFORE BEING REMOVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3: TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.00/+1.0/090 | OD | Physician | 08/14/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17715582&pc=QCB | 2023826-2023-03883 | 2023-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.50/+0.50/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT WAS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK . A6: UNK. CLAIM # (B)(4). | -08.50/+0.50/091 | OS | Physician | 08/18/2023 | 2023 | 41 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17715624&pc=QCB | 2023826-2023-03870 | 2023-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.50/+1.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK . A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.50/+1.0/088 | OS | Physician | 08/14/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17715638&pc=QCB | 2023826-2023-03885 | 2023-08-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50/+1.5/175 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -08.50/+1.5/175 | OS | Physician | 08/21/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17715893&pc=QCB | 2023826-2023-03976 | 2023-08-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.5/2.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SAME LENGTH LENS AND THIS RESOLVED THE PROBLEM. THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) CLAIM(B)(4). | -12.5/2.0/091 | OD | Physician | 09/05/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17716164&pc=QCB | 2023826-2023-03915 | 2023-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.0/115 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/+1.0/115 | OD | Physician | 08/14/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17716166&pc=QCB | 2023826-2023-03916 | 2023-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.0/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. DUE TO LOW VAULT, THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/+1.0/091 | OS | Physician | 08/14/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17716168&pc=QCB | 2023826-2023-03919 | 2023-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+2.0/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. DUE TO LOW VAULT, THE LENS WAS EXCHANGED ON (B)(6) 2023, FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.5/+2.0/092 | OD | Physician | 08/14/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17716169&pc=QCB | 2023826-2023-03914 | 2023-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+2.5/089 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -8.0/+2.5/089 | OD | Physician | 08/14/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17716171&pc=QCB | 2023826-2023-03920 | 2023-04-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.0/069 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. DUE TO LOW VAULT, THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+1.0/069 | OS | Physician | 08/14/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17716242&pc=QCB | 2023826-2023-03967 | 2023-08-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/2.0/169 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -11.0/2.0/169 | OS | Physician | 09/05/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17716243&pc=QCB | 2023826-2023-03968 | 2023-08-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -14.0/2.0/092 | OD | Physician | 09/07/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17716244&pc=QCB | 2023826-2023-03966 | 2023-08-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/3.0/012 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -11.0/3.0/012 | OD | Physician | 09/05/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17716245&pc=QCB | 2023826-2023-03965 | 2023-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -130/2.0/092(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -130/2.0/092 | OD | Physician | 09/05/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17716246&pc=QCB | 2023826-2023-03964 | 2023-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/1.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(6). | -7.0/1.5/094 | OS | Physician | 09/05/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17716247&pc=QCB | 2023826-2023-03963 | 2023-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/1.0/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-HEIGHT: UNK. CLAIM #(B)(4). | -7.0/1.0/074 | OD | Physician | 09/05/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17716249&pc=QCB | 2023826-2023-03969 | 2023-08-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.0/3.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) CLAIM(B)(4). | -5.0/3.0/094 | OD | Physician | 09/07/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17716978&pc=QCB | 2023826-2023-03863 | 2023-08-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.5/089 (SPHERE/CYLINDER/AXIS) WAS TOO SMALL AND THE SURGEON WENT WITH A 13.2MM LENS AFTER OPENING THE LENS. UNSURE IF THERE WAS PATIENT CONTACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. D6A: UNK. D6B: UNK . H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.50/+1.5/089 | NA | Other Health Care Professional | 08/15/2023 | 2023 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17720815&pc=QCB | 2023826-2023-03958 | 2023-07-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -13.0/1.5/070 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -13.0/1.5/070 | OS | Physician | 08/15/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17725062&pc=QCB | 2023826-2023-03960 | 2023-07-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5/2.0/092 (SPHERE/CYLINDER/AXIS) TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS INTRAOPERATIVELY REMOVED AND REPLACED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4). | -11.5/2.0/092 | OS | Physician | 08/16/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17725786&pc=QCB | 2023826-2023-03961 | 2023-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -12.5/4.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS THAT WAS IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: H3 - PRODUCT EVALUATION FOUND LENS RETURNED MICROCENTRIFUGE VIAL IN LIQUID WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.5/4.0/092 | OD | Physician | 08/16/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17725936&pc=QCB | 2023826-2023-03962 | 2023-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -11.0/3.5/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS THAT WAS IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: H3 - PRODUCT EVALUATION FOUND LENS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE, RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.0/3.5/094 | OS | Physician | 08/16/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17726008&pc=QCB | 2023826-2023-03970 | 2023-04-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/2.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO LOW VAULT WITHOUT ROTATION. ON (B)(6) 2023 A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -9.0/2.5/074 | OS | Physician | 09/07/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17726009&pc=QCB | 2023826-2023-03972 | 2023-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/4.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -10.5/4.0/090 | OD | Physician | 09/07/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17726010&pc=QCB | 2023826-2023-03974 | 2023-08-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -11.0/2.0/092 | OD | Physician | 09/08/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17726011&pc=QCB | 2023826-2023-03971 | 2023-08-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.0/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - CORRECTED TO: "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.0/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN." IN THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -10.5/1.0/073 | OD | Physician | 09/07/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17726012&pc=QCB | 2023826-2023-03973 | 2023-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -9.0/2.0/092 | OS | Physician | 09/08/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17726148&pc=QCB | 2023826-2023-03881 | 2023-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.00/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023, DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2023, BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH A LONGER LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.00/+2.0/090 | OD | Physician | 08/17/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17726149&pc=QCB | 2023826-2023-03880 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.00/+3.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) (IMPLANTED HORIZONTALLY), ON (B)(6) 2023. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-03879, BUT THE PROBLEM WAS NOT RESOLVED, THERE WAS NO DIFFERENCE IN ARCH HEIGHT AFTER REPLACEMENT, THE PATIENT IS UNDER FOLLOW-UP OBSERVATION AT PRESENT. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: B5- IT WAS LATER REPORTED, "AFTER THE REPLACEMENT OF THE 12.6MM HORIZONTAL IMPLANTED LENS, THE VAULT WAS 50UM IN THE LAST REVIEW. THE PATIENT HAD NO OTHER COMPLAINTS AND HAD GOOD VISION. AFTER COMMUNICATION, HE DECIDED TO CONTINUE TO OBSERVE FOR A PERIOD OF TIME". CLAIM# (B)(4). Manufacturer Narrative: A4: UNK, A5: UNK, A6: UNK, B3: UNK, H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4) | -13.00/+3.0/088 | OS | Physician | 08/17/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17726161&pc=QCB | 2023826-2023-03878 | 2023-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -05.00/+1.0/003 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), (IMPLANTED VERTICALLY), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED HORIZONTALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | -05.00/+1.0/003 | OD | Physician | 08/17/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17726162&pc=QCB | 2023826-2023-03886 | 2023-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+5.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENS BUT THE PROBLEM WAS NOT RESOLVED. THE REPLACEMENT LENS HAD ROTATED AND WAS UNSTABLE. SEE MFR. REPORT # 2023826-2023-03887 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 19-OCT-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH SEVERE DAMAGE. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND RESIDUE FIBERS ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS WERE UNABLE TO BE PERFORMED DUE TO SIGNIFICANT LENS DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -11.50/+5.0/092 | OS | Physician | 08/18/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17726196&pc=QCB | 2023826-2023-03879 | 2023-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.00/+3.0/175 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) (IMPLANTED VERTICALLY), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023, DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED HORIZONTALLY) BUT THE PROBLEM WAS NOT RESOLVED. THERE WAS NO DIFFERENCE IN ARCH HEIGHT AFTER REPLACEMENT, THE PATIENT IS UNDER FOLLOW-UP OBSERVATION AT PRESENT. SEE MFR. REPORT # 2023826-2023-03880, FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | -13.00/+3.0/175 | OS | Physician | 08/17/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17730452&pc=QCB | 2023826-2023-03923 | 2022-09-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.0/083 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -10.0/+1.0/083 | OS | Physician | 08/14/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17730457&pc=QCB | 2023826-2023-03922 | 2022-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.0/085 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -11.0/+1.0/085 | OD | Physician | 08/14/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17730458&pc=QCB | 2023826-2023-03924 | 2023-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+2.0/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN AND IT IS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: (B)(4). | -13.5/+2.0/092 | OS | Physician | 08/14/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17730460&pc=QCB | 2023826-2023-03930 | 2023-07-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+2.0/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM (B)(4). | -8.0/+2.0/094 | OS | Physician | 08/14/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17730463&pc=QCB | 2023826-2023-03926 | 2023-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+2.0/088 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN AND IT IS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -6.5/+2.0/088 | OD | Physician | 08/14/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17730464&pc=QCB | 2023826-2023-04019 | 2023-08-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS WITH DIOPTER -15.5/+4.5/162 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN AND IT IS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: (B)(4). | -15.5/+4.5/162 | OS | Physician | 08/18/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733289&pc=QCB | 2023826-2023-04027 | 2023-07-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_ 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/1.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ON (B)(6) 2023 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -11.0/1.5/092 | OD | Physician | 09/11/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733290&pc=QCB | 2023826-2023-04025 | 2023-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/3.5/113 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK, (B)(4). | -13.0/3.5/113 | OD | Physician | 09/11/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733369&pc=QCB | 2023826-2023-03977 |  | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A TORIC ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENTS EYE AND THE LENS HAD ROTATED. THE LENS WAS SUCCESSFULLY REPOSITIONED AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+4.0/089, INTO THE PATIENTS LEFT EYE (OS). THE LENS HAD ROTATED AND WAS SUCCESSFULLY REPOSITIONED. THE LENS REMAINED IMPLANTED. CLAIM # (B)(4). Manufacturer Narrative: A2: UNK. A4: UNK. A5: UNK. A6: UNK. B3: UNK. D4: EXPIRATION DATE UNK. D6A: UNK. D6B: UNK. H4: UNK. CLAIM # (B)(4). | -10.00/+4.0/089 | OS | Physician | 08/23/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733378&pc=QCB | 2023826-2023-03887 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+5.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS HAD ROTATED AND WAS UNSTABLE. THE LENS REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-03886. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK, A5: UNK, A6: UNK, B3: UNK, D6B: UNK, H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4). | -11.50/+5.0/094 | OS | Physician | 08/18/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733438&pc=QCB | 2023826-2023-03927 | 2023-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+1.5/ 072 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.0/+1.5/ 072 | OD | Physician | 08/14/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733439&pc=QCB | 2023826-2023-04010 | 2023-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+3.0/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+3.0/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -13.5/+3.0/092 | OD | Physician | 08/15/2023 | 2023 | 9 MO | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733441&pc=QCB | 2023826-2023-04008 | 2023-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+0.5/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#(B)(4). | -7.5/+0.5/095 | OD | Physician | 08/15/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733442&pc=QCB | 2023826-2023-03928 | 2023-07-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+1.0/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.5/+1.0/090 | OD | Physician | 08/14/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733445&pc=QCB | 2023826-2023-04015 | 2023-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+2.0/073 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -8.0/+2.0/073 | OS | Physician | 08/15/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733446&pc=QCB | 2023826-2023-04016 | 2023-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.5/131 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+2.5/131 | OD | Physician | 08/18/2023 | 2023 | 43 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733447&pc=QCB | 2023826-2023-04018 | 2023-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+1.0/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN AND IT IS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21 YEARS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | -8.0/+1.0/092 | OS | Physician | 08/18/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733449&pc=QCB | 2023826-2023-04017 | 2023-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/+0.5/106 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN AND IT IS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#(B)(4). | -5.5/+0.5/106 | OD | Physician | 08/18/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733484&pc=QCB | 2023826-2023-03891 | 2023-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -10.0/2.0/010 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL ACD<3.0. (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -10.0/2.0/010 | OD | Physician | 08/26/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733485&pc=QCB | 2023826-2023-03892 | 2023-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -14.5/3.0/166 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL ACD<3.0. CLAIM# (B)(4). | -14.5/3.0/166 | OS | Physician | 08/26/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17733486&pc=QCB | 2023826-2023-03897 | 2023-08-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -8.0/2.0/075 (SPHERE/CYLINDER/AXIS) WAS INTRAOPERATIVELY IMPLANTED, REMOVED, AND REPLACED VERTICALLY INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PROBLEM WAS RESOLVED. REPORTEDLY, "ALTHOUGH BOTH ARE 12.6 MODELS. HOWEVER, THE HEIGHT OF THE HORIZONTAL ARCH IS TOO IGHT, AND IT IS REPLACED WITH VERTICAL PLACEMENT." CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK, H6: OFF-LABEL AGE<21, (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). | -8.0/2.0/075 | OD | Physician | 08/28/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17741571&pc=QCB | 2023826-2023-03993 | 2023-08-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM513.7 IMPLANTABLE COLLAMER LENS OF A -7.0/1.5/73 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED UPSIDE DOWN INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. INTRAOPERATIVELY IRISPROLAPS AT MAIN INCISION. THE LENS WAS INTRAOPERATIVELY IMPLANTED AND REMOVED. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "VISUAL ACCUITY IS GOING TO GET BACK TO NORMAL. PUPIL IS ALREADY AROUND AGAIN AND REACTIVE. SLIGHT THINNING OF THE IRIS STROMA IN THE AREA OF FORMER PROLAPS CAN BE DETECTED. AT THE MOMENT THE PATIENT WILL BE MONITORED AND NO PERMANENT INJURY WILL BE EXPECTED. A DECISION ABOUT A FURTHER ICL IMPLANTATION WILL BE DONE IN THE NEXT MONTHS." Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -7.0/1.5/73 | OS | Physician | 08/29/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17750465&pc=QCB | 2023826-2023-03994 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -17.50/1.0/167 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). HIGH VAULT AND ANGLE CLOSURES ARE OBSERVED. Manufacturer Narrative: A4-A6:UNK H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -17.50/1.0/167 | OD | Physician | 08/29/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17750856&pc=QCB | 2023826-2023-04032 | 2021-12-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_ 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/2.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -12.5/2.5/088 | OS | Physician | 09/11/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17750857&pc=QCB | 2023826-2023-04031 | 2021-12-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.5/1.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -15.5/1.5/095 | OD | Physician | 09/11/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17750859&pc=QCB | 2023826-2023-04028 | 2023-07-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_ 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/3.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -9.0/3.0/094 | OS | Physician | 09/11/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17750860&pc=QCB | 2023826-2023-04030 | 2023-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_ 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/2.0/081 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK, (B)(4). | -10.0/2.0/081 | OS | Physician | 09/11/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17750868&pc=QCB | 2023826-2023-04045 | 2023-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/2.0/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -14.5/2.0/073 | OS | Physician | 09/11/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17750870&pc=QCB | 2023826-2023-04044 | 2023-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/2.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -12.5/2.0/093 | OD | Physician | 09/11/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17750872&pc=QCB | 2023826-2023-04042 | 2023-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/1.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -14.5/1.0/094 | OS | Physician | 09/11/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17750887&pc=QCB | 2023826-2023-04046 | 2023-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -18.0/1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - CORRECTED TO : " THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -18.0/1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN." IN THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -18.0/1.0/090 | OD | Physician | 09/11/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17750889&pc=QCB | 2023826-2023-04050 | 2023-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/2.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(4) 2023. ON (B)(4) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK, (B)(4). | -13.0/2.0/095 | OS | Physician | 09/12/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17750914&pc=QCB | 2023826-2023-04051 | 2023-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.0/2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK . CLAIM # (B)(4). | -16.0/2.0/092 | OS | Physician | 09/12/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17751158&pc=QCB | 2023826-2023-03995 | 2022-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -17.50/0.5/025 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). HIGH VAULT AND ANGLE CLOSURES ARE OBSERVED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -17.50/0.5/025 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND SHORTER LENGTH LENS DUE TO HIGH VAULT, ELEVATED IOP AND ANGLE CLOSURES. IT IS NOT KNOWN IF THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENTS IS UNKNOWN. D6A - (B)(6) 2022; D6B - (B)(6) 2023. CLAIM # (B)(4). Manufacturer Narrative: B3: DATE OF EVENT: 13-SEPT-2022. CLAIM#: (B)(4). | -17.50/0.5/025 | OS | Physician | 12/04/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17751257&pc=QCB | 2023826-2023-04049 | 2023-08-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/1.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -14.0/1.5/089 | OD | Physician | 09/12/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17761002&pc=QCB | 2023826-2023-04002 | 2023-08-31 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -12.00/1.5/072 (SPHERE/CYLINDER/AXIS) LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT CONTACT OR INJURY. CAUSE OF THE EVENT INS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR LOTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -12.00/1.5/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) (B)(6)2023. THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS REMOVED INTRAOPERTIVELY. THERE WAS NO PATIENT INJURY. ON (B)(6) 2023 A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED. THE PROBLEM WAS RESOLVED. CLAIM #: (B)(4). | -12.00/1.5/072 | OS | Physician | 09/01/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17761234&pc=QCB | 2023826-2023-04074 | 2022-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -1.50/2.5/059 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS THERE WAS LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND IT WAS INDICATED THE LEFT EYE WAS NOT STABLE AND THE PATIENT COULD NOT GET USED TO THE LENS AND THE LENS WAS EXPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -1.50/2.5/059 | OS | Physician | 08/31/2023 | 2023 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17769891&pc=QCB | 2023826-2023-04113 | 2023-06-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11.00/1.5/106 (SPHERE/CYLINDER/AXIS) LENS TORE WHILE LOADING; NO PATIENT CONTACT. NO FURTHER INFORMATION HAS BEEN PROVIDED. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -11.00/1.5/106 | NA | Other Health Care Professional | 08/22/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17769892&pc=QCB | 2023826-2023-04114 | 2023-06-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -12.50/3.0/109 (SPHERE/CYLINDER/AXIS) LENS TORE WHILE LOADING; NO PATIENT CONTACT. NO FURTHER INFORMATION HAS BEEN PROVIDED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -12.50/3.0/109 | NA | Other Health Care Professional | 08/22/2023 | 2023 | 47
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17770212&pc=QCB | 2023826-2023-04057 | 2023-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/1.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -7.0/1.0/094 | OD | Physician | 09/13/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17770214&pc=QCB | 2023826-2023-04054 | 2023-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/3.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -12.0/3.5/090 | OS | Physician | 09/12/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17770215&pc=QCB | 2023826-2023-04058 | 2023-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/2.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -12.0/2.5/089 | OS | Physician | 09/13/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17770217&pc=QCB | 2023826-2023-04053 | 2023-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/2.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - CORRECTED TO: "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/2.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON 08-AUG-2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN." IN THE INTIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: A4-A6: UNK, H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -11.5/2.5/090 | OD | Physician | 09/12/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17770269&pc=QCB | 2023826-2023-04105 | 2023-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -17.50/4.0/115 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. IT WAS REPORTED THAT MEDICAL OR SURGICAL INTERVENTION WAS NECESSARY AND REPOSITIONING WAS REPORTED. IN THE REPORTERS OPINION PATIENT-RELATED FACTOR AND USER ERROR WAS THE CAUSE OF THE EVENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: 12.6MM VTICMO12.6 SHOULD BE CORRECTED 13.2MM VTICMO13.2. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -17.50/4.0/115 | OD | Physician | 08/24/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17770611&pc=QCB | 2023826-2023-03701 | 2023-05-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -8.0/2.0/174 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) AT 60 DEGREES ON (B)(6)2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE SAME MODEL AND SIZE AND POWER LENS DUE TO LOW VAULT. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE ON A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC SPLIT AND HAPTIC BROKEN. THE LENS HAS BEEN RETURNED SIGNIFICANTLY DAMAGED, AS THE BOX HAS BEEN TICKED ACCORDINGLY, THUS DIMENSIONAL INSPECTION COULDN'T HAVE BEEN PERFORMED. CLAIM # (B)(4). Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM), (B)(4). | -8.0/2.0/174 | OS | Physician | 08/23/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17776290&pc=QCB | 2023826-2023-03703 | 2023-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -11.50/1.0/075 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS IMPLANTED UPSIDE DOWN. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL, SIZE AND POWER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -11.50/1.0/075 | OS | Physician | 08/23/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17776747&pc=QCB | 2023826-2023-03702 | 2023-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.0/0.5/66 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON 12-JUN-2023. THE LENS WAS IMPLANTED UPSIDE DOWN. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -11.0/0.5/66 | OS | Physician | 08/23/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17777435&pc=QCB | 2023826-2023-04148 | 2023-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.5/3.0/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2023, THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS BUT THIS DID NOT RESOLVE THE PROBLEM. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -9.5/3.0/087 | OD | Physician | 09/14/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17777437&pc=QCB | 2023826-2023-04149 | 2023-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.0/2.0/074 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS BUT THIS DID NOT RESOLVE THE PROBLEM. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.0/2.0/074 | OS | Physician | 09/14/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17777470&pc=QCB | 2023826-2023-04128 | 2023-03-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -7.5/1.0/091 | OS | Physician | 09/14/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17777887&pc=QCB | 2023826-2023-04106 | 2023-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -13.50/2.5/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.50/2.5/089 | OS | Physician | 08/28/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17778166&pc=QCB | 2023826-2023-04151 | 2023-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+2.0/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -12.5/+2.0/091 | OS | Physician | 08/21/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17778168&pc=QCB | 2023826-2023-04152 | 2023-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+2.0/073 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -8.5/+2.0/073 | OD | Physician | 08/21/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17778169&pc=QCB | 2023826-2023-04153 | 2023-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+1.0/014 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.5/+1.0/014 | OS | Physician | 08/21/2023 | 2023 | 45 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17778200&pc=QCB | 2023826-2023-04111 | 2023-07-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -12.0/2.5/094 (SPHERE/CYLINDER/AXIS) TORE/BROKE BEFORE/DURING INJECTION/DELIVERY INTO THE PATIENTS RIGHT EYE (OD). INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK. ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT, (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - MEDICAL DEVICE PROBLEM CODE 1494: OFF-LABEL USE (UNDER 21 YRS OF AGE AT DATE OF IMPLANTATION) ADDED TO INITIAL MDR. CLAIM#: (B)(4). | -12.0/2.5/094 | OD | Physician | 08/30/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17783548&pc=QCB | 2023826-2023-04170 | 2023-08-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.5/2.5/128 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE. INTRAOPERATIVELY THE LENS WAS REMOVED WITH NO PATIENT INJURY. ON (B)(6) 2023 A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.5/2.5/128 | OD | Physician | 09/11/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17785352&pc=QCB | 2023826-2023-04157 | 2023-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.5/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+1.5/091 | OS | Physician | 08/22/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17785353&pc=QCB | 2023826-2023-04158 | 2023-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.5/070 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -9.5/+1.5/070 | OD | Physician | 08/22/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17785354&pc=QCB | 2023826-2023-04159 | 2023-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.5/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -9.5/+1.5/092 | OS | Physician | 08/22/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17785356&pc=QCB | 2023826-2023-04160 | 2023-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+3.0/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+3.0/095 | OD | Physician | 08/22/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17786027&pc=QCB | 2023826-2023-04133 | 2023-06-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_ 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/2.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -10.0/2.0/095 | OS | Physician | 09/15/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17786028&pc=QCB | 2023826-2023-04132 | 2023-07-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_ 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.0/2.5/175 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -4.0/2.5/175 | OS | Physician | 09/15/2023 | 2023 | 48 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17786030&pc=QCB | 2023826-2023-04135 | 2023-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_ 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/2.5/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - CORRECTED TO "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_ 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/2.5/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN." IN THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -7.5/2.5/097 | OD | Physician | 09/15/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17786031&pc=QCB | 2023826-2023-04134 | 2022-11-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_ 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/2.0/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL/ LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AND HAPTIC BROKEN TO THE LENS. LENS RETURNED DAMAGED. UNABLE TO PERFORM DIMENSIONAL INSPECTION. CLAIM#(B)(4). | -10.0/2.0/083 | OS | Physician | 09/15/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17786415&pc=QCB | 2023826-2023-04090 | 2023-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -9.5/2.5/077 WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. REPORTEDLY, "PATIENT DOESN'T WANT TO HAVE AN EVO ICL IN ANOTHER LENGTH. SHE WERES CONTACT LENSES AGAIN." CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATE LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: H6 - HEALTH EFFECT - IMPACT CODE - 4627 CORRECTED IN THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -9.5/2.5/077 | OS | Physician | 12/12/2023 | 2023 | 43 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17786645&pc=QCB | 2023826-2023-04120 | 2023-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -12.5/4.5/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. LENS ROTATION WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). | -12.5/4.5/095 | OD | Physician | 08/25/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17786646&pc=QCB | 2023826-2023-04119 | 2023-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -9.50/5.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. LENS ROTATION WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H10: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -9.50/5.0/091 | OS | Physician | 08/25/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17792264&pc=QCB | 2023826-2023-04077 | 2023-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.50/+2.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LOW VAULTING AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -15.50/+2.5/092 | OS | Physician | 08/26/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17794703&pc=QCB | 2023826-2023-04165 | 2023-03-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/+2.0/086 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -6.0/+2.0/086 | OS | Physician | 08/23/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17794704&pc=QCB | 2023826-2023-04162 | 2023-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+2.5/126 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -11.5/+2.5/126 | OD | Physician | 08/23/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17794707&pc=QCB | 2023826-2023-04168 | 2023-05-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+2.0/109 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.5/+2.0/109 | OS | Physician | 08/23/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17795157&pc=QCB | 2023826-2023-04093 | 2023-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -11.5 DIOPTER WAS IMPLANTED UPSIDE DOWN INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SAME SIZE LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT-RACE:UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#(B)(4). | -11.5 diopte | OS | Physician | 09/07/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17795168&pc=QCB | 2023826-2023-04163 | 2023-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.5/086 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT. THAT SAME DAY, A SHORTER LENGTH LENS WAS IMPLANTED, ALTHOUGH IT IS UNCLEAR WHETHER THIS WAS PERFORMED WITHIN THE SAME SURGERY THAT THE LENS WAS REMOVED. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.5/+1.5/086 | OS | Physician | 08/23/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17795402&pc=QCB | 2023826-2023-04166 | 2023-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+4.0/088 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN AND IT IS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -9.5/+4.0/088 | OS | Physician | 08/23/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17797232&pc=QCB | 2023826-2023-04176 | 2023-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+2.5/070 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -9.0/+2.5/070 | OD | Physician | 08/24/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17797233&pc=QCB | 2023826-2023-04174 | 2023-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+2.0/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -14.0/+2.0/092 | OD | Physician | 08/24/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17797234&pc=QCB | 2023826-2023-04175 | 2023-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0/+2.0/097 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -15.0/+2.0/097 | OD | Physician | 08/24/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17797241&pc=QCB | 2023826-2023-04179 | 2023-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+3.5/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | -9.5/+3.5/094 | OS | Physician | 08/24/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17797242&pc=QCB | 2023826-2023-04182 | 2023-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/151 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN AND IT IS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+1.0/151 | OD | Physician | 08/24/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17797243&pc=QCB | 2023826-2023-04180 | 2023-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.5/074 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -10.5/+1.5/074 | OD | Physician | 08/24/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17797247&pc=QCB | 2023826-2023-04183 | 2023-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+2.5/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -13.0/+2.5/092 | OS | Physician | 08/24/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17797248&pc=QCB | 2023826-2023-04185 | 2023-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+1.0/053 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -7.5/+1.0/053 | OS | Physician | 08/25/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17797771&pc=QCB | 2023826-2023-04186 | 2023-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+2.5/085 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -12.0/+2.5/085 | OS | Physician | 08/26/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17797774&pc=QCB | 2023826-2023-04190 | 2023-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+2.5/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) CLAIM# (B)(4). | -9.0/+2.5/093 | OS | Physician | 08/26/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17797777&pc=QCB | 2023826-2023-04189 | 2023-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+1.5/095 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -13.5/+1.5/095 | OS | Physician | 08/26/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17797778&pc=QCB | 2023826-2023-04191 | 2023-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+2.5/010 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -11.5/+2.5/010 | OD | Physician | 08/26/2023 | 2023 | 16 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17799578&pc=QCB | 2023826-2023-04171 | 2023-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -5.50/+5.50/101 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS ROTATION. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND THERE WAS DEVICE CAUSALITY AND NOTED "PATIENT HAD GOOD VA 1/7 POST-OP AND TORIC ALIGNMENT CORRESPONDED WITH IOD, ROTATION HAS OCCURRED BETWEEN DAY 1 AND DAY 5 POST-OP DESPITE 2CT VAULT". Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: B5: IT WAS LATER REPORTED THAT ON (B)(6) 2023, THE LENS WAS REPOSITIONED AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). | -5.50/+5.50/101 | OD | Physician | 08/30/2023 | 2023 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17800581&pc=QCB | 2023826-2023-04096 | 2023-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF A -10.5/2.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS A PATIENT-RELATED-FACTOR. REPORTEDLY, "EXCESSIVE VAULT ESPECIALLY MID-PERIPHERY NEARLY TOUCHING CORNEA." Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL ACD<3.0. PATIENT-RELATED-FACTOR. CLAIM# (B)(4). Manufacturer Narrative: B5- LENS WAS EXPLANTED ON (B)(6) 2023 AND REPLACEMENT LENS WAS IMPLANTED ON SAME DATE DIFFERENT SURGERY WITH A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. CLAIM# (B)(4). | -10.5/2.0/093 | OD | Physician | 09/11/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17801713&pc=QCB | 2023826-2023-04078 | 2023-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.00/+2.0/008 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID AND RESIDUE/DEBRIS. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK . H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -11.00/+2.0/008 | OD | Physician | 08/28/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17801787&pc=QCB | 2023826-2023-04079 | 2023-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.50/+3.0/178 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) (IMPLANTED VERTICALLY), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED HORIZONTALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK . H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -15.50/+3.0/178 | OS | Physician | 12/04/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17803928&pc=QCB | 2023826-2023-04094 | 2023-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -10.0/1.5/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6)2023, LENS OPACITY ASC WAS OBSERVED. ON (B)(6) 2023, THE LENS WAS REMOVED. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY, "THE OPERATION EYE IS STILL RECOVERING , AND THE DOCTOR WILL DECIDE THE NEXT TREATMENT PLAN AFTER THE OPERATION EYE'S CONDITION IS STABLE. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: OFF-LABEL AGE<21. CLAIM#: (B)(4). | -10.0/1.5/090 | OS | Physician | 09/12/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17803991&pc=QCB | 2023826-2023-04097 | 2023-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF A -17.5/1.5/89 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "ISSUE WITH ZONULES." Manufacturer Narrative: A4-A6:UNK, H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -17.5/1.5/89 | OD | Physician | 09/11/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17805044&pc=QCB | 2023826-2023-04121 | 2023-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -11.50/1.0/092 (SPHERE/CYLINDER/AXIS) "DOCTOR NOTICED LENS WAS TORN". THE LENS WAS REMOVED ON AN UNKNOWN DATE AND REPLACED WITH BACKUP LENS ALSO UNKNOWN WHEN THIS OCCURRED. NO PATIENT INJURY WAS REPORTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -11.50/1.0/092 | NA | Other Health Care Professional | 08/25/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17813908&pc=QCB | 2023826-2023-04083 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/081 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS EYE. THE LENS WAS REMOVED DUE TO SMALL VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. D4 - UNK. H4 - UNK. H6 - WORK ORDER SEARCH IS N/A AS SERIAL NUMBER IS UNK. CLAIM# (B)(4). | -10.00/+2.0/081 | NA | Physician | 08/31/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17813923&pc=QCB | 2023826-2023-04082 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.00/+1.5/010 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS EYE. THE LENS WAS REMOVED DUE TO LARGE VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2: UNK. A3: UNK. A4: UNK. A5: UNK. A6: UNK. B3: UNK. D6A: UNK. D6B:UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED. EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. D4 - UNK; H4 - UNK; H6 - WORK ORDER SEARCH IS N/A AS SERIAL NUMBER IS UNK. CLAIM# (B)(4). | -16.00/+1.5/010 | NA | Physician | 08/31/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17813924&pc=QCB | 2023826-2023-04084 | 2023-07-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/096 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. D6B:UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.00/+2.0/096 | OS | Physician | 09/01/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17813934&pc=QCB | 2023826-2023-04081 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.50/+2.5/169 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS EYE. THE LENS WAS REMOVED DUE TO SMALL VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2: UNK. A3: UNK. A4: UNK. A5: UNK. A6: UNK . B3: UNK. D6A: UNK. D6B:UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. H6 - WORK ORDER SEARCH IS N/A AS SERIAL NUMBER IS UNK. CLAIM# (B)(4). | -08.50/+2.5/169 | NA | Physician | 08/31/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814298&pc=QCB | 2023826-2023-04139 | 2023-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -18.0/2.0/070 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "THE LENS POSITION IS ADJUSTED FROM HORIZONTAL TO VERTICAL." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -18.0/2.0/070 | OD | Physician | 09/17/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814299&pc=QCB | 2023826-2023-04138 | 2023-05-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/2.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. CLAIM #(B)(4). | -9.0/2.0/090 | OD | Physician | 09/17/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814300&pc=QCB | 2023826-2023-04141 | 2023-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/1.5/079(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -11.0/1.5/079 | OS | Physician | 09/17/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814301&pc=QCB | 2023826-2023-04137 | 2023-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/1.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6 DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -8.0/1.0/095 | OD | Physician | 09/17/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814302&pc=QCB | 2023826-2023-04140 | 2023-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.0/109 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -10.5/1.0/109 | OD | Physician | 09/17/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814304&pc=QCB | 2023826-2023-04142 | 2023-01-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.50/1.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. (B)(4). | -9.50/1.0/089 | OS | Physician | 09/17/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814306&pc=QCB | 2023826-2023-04147 | 2023-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/0.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4) . | -8.0/0.5/095 | OD | Physician | 09/18/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814307&pc=QCB | 2023826-2023-04247 | 2023-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/1.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -9.0/1.5/093 | OD | Physician | 09/18/2023 | 2023 | 43 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814308&pc=QCB | 2023826-2023-04145 | 2023-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/2.0/088(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON 25-AUG-2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK, H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -14.0/2.0/088 | OS | Physician | 09/18/2023 | 2023 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814309&pc=QCB | 2023826-2023-04146 | 2023-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/2.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | -14.0/2.0/088 | OD | Physician | 09/18/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814312&pc=QCB | 2023826-2023-04245 | 2023-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/1.0/108 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -9.5/1.0/108 | OD | Physician | 09/18/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814313&pc=QCB | 2023826-2023-04246 | 2023-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/2.0/170 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -11.5/2.0/170 | OD | Physician | 09/18/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814314&pc=QCB | 2023826-2023-04249 | 2023-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_ 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/2.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO LOW VAULT WITHOUT ROTATION. ON (B)(6) 2023 A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. (B)(4). | -13.0/2.0/090 | OS | Physician | 09/19/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814317&pc=QCB | 2023826-2023-04254 | 2023-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_ 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/2.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. (B)(4). | -11.5/2.0/090 | OD | Physician | 09/19/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814324&pc=QCB | 2023826-2023-04252 | 2023-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.50/1.5/014 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. (B)(4). | -7.50/1.5/014 | OD | Physician | 09/19/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814329&pc=QCB | 2023826-2023-04256 | 2023-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/1.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME LENGTH/MODEL LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "ACCORDING TO THE ASSESSMENT OF POST-OP ARCH, CHANGE THE ICL POSITION FROM HORIZONTAL TO VERTICAL." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.0/1.5/094 | OD | Physician | 09/20/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814429&pc=QCB | 2023826-2023-04199 | 2023-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.5/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -8.5/+1.5/092 | OS | Physician | 08/28/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814430&pc=QCB | 2023826-2023-04207 | 2023-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.5/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#(B)(4). | -11.5/+1.5/091 | OS | Physician | 08/28/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814431&pc=QCB | 2023826-2023-04202 | 2023-08-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.5/083 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) . CLAIM# (B)(4). | -8.5/+1.5/083 | OD | Physician | 08/28/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814433&pc=QCB | 2023826-2023-04196 | 2023-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.0/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | -10.0/+1.0/090 | OS | Physician | 08/28/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814434&pc=QCB | 2023826-2023-04193 | 2023-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+2.0/155 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -9.0/+2.0/155 | OD | Physician | 08/28/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814595&pc=QCB | 2023826-2023-04274 | 2023-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -16.5/2.5/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REMOVED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. A REPLACEMENT LENS OF LONGER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -16.5/2.5/091 | OS | Physician | 08/31/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814963&pc=QCB | 2023826-2023-04204 | 2023-07-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VICMO13.7 IMPLANTABLE COLLAMER LENS WITH DIOPTER -6.5 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN AND IT IS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 -8.5/2.0/110 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED ON (B)(6) 2023 INTO THE RIGHT EYE (OD). LOW VAULT WAS OBSERVED. LENS WAS IMPLANTED, REMOVED AND REPLACED INTRAOPERATIVELY ON (B)(6) 2023 WITH A LONGER LENGTH LENS AND PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -8.5/2.0/110 | OD | Physician | 08/28/2023 | 2023 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814964&pc=QCB | 2023826-2023-04195 | 2023-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.5/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+1.5/090 | OD | Physician | 08/28/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814966&pc=QCB | 2023826-2023-04205 | 2023-06-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+2.0/012 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.0/+2.0/012 | OD | Physician | 08/28/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17814967&pc=QCB | 2023826-2023-04201 | 2023-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+2.0/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT. IN A SEPARATE SURGERY THAT DAY A SHORTER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.5/+2.0/093 | OS | Physician | 08/28/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17815158&pc=QCB | 2023826-2023-04266 | 2023-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -7.0/1.5/014 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE SAME MODEL AND SIZE BUT DIFFERENT POWER LENS DUE TO LOW VAULT. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | -7.0/1.5/014 | OS | Physician | 12/04/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17815183&pc=QCB | 2023826-2023-04267 | 2023-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -14.0/1.0/049 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION - LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, DRY WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -14.0/1.0/049 | OD | Physician | 08/28/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17818025&pc=QCB | 2023826-2023-04264 | 2023-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.0/2.0/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND REPLACED WITH A SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.0/2.0/89 | OD | Physician | 09/13/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17818335&pc=QCB | 2023826-2023-04231 | 2023-08-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -8.0/2.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. EXCESSIVE VAULT AND REFRACTIVE SURPRISE WAS OBSERVED. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY, "VAULT IS TOO HIGH AND REST REFRACTION: S+1.0 C.75 X 18". Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: UDI: SHOULD BE CORRECTED TO: (B)(4). CLAIM# (B)(4). | -8.0/2.0/091 | OD | Physician | 09/14/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17818351&pc=QCB | 2023826-2023-04232 | 2023-08-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -9.0/1.5/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT AND REFRACTIVE SURPRISE WAS OBSERVED. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY, "VAULT IS TOO HIGH AND REST REFRACTION: S+0.5 C.-1.25 X 171" Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -9.0/1.5/087 | OS | Physician | 09/14/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17822283&pc=QCB | 2023826-2023-04305 | 2023-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/1.5/070 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. (B)(4). | -13.5/1.5/070 | OS | Physician | 09/26/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17822284&pc=QCB | 2023826-2023-04263 | 2023-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_ 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/3.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CLAIM #(B)(4). | -9.0/3.0/093 | OS | Physician | 09/22/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17822285&pc=QCB | 2023826-2023-04261 | 2023-05-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.0/4.0/079 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -16.0/4.0/079 | OS | Physician | 09/22/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17822287&pc=QCB | 2023826-2023-04259 | 2023-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/2.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. (B)(4) . | -7.5/2.0/094 | OS | Physician | 09/22/2023 | 2023 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17822288&pc=QCB | 2023826-2023-04260 | 2023-08-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_ 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/1.5/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: A4-A6: UNK, (B)(4) | -11.5/1.5/083 | OS | Physician | 09/22/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17822289&pc=QCB | 2023826-2023-04262 | 2023-03-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.0/2.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. (B)(4). | -16.0/2.5/094 | OS | Physician | 09/22/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17822291&pc=QCB | 2023826-2023-04306 | 2023-03-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/1.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON 21-MAR-2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK, H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM), (B)(4). | -9.0/1.0/092 | OS | Physician | 09/26/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17822371&pc=QCB | 2023826-2023-04293 | 2023-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+2.0/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -9.0/+2.0/090 | OS | Physician | 08/31/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17822373&pc=QCB | 2023826-2023-04290 | 2023-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+2.0/073 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -8.5/+2.0/073 | OS | Physician | 08/30/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17822375&pc=QCB | 2023826-2023-04294 | 2023-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.5/069 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -10.5/+1.5/069 | OS | Physician | 08/31/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17822378&pc=QCB | 2023826-2023-04297 | 2023-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.0/075 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -10.0/+2.0/075 | OS | Physician | 08/31/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17822379&pc=QCB | 2023826-2023-04295 | 2023-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.0/+2.0/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LOW VAULT. THAT SAME DAY THE LENS WAS REPLACED WITH A LONGER LENGTH LENS ALTHOUGH IT IS UNCLEAR WHETHER THIS WAS WITHIN THE SAME PROCEDURE THAT THE LENS WAS REMOVED. THE PROBLEM WAS RESOLVED. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- PER THE REPORTER'S EMAIL ON (B)(6) 2023 THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. H6 HEALTH EFFECT - IMPACT CODE: CHANGED TO 4629. CLAIM# (B)(4). Manufacturer Narrative: B5 CORRECTION: PER THE REPORTER'S EMAIL ON (B)(6) 2023 THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. G3 - CORRECTED TO 05-NOV-2023 IN THE PREVIOUS MDR (SUPPLEMENTAL #1). CLAIM# (B)(4). | -17.0/+2.0/093 | OS | Physician | 08/31/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17822380&pc=QCB | 2023826-2023-04296 | 2023-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+2.0/001 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -8.5/+2.0/001 | OD | Physician | 08/31/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17822381&pc=QCB | 2023826-2023-04298 | 2023-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/+2.5/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) AND (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | -5.5/+2.5/092 | OS | Physician | 08/31/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17827767&pc=QCB | 2023826-2023-04300 | 2023-06-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+2.0/068 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT. IN A SEPARATE SURGERY THAT DAY A SHORTER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -12.0/+2.0/068 | OD | Physician | 09/01/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17827770&pc=QCB | 2023826-2023-04301 | 2023-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+1.5/102 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT. IN A SEPARATE SURGERY THAT DAY A SHORTER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.5/+1.5/102 | OD | Physician | 09/01/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17830163&pc=QCB | 2023826-2023-04310 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-09-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -11.00/1.5/096 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) AS A BACK UP LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.00/1.5/096 | OS | Physician | 09/06/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17830164&pc=QCB | 2023826-2023-04308 |  | Injury | STAAR SURGICAL COMPANY | 2023-09-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -11.00/1.5/096 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). IT WAS REPORTED THAT THE LENS WAS EXPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.00/1.5/096 | OS | Physician | 09/06/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17830165&pc=QCB | 2023826-2023-04309 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-09-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO_12.1, -15.50/2.0/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) AS A BACK UP LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -15.50/2.0/095 | OS | Physician | 09/06/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17831119&pc=QCB | 2023826-2023-04219 | 2023-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/054 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2023 AND THE PROBLEM WAS RESOLVED. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT POWER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND FIBRE ON THE SURFACE OF THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.0/2.0/054 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THERE WAS LENS ROTATION NOT ASSOCIATED TO LOW VAULT. ON (B)(6) 2023, THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. "AFTER SURGERY THE POSITION OF THE LENS OF THE PATIENT'S RIGHT EYE WAS ROTATED SEVERAL TIMES AND THE ROTATION POSITION WAS 19 COUNTERCLOCKWISE VERTICALLY; SO THE INTRAOCULAR LENS WAS REPLACED IN THIS DIRECTION". ON (B)(6) 2023, THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT AXIS LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FIBRE ON THE SURFACE OF THE PRODUCT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.00/+2.0/054 | OD | Physician | 09/04/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17831196&pc=QCB | 2023826-2023-04211 | 2023-09-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.50/+1.0/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -08.50/+1.0/080 | OS | Physician | 09/04/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17831269&pc=QCB | 2023826-2023-04212 | 2023-09-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -06.50/+1.5/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2: UNK. A4: UNK. A5: UNK. A6: UNK. D6B: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -06.50/+1.5/090 | OD | Physician | 09/04/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17831295&pc=QCB | 2023826-2023-04218 | 2023-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.00/+2.5/111 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2023 AND THE PROBLEM WAS RESOLVED. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT POWER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -16.0/2.5/111 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THERE WAS LENS ROTATION NOT ASSOCIATED TO LOW VAULT. ON (B)(6) 2023 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT AXIS LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. CLAIM#: (B)(6). | -16.00/+2.5/111 | OS | Physician | 09/04/2023 | 2023 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17839066&pc=QCB | 2023826-2023-04235 |  | Injury | STAAR SURGICAL COMPANY | 2023-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | UNK | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). REPORTEDLY, "THE PATIENT HAS AN UNEXPTECTED HYPEROPIC OUTCOME." Manufacturer Narrative: A4-A6:UNK. B3: UNK. D4 (EXPERIATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK. H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM# (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VITCM5_12.6 -10.50/1.0/177 (SPHERE/CYLINDER/AXIS) UPSIDE DOWN INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 WITH A SAME LENGTH DIFFERENT DIOPTER LENS AND PROBLEM WAS RESOLVED. REPORTEDLY "ALL RESOLVED AND PATIENT HAS GOOD OUTCOME". CAUSE OF EVENT IS USER ERROR; DEVICE DID NOT FAIL TO PERFORM AS INTENDED. H6: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -10.50/1.0/177 | OS | Physician | 09/15/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17839072&pc=QCB | 2023826-2023-04234 |  | Injury | STAAR SURGICAL COMPANY | 2023-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | UNK | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). REPORTEDLY, "THE PATIENT HAS AN UNEXPECTED HYPEROPIC OUTCOME." Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VITCM5_12.6 -10.50/0.5/067 (SPHERE/CYLINDER/AXIS) UPSIDE DOWN INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 WITH A SAME LENGTH DIFFERENT DIOPTER LENS AND PROBLEM WAS RESOLVED. REPORTEDLY "ALL RESOLVED AND PATIENT HAS GOOD OUTCOME". CAUSE OF EVENT IS USER ERROR; DEVICE DID NOT FAIL TO PERFORM AS INTENDED. H6 - INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK. B3: UNK. D4 (EXPIRATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK. H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM# (B)(4). | -10.50/0.5/067 | OD | Physician | 09/15/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17839651&pc=QCB | 2023826-2023-04283 |  | Injury | STAAR SURGICAL COMPANY | 2023-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -14.50/2.0/069 (SPHERE/CYLINDER/AXIS) WAS REMOVED FROM THE PATIENTS EYE ON (B)(6) 2023. A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED ON THE SAME DATE, IT IS UNKNOWN IF THE LENS WAS REPLACED WITHIN THE SAME SURGERY AS REMOVAL OF THE INITIAL LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -14.50/2.0/069 | NA | Other Health Care Professional | 09/05/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17840477&pc=QCB | 2023826-2023-04236 | 2022-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.0/4.0/88 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2023, THE LENS WAS EXCHANGED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -9.0/4.0/88 | OS | Physician | 09/17/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17840478&pc=QCB | 2023826-2023-04237 | 2023-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -13.5/2.0/070 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED UPSIDE DOWN INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A REPLACEMENT LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. NO PATIENT INJURY IS REPORTED. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -13.5/2.0/070 | OS | Physician | 09/17/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17846204&pc=QCB | 2023826-2023-04216 | 2023-07-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.00/+2.0/162 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CORRECTED DATA: H3: DEVICE EVALUATED BY MANUFACTURER? NO: H10: LENS EVALUATION NOT REQUIRED. CLAIM#: (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -07.00/+2.0/162 | OD | Physician | 09/06/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17847275&pc=QCB | 2023826-2023-04321 | 2023-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.5/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -10.5/+1.5/091 | OD | Physician | 09/03/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17847292&pc=QCB | 2023826-2023-04320 | 2023-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+2.0/086 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/+2.0/086 | OS | Physician | 09/03/2023 | 2023 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17848948&pc=QCB | 2023826-2023-04345 | 2023-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.50/2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -10.50/2.0/092 | OS | Physician | 09/27/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17848949&pc=QCB | 2023826-2023-04350 | 2023-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.50/1.0/069 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -11.50/1.0/069 | OD | Physician | 09/27/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17848950&pc=QCB | 2023826-2023-04351 | 2023-03-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/1.0/014 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "THE FIRST TICL IS IN THE VERTICAL POSITION." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -9.0/1.0/014 | OD | Physician | 09/27/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17848952&pc=QCB | 2023826-2023-04352 | 2023-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/2.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -10.0/2.0/094 | OS | Physician | 09/27/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17848958&pc=QCB | 2023826-2023-04343 | 2023-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/2.5/004 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "BOTH TICLS ARE VERTICAL; THE FIRST AND SECOND TICLS ARE BOTH VERTICAL." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -11.5/2.5/004 | OD | Physician | 09/27/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17848965&pc=QCB | 2023826-2023-04344 | 2023-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/2.5/009 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON 09-FEB-2023. ON 08-AUG-2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "BOTH TICLS ARE VERTICAL; THE FIRST AND SECOND TICLS ARE BOTH VERTICAL." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK CLAIM #(B)(4). | -12.0/2.5/009 | OS | Physician | 09/27/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17848967&pc=QCB | 2023826-2023-04336 | 2023-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/5.0/142 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -7.5/5.0/142 | OD | Physician | 09/27/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17848968&pc=QCB | 2023826-2023-04340 | 2023-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.50/1.0/095(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -8.50/1.0/095 | OS | Physician | 09/27/2023 | 2023 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17848970&pc=QCB | 2023826-2023-04338 | 2023-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/1.0/154 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "THE FIRST TICL IS IN THE VERTICAL POSITION." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -8.5/1.0/154 | OS | Physician | 09/27/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17848971&pc=QCB | 2023826-2023-04337 | 2023-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.5/004 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "THE FIRST TICL IS IN THE VERTICAL POSITION." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -10.0/1.5/004 | OD | Physician | 09/27/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17853892&pc=QCB | 2023826-2023-04356 | 2023-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "REPLACEMENT LENS WAS VERTICAL IMPLANTED." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -10.0/1.5/088 | OD | Physician | 09/28/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17857764&pc=QCB | 2023826-2023-04404 | 2023-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.5/3.0/147 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2023 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT AXIS LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -14.5/3.0/147 | OD | Physician | 09/27/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17858003&pc=QCB | 2023826-2023-04376 | 2023-09-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -5/1.5/112 (SPHERE/CYLINDER/AXIS) LENS TORE DURING LOADING AFTER OPENING VIAL. THE LENS WAS NOT IMPLANTED. THE CAUSE OF THE EVENT WAS THE DEVICE. REPORTEDLY, "PROBABLY THE LOADING FORCEPS TORE THE LENS." THE PROBLEM WAS NOT RESOLVED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | -5/1.5/112 | NA | Physician | 09/27/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17858004&pc=QCB | 2023826-2023-04377 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -7.50/1.0/002 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4-A6:UNK. CLAIM# (B)(4). | -7.50/1.0/002 | OD | Physician | 09/27/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17858048&pc=QCB | 2023826-2023-04325 | 2023-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.50/+1.5/107 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -14.50/+1.5/107 | OD | Physician | 09/06/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17858050&pc=QCB | 2023826-2023-04327 | 2023-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.5/+2.0/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT. IN A SEPARATE SURGERY THAT DAY A SHORTER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -15.5/+2.0/095 | OD | Physician | 09/06/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17858051&pc=QCB | 2023826-2023-04324 | 2023-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.0/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -11.0/+2.0/090 | OD | Physician | 09/06/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17858056&pc=QCB | 2023826-2023-04330 | 2023-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+2.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/+2.0/091 | OD | Physician | 09/07/2023 | 2023 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17858057&pc=QCB | 2023826-2023-04333 | 2023-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/073 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -9.5/+1.0/073 | OD | Physician | 09/08/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17858059&pc=QCB | 2023826-2023-04326 | 2023-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.5/+2.0/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -15.5/+2.0/089 | OS | Physician | 09/06/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17858060&pc=QCB | 2023826-2023-04334 | 2023-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6:DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -11.0/+1.0/091 | OD | Physician | 09/09/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17858063&pc=QCB | 2023826-2023-04335 | 2023-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.0/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -10.5/1.0/090 | OS | Physician | 09/09/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17861995&pc=QCB | 2023826-2023-04062 | 2023-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.5/1.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2023, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS REPLACED FOR SAME LENGTH SPHERICAL LENS AND THE PROBLEM WAS RESOLVED. REPORTEDLY, PATIENT IS HAPPY WITH VISION. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.5/1.0/075 | OS | Physician Assistant | 09/11/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17862005&pc=QCB | 2023826-2023-04060 | 2023-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED INTRAOPERATIVELY A 12.6MM VTICM5_12.6, -10.5/3.0/071 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LENS WAS IMPLANTED UPSIDE DOWN. PATIENT CONTACT (LENS TOUCHED THE EYE). NO PATIENT INJURY. PROBLEM NOT RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. REPORTEDLY, LENS TORE DURING REMOVAL FROM EYE. CORNEAL EDEMA OBSERVED BUT HAS SINCE CLEARED WITH STEROIDS AND HYPEROSMOTIC EYE DROPS. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.5/3.0/071 | OS | Physician | 09/11/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17862006&pc=QCB | 2023826-2023-04061 | 2023-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.0/1.0/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2023, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS REPLACED FOR SAME LENGTH SPHERICAL LENS AND THE PROBLEM WAS RESOLVED. REPORTEDLY, PATIENT IS HAPPY WITH VISION. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.0/1.0/109 | OD | Physician Assistant | 09/11/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17862013&pc=QCB | 2023826-2023-04222 | 2023-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Patient Device Interaction Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -11.0/1.5/179 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. LENS WAS REPOSITIONED ON (B)(6) 2023 BUT THIS DID NOT RESOLVE THE LENS ROTATION AND LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENGTH SPHERICAL LENS ON (B)(6) 2023. THIS DID NOT RESOLVE THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR, LENS TOO SHORT CAUSE ROTATION. REPORTEDLY, PATIENT STATES NEAR VISION NOT SO CLEAR. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE CODE: 1494: THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM#: (B)(4). | -11.0/1.5/179 | OD | Physician | 09/15/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17862439&pc=QCB | 2023826-2023-04402 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -5.50/0.50/101 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO "VAULT IS 0, AND WE NEED 13.7 ICL". THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM#: (B)(4). | -5.50/0.50/101 | OD | Physician | 09/29/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17862452&pc=QCB | 2023826-2023-04403 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -4.00/0.50/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM# (B)(4). | -4.00/0.50/95 | OS | Physician | 09/29/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17863303&pc=QCB | 2023826-2023-04409 | 2023-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -10.0/1.5/091 | OS | Physician | 10/02/2023 | 2023 | 46 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17863304&pc=QCB | 2023826-2023-04408 | 2023-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.50/0.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -6.50/0.5/087 | OD | Physician | 10/02/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17863324&pc=QCB | 2023826-2023-04410 | 2023-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/2.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -8.5/2.0/095 | OD | Physician | 10/02/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17865101&pc=QCB | 2023826-2023-04383 | 2023-08-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -16.0/2.0/095 (SPHERE/CYLINDER/AXIS) TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS NOT IMPLANTED, BUT THERE WAS PATIENT CONTACT ON (B)(6) 2023 . ON (B)(6) 2023, A REPLACEMENT LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. NO PATIENT INJURY WAS REPORTED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -16.0/2.0/095 | OS | Physician | 09/28/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17872496&pc=QCB | 2023826-2023-04268 | 2023-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -10.0/1.5/89 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. LENS ROTATION WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H10: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -10.0/1.5/89 | OS | Physician | 09/08/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17873815&pc=QCB | 2023826-2023-04428 | 2023-08-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/+2.5/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -14.5/+2.5/091 | OS | Physician | 09/04/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17878139&pc=QCB | 2023826-2023-04426 | 2023-08-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -15.0/+2.5/071 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE LENS HAD TORE DURING LOADING AND NOTED "DURING PULLING IN CARTRIDGE". INTRAOPERATIVELY THE LENS WAS REMOVED. ON (B)(6) 2023 A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED "THE LENS WAS TORN DURING LOADING, TRIED WITH THE OTHER SIDE BUT ALSO SAME TORN". Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -15.0/+2.5/071 | OS | Physician | 09/26/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17882132&pc=QCB | 2023826-2023-04446 | 2023-06-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -14.0/+2.5/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS ROTATION. ON (B)(6) 2023 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND REPLACED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -14.0/+2.5/89 | OD | Physician | 09/27/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17882722&pc=QCB | 2023826-2023-04421 | 2023-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -12.0/4.0/075 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON 09-JUN-2023 THE LENS WAS EXPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: RETURN DATE: (B)(6) 2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BROKEN WITH RESIDUE/DEBRIS ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.0/4.0/075 | OS | Physician | 09/12/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17883067&pc=QCB | 2023826-2023-04420 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -4.50/1.5/091 (SPHERE/CYLINDER/AXIS) WAS EXPLANTED FROM THE PATIENTS EYE ON (B)(6) 2023 FOR AN UNKNOWN REASON. A REPLACEMENT LENS WAS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6: HEALTH EFFECT - IMPACT CODE: 4627 CORRECTED TO 4629. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -4.50/1.5/091 | NA | Other Health Care Professional | 12/11/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17883088&pc=QCB | 2023826-2023-04424 | 2023-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.5/1.0/072 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.5/1.0/072 | OD | Physician | 09/13/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17885103&pc=QCB | 2023826-2023-04314 | 2023-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: A2: AGE AT TIME OF EVENT, DATE OF BIRTH: (B)(6) 1994 . B3: DATE OF EVENT: 04-SEP-2023 . CLAIM # (B)(4). | -11.50/+1.5/082 | OD | Physician | 09/13/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17885125&pc=QCB | 2023826-2023-04312 | 2023-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LOW VAULTING AND LENS ROTATION.THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: B4: DATE OF THIS REPORT: 11-SEP-2023. CLAIM # (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. CLAIM # (B)(4). | -15.00/+2.0/090 | OD | Physician | 09/11/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17885159&pc=QCB | 2023826-2023-04317 | 2023-07-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+1.5/094 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS LATER THAT SAME DAY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.00/+1.5/094 | OS | Physician | 09/14/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17894897&pc=QCB | 2023826-2023-04225 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1, -11.0/2.0/081 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS TORE IN THE INJECTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A1 - UNK. A2 - UNK. A3 - UNK. A4 - UNK. A5 - UNK. A6 - UNK. B3 - UNK. D6A - UNK. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -11.0/2.0/081 | NA | Physician | 09/21/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17898593&pc=QCB | 2023826-2023-04251 | 2023-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/2.5/012 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "THE SECOND LENS WAS IMPLANTED VERTICALLY AND GAVE THE PATIENT FULL CORRECTION, SO THE DEGREE WAS CHANGED." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -8.5/2.5/012 | OS | Physician | 09/19/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17898649&pc=QCB | 2023826-2023-04453 | 2023-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/3.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -11.5/3.5/087 | OS | Physician | 10/06/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17898669&pc=QCB | 2023826-2023-04451 | 2023-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/2.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -10.5/2.0/095 | OD | Physician | 10/06/2023 | 2023 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17898674&pc=QCB | 2023826-2023-04452 | 2023-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/1.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -9.0/1.5/089 | OS | Physician | 10/06/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17902261&pc=QCB | 2023826-2023-04472 | 2023-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -8.5/2.0/84 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING, THE LENS WAS REPOSITIONED ON (B)(6) 2023 BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND ON THIS SAME DATE REPLACED WITH A LONGER LENGTH LENS WHICH RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - TRACE CATARACT WAS OBSERVED. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -8.5/2.0/84 | OS | Physician | 09/21/2023 | 2023 | 50 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17906060&pc=QCB | 2023826-2023-04482 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | UNK | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 OF A -6.50/1.0/160 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. CAUSE OF THE EVENT IS A PATIENT RELATED FACTOR. Manufacturer Narrative: CLAIM# (B)(4). | -6.50/1.0/160 | OS | Physician | 09/13/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17908778&pc=QCB | 2023826-2023-04471 | 2023-06-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+2.0/071 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - "INTO THE PATIENT'S EYE LEFT EYE (OS) ON (B)(6) 2023" SHOULD BE CORRECTED TO "INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 " IN THE INITIAL MDR CLAIM #: (B)(4). Manufacturer Narrative: H6- DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -12.0/+2.0/071 | OS | Physician | 09/12/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17908779&pc=QCB | 2023826-2023-04457 | 2023-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.5/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.0/+2.5/093 | OS | Physician | 09/11/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17908786&pc=QCB | 2023826-2023-04458 | 2023-06-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+1.0/166 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN AND IT WAS REPORTED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CLAIM# (B)(4). | -6.5/+1.0/166 | OD | Physician | 09/11/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17908791&pc=QCB | 2023826-2023-04465 | 2023-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+2.0/073 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 . THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS TWO SIZES LONGER IN LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -12.5/+2.0/073 | OS | Physician | 09/12/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17908792&pc=QCB | 2023826-2023-04461 | 2023-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/071 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT. IN A SEPARATE SURGERY THAT SAME DAY A SHORTER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) (B)(4). | -9.5/+1.0/071 | OD | Physician | 09/12/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17908793&pc=QCB | 2023826-2023-04462 | 2023-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/+2.0/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT. THAT SAME DAY, A SHORTER LENGTH LENS WAS IMPLANTED, ALTHOUGH IT IS UNCLEAR WHETHER THIS WAS PERFORMED WITHIN THE SAME SURGERY THAT THE LENS WAS REMOVED. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: UPDATE PER EMAIL RECEIVED ON 23-NOV-2023: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.0 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. H6: HEALTH EFFECT - IMPACT CODE: CHANGED TO 4629. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -5.5/+2.0/093 | OS | Physician | 09/12/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17908795&pc=QCB | 2023826-2023-04466 | 2023-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/003 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. A LOW VAULT WAS NOTED. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS OF THE SAME LENGTH AND IT WAS IMPLANTED VERTICALLY. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(6). | -10.0/+1.5/003 | OD | Physician | 09/12/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17908796&pc=QCB | 2023826-2023-04469 | 2023-07-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS WITH DIOPTER -14.50/+4.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -14.50/+4.0/091 | OD | Physician | 09/12/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17908930&pc=QCB | 2023826-2023-04490 | 2023-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -18.0/1.0/105 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM (B)(4). | -18.0/1.0/105 | OD | Physician | 10/07/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17908932&pc=QCB | 2023826-2023-04488 | 2023-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.5/3.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM (B)(4). | -16.5/3.5/087 | OD | Physician | 10/09/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17909624&pc=QCB | 2023826-2023-04459 | 2023-06-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+0.5/011 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN AND IT WAS REPORTED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: (B)(4). | -8.5/+0.5/011 | OS | Physician | 09/11/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17909625&pc=QCB | 2023826-2023-04464 | 2023-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.0/005 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.5/+2.0/005 | OS | Physician | 09/12/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17909626&pc=QCB | 2023826-2023-04470 | 2023-06-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+3.0/100 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -9.5/+3.0/100 | OD | Physician | 09/12/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17909655&pc=QCB | 2023826-2023-04489 | 2023-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.5/3.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -15.5/3.5/092 | OS | Physician | 10/09/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17914243&pc=QCB | 2023826-2023-04492 | 2023-09-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -13.50/+1.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE. INTRAOPERATIVELY THE LENS WAS REMOVED WITH NO PATIENT INJURY. A REPLACEMENT LENS WAS IMPLANTED AT A LATER DATE. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -13.50/+1.0/090 | OS | Physician | 09/18/2023 | 2023 | 1995 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17914247&pc=QCB | 2023826-2023-04484 | 2023-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -13.50/+2.50/68 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2023 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND THE DEVICE DID FAIL TO PERFORM AS INTENDED AND NOTED "DESPITE A VAULT OF 1.5 CT ICL HAS ROTATED". Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -13.50/+2.50/68 | OD | Physician | 09/19/2023 | 2023 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17914248&pc=QCB | 2023826-2023-04491 | 2023-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -14.00/+2.00/71 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2023 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND THE DEVICE DID FAIL TO PERFORM AS INTENDED AND NOTED "EVEN THOUGH LENS VAULT IS GOOD THIS LENS HAS ROTATED TWICE." Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -14.00/+2.00/71 | OS | Physician | 09/19/2023 | 2023 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17916946&pc=QCB | 2023826-2023-04395 | 2023-06-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00/+2.5/002 (SPHERE/CYLINDER/AXIS), GOT CAUGHT IN THE INJECTOR DURING LOADING AND TORE. THERE WAS NO PATIENT CONTACT. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | -12.00/+2.5/002 | NA | Physician | 09/18/2023 | 2023 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17916954&pc=QCB | 2023826-2023-04394 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -07.50/+1.0/006 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: H6- MEDICAL IMPACT CODE: "4629" SHOULD BE REMOVED AND "4627" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. B3: UNK. CLAIM # (B)(4). | -07.50/+1.0/006 | OS | Physician | 09/18/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17917356&pc=QCB | 2023826-2023-04476 | 2023-09-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -7.5/1.5/069 (SPHERE/CYLINDER/AXIS) LENS TORE DURING LOADING AFTER OPENING THE VIAL. ON (B)(6) 2023, THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED, AND REPLACED WITH A REPLACEMENT LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -7.5/1.5/069 (SPHERE/CYLINDER/AXIS) LENS TORE DURING LOADING AFTER OPENING THE VIAL. ON (B)(6) 2023, THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED. ON (B)(6) 2023 A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. (B)(4). | -7.5/1.5/069 | NA | Physician | 10/03/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17917562&pc=QCB | 2023826-2023-04522 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THERE WAS AN ISSUE WITH A 12.6MM VTICMO12.6 -15.00/2.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -15.00/2.0/091 | NA | Physician | 09/13/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17917663&pc=QCB | 2023826-2023-04423 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTED INDICATED THAT A 12.1MM VTICMO12.1 -14.50/2.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS EXPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: B5 - THE REPORTED INDICATED THAT A 12.1MM VTICMO12.1 -14.50/2.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS EXCHANGED. H6 - HEALTH EFFECT - IMPACT CODE: 4627 CORRECTED TO 4629 IN INTIAL MDR. CLAIM # (B)(4). | -14.50/2.0/091 | NA | Physician | 11/22/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17924925&pc=QCB | 2023826-2023-04524 | 2023-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/1.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -8.5/1.5/088 | OS | Physician | 10/09/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17927858&pc=QCB | 2023826-2023-04530 | 2023-09-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6; -16.5/1.5/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS HAD TORE DURING LOADING, THIS OCCURRED ON (B)(6) 2023. THE LENS WAS NOT IMPLANTED. ON (B)(6) 2023 A REPLACEMENT LENS WAS LATER IMPLANTED OF THE SAME LENGTH AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS "LENS BROKEN DURING LOADING". Manufacturer Narrative: TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -16.5/1.5/92 | NA | Physician | 10/09/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17928015&pc=QCB | 2023826-2023-04495 | 2023-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+1.5/078 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.0/+1.5/078 | OD | Physician | 09/13/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17928016&pc=QCB | 2023826-2023-04493 | 2023-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+1.0/092 | OS | Physician | 09/13/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17928019&pc=QCB | 2023826-2023-04494 | 2023-08-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/+0.5/85 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -6.0/+0.5/85 | OD | Physician | 09/13/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17932900&pc=QCB | 2023826-2023-04398 | 2023-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.00/+4.0/110 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED ON (B)(6) 2023 AND THE PROBLEM WAS RESOLVED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.00/+4.0/110 | OD | Physician | 09/20/2023 | 2023 | 43 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17932925&pc=QCB | 2023826-2023-04397 | 2023-08-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.00/+1.5/151 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -07.00/+1.5/151 | OD | Physician | 09/18/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17932930&pc=QCB | 2023826-2023-04396 | 2023-08-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/092 (SPHERE/CYLINDER/AXIS), GOT CAUGHT IN THE INJECTOR DURING LOADING AND TORE. THERE WAS NO PATIENT CONTACT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.00/+2.0/092 | NA | Physician | 09/18/2023 | 2023 | 44 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17932972&pc=QCB | 2023826-2023-04393 | 2023-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.00/+1.0/112 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | -08.00/+1.0/112 | OD | Physician | 09/17/2023 | 2023 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17933032&pc=QCB | 2023826-2023-04392 | 2023-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/071 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO THE LENS WAS STUCK IN THE CARTRIDGE AND TORE/BROKE DURING INJECTION /DELIVERY INTO THE EYE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | -10.00/+1.0/071 | OS | Physician | 09/17/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17936184&pc=QCB | 2023826-2023-04538 | 2023-07-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.50/1.0/110 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.50/1.0/110 | OS | Physician | 09/18/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17936339&pc=QCB | 2023826-2023-04534 | 2023-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/1.5/130 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK CLAIM #(B)(4). | -12.0/1.5/130 | OS | Physician | 10/11/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17936340&pc=QCB | 2023826-2023-04532 | 2023-08-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/2.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "VAULT WAS ADJUSTED BY CHANGING THE POSITION OF THE LENS FIXATION." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - CORRECTED TO: "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_ 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/2.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "VAULT WAS ADJUSTED BY CHANGING THE POSITION OF THE LENS FIXATION." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN." IN THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -12.0/2.5/090 | OD | Physician | 10/12/2023 | 2023 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17936347&pc=QCB | 2023826-2023-04539 | 2023-01-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5/1.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) SHOULD ALSO HAVE BEEN INCLUDED WITH PREVIOUS MDR. CLAIM#: (B)(4). Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -9.5/1.0/093 | OD | Physician | 09/18/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17936871&pc=QCB | 2023826-2023-04507 | 2023-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+1.0/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.0/+1.0/095 | OD | Physician | 09/15/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17936872&pc=QCB | 2023826-2023-04506 | 2023-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+3.0/173 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -8.5/+3.0/173 | OS | Physician | 09/15/2023 | 2023 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17936999&pc=QCB | 2023826-2023-04503 | 2023-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+2.5/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -12.5/+2.5/092 | OD | Physician | 09/14/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17937000&pc=QCB | 2023826-2023-04501 | 2023-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.0/013 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: OFF-LABEL USE (ACD < 3.0MM) AND (OVER 45YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -9.0/+1.0/013 | OS | Physician | 09/14/2023 | 2023 | 47 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17937003&pc=QCB | 2023826-2023-04500 | 2023-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+2.0/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS OF THE SAME LENGTH AND IT WAS IMPLANTED VERTICALLY. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -13.0/+2.0/093 | OS | Physician | 09/14/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17937006&pc=QCB | 2023826-2023-04504 | 2023-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+2.5/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -12.5/+2.5/092 | OS | Physician | 09/14/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17944985&pc=QCB | 2023826-2023-04540 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-10-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -14.00/2.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. PER REPORTER THE LENS IS OVER VAULTED AND REPORTS A PLANNED EXPLANT/EXCHANGE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.00/2.0/090 | OD | Other Health Care Professional | 09/18/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17946044&pc=QCB | 2023826-2023-04513 | 2023-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.0/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS TWO SIZES LONGER IN LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.0/+2.0/090 | OD | Physician | 09/17/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17946048&pc=QCB | 2023826-2023-04514 | 2023-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.5/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) . CLAIM# (B)(4). | -10.5/+1.5/092 | OD | Physician | 09/17/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17946060&pc=QCB | 2023826-2023-04509 | 2023-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/105 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+1.0/105 | OS | Physician | 09/17/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17946061&pc=QCB | 2023826-2023-04510 | 2023-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+2.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.0/+2.0/093 | OD | Physician | 09/17/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17946069&pc=QCB | 2023826-2023-04511 | 2023-05-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+2.5/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) . CLAIM# (B)(4). | -7.0/+2.5/093 | OD | Physician | 09/17/2023 | 2023 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17946078&pc=QCB | 2023826-2023-04512 | 2023-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+2.0/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+2.0/090 | OS | Physician | 09/17/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17946685&pc=QCB | 2023826-2023-04517 | 2023-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.5/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+1.5/094 | OD | Physician | 09/18/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17946686&pc=QCB | 2023826-2023-04519 | 2023-08-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/+2.0/079 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. AN EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS OF THE SAME LENGTH AND IT WAS IMPLANTED VERTICALLY. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -6.0/+2.0/079 | OS | Physician | 09/18/2023 | 2023 | 45 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17946687&pc=QCB | 2023826-2023-04508 | 2023-01-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.0/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) AND (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -10.0/+1.0/090 | OS | Physician | 09/18/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17947869&pc=QCB | 2023826-2023-04580 | 2023-07-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+1.50/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT) CLAIM# (B)(4). | -8.0/+1.50/091 | OD | Physician | 09/18/2023 | 2023 | 48 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17947870&pc=QCB | 2023826-2023-04582 | 2023-08-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.50/+1.00/004 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -5.50/+1.00/004 | OS | Physician | 09/18/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17947871&pc=QCB | 2023826-2023-04583 | 2023-06-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.50/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.50/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -9.0/+1.50/091 | OD | Physician | 09/18/2023 | 2023 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17947872&pc=QCB | 2023826-2023-04581 | 2023-07-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.5/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -10.0/+2.5/089 | OS | Physician | 09/18/2023 | 2023 | 48 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17948872&pc=QCB | 2023826-2023-04429 | 2023-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6. -13.5/2.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS TORE/BROKE DURING INJECTION/DELIVER INTO THE EYE. THERE WAS PATIENT CONTACT(LENS TOUCHED THE EYE). NO PATIENT INJURY. THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. REPORTEDLY, PATIENT AND SURGEON ARE HAPPY. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.5/2.5/093 | OD | Physician | 09/25/2023 | 2023 | 48 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17948873&pc=QCB | 2023826-2023-04430 | 2023-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2. -12.5/2.5/071 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS TORE/BROKE DURING INJECTION/DELIVER INTO THE EYE. THERE WAS PATIENT CONTACT(LENS TOUCHED THE EYE). NO PATIENT INJURY. THE LENS WAS EXCHANGED SAME LENGTH LENS ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. REPORTEDLY, PATIENT AND SURGEON ARE HAPPY. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.5/2.5/071 | OS | Physician | 09/25/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17948874&pc=QCB | 2023826-2023-04431 | 2023-09-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -9.5/1.5/1100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING INJECTION/DELIVERY INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REPLACED INTRAOPERATIVELY WITH THE BACK UP LENS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS OTHER: LENS QUALITY (YELLOWING FOUND AROUND EDGES OF LENS). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.5/1.5/1100 | OD | Physician | 09/25/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17948875&pc=QCB | 2023826-2023-04433 | 2023-08-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -4.0/0.5/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT IS HAPPY. Manufacturer Narrative: CLAIM# (B)(4). | -4.0/0.5/100 | OS | Physician | 09/26/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17948878&pc=QCB | 2023826-2023-04434 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICM5_12.6. -11.00/1.0/005 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD). THE LENS TORE/BROKE DURING INJECTION/DELIVER INTO THE EYE. THERE WAS PATIENT CONTACT(LENS TOUCHED THE EYE). NO PATIENT INJURY. THE BACK UP LENS WAS IMPLANTED INTRAOPERATIVLY. THIS RESOLVED THE PROBLEM. REPORTEDLY, SURGEON MAY HAVE CAUSED PROBLEM LOADING LENS. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.00/1.0/005 | OD | Physician | 09/27/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17948879&pc=QCB | 2023826-2023-04432 | 2023-08-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -5.5/0.5/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT IS HAPPY. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. CLAIM# (B)(4). | -5.5/0.5/102 | OD | Physician | 09/28/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17948881&pc=QCB | 2023826-2023-04436 | 2023-09-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -10.0/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A LONGER LENGTH LENS ON (B)(6) 2023 FROM PATIENT'S LEFT RIGHT EYE (OS) DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -10.0/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY ON (B)(6) 2023 FROM PATIENT'S LEFT RIGHT EYE (OS) DUE TO LOW VAULT. ON AN UNKNOWN DATE THE LENS WAS REPLACED WITH A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. D6B - CORRECTED TO UNK IN THE INITIAL MDR. CLAIM (B)(4). | -10.0/1.0/089 | OS | Physician | 11/13/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17948882&pc=QCB | 2023826-2023-04437 | 2023-08-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.0/1.0/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANGED DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT ON (B)(6) 2023. IN A SECOND SURGERY ON (B)(6) 2023 A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS OTHER: SIZING ISSUE. Manufacturer Narrative: A2 DOB: (B)(6) 2023 CORRECTED TO (B)(6) 1995. B3 - DATE OF EVENT: 31-MAR-2023 CORRECTED TO 31-AUG-2023. H6 - HEALTH EFFECT - IMPACT CODE: 4627-4629 CORRECTED TO 4629 IN ITIALL MDR. CLAIM # (B)(4). Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - HEALTH EFFECT CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGELS. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.0/1.0/109 | OD | Physician | 09/27/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17948883&pc=QCB | 2023826-2023-04435 | 2023-09-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 12.6MM VTICMO12.6. -13.0/2.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT(LENS TOUCHED THE EYE). NO PATIENT INJURY. A REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H1 - TYPE OF EVENT - SERIOUS CORRECTED TO MALFUNCTION. CLAIM # (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.0/2.5/087 | OS | Physician | 11/23/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17948885&pc=QCB | 2023826-2023-04440 | 2023-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.5/4.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LOW VAULT OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). | -11.5/4.0/075 | OS | Physician | 09/28/2023 | 2023 | 49 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17948894&pc=QCB | 2023826-2023-04439 | 2023-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.5/4.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 202322. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. CLAIM# (B)(4). | -12.5/4.5/094 | OD | Physician | 09/28/2023 | 2023 | 49 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17948897&pc=QCB | 2023826-2023-04443 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -2.5/1.5/165 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS. HAS AN EVENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM# (B)(4). | -2.5/1.5/165 | OS | Physician | 09/28/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17951047&pc=QCB | 2023826-2023-04483 | 2023-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -16.0/2.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY, "DURING THE RE-IMPLANTATION PROCESS, THE ICL WAS BROKEN DURING THE INJECTION, AND THE OPERATION WAS PERFORMED AFTER RE-ORDERING THE LENS. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -16.0/2.0/094 | OS | Physician | 10/07/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17951385&pc=QCB | 2023826-2023-04608 | 2023-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/2.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -10.0/2.5/092 | OD | Physician | 10/16/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17951514&pc=QCB | 2023826-2023-04609 | 2023-09-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -10.5/1.5/088 | OD | Physician | 10/16/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17951782&pc=QCB | 2023826-2023-04612 | 2023-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.5/3.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -15.5/3.0/090 | OD | Physician | 10/16/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17952359&pc=QCB | 2023826-2023-04610 | 2023-09-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/1.5/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). Manufacturer Narrative: D4 - CORRECTED TO: SERIAL # (B)(6) IN THE INITIAL MDR. CLAIM # (B)(4). | -9.5/1.5/073 | OS | Physician | 10/16/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17952382&pc=QCB | 2023826-2023-04607 | 2023-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.50/2.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -9.50/2.5/088 | OS | Physician | 10/16/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17953777&pc=QCB | 2023826-2023-04617 | 2023-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -12.0/1.0/091 | OD | Physician | 10/16/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17953779&pc=QCB | 2023826-2023-04613 | 2023-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_ 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-HEIGHT: UNK. DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) CLAIM #((B)(4). | -6.0/1.0/090 | OD | Physician | 10/16/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17953780&pc=QCB | 2023826-2023-04616 | 2023-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/1.0/152 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -7.0/1.0/152 | OD | Physician | 10/16/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17953781&pc=QCB | 2023826-2023-04618 | 2023-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/2.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -12.0/2.0/090 | OS | Physician | 10/16/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17953943&pc=QCB | 2023826-2023-04619 | 2023-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/2.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -14.0/2.5/092 | OD | Physician | 10/16/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17953987&pc=QCB | 2023826-2023-04590 | 2023-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.5/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM#: (B)(4). | -11.0/+1.5/093 | OS | Physician | 09/20/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17953989&pc=QCB | 2023826-2023-04589 | 2023-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+2.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: (B)(4). | -13.0/+2.0/091 | OD | Physician | 09/20/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17953990&pc=QCB | 2023826-2023-04587 | 2023-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+4.0/089 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: (B)(4). | -11.5/+4.0/089 | OD | Physician | 09/20/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17953992&pc=QCB | 2023826-2023-04588 | 2023-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+4.5/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+4.5/094 | OS | Physician | 09/20/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17954021&pc=QCB | 2023826-2023-04624 | 2023-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+1.0/109 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.0/+1.0/109 | OD | Physician | 09/22/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17954110&pc=QCB | 2023826-2023-04620 | 2023-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.5/039 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -10.5/1.5/039 | OS | Physician | 10/16/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17954114&pc=QCB | 2023826-2023-04586 | 2023-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) CLAIM# (B)(4). | -10.0/+1.0/093 | OD | Physician | 09/19/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17954115&pc=QCB | 2023826-2023-04591 | 2023-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+3.5/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -12.0/+3.5/091 | OD | Physician | 09/20/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17954291&pc=QCB | 2023826-2023-04543 | 2023-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.5/5.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS OF DIFFERENT POWER. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.5/5.0/090 | OD | Physician | 09/20/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17954511&pc=QCB | 2023826-2023-04629 | 2023-08-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -16.5/1.5/074 (SPHERE/CYLINDER/AXIS) TORE DURING LOADING. THERE WAS NO PATIENT CONTACT. A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | -16.5/1.5/074 | NA | Physician | 09/22/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17961774&pc=QCB | 2023826-2023-04643 | 2023-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED, THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.0/123 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS, DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6: DEVICE CODE 1494: OFF LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -10.5/1.0/123 | OD | Physician | 10/17/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17961775&pc=QCB | 2023826-2023-04641 | 2023-08-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/2.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -14.0/2.0/094 | OD | Physician | 10/17/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17961819&pc=QCB | 2023826-2023-04560 | 2023-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6-HEALTH EFFECT- IMPACT CODE: "4629" SHOULD BE REMOVED AND CODE "4627" SHOULD BE ADDED. CLAIM# (B)(4). Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. CLAIM # (B)(4). | -11.00/+1.5/094 | OD | Physician | 09/28/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17961821&pc=QCB | 2023826-2023-04555 | 2023-09-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS LEFT EYE (OS). THE LENS HAD EXCESSIVE VAULTING AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -05.50/+0.5/104 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. D2: COMMON DEVICE NAME: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | -05.50/+0.5/104 | OS | Physician | 09/24/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17961824&pc=QCB | 2023826-2023-04554 | 2023-09-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS RIGHT EYE (OD). THE LENS HAD EXCESSIVE VAULTING AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -04.50/+0.5/087 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. CLAIM # (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | -04.50/+0.5/087 | OD | Physician | 09/24/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17962121&pc=QCB | 2023826-2023-04571 | 2023-08-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -11.0/4.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LOW VAULT WITHOUT ROTATION IS OBSERVED. CAUSE OF THE EVENT IS USER ERROR WITH THE DEVICE NOT FAILING TO PERFORM AS INTENDED. REPORTEDLY, "ANATOMICAL DIFFERENCE OF THE PATIENT. ERROR IN MEASUREMENT OF WTW." THE PROBLEM WAS NOT RESOLVED. THE STATUS OF THE EYE IS, "CALM EYE WITHOUT SYMPTOMS." Manufacturer Narrative: H6: USER ERROR. CLAIM# (B)(4). | -11.0/4.0/093 | OS | Physician | 10/12/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17965981&pc=QCB | 2023826-2023-04549 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -8.00/1.0/054 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS EYE. EXCESSIVE VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -8.00/1.0/054 | NA | Other Health Care Professional | 09/19/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17966127&pc=QCB | 2023826-2023-04550 | 2023-09-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6 VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -8.0/5.5/090 (SPHERE/CYLINDER/AXIS) LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY ON (B)(6) 2023. A REPLACEMENT LENS WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: A4, A5, A6-UNK. H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -8.0/5.5/090 | NA | Physician | 09/19/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17966723&pc=QCB | 2023826-2023-04650 | 2023-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -15.0/1.5/111 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LET EYE (OS) ON (B)(6)2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM) CLAIM# (B)(4). | -15.0/1.5/111 | OS | Physician | 10/12/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17969209&pc=QCB | 2023826-2023-04662 | 2023-09-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/1.5/112 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED A LOW VAULT WITHOUT ROTATION. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #7(B)(4). | -12.5/1.5/112 | OD | Physician | 10/18/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17969210&pc=QCB | 2023826-2023-04666 | 2023-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.50/0.5/075 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -6.50/0.5/075 | OS | Physician | 10/18/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17969255&pc=QCB | 2023826-2023-04663 | 2023-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/1.5/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -11.0/1.5/073 | OD | Physician | 10/18/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17969324&pc=QCB | 2023826-2023-04667 | 2023-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/4.0/016 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "THE ICL IS SHIFTED FROM A VERTICAL POSITION TO A HORIZONTAL POSITION." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -13.0/4.0/016 | OD | Physician | 10/18/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17971145&pc=QCB | 2023826-2023-04633 | 2023-07-10 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 13.7MM, VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF -7.5/2.5/087 (SPHERE/CYLINDER/AXIS) TORE DURING LOADING. THERE WAS NO PATIENT CONTACT. ON (B)(6) 2023 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). | -7.5/2.5/087 | NA | Physician | 09/25/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17971327&pc=QCB | 2023826-2023-04635 | 2023-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -6.00/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO LOW VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -6.00/2.0/089 | OS | Physician | 09/26/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17971367&pc=QCB | 2023826-2023-04637 | 2023-05-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -7.50/1.5/114 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). | -7.50/1.5/114 | OD | Physician | 09/27/2023 | 2023 | 47 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17971382&pc=QCB | 2023826-2023-04638 | 2023-05-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -5.50/2.0/104 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -5.50/2.0/104 | OS | Physician | 09/27/2023 | 2023 | 47 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17971403&pc=QCB | 2023826-2023-04573 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -9.0/0.5/096 (SPHERE/CYLINDER/AXIS) LENS TORE DURING LOADING AND WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A REPLACEMENT LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -9.0/0.5/096 | OD | Physician | 10/13/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17971798&pc=QCB | 2023826-2023-04639 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF -11.0/3.0/062 (SPHERE/CYLINDER/AXIS) HAD TO BE RELOADED DURING INSERTION AND IT WAS OBSERVED THAT A PORTION OF THE HAPTIC WAS TORN. THE SURGEON DECIDED TO USE THE LENS BECAUSE THE OPTIC WAS NOT AFFECTED. THE LENS WAS IMPLANTED IN THE PATIENTS LEFT EYE (OS) AND REMAINS IMPLANTED. NO PATIENT INJURY HAS BEEN REPORTED. Manufacturer Narrative: D6A:UNK. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -11.0/3.0/062 | OS | Physician | 09/27/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17971999&pc=QCB | 2023826-2023-04652 | 2023-09-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -16.00/2.0/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). ON (B)(6) 2023 THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY DUE TO THE LENS BEING IMPLANTED UPSIDE DOWN. NO PATIENT INJURY REPORTED. A REPLACEMENT LENS OF THE SAME MODEL AND SIZE AND SIMILAR POWER WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). | -16.00/2.0/086 | OD | Physician | 09/20/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17972000&pc=QCB | 2023826-2023-04651 |  | Injury | STAAR SURGICAL COMPANY | 2023-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -10.50/5.5/020 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). LENS ROTATION WAS REPORTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - HEALTH EFFECTS - IMPACT CODE: 2199 CORRECTED TO 4627. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -10.50/5.5/020 | OS | Physician | 12/22/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17972048&pc=QCB | 2023826-2023-04654 | 2018-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -16.50/2.0/079 (SPHERE/CYLINDER/AXIS) LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS OF THE SAME SIZE AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -16.50/2.0/079 | OS | Physician | 09/21/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17978592&pc=QCB | 2023826-2023-04561 | 2023-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/+2.5/071 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -11.00/+2.5/071 | OS | Physician | 09/28/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17978602&pc=QCB | 2023826-2023-04562 | 2023-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/012 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2023 BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3: TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -09.50/+1.0/012 | OS | Physician | 09/27/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17981335&pc=QCB | 2023826-2023-04702 | 2023-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/1.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -12.5/1.0/093 | OS | Physician | 10/18/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17981336&pc=QCB | 2023826-2023-04703 | 2023-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/3.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -9.5/3.0/088 | OS | Physician | 10/18/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17982426&pc=QCB | 2023826-2023-04713 | 2023-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF - 2.00/3.5/097 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 EXCHANGED FOR A SAME MODEL/SIZE LENS BUT DIFFERENT POWER DUE TO LOW VAULT. THE EXCHANGE LENS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: H10: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL OBSERVATION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | - 2.00/3.5/097 | OD | Physician | 09/22/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17986863&pc=QCB | 2023826-2023-04598 | 2023-08-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6, -9.50/1.0/062 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. LOW VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. DEVICE DID NOT FAIL TO PERFORM AS INTENDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4) Manufacturer Narrative: ADDITIONAL INFORMATION: B5: REPORTEDLY, "ICL VAULTS ARE TOO LOW AND I'M SEEING SOME EARLY CATARACT DEVELOPMENT, SO WILL NEED TO DO AN EXCHANGE". CLAIM# (B)(4). | -9.50/1.0/062 | OS | Physician | 10/02/2023 | 2023 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17986864&pc=QCB | 2023826-2023-04605 | 2023-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -7.0/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. A SAME LENGTH LENS WAS EXCHANGED ON 01/AUG/2023 DUE TO A FOOTPLATE WAS FOUND MISSING FROM THE LENS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT REPORTED AS DEVICE. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: THE DHR REVIEW INDICATES THAT THE PRODUCT HAS NOT BEEN MANUFACTURED WITHIN THE ESTABLISHED PROCESS PARAMETERS AND THAT THERE IS INDICATION THAT THE MANUFACTURING AND/OR PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. A DEFINITIVE ROOT CAUSE FOR THIS EVENT WAS UNABLE TO BE DETERMINED. AS PREVENTATIVE ACTION, (B)(4) AND (B)(4) HAVE BEEN UPDATED TO REDUCE THE LIKELIHOOD OF DAMAGE CAUSED BY INADVERTENT CONTACT BETWEEN THE LENS AND THE DRY GLASS SURFACE OF THE 2ML VIAL. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: 1494: OFF-LABEL. AGE <21 AT TIME OF IMPLANTATION. CLAIM#(B)(4). | -7.0/2.0/090 | OS | Physician | 10/06/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17986869&pc=QCB | 2023826-2023-04602 | 2023-06-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -15.5/4.0/175 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS REPOSITIONED AND DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS, THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH REDIUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -15.5/4.0/175 | OS | Physician | 10/04/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17988695&pc=QCB | 2023826-2023-04664 | 2023-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/1.5/070 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -12.0/1.5/070 | OS | Physician | 10/18/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17989284&pc=QCB | 2023826-2023-04739 | 2023-08-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -11.0/3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023.THE LENS WAS IMPLANTED, REMOVED AND REPLACED INTRAOPERATIVELY AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -11.0/3.0/093 | OS | Physician | 09/28/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17990984&pc=QCB | 2023826-2023-04740 | 2023-08-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -6.0/2.0/164 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: D4 - EXPIRATION DATE: 31-DEC-2023 CORRECTED TO 28-FEB-2026 IN THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL AND RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | -6.0/2.0/164 | OD | Physician | 12/04/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17991136&pc=QCB | 2023826-2023-04741 | 2023-08-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -6.0/3.5/002 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -6.0/3.5/002 | OS | Physician | 12/04/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17992483&pc=QCB | 2023826-2023-04672 | 2023-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.5/112 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) . CLAIM# (B)(4). | -11.0/+1.5/112 | OD | Physician | 09/12/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17992490&pc=QCB | 2023826-2023-04671 | 2023-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+2.0/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -8.5/+2.0/095 | OD | Physician | 09/25/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17999176&pc=QCB | 2023826-2023-04700 | 2023-07-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00/+2.5/095 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO THE LENS WAS STUCK IN THE CARTRIDGE OR INJECTION SYSTEM. THE LENS WAS DAMAGED AND THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/+2.5/095 | NA | Physician | 10/06/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17999351&pc=QCB | 2023826-2023-04697 | 2023-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+3.0/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LOW VAULTING AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | -11.50/+3.0/086 | OD | Physician | 10/02/2023 | 2023 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17999361&pc=QCB | 2023826-2023-04698 | 2023-09-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.00/+2.5/112, TORE DURING LOADING AND BECAME STUCK IN THE CARTRIDGE OR INJECTION SYSTEM. THERE WAS PATIENT CONTACT BUT NO INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/ LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -17.00/+2.5/112 | NA | Physician | 10/02/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17999693&pc=QCB | 2023826-2023-04746 | 2023-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/1.0/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -9.0/1.0/072 | OS | Physician | 10/22/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17999706&pc=QCB | 2023826-2023-04745 | 2023-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.0/2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -16.0/2.0/092 | OD | Physician | 10/22/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17999711&pc=QCB | 2023826-2023-04744 | 2023-08-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/1.0/148 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -9.0/1.0/148 | OS | Physician | 10/21/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17999712&pc=QCB | 2023826-2023-04751 | 2023-07-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/2.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.0/2.5/090 | OS | Physician | 10/23/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17999716&pc=QCB | 2023826-2023-04752 | 2023-09-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.0/070 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK . H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -10.5/1.0/070 | OS | Physician | 10/23/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=17999717&pc=QCB | 2023826-2023-04747 | 2023-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/2.5/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -14.0/2.5/080 | OS | Physician | 10/23/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18001276&pc=QCB | 2023826-2023-04673 | 2023-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.5/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -10.5/+2.5/095 | OD | Physician | 09/26/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18001277&pc=QCB | 2023826-2023-04681 | 2023-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+3.0/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -11.5/+3.0/089 | OS | Physician | 09/26/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18001278&pc=QCB | 2023826-2023-04675 | 2023-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+2.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -12.5/+2.0/091 | OD | Physician | 09/26/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18001279&pc=QCB | 2023826-2023-04680 | 2023-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+3.0/089 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -12.0/+3.0/089 | OD | Physician | 09/26/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18007305&pc=QCB | 2023826-2023-04778 | 2023-09-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -11.00/2.0/118 (SPHERE/CYLINDER/AXIS) LENS TEAR/BREAK DURING LOADING. THE LENS WAS NOT IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -11.00/2.0/118 | OD | Physician | 09/25/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18007459&pc=QCB | 2023826-2023-04779 | 2023-09-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -14.50/2.0/092 (SPHERE/CYLINDER/AXIS) LENS TEAR/BREAK DURING LOADING. THE LENS WAS NOT IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -14.50/2.0/092 | OD | Physician | 09/25/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18007606&pc=QCB | 2023826-2023-04780 | 2023-12-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 OF -9.00/1.0/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE/BROKE DURING LOADING ON (B)(6) 2023. THE LENS WAS NOT IMPLANTED INTO THE PATIENTS LEFT EYE (OS). A REPLACEMENT LENS OF THE SAME MODEL/SIZE AND POWER WAS IMPLANTED ON (B)(6) 2023 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -9.00/1.0/095 | OS | Physician | 09/26/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18007846&pc=QCB | 2023826-2023-04781 | 2023-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -12.00/2.0/87 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED AND THE PROBLEM WAS NOT RESOLVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H10: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -12.00/2.0/87 | OS | Physician | 09/26/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18008531&pc=QCB | 2023826-2023-04809 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) AND SEEMS TO HAVE HIGH VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#:(B)(4). Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -18.00/5.5/170 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) AND SEEMS TO HAVE HIGH VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -18.00/5.5/170 | OS | Physician | 09/29/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18015704&pc=QCB | 2023826-2023-04677 | 2023-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+1.5/108 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LOW VAULT. A LENS TWO SIZES LONGER IN LENGTH WAS IMPLANTED ON (B)(6) 2023 AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -6.5/+1.5/108 | OD | Physician | 09/27/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18015705&pc=QCB | 2023826-2023-04685 | 2023-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+3.0/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) CLAIM# (B)(4). | -12.5/+3.0/090 | OD | Physician | 09/27/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18015707&pc=QCB | 2023826-2023-04679 | 2023-08-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/174 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS OF THE SAME LENGTH AND IT WAS IMPLANTED HORIZONTALLY DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITONAL DATA: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -9.5/+1.0/174 | OS | Physician | 09/27/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18015728&pc=QCB | 2023826-2023-04682 | 2020-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. LATER ON (B)(6) 2023 A SHORTER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -6.5/1.5/082 | OD | Physician | 09/27/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18016305&pc=QCB | 2023826-2023-04676 | 2023-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+2.0/169 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -7.0/+2.0/169 | OS | Physician | 09/27/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18016634&pc=QCB | 2023826-2023-04784 | 2023-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -7.00/1.0/014 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -7.00/1.0/014 | OS | Physician | 09/27/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18017064&pc=QCB | 2023826-2023-04753 | 2023-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/1.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK . H6-DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -7.5/1.0/093 | OD | Physician | 10/23/2023 | 2023 | 48 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18017065&pc=QCB | 2023826-2023-04815 | 2023-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/3.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -14.5/3.0/093 | OD | Physician | 10/24/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18017066&pc=QCB | 2023826-2023-04817 | 2023-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/1.5/071 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "THE LENS EXCHANGED WAS IMPLANTED VERTICALLY. AS REPLACEMENT LENS WERE NOT IN STOCK, THE ICL LENS DEGREE WAS ADJUSTED." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -9.5/1.5/071 | OD | Physician | 10/24/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18017067&pc=QCB | 2023826-2023-04818 | 2023-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/2.5/109 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "THE LENS EXCHANGED WAS IMPLANTED VERTICALLY. AS REPLACEMENT LENS WERE NOT IN STOCK, THE ICL LENS DEGREE WAS ADJUSTED." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -9.0/2.5/109 | OS | Physician | 10/24/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18017068&pc=QCB | 2023826-2023-04816 | 2023-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/1.0/114 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -13.5/1.0/114 | OS | Physician | 10/24/2023 | 2023 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18017242&pc=QCB | 2023826-2023-04844 | 2023-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -13.0/2.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS ROTATION. ON (B)(6) 2023 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT IMPLANTED FROM HORIZONTAL TO VERTICAL POSITION AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -13.0/2.0/093 | OD | Physician | 10/15/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18017386&pc=QCB | 2023826-2023-04786 | 2023-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -8.00/1.5/086 (SPHERE/CYLINDER AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. LENS ROTATION WAS OBSERVED. LENS WAS REPOSITIONED AND THE PROBLEM WAS NOT RESOLVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS OF THE SAME SIZE AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -8.00/1.5/086 | OS | Physician | 09/28/2023 | 2023 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18023231&pc=QCB | 2023826-2023-04684 | 2023-07-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+2.5/170 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -13.5/+2.5/170 | OS | Physician | 09/27/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18023249&pc=QCB | 2023826-2023-04687 | 2023-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.0/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS TWO SIZES LONGER IN LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+2.0/091 | OS | Physician | 09/27/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18023250&pc=QCB | 2023826-2023-04686 | 2023-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.0/099 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS TWO SIZES LONGER IN LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+2.0/099 | OD | Physician | 09/27/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18023251&pc=QCB | 2023826-2023-04691 | 2023-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.0/+1.5/168 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -17.0/+1.5/168 | OD | Physician | 09/28/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18023282&pc=QCB | 2023826-2023-04692 | 2023-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -18.0/+1.0/010 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -18.0/+1.0/010 | OS | Physician | 09/28/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18024739&pc=QCB | 2023826-2023-04695 | 2023-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.5/+1.5/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -3.5/+1.5/090 | OS | Physician | 09/28/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18024740&pc=QCB | 2023826-2023-04696 | 2023-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+3.0/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/+3.0/094 | OD | Physician | 09/28/2023 | 2023 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18024747&pc=QCB | 2023826-2023-04694 | 2023-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.0/+1.5/097 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -4.0/+1.5/097 | OD | Physician | 09/28/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18024795&pc=QCB | 2023826-2023-04689 | 2023-08-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.5/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) AND (UNDER 21YRS OF AGE AT DATE OF IMPLANT(B)(4). | -11.0/+1.5/090 | OD | Physician | 09/28/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18024796&pc=QCB | 2023826-2023-04690 | 2023-08-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+2.0/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) AND (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -11.5/+2.0/090 | OS | Physician | 09/28/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18025192&pc=QCB | 2023826-2023-04693 | 2023-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+1.5/109 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -7.5/+1.5/109 | OS | Physician | 09/28/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18026421&pc=QCB | 2023826-2023-04824 | 2023-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/4.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -10.0/4.5/093 | OS | Physician | 10/24/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18026422&pc=QCB | 2023826-2023-04823 | 2023-10-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/1.0/172 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -9.5/1.0/172 | OD | Physician | 10/24/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18026424&pc=QCB | 2023826-2023-04819 | 2023-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/0.5/174 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -10.5/0.5/174 | OD | Physician | 10/23/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18026425&pc=QCB | 2023826-2023-04820 | 2023-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/1.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -9.5/1.5/095 | OD | Physician | 10/23/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18026427&pc=QCB | 2023826-2023-04827 | 2023-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -18.0/1.0/084 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -18.0/1.0/084 | OS | Physician | 10/24/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18026429&pc=QCB | 2023826-2023-04826 | 2023-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -18.0/1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: (B)(4). CLAIM # (B)(4). | -18.0/1.0/091 | OD | Physician | 10/24/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18026431&pc=QCB | 2023826-2023-04832 | 2023-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/1.0/071(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -9.0/1.0/071 | OS | Physician | 10/25/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18026432&pc=QCB | 2023826-2023-04831 | 2023-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/1.5/095(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -9.5/1.5/095 | OD | Physician | 10/25/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18026433&pc=QCB | 2023826-2023-04830 | 2023-10-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/1.5/091(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -5.5/1.5/091 | OS | Physician | 10/25/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18026434&pc=QCB | 2023826-2023-04828 | 2023-07-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/1.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -12.5/1.5/093 | OD | Physician | 10/24/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18026499&pc=QCB | 2023826-2023-04735 | 2023-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.0/15/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -11.0/15/089 | OD | Physician | 10/18/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18030400&pc=QCB | 2023826-2023-04858 | 2023-10-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: HOLD VMTHE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -2.5/2/118 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THERE IS EXCESSIVE VAULTING. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -2.5/2/118 | OD | Physician | 10/06/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18035691&pc=QCB | 2023826-2023-04859 | 2023-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.5/+3.5/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -15.5/+3.5/090 | OD | Physician | 10/02/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18035692&pc=QCB | 2023826-2023-04863 | 2023-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.5/+5.5/086 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM#: (B)(4). | -15.5/+5.5/086 | OS | Physician | 10/04/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18036028&pc=QCB | 2023826-2023-04834 | 2023-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0/2.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -15.0/2.5/089 | OS | Physician | 10/25/2023 | 2023 | 41 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18036029&pc=QCB | 2023826-2023-04835 | 2023-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_ 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/3.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -11.0/3.5/094 | OD | Physician | 10/25/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18036030&pc=QCB | 2023826-2023-04836 | 2023-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_ 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/3.0/095(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -11.0/3.0/095 | OS | Physician | 10/25/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18036032&pc=QCB | 2023826-2023-04833 | 2023-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/2.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -10.5/2.5/095 | OD | Physician | 10/25/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18036033&pc=QCB | 2023826-2023-04838 | 2023-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_ 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/1.0/084(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -6.0/1.0/084 | OS | Physician | 10/25/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18036036&pc=QCB | 2023826-2023-04839 | 2023-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_ 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/3.5/090(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -11.5/3.5/090 | OD | Physician | 10/25/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18036039&pc=QCB | 2023826-2023-04840 | 2023-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_ 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/4.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -12.5/4.0/090 | OS | Physician | 10/25/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18036040&pc=QCB | 2023826-2023-04892 | 2023-10-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.50/1.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -6.50/1.0/088 | OS | Physician | 10/26/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18036043&pc=QCB | 2023826-2023-04890 | 2023-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/1.5/172 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -12.0/1.5/172 | OD | Physician | 10/26/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18036045&pc=QCB | 2023826-2023-04893 | 2023-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "THE PATIENT'S SCLEROCRATIC HEIGHT IS TOO HIGH, AND THE CLINICIAN CONFIRMS THAT THE LENS NEEDS TO BE REPLACED FROM A TRANSVERSE TO A VERTICAL POSITION." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -10.5/1.5/074 | OS | Physician | 10/26/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18048065&pc=QCB | 2023826-2023-04716 | 2023-06-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7, -9.0/1.5/138 (SPHERE/CYLINDER/AXIS) THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS RIGHT EYE ON (B)(6) 2023. THERE WAS PATIENT CONTACT(LENS TOUCHED THE EYE). NO PATIENT INJURY. A NEW SAME LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.0/1.5/138 | OD | Physician | 10/12/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18048072&pc=QCB | 2023826-2023-04722 | 2023-10-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7. -14.0/1.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE. THE LENS WAS OBSERVED BEING UPSIDE DOWN IN PATIENTS EYE. THE LENS WAS REMOVED AND A BACK UP LENS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5 - REPORTEDLY, STATUS OF THE EYE IS "HEALING WELL". (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION ADDED TO FORM. CLAIM# (B)(4). | -14.0/1.0/087 | NA | Physician | 12/21/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18048074&pc=QCB | 2023826-2023-04717 | 2023-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -10.0/3.5/002 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS REPOSITIONED. THIS DID NOT RESOLVE THE PROBLEM. A LONGER LENGTH LENS WAS EXCHANGED ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -10.0/3.5/002 | OD | Physician | 10/10/2023 | 2023 | 42 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18048079&pc=QCB | 2023826-2023-04724 |  | Injury | STAAR SURGICAL COMPANY | 2023-11-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -11.0/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 UNK. A5 UNK. A6 UNK. B3 UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.0/2.5/091 | OD | Physician | 10/13/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18048332&pc=QCB | 2023826-2023-04871 | 2023-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+1.5/006 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -14.0/+1.5/006 | OS | Physician | 10/07/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18050166&pc=QCB | 2023826-2023-04909 | 2023-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -8.0/1.0/091 | OD | Physician | 10/29/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18050182&pc=QCB | 2023826-2023-04911 | 2023-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/2.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -8.0/2.5/089 | OD | Physician | 10/29/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18050183&pc=QCB | 2023826-2023-04910 | 2023-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/2.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL/ LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "ACCORDING TO THE ASSESSMENT OF POST-OP ARCH, CHANGE THE ICL POSITION FROM HORIZONTAL TO VERTICAL." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # : (B)(4). | -9.5/2.5/094 | OD | Physician | 10/29/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18050427&pc=QCB | 2023826-2023-04912 | 2023-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/2.0/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -7.0/2.0/097 | OS | Physician | 10/29/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18052033&pc=QCB | 2023826-2023-04846 |  | Injury | STAAR SURGICAL COMPANY | 2023-11-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF A -14.0/3.0/91 (SPHERE/CYLINDER/AXIS) LENS WAS INTRAOPERATIVELY IMPLANTED AND REMOVED ON (B)(6) 2023 DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE PATIENT'S LEFT EYE (OS). ON (B)(6) 2023, A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THERE WAS PATIENT CONTACT AND PATIENT INJURY. REPORTEDLY, "LENS GOT STUCKED IN THE PLUNGER." STATUS OF THE EYE IS QUIET. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+3.0/91 (SPHERE/CYLINDER/AXIS) INTO THE LEFT EYE (OS) ON (B)(6) 2023. REPORTEDLY THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. ON (B)(6) 2023 A LENS OF THE SAME PARAMETERS WAS IMPLANTED, AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE AS "QUIET". H6: HEALTH EFFECT- IMPACT CODE: "4627" SHOULD BE REMOVED AND CODE "2199" SHOULD BE ADDED. CLAIM#: (B)(4). Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -14.0/3.0/91 | OS | Physician | 10/24/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18052035&pc=QCB | 2023826-2023-04848 | 2023-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.0/2.0/171 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -10.0/2.0/171 | OD | Physician | 10/24/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18052095&pc=QCB | 2023826-2023-04757 | 2023-09-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.00/+2.0/084 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS WAS IMPLANTED UPSIDE DOWN IN THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -06.00/+2.0/084 | OS | Physician | 10/08/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18052236&pc=QCB | 2023826-2023-04928 |  | Injury | STAAR SURGICAL COMPANY | 2023-11-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -9.5/3.0/089 (SPHERE/CYLINER/AXIS) IMPLANTABLE COLLAMER LENS WAS REMOVED BECAUSE IT FLIPPED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -9.5/3.0/089 | NA | Physician | 10/04/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18052310&pc=QCB | 2023826-2023-04849 | 2023-10-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM513.7 IMPLANTABLE COLLAMER LENS OF -7.5/2.5/74 TORE/BROKE BEFORE OR DURING LOADING. IT IS UNKNOWN TO WHEN THE TEAR OCCURED. NO LENS WAS IMPLANTED. Manufacturer Narrative: A4-A6:UNK; H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -7.5/2.5/74 | NA | Physician | 10/24/2023 | 2023 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18052451&pc=QCB | 2023826-2023-04930 |  | Injury | STAAR SURGICAL COMPANY | 2023-11-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 -7.50/1.5/107 (SPHERE/CYLINER/AXIS) IMPLANTABLE COLLAMER LENS FLIPPED ON INSERTION. THE LENS WAS EXPLANTED AND THE BACKUP LENS WAS INSERTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: D6A:UNK. D6B:UNK. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -7.50/1.5/107 | NA | Physician | 10/06/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18057431&pc=QCB | 2023826-2023-04942 | 2023-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED 12.1MM VTICMO12.1; -8.0/2.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS HAD ROTATED. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5 - LENS WAS REPOSITIONED ON (B)(6) 2023 BUT PROBLEM WAS NOT RESOLVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). | -8.0/2.0/093 | OD | Physician | 10/29/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18062030&pc=QCB | 2023826-2023-04852 | 2023-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | UNK | Event Description: THE REPORTER INDICATED THAT A 12.6MM IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS REMOVED AND LATER REPLACED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "THE SURGEON AND PATIENT ARE HAPPY." Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM5_12.6; -14.0/2.5/110 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND LATER REPLACED WITH A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. CLAIM#: (B)(4). Manufacturer Narrative: A4-A6:UNK. B3: UNK. D4 (EXPERIATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK. H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM# (B)(4). | -14.0/2.5/110 | OS | Physician | 10/25/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18062032&pc=QCB | 2023826-2023-04851 | 2023-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | UNK | Event Description: THE REPORTER INDICATED THAT A 12.6MM IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). THE LENS WAS REMOVED AND LATER REPLACED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "THE SURGEON AND PATIENT ARE HAPPY." Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM5_12.6; -12.0/2.5/067 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND LATER REPLACED WITH A SHORTER LENGTH LENS THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS PATIENT RELATED FACTOR AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. CLAIM#: (B)(4). Manufacturer Narrative: A4-A6:UNK. B3: UNK. D4 (EXPERIATION DATE); UNK NO SERIAL NUMBER REPORTED. D6A: UNK. H4 (DEVICE MANUFACTURING DATE): NO SERIAL NUMBER REPORTED. CLAIM# (B)(4). | -12.0/2.5/067 | OD | Physician | 10/25/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18062372&pc=QCB | 2023826-2023-04933 | 2023-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.0/3.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK . CLAIM # (B)(4). | -17.0/3.0/090 | OD | Physician | 10/31/2023 | 2023 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18062375&pc=QCB | 2023826-2023-04935 | 2023-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -9.5/1.0/091 | OS | Physician | 10/31/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18062377&pc=QCB | 2023826-2023-04958 | 2023-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/1.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -11.5/1.5/092 | OD | Physician | 11/01/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18062379&pc=QCB | 2023826-2023-04936 | 2023-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.5/092(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ON (B)(6) 2023 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -10.0/1.5/092 | OD | Physician | 10/31/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18062385&pc=QCB | 2023826-2023-04938 | 2023-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/2.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK . H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -11.0/2.0/094 | OS | Physician | 10/31/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18062386&pc=QCB | 2023826-2023-04940 | 2023-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_ 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/4.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -14.0/4.5/086 | OD | Physician | 11/01/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18062392&pc=QCB | 2023826-2023-04941 | 2023-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_ 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.50/2.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -12.50/2.0/086 | OS | Physician | 11/01/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18062393&pc=QCB | 2023826-2023-04959 | 2023-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/3.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL/LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) . CLAIM # (B)(4). | -13.0/3.0/092 | OS | Physician | 11/01/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18062743&pc=QCB | 2023826-2023-04963 | 2023-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF -7.50/1.00/059 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023 AS A REPLACEMENT LENS TO ADDRESS LOW VAULT. PER REPORTER THE PROBLEM DID NOT RESOLVE, " VAULTING STILL LOW AFTER ICL EXCHANGE". LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -7.50/1.00/059 | OD | Physician | 10/09/2023 | 2023 | 47 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18062744&pc=QCB | 2023826-2023-04962 | 2023-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -7.50/1.00/064 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE DID NOT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -7.50/1.00/064 | OD | Physician | 10/09/2023 | 2023 | 47 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18062808&pc=QCB | 2023826-2023-04966 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-11-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -14.50/0.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 AS A REPLACEMENT LENS TO ADDRESS LOW VAULT. PER REPORTER THE PROBLEM DID NOT RESOLVE, ICL HEIGHT IS LOW AFTER THE REPLACEMENT. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -14.50/0.5/093 | OS | Physician | 10/10/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18062809&pc=QCB | 2023826-2023-04965 | 2023-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -14.50/0.5/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO LOW VAULT. ON (B)(6) 2023 A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED BUT IT DID NOT RESOLVE THE PROBLEM, THE VAULT IS STILL LOW AFTER REPLACEMENT. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -14.50/0.5/091 | OS | Physician | 10/10/2023 | 2023 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18066235&pc=QCB | 2023826-2023-04857 |  | Injury | STAAR SURGICAL COMPANY | 2023-11-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -18.0/5.5/030(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 AND ON (B)(6) 2022, THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. THE PROBLEMS WERE NOT RESOLVED. REPORTEDLY, "THE LENS WAS OK FOR MANY MONTHS THEN IT ROTATED," AND "IT WAS STEADY FOR OVER 2 YEARS AND IT ROTATED IT AGAIN." Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B5: LENS WAS IMPLANTED INTO A PATIENT WITH PRE-EXISTING KERATOCONUS. H6: OFF-LABEL: PRE-EXISTING KERATOCONUS. WORK ORDER SEARCH WAS PERFORMED BUT WAS NOT REQUIRED. CLAIM# (B)(4). | -18.0/5.5/030 | OS | Physician | 10/30/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18067653&pc=QCB | 2023826-2023-04873 | 2023-09-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.5/088 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -10.5/+2.5/088 | OS | Physician | 10/09/2023 | 2023 | 46 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18067654&pc=QCB | 2023826-2023-04874 | 2023-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.0/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+1.0/094 | OD | Physician | 10/09/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18067655&pc=QCB | 2023826-2023-04876 | 2023-07-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.0/115 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) CLAIM# (B)(4). | -10.5/+2.0/115 | OS | Physician | 10/09/2023 | 2023 | 17 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18067656&pc=QCB | 2023826-2023-04875 | 2023-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.0/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+1.0/094 | OS | Physician | 10/09/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18067657&pc=QCB | 2023826-2023-04872 | 2023-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+1.0/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LOW VAULT. IN A SEPARATE SURGERY THAT DAY A LONGER LENGTH LENS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM, HOWEVER IT IS UNCLEAR IF THIS WAS PERFORMED WITHIN THE SAME SURGERY. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -12.0/+1.0/095 | OD | Physician | 10/09/2023 | 2023 | 46 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18067935&pc=QCB | 2023826-2023-04763 | 2023-07-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -09.00/+1.0/085 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO THE LENS TORE DURING LOADING, DUE TO THE LENS CATCHING ON THE INJECTOR. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -09.00/+1.0/085 | NA | Physician | 10/13/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18075204&pc=QCB | 2023826-2023-04918 | 2023-09-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.5/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM, UNDER 21 YRS OF AGE AT DATE OF IMPLANT. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -14.5/2.5/091 | OD | Physician | 10/16/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18075205&pc=QCB | 2023826-2023-04923 | 2023-09-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 12.6MM VTICMO12.6. -14.5/1.0/168 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT(LENS TOUCHED THE EYE). NO PATIENT INJURY. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.5/1.0/168 | OD | Physician | 10/18/2023 | 2023 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18075209&pc=QCB | 2023826-2023-04926 | 2023-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.5/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. THE LENS WAS REPOSITIONED BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE: LENS TOO SMALL THAN REQUIRED. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: LENS WAS REPOSITIONED ON (B)(6) 2023 BUT DID NOT RESOLVED THE PROBLEM. CLAIM# (B)(4). | -8.5/1.5/089 | OD | Physician | 10/20/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18075210&pc=QCB | 2023826-2023-04924 | 2023-08-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -9.0/2.0/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.0/2.0/072 | OS | Physician | 10/18/2023 | 2023 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18075224&pc=QCB | 2023826-2023-04925 | 2023-09-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 13.2MM VTICM5_13.2. -10.0/2.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. DUE TO THE LENS TEAR/BRAKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT(LENS TOUCHED THE EYE). NO PATIENT INJURY. A LONGER LENGTH LENS WAS IMPLANTED ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.0/2.0/095 | OD | Physician | 10/19/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18081083&pc=QCB | 2023826-2023-04960 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS OF -9.50/2.5/065 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). A SCHEDULED LENS REMOVAL FOR AN UNKNOWN REASON WAS REPORTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: D6A: UNK. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -9.50/2.5/065 | OS | Other Health Care Professional | 10/09/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18081125&pc=QCB | 2023826-2023-04961 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS OF -10.00/2.5/107 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). A SCHEDULED LENS REMOVAL FOR AN UNKNOWN REASON WAS REPORTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: D6A:UNK. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#:(B)(4). | -10.00/2.5/107 | OD | Other Health Care Professional | 10/09/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18081562&pc=QCB | 2023826-2023-04878 | 2023-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+4.5/122 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -12.0/+4.5/122 | OS | Physician | 10/11/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18081563&pc=QCB | 2023826-2023-04886 | 2023-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.5/115 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/+1.5/115 | OD | Physician | 10/12/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18081564&pc=QCB | 2023826-2023-04885 | 2023-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.5/085 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -8.5/+1.5/085 | OD | Physician | 10/12/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18081566&pc=QCB | 2023826-2023-04877 | 2023-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.5/071 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -9.5/+1.5/071 | OS | Physician | 10/11/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18086566&pc=QCB | 2023826-2023-04883 | 2023-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/110 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXPLANTED DUE TO LOW VAULT. LATER ON (B)(6) 2023 A LONGER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.0/+1.5/110 | OD | Physician | 10/12/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18086567&pc=QCB | 2023826-2023-04887 | 2023-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+2.5/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.5/+2.5/092 | OD | Physician | 10/12/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18086568&pc=QCB | 2023826-2023-04881 | 2023-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.5/130 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: (B)(4). | -9.5/+1.5/130 | OS | Physician | 10/12/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18086569&pc=QCB | 2023826-2023-05010 | 2023-09-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+2.5/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -13.0/+2.5/093 | OD | Physician | 10/13/2023 | 2023 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18086712&pc=QCB | 2023826-2023-05028 | 2023-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8.50/+1.50/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE. ON THIS SAME DATE THE LENS WAS EXPLANTED WITH NO PATIENT INJURY. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -8.50/+1.50/091 | OD | Physician | 10/17/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18089677&pc=QCB | 2023826-2023-04967 | 2023-09-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON 20-SEP-2023 A 13.2MM, VTICM5_13.2, -8.50/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS GOT STUCK IN THE PLUNGER AND WAS NOTICED TO BE TORN/BROKEN WHEN DELIVERED IN THE EYE WHILE MANIPULATING THE HAPTICS. ON (B)(6) 2023 THE LENS WAS EXCHANGED AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -8.50/2.0/089 | NA | Physician | 10/12/2023 | 2023 | 48 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18097758&pc=QCB | 2023826-2023-04979 | 2023-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -9.5/2.5/091 | OD | Physician | 11/02/2023 | 2023 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18097761&pc=QCB | 2023826-2023-04981 | 2023-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/1.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -9.5/1.0/093 | OD | Physician | 11/01/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18097762&pc=QCB | 2023826-2023-04985 | 2023-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/2.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "THE REPLACEMENT LENS WAS VERTICALLY IMPLANTED." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -11.0/2.5/095 | OD | Physician | 11/02/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18097764&pc=QCB | 2023826-2023-04984 | 2023-05-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/2.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -12.0/2.5/089 | OD | Physician | 11/01/2023 | 2023 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18097765&pc=QCB | 2023826-2023-04988 | 2023-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/3.0/111(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -10.5/3.0/111 | OS | Physician | 11/02/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18097766&pc=QCB | 2023826-2023-04987 | 2023-10-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/3.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL/LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "THE SECOND TICL IS IN A VERTICAL POSITION." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -14.0/3.0/089 | OS | Physician | 11/02/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18097767&pc=QCB | 2023826-2023-04986 | 2023-10-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0/2.5/096 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME MODEL/LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "THE SECOND TICL IS IN A VERTICAL POSITION." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -15.0/2.5/096 | OD | Physician | 11/02/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18097771&pc=QCB | 2023826-2023-04991 | 2023-06-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/2.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -11.0/2.0/093 | OD | Physician | 11/02/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18097773&pc=QCB | 2023826-2023-05030 | 2023-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/3.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -7.5/3.0/091 | OS | Physician | 11/06/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18097775&pc=QCB | 2023826-2023-05035 | 2023-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/2.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(6). | -9.5/2.5/094 | OD | Physician | 11/06/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18097777&pc=QCB | 2023826-2023-05033 | 2023-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/1.0/081 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -9.0/1.0/081 | OS | Physician | 11/05/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18097778&pc=QCB | 2023826-2023-05034 | 2023-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/1.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -13.5/1.5/086 | OS | Physician | 11/06/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18097780&pc=QCB | 2023826-2023-05037 | 2023-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/2.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -12.5/2.5/094 | OS | Physician | 11/07/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18105875&pc=QCB | 2023826-2023-04793 | 2023-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -10.00/1.5/080 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS ROTATION WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS MODEL OF THE SAME SIZE. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: H10: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH. NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -10.00/1.5/080 | OD | Physician | 10/11/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18106554&pc=QCB | 2023826-2023-05040 | 2023-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/2.0/178 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -11.0/2.0/178 | OD | Physician | 11/07/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18106555&pc=QCB | 2023826-2023-05038 | 2023-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/2.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -11.5/2.0/094 | OD | Physician | 11/07/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18106556&pc=QCB | 2023826-2023-05039 | 2023-08-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/2.0/071 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -10.5/2.0/071 | OS | Physician | 11/07/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18106561&pc=QCB | 2023826-2023-05041 | 2023-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/2.0/178 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -8.0/2.0/178 | OD | Physician | 11/07/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18106578&pc=QCB | 2023826-2023-05042 | 2023-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/2.5/179 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -11.0/2.5/179 | OS | Physician | 11/07/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18107024&pc=QCB | 2023826-2023-05062 | 2021-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -12.5/4.0/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -12.5/4.0/088 | OS | Physician | 11/05/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18107626&pc=QCB | 2023826-2023-05012 | 2023-05-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/044 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+1.0/044 | OS | Physician | 10/15/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18107629&pc=QCB | 2023826-2023-05017 | 2023-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -18.0/+1.5/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -18.0/+1.5/093 | OD | Physician | 10/15/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18107631&pc=QCB | 2023826-2023-05013 | 2023-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+2.0/093 | OD | Physician | 10/15/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18107632&pc=QCB | 2023826-2023-05016 | 2023-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -18.0/+1.5/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -18.0/+1.5/092 | OS | Physician | 10/15/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18113113&pc=QCB | 2023826-2023-04794 | 2023-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -11.00/2.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED AND THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS MODEL OF THE SAME SIZE. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: H10: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL OBSERVATION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -11.00/2.5/092 | OS | Physician | 10/11/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18115881&pc=QCB | 2023826-2023-04947 | 2023-08-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 OF -9.00/2.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS FOUND TO BE TORN/BROKE DURING INJECTION DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THE LENS WAS NOT IMPLANTED. NO PATIENT CONTACT; THE LENS DID NOT TOUCH THE EYE. NO PATIENT INJURY. A SPHERICAL LENS WAS IMPLANTED AS A REPLACEMENT LENS OF THE SAME SIZE AND LRI WAS PERFORMED FOR ASTIGMATISM. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -9.00/2.0/093 | OD | Physician | 10/11/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18116357&pc=QCB | 2023826-2023-05026 | 2023-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.5/071 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS TWO SIZES LONGER IN LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -9.5/+1.5/071 | OD | Physician | 10/16/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18116361&pc=QCB | 2023826-2023-05020 | 2023-09-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.0/087 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.5/+1.0/087 | OS | Physician | 10/16/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18116369&pc=QCB | 2023826-2023-05024 | 2023-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+3.0/070 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -10.0/+3.0/070 | OD | Physician | 10/16/2023 | 2023 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18116371&pc=QCB | 2023826-2023-05027 | 2023-09-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -10.0/+1.5/089 | OS | Physician | 10/16/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18116442&pc=QCB | 2023826-2023-05025 | 2023-09-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+2.5/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6)-2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -7.5/+2.5/094 | OS | Physician | 10/16/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18117391&pc=QCB | 2023826-2023-04766 | 2023-08-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -16.50/+1.0/166 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -16.50/+1.0/166 | OD | Physician | 10/17/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18117392&pc=QCB | 2023826-2023-04771 | 2023-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. CLAIM # (B)(4). | -10.00/+1.5/094 | OD | Physician | 10/17/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18120273&pc=QCB | 2023826-2023-05076 | 2023-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+1.0/086 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.5/+1.0/086 | OS | Physician | 10/17/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18120276&pc=QCB | 2023826-2023-05071 | 2003-06-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+2.0/087 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -8.5/+2.0/087 | OD | Physician | 10/17/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18120279&pc=QCB | 2023826-2023-05073 | 2023-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.5/075 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS OF THE SAME LENGTH AND IMPLANTED VERTICALLY DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -11.0/+1.5/075 | OS | Physician | 10/17/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18120284&pc=QCB | 2023826-2023-05069 | 2023-06-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+2.5/111 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -9.0/+2.5/111 | OS | Physician | 10/17/2023 | 2023 | 17 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18120292&pc=QCB | 2023826-2023-05072 | 2023-07-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+2.0/087 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS OF THE SAME LENGTH AND IMPLANTED VERTICALLY DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.5/+2.0/087 | OD | Physician | 10/17/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18120829&pc=QCB | 2023826-2023-05043 | 2023-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.5/2.0/100 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -3.5/2.0/100 | OS | Physician | 11/07/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18121319&pc=QCB | 2023826-2023-05068 | 2023-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+1.5/085 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) . CLAIM# (B)(4). | -6.5/+1.5/085 | OD | Physician | 10/17/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18122722&pc=QCB | 2023826-2023-05095 | 2023-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+1.0/114 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. IN A SEPARATE SURGERY THAT DAY A LENS OF THE SAME LENGTH WAS IMPLANTED VERTICALLY AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -14.0/+1.0/114 | OD | Physician | 10/19/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18122724&pc=QCB | 2023826-2023-05077 | 2023-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+3.0/090 INTO THE PATIENT'S RIGHT EYE (OD) (B)(6)2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -7.0/+3.0/090 | OD | Physician | 10/18/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18122727&pc=QCB | 2023826-2023-05079 | 2023-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/+1.0/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -6.0/+1.0/092 | OD | Physician | 10/18/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18122752&pc=QCB | 2023826-2023-04976 | 2022-05-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -15.0/3.5/94 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -15.0/3.5/94 | OD | Physician | 11/02/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18122760&pc=QCB | 2023826-2023-05093 | 2023-10-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS WITH DIOPTER -8.5/+1.0/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.5/+1.0/095 | OD | Physician | 10/19/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18129698&pc=QCB | 2023826-2023-04995 | 2023-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -11.5/2.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2023. BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -11.5/2.0/092 | OS | Physician | 10/23/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18129701&pc=QCB | 2023826-2023-05002 | 2023-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -9.5/3.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS REPLACED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME. CLAIM# (B)(4). | -9.5/3.5/080 | OS | Physician | 10/25/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18129707&pc=QCB | 2023826-2023-04999 | 2023-09-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5 12.6, -10.0/1.0/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON 14/AUG/2023. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. THE LENS WAS REPOSITIONED ON 23/OCT/2023. THIS DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.0/1.0/111 | OD | Physician | 10/24/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18129715&pc=QCB | 2023826-2023-05006 | 2023-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5 13.2, -11.0/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS OBSERVED WITH ROTATION. LENS WAS REPOSITIONED. THIS DID NOT RESOLVE THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5 13.2, -11.0/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO PATIENTS LEFT EYE (OD) ON (B)(6) 2023. LENS WAS REPOSITIONED ON (B)(6) 2023 DUE TO ROTATION. THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS. THIS RESOVLED THE PROBLEM. CAUSE OF EVENT IS PATIENT RELATED FACTOR. REPORTEDLY, PATIENT IS HAPPY WIHT VISION. Manufacturer Narrative: A1 - UNK. A2 - UNK. A3 - UNK. A4 - UNK. A5 - UNK. A6 - UNK. B3 - UNK. D6A - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION ADDED TO FORM. CLAIM# (B)(4). | -11.0/2.0/091 | OS | Physician | 10/27/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136062&pc=QCB | 2023826-2023-05081 | 2023-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -17.5/3.5/104 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -17.5/3.5/104 | OD | Physician | 11/05/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136063&pc=QCB | 2023826-2023-05083 | 2023-10-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -16.0/2.5/90 (SPHERE/CYLINDER/AXIS) TORE DURING LOADING AFTER OPENING THE VIAL. CAUSE OF THE EVENT WAS USER ERROR. REPORTEDLY, "DOCTOR DAMAGED THE LENS WHILE LOADING THE LENS." THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4-A6: UNK. H6: USER ERROR. CLAIM# (B)(4). | -16.0/2.5/90 | NA | Physician | 11/07/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136580&pc=QCB | 2023826-2023-05046 | 2023-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00/+4.5/080 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LENS WAS IMPLANTED UPSIDE DOWN. THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK, A5: UNK, A6: UNK. H6 WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/+4.5/080 | OS | Physician | 10/23/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136581&pc=QCB | 2023826-2023-05050 | 2023-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.50/+2.5/105 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A DIFFERENT MODEL/POWER BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK . H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -16.50/+2.5/105 | OD | Physician | 10/25/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18136622&pc=QCB | 2023826-2023-05113 | 2023-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.5/+3.00/104 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 2.8MM). CLAIM# (B)(4). | -6.5/+3.00/104 | OD | Physician | 11/06/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18143049&pc=QCB | 2023826-2023-05117 | 2023-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+4.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) AND (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -7.5/+4.0/093 | OD | Physician | 10/20/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18143050&pc=QCB | 2023826-2023-05123 | 2023-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+2.0/034 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -11.5/+2.0/034 | OD | Physician | 10/20/2023 | 2023 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18143051&pc=QCB | 2023826-2023-05116 | 2023-08-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+2.5/111 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. IN A SEPARATE SURGERY THAT DAY A SHORTER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -9.5/+2.5/111 | OD | Physician | 10/20/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18143052&pc=QCB | 2023826-2023-05143 | 2023-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+1.0/089 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) . CLAIM# (B)(4). | -13.0/+1.0/089 | OD | Physician | 10/21/2023 | 2023 | 16 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18143055&pc=QCB | 2023826-2023-05121 | 2023-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTCMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.5/+0.5/111 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS TWO SIZES SHORTER IN LENGTH DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -4.5/+0.5/111 | OD | Physician | 10/20/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18143063&pc=QCB | 2023826-2023-05118 | 2023-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+4.0/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) AND (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -7.5/+4.0/092 | OS | Physician | 10/20/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18145673&pc=QCB | 2023826-2023-05067 | 2023-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/3.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -11.5/3.5/093 | OD | Physician | 11/09/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18145675&pc=QCB | 2023826-2023-05065 | 2023-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/1.0/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -11.5/1.0/082 | OS | Physician | 11/08/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18145677&pc=QCB | 2023826-2023-05135 | 2023-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_ 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.50/3.0/001 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "REPLACED THE LENS, ADJUSTED THE VAULT, AND THE RESULT IS GOOD. THERE IS NO CHANGE TO THE LENS SIZE BECAUSE THE VAULT WAS ADJUSTED BY ADJUSTING THE FIXED POSITION OF THE LENS." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6: TYPE OF INVESTIGATION: 4110 LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -5.50/3.0/001 | OS | Physician | 11/10/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18145678&pc=QCB | 2023826-2023-05134 | 2023-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_ 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/3.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "REPLACED THE LENS, ADJUSTED THE VAULT, AND THE RESULT IS GOOD." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -6.0/3.0/088 | OD | Physician | 11/10/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18145681&pc=QCB | 2023826-2023-05132 | 2023-10-31 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2IMPLANTABLE COLLAMER LENS OF DIOPTER -14.50/5.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) CLAIM (B)(4). | -14.50/5.0/095 | OD | Physician | 11/10/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18146004&pc=QCB | 2023826-2023-05142 | 2023-10-31 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -2.00/+4.0/014 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS THE PATIENT HAS EXCESSIVE VAULTING. LENS REMAINS IMPLANTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -2.00/+4.0/014 | OS | Physician | 11/01/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18151101&pc=QCB | 2023826-2023-05086 | 2023-10-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM513.7 IMPLANTABLE COLLAMER LENS OF -12.0/2.0/94 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS A REPLACEMENT INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023, DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE IS "OK." CAUSE OF THE EVENT IS UNKNOWN. IT WAS REPORTED THAT THE RESULTING VAULT IS STILL VERY LOW. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -12.0/2.0/94 | OD | Physician | 11/08/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18151472&pc=QCB | 2023826-2023-05085 | 2023-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -12.0/2.0/100 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS NOT RESOLVED. STATUS OF THE EYE IS "OK." CAUSE OF THE EVENT IS UNKNOWN. IT WAS REPORTED THAT THE RESULTING VAULT IS STILL VERY LOW. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -12.0/2.0/100 | OD | Physician | 11/08/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18156104&pc=QCB | 2023826-2023-04950 | 2023-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -11.50/5.5/003 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LOW VAULT WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -11.50/5.5/003 | OS | Physician | 10/17/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18156343&pc=QCB | 2023826-2023-04951 | 2023-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -10.50/4.5/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LOW VAULT WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -10.50/4.5/090 | OD | Physician | 10/17/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18156501&pc=QCB | 2023826-2023-05136 | 2023-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_ 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.50/3.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "REPLACED THE LENS, ADJUSTED THE VAULT, AND THE RESULT IS GOOD. " THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -5.50/3.0/091 | OS | Physician | 11/10/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18156504&pc=QCB | 2023826-2023-05141 | 2023-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/1.5/069 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -9.5/1.5/069 | OD | Physician | 11/14/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18156505&pc=QCB | 2023826-2023-05140 | 2023-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/1.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -9.5/1.5/087 | OD | Physician | 11/13/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18156509&pc=QCB | 2023826-2023-05138 | 2023-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/1.0/071 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "BOTH IN VERTICAL POSITION." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK CLAIM (B)(4) | -9.0/1.0/071 | OD | Physician | 11/12/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18156839&pc=QCB | 2023826-2023-05145 | 2023-10-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -12.50/1.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS GOT STUCK IN THE CARTRIDGE OR INJECTION SYSTEM, UNABLE TO IMPLANT. THERE WAS PATIENT CONTACT WITH NO PATIENT INJURY. A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON NOTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+1.0/90 (SPHERE/CYLINDER/AXIS) WAS STUCK IN THE CARTRIDGE/ INJECTOR SYSTEM ON 19/OCT/2023. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. THE LENS WAS NOT IMPLANTED. A LENS OF THE SAME PARAMETERS WAS IMPLANTED INTO THE RIGHT EYE (OD) AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "HAPPY WITH VISION AND COMFORT TODAY. ODD CELLS IN ANTERIOR CHAMBER. WOUNDS SEALED". THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS THE DEVICE. CAUSE DETAILS PROVIDED: "PRIMARY LENS CAUGHT IN INJECTOR, UNABLE TO IMPLANT". B3- "10/19/2023" SHOULD BE ADDED. CLAIM# (B)(4). | -12.50/1.0/090 | OD | Physician | 10/19/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18166325&pc=QCB | 2023826-2023-04945 |  | Injury | STAAR SURGICAL COMPANY | 2023-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -14.00/1.0/014 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS EYE ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REPLACED FOR A DIFFERENT SIZE LENS FOR AN UNKNOWN REASON. NO FURTHER INFORMATION HAS BEEN PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -14.00/1.0/014 | NA | Physician | 10/19/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18166419&pc=QCB | 2023826-2023-01137 | 2023-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -9.0/3.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE PROBLEM DID NOT RESOLVE. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS THAT WAS IMPLANTED VERTICALLY AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -9.0/3.0/093 | OD | Physician | 02/28/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18170880&pc=QCB | 2023826-2023-05173 | 2023-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.0/+2.0/105 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. AN EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS OF THE SAME LENGTH AND IT WAS IMPLANTED VERTICALLY. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -4.0/+2.0/105 | OS | Physician | 10/23/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18170882&pc=QCB | 2023826-2023-05171 | 2023-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.0/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+2.0/094 | OS | Physician | 10/22/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18170883&pc=QCB | 2023826-2023-05172 | 2023-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/+1.0/087 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. AN EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS OF THE SAME LENGTH AND IT WAS IMPLANTED VERTICALLY. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -6.0/+1.0/087 | OD | Physician | 10/23/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18172184&pc=QCB | 2023826-2023-05179 | 2023-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+1.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -7.5/+1.0/093 | OD | Physician | 10/23/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18172186&pc=QCB | 2023826-2023-05178 | 2023-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+3.5/071 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -12.0/+3.5/071 | OS | Physician | 10/23/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18172188&pc=QCB | 2023826-2023-05177 | 2023-05-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.5/072 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.5/+1.5/072 | OS | Physician | 10/23/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18172204&pc=QCB | 2023826-2023-05176 | 2022-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+3.5/179 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.0/+3.5/179 | OS | Physician | 10/23/2023 | 2023 | 41 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18172209&pc=QCB | 2023826-2023-05175 | 2023-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.5/035 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#(B)(4). | -11.5/+1.5/035 | OD | Physician | 10/23/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18172335&pc=QCB | 2023826-2023-05152 | 2023-03-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -8.5/5.0/075 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE PROBLEM DID NOT RESOLVE. LENS REMAINS IMPLANTED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS DUE TO PATIENT-RELATED FACTORS, HOWEVER IT WAS ALSO REPORTED THAT IT IS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENS. THIS DID NOT RESOLVE THE PROBLEM. CLAIM# (B)(4). | -8.5/5.0/075 | OS | Physician | 10/24/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18174968&pc=QCB | 2023826-2023-05107 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.5/1.0/037 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). LENS STUCK IN CARTRIDGE OR JINJECTION SYSTEM. BACK UP LENS IMPLANTED. PROBLEM RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A3 - UNK. A4 - UNK. A5 - UNK. A6 - UNK. B3 - UNK. D6A - UNK H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -8.5/1.0/037 | OS | Physician | 11/01/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18174982&pc=QCB | 2023826-2023-05099 | 2023-10-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -9.5/2.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. LOW VAUT WAS OBSERVED. THE LENS REMAINS IMPLANTED. REPORTEDLY: PATIENT IS SUPER HAPPY. PATIENT WILL BE MONITORED AND NO EXCHANGE IS PLANNED. Manufacturer Narrative: D6A - (B)(6)2023. A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.5/2.0/094 | OS | Physician | 10/30/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18174987&pc=QCB | 2023826-2023-05100 | 2023-10-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5 13.2, -09.5/5.5/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE BEFORE/DURING LOADING. NO PATIENT CONTACT. BACK UP LENS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/5.5/112 (SPHERE/CYLINDER/AXIS) TORE/BROKE BEFORE/DURING LOADING ON (B)(6) 2023. THERE WAS NO PATIENT CONTACT. A REPLACEMENT LENS OF THE SAME MODEL AND LENGTH WAS IMPLANTED THE SAME DAY. THIS RESOLVED THE PROBLEM. NO INJURY REPORTED. CAUSE REPORTED AS UNKNOWN. CLAIM #(B)(4) Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -09.5/5.5/112 | NA | Physician | 12/11/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18174989&pc=QCB | 2023826-2023-05110 | 2023-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.5/2.5/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS FROM THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. A SHORTER LENGTH LENS WAS IMPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE: LENS TOO BIG. Manufacturer Narrative: CLAIM# (B)(4). | -12.5/2.5/072 | OS | Physician | 11/03/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18175006&pc=QCB | 2023826-2023-05109 | 2023-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -13.0/3.0/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS FROM THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. A SHORTER LENGTH LENS WAS IMPLANTED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE: LENS TOO BIG. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. CLAIM # (B)(4). | -13.0/3.0/102 | OD | Physician | 11/03/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18175017&pc=QCB | 2023826-2023-05108 | 2023-09-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -8.0/1.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING INJECTION/DELIVERY INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -8.0/1.5/083 | OS | Physician | 11/02/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18175085&pc=QCB | 2023826-2023-05102 | 2023-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -11.0/0.5/167 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -11.0/0.5/167 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION ADDED TO FORM. CLAIM#: (B)(4). | -11.0/0.5/167 | OD | Physician | 10/31/2023 | 2023 | 45 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18175114&pc=QCB | 2023826-2023-05098 | 2023-10-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -8.5/1.0/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LOW VAUT WAS OBSERVED. THE LENS REMAINS IMPLANTED. PATIENT WILL BE MONITORED AND NO EXCHANGE IS PLANNED. Manufacturer Narrative: D6A - (B)(6) 2023. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -8.5/1.0/102 | OD | Physician | 10/30/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18175477&pc=QCB | 2023826-2023-05190 | 2023-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+3.0/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.0/+3.0/090 | OD | Physician | 10/24/2023 | 2023 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18175485&pc=QCB | 2023826-2023-05193 | 2023-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0/+1.5/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT REPORT A CAUSE FOR THIS EVENT. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -15.0/+1.5/095 | OD | Physician | 10/24/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18175508&pc=QCB | 2023826-2023-05196 | 2023-06-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+2.5/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT REPORT A CAUSE FOR THIS EVENT. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -14.0/+2.5/090 | OS | Physician | 10/24/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18175510&pc=QCB | 2023826-2023-05194 | 2023-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0/+1.5/087 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT REPORT A CAUSE FOR THIS EVENT. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -15.0/+1.5/087 | OS | Physician | 10/24/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18175530&pc=QCB | 2023826-2023-05195 | 2023-06-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.50/+2.0/085 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT REPORT A CAUSE FOR THIS EVENT. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -12.50/+2.0/085 | OD | Physician | 10/24/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18179067&pc=QCB | 2023826-2023-05199 | 2023-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.5/+1.0/015 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -3.5/+1.0/015 | OD | Physician | 10/25/2023 | 2023 | 57 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18179069&pc=QCB | 2023826-2023-05202 | 2023-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.5/+3.0/110 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -16.5/+3.0/110 | OD | Physician | 10/25/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18179070&pc=QCB | 2023826-2023-05204 | 2023-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.0/074 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4) | -8.5/+1.0/074 | OD | Physician | 10/25/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18179071&pc=QCB | 2023826-2023-05203 | 2023-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.0/073 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON 25-AUG-2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -8.5/+1.0/073 | OS | Physician | 10/25/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18179627&pc=QCB | 2023826-2023-05184 | 2023-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2 MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -13.5/1.5/085 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- ONE SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -13.5/1.5/085 | OS | Physician | 10/23/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18179726&pc=QCB | 2023826-2023-05221 | 2023-10-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -15.00/1.5/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE EYE ON (B)(6) 2023. THE LENS TORE/BROKE DURING LOADING WHILE INJECTING INTO THE EYE. THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY ON (B)(6) 2023. A REPLACEMENT LENS WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -15.00/1.5/091 | NA | Physician | 10/25/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18180089&pc=QCB | 2023826-2023-05215 | 2023-10-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/+1.0/071 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -6.0/+1.0/071 | OD | Physician | 10/26/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18180090&pc=QCB | 2023826-2023-05208 | 2023-09-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+2.5/088 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THAT SAME DAY A SHORTER LENGTH LENS WAS IMPLANTED HOWEVER IT IS UNCLEAR IF THIS WAS PERFORMED INTRAOPERATIVELY. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.5/+2.5/088 | OD | Physician | 10/25/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18180591&pc=QCB | 2023826-2023-05244 | 2023-05-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS WITH DIOPTER -6.0/+1.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON 13-MAY-2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM (B)(4). | -6.0/+1.0/091 | OD | Physician | 10/26/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18180593&pc=QCB | 2023826-2023-05238 | 2023-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+2.50/175 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -8.0/+2.50/175 | OD | Physician | 10/26/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18180596&pc=QCB | 2023826-2023-05242 | 2023-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+2.5/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.0/+2.5/094 | OS | Physician | 10/26/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18180597&pc=QCB | 2023826-2023-05243 | 2023-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.0/047 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.5/+1.0/047 | OD | Physician | 10/26/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18188085&pc=QCB | 2023826-2023-05126 | 2023-08-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.00/+3.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A DIFFERENT MODEL/DIOPTER BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -11.00/+3.0/091 | OD | Physician | 10/29/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18188445&pc=QCB | 2023826-2023-05280 | 2023-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.0/2.5/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. STATUS OF THE EYE IS "OK." Manufacturer Narrative: CLAIM# (B)(4). | -12.0/2.5/087 | OD | Physician | 10/31/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18188448&pc=QCB | 2023826-2023-05284 | 2023-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VICMO12.6 IMPLANTABLE COLLAMER LENS OF -8.0/1.0/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. STATUS OF THE EYE IS "OK." Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL - ACD<3.0. CLAIM# (B)(4). | -8.0/1.0/095 | OD | Physician | 10/31/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18188486&pc=QCB | 2023826-2023-05287 | 2023-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/2.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -9.5/2.0/091 | OS | Physician | 11/01/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18188682&pc=QCB | 2023826-2023-05256 |  | Injury | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -10.50/1.5/137 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. IT WAS REPORTED THAT THE LENS ROTATED IN THE EYE AND WAS THE WRONG SIZE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.50/1.5/137 | NA | Physician | 11/01/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18188685&pc=QCB | 2023826-2023-05257 | 2023-11-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -8.00/2.5/105 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON NOTICED A DEFECT IN THE LENS AND REMOVED AND REPLACED IT WITH A BACK UP LENS. THERE WAS NO PATIENT INJURY. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF A -8.0/2.5/105 (SPHERE/CYLINDER/AXIS) UPSIDE DOWN INTO THE PATIENT'S RIGHT EYE (OS) ON (B)(6) 2023. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITHIN THE INITIAL SURGERY FOR A SAME MODEL AND LENGTH LENS. THIS RESOLVED THE PROBLEM. NO PATIENT INJURY. CAUSE REPORTED AS "PATIENT RELATED FACTOR" AND "USER ERROR" WITH 'NO' THE DEVICE DID NOT FAIL TO PERFORM. REPORTEDLY, "DURING INSERTION OF A POORLY COOPERATING PT ICL WAS INJECTED UPSIDE DOWN AND NEED TO BE REMOVED WITH USE OF BACK-UP ICL". D6A - (B)(6) 2023. D6B - (B)(6) 2023. CLAIM # (B)(4). | -8.00/2.5/105 | OD | Physician | 12/11/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18188686&pc=QCB | 2023826-2023-05290 | 2023-11-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -15.5/+1.5/072 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS HAD TORE DURING IMPLANTATION/DELIVERY INTO THE EYE. INTRAOPERATIVELY THE LEN WAS REMOVED WITH NO PATIENT INJURY. ON (B)(6) 2023 A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE - THERE IS NOT ENOUGH SAFETY AND EFFECTIVENESS DATA TO SUPPORT IMPLANTATION WITH PATIENTS WITH PREEXISTING PROGRESSIVE MYOPIA. CLAIM#: (B)(4). | -15.5/+1.5/072 | OS | Physician | 11/20/2023 | 2023 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18196744&pc=QCB | 2023826-2023-05262 | 2023-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+3.0/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULT. IN A SEPARATE SURGERY THAT SAME DAY A SHORTER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -8.0/+3.0/092 | OS | Physician | 10/27/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18196745&pc=QCB | 2023826-2023-05260 | 2023-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.0/082 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.0/+2.0/082 | OS | Physician | 10/27/2023 | 2023 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18196746&pc=QCB | 2023826-2023-05261 | 2023-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+2.0/168 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED FOR A LENS OF THE SAME LENGTH AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) AND (ACD < 3.0MM). CLAIM# (B)(4). | -12.0/+2.0/168 | OS | Physician | 10/27/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18196844&pc=QCB | 2023826-2023-05129 | 2023-10-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+3.0/080 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6)2023 AND THE LENS WAS STUCK IN CARTRIDGE OR INJECTION SYSTEM. THERE WAS PATIENT CONTACT BUT NO INJURY. THE LENS WAS DAMAGED. THE LENS WAS REPLACED WITH A SHORTER LENS. Manufacturer Narrative: A4: UNK A5: UNK A6: UNK H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4). | -11.50/+3.0/080 | OD | Physician | 11/02/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18196875&pc=QCB | 2023826-2023-05156 |  | Injury | STAAR SURGICAL COMPANY | 2023-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 -12.50/2.0/112 (SPHERE/CYLINER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE. A LENS EXCHANGE DUE TO A SIZING ISSUE WAS REPORTER. LENS WAS EXPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -12.50/2.0/112 | NA | Other Health Care Professional | 10/26/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18196919&pc=QCB | 2023826-2023-05266 | 2023-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+1.0/109 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.0/+1.0/109 | OD | Physician | 10/29/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18196920&pc=QCB | 2023826-2023-05270 | 2023-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.5/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.5/+1.5/091 | OS | Physician | 10/29/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18196921&pc=QCB | 2023826-2023-05272 | 2023-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.0/008 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -10.0/+1.0/008 | OD | Physician | 10/29/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18196924&pc=QCB | 2023826-2023-05265 | 2023-08-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+0.5/038 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#(B)(4). | -10.5/+0.5/038 | OS | Physician | 10/29/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18196925&pc=QCB | 2023826-2023-05263 | 2023-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+2.0/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) AND (ACD < 3.0MM). CLAIM#(B)(4). | -9.5/+2.0/089 | OS | Physician | 10/29/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18196926&pc=QCB | 2023826-2023-05267 | 2023-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+1.0/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.5/+1.0/094 | OS | Physician | 10/29/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18196928&pc=QCB | 2023826-2023-05271 | 2023-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+3.5/108 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -13.5/+3.5/108 | OD | Physician | 10/29/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18196942&pc=QCB | 2023826-2023-05306 | 2023-09-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6; -11.5/+2.0/101 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS HAD TORE DURING LOADING. THIS OCCURRED ON 29-SEP-2023. ON THIS SAME DATE A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/+2.0/101 | NA | Physician | 11/14/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18196965&pc=QCB | 2023826-2023-05304 |  | Injury | STAAR SURGICAL COMPANY | 2023-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -8.50/1.5/106 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.50/1.5/106 | OD | Physician | 10/27/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18196980&pc=QCB | 2023826-2023-05305 | 2023-10-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -18.00/1.0/154 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT WAS REPORTER, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -18.00/1.0/154 | OD | Physician | 10/27/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18201847&pc=QCB | 2023826-2023-05274 | 2023-09-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.0/071 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -10.0/+2.0/071 | OS | Physician | 10/31/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18201850&pc=QCB | 2023826-2023-05279 | 2023-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.0/179 INTO THE PATIENT'S LEFT EYE (OS) ON 24-SEP-2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS TWO SIZES LONGER IN LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -10.0/+1.0/179 | OS | Physician | 10/31/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18210918&pc=QCB | 2023826-2023-05104 | 2023-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -17.50/1.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED VERTICALLY FOR A SAME LENGTH LENS, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -17.50/1.5/094 | OD | Physician | 11/06/2023 | 2023 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18210925&pc=QCB | 2023826-2023-05245 | 2023-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.0/3.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXPLANTED, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT DID NOT CONSIDER REPLACEMENT OR ALIGNMENT SURGERY. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.0/3.0/087 | OS | Physician | 11/06/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18212229&pc=QCB | 2023826-2023-05090 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-11-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF A -16.50/2.5/057 (SPHERE/CYLINDER/AXIS) BECAME STUCK IN CARTRIDGE OR INJECTION SYSTEM. THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | -16.50/2.5/057 | OS | Physician | 11/10/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18213259&pc=QCB | 2023826-2023-05331 | 2023-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.0/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -10.5/+2.0/092 | OD | Physician | 11/02/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18213261&pc=QCB | 2023826-2023-05333 | 2023-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+2.5/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -8.5/+2.5/091 | OS | Physician | 11/02/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18213263&pc=QCB | 2023826-2023-05329 | 2023-10-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+3.5/072 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -9.0/+3.5/072 | OD | Physician | 11/02/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18213265&pc=QCB | 2023826-2023-05332 | 2023-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+2.0/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. AN EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS OF THE SAME LENGTH AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -12.5/+2.0/093 | OS | Physician | 11/02/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18213268&pc=QCB | 2023826-2023-05325 | 2023-09-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.0/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -8.5/+1.0/095 | OD | Physician | 11/02/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18213270&pc=QCB | 2023826-2023-05328 | 2023-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.5/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -9.0/+1.5/094 | OS | Physician | 11/02/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18220800&pc=QCB | 2023826-2023-05163 | 2023-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.0/2.5/073 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, AN EXCHANGE WAS MADE FOR A LARGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL: AGE<21. (B)(4). | -12.0/2.5/073 | OS | Physician | 11/12/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18220821&pc=QCB | 2023826-2023-05164 | 2023-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -8.5/1.0/070 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6)2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS BOTH UNKNOWN AND A PATIENT RELATED FACTOR. STATUS OF THE EYE IS REPORTED AS, "OK." Manufacturer Narrative: H6: PATIENT RELATED FACTOR: CLAIM# (B)(4). | -8.5/1.0/070 | OS | Physician | 04/24/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18222774&pc=QCB | 2023826-2023-05374 | 2023-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5/1.5/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2023 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE PROBLEM DID NOT RESOLVE. ON (B)(6) 2023 THE LENS WAS EXPLANTED AND LATER REPLACED WITH A LONGER LENGTH LENS AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.5/1.5/095 | OS | Physician | 10/30/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18223148&pc=QCB | 2023826-2023-05379 | 2022-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -6.5/+2.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LENS ROTATION. ON (B)(6) 2022 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT DIFFERENT AXIS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.5/+2.0/089 | OS | Physician | 11/24/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18232718&pc=QCB | 2023826-2023-05388 | 2023-10-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT WHILE ATTEMPTING TO IMPLANT A 13.2MM VTICM5_13.2; -9.00/+3.0/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023; THE LENS HAD TORE DURING LOADING. ON THIS SAME DATE A REPLACEMENT LENS OF THE SAME LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.00/+3.0/095 | OD | Physician | 11/14/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18232719&pc=QCB | 2023826-2023-05381 | 2023-11-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6; -13/+2.5/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. THE LENS WAS NOT IMPLANTED. IT WAS ADDITIONALLY NOTED "LENS TORN DURING INJECTION/TEAR ALONG THE SMALL AXIS AT PROXIMAL HAPTIC LEVEL. A NEW LENS HAS BEEN ORDERED AND SURGERY IS PLANNED AT (B)(6 )2023. Manufacturer Narrative: ADDITONAL DATA: B5- THE REPORTER INDICATED THAT THE SURGEON NOTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+2.5/102 (SPHERE/CYLINDER/AXIS) TORE/BROKE DURING INJECTION/DELIVERY ON (B)(6) 2023. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. THE LENS WAS NOT IMPLANTED. ON (B)(6) 2023 A LENS OF THE SAME PARAMETERS WAS IMPLANTED INTO THE LEFT EYE (OS) AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "NO PAIN, EXCELLENT VISION". THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. CAUSE DETAILS PROVIDED: "LENS TORN DURING INJECTION / TEAR ALONG THE SMALL AXIS AT PROXIMAL HAPTIC LEVEL". (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4) | -13/+2.5/102 | OS | Physician | 11/10/2023 | 2023 | 1993 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18233719&pc=QCB | 2023826-2023-05365 | 2023-10-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.0/0.5/089 (SPHERE/CYLINDER/AXIS) LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED, AND REPLACED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. NO PATIENT INJURY WAS REPORTED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | -10.0/0.5/089 | OS | Physician | 11/17/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18234527&pc=QCB | 2023826-2023-05291 | 2023-10-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.00/+2.5/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON 26-OCT-2023. THERE WAS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -08.00/+2.5/091 | OD | Physician | 11/06/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18234536&pc=QCB | 2023826-2023-05302 | 2022-12-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -05.50/+3.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -05.50/+3.5/090 | OD | Physician | 11/05/2023 | 2023 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18234548&pc=QCB | 2023826-2023-05299 | 2023-09-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.50/+1.5/084 | OD | Physician | 11/07/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18234552&pc=QCB | 2023826-2023-05301 | 2022-12-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -05.50/+2.0/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2022 BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -05.50/+2.0/083 | OS | Physician | 11/05/2023 | 2023 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18234737&pc=QCB | 2023826-2023-05336 | 2023-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/2.0/163 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. REPORTEDLY, "THE SECOND LENS IS INSERTED HORIZONTALLY INTO THE EYE." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -9.5/2.0/163 | OD | Physician | 11/17/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18234738&pc=QCB | 2023826-2023-05160 | 2023-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/3.0/176 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(6). | -8.5/3.0/176 | OS | Physician | 11/15/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18234740&pc=QCB | 2023826-2023-05157 | 2023-11-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/2.0/001 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -11.5/2.0/001 | OD | Physician | 11/15/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18234741&pc=QCB | 2023826-2023-05335 | 2021-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/1.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) CLAIM #(B)(4). | -9.5/1.0/089 | OS | Physician | 11/16/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18234742&pc=QCB | 2023826-2023-05334 | 2021-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/1.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -11.0/1.0/092 | OD | Physician | 11/16/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18234743&pc=QCB | 2023826-2023-05158 | 2023-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/3.0/180 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "T1610400 WAS IMPLANTED IN A VERTICAL POSITION T1367651 IS A HORIZONTAL ENTRY." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) CLAIM #(B)(4). | -8.5/3.0/180 | OD | Physician | 11/15/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18234744&pc=QCB | 2023826-2023-05337 | 2023-11-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/1.0/014 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON 05-NOV-2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -12.5/1.0/014 | OS | Physician | 11/17/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18234746&pc=QCB | 2023826-2023-05341 | 2023-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/1.0/109 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -11.0/1.0/109 | OD | Physician | 11/20/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18234747&pc=QCB | 2023826-2023-05343 | 2023-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/2.0/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -10.5/2.0/072 | OS | Physician | 11/19/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18234748&pc=QCB | 2023826-2023-05342 | 2023-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/1.0/075 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -11.0/1.0/075 | OS | Physician | 11/20/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18234749&pc=QCB | 2023826-2023-05338 | 2023-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/0.5/130 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON 08-JUN-2023. ON 23-OCT-2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -8.5/0.5/130 | OD | Physician | 11/17/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18234750&pc=QCB | 2023826-2023-05344 | 2023-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/1.0/131 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -8.0/1.0/131 | OD | Physician | 11/19/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18234751&pc=QCB | 2023826-2023-05345 | 2023-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.50/1.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -11.50/1.5/093 | OD | Physician | 11/19/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18234752&pc=QCB | 2023826-2023-05348 | 2023-10-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/2.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(3) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -13.5/2.0/091 | OD | Physician | 11/20/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18234753&pc=QCB | 2023826-2023-05350 | 2018-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/3.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "THE REPLACEMENT LENS WAS VERTICALLY IMPLANTED." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -11.5/3.0/087 | OS | Physician | 11/21/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18234756&pc=QCB | 2023826-2023-05352 | 2023-11-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/2.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON 02-MAY-2021. ON 20-NOV-2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -8.0/2.5/092 | OS | Physician | 11/21/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18234757&pc=QCB | 2023826-2023-05354 | 2023-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/1.0/070 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -9.0/1.0/070 | OS | Physician | 11/21/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18234769&pc=QCB | 2023826-2023-05351 | 2023-11-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -10.5/2.0/092 | OD | Physician | 11/21/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18234771&pc=QCB | 2023826-2023-05353 | 2023-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/3.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -10.5/3.0/091 | OS | Physician | 11/21/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18249826&pc=QCB | 2023826-2023-05295 | 2023-11-09 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.50/+1.0/002 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS). THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.50/+1.0/002 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6)2023. THE LENS WAS REPORTED AS LOW VAULTING AND PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. CLAIM # (B)(4) Manufacturer Narrative: A2: UNK, A4: UNK, A5: UNK, A6: UNK, B3: UNK, D4: EXPIRATION DATE UNK, D6A: UNK, D6B: UNK, H4: UNK. (B)(4). | -13.50/+1.0/002 | OS | Physician | 11/09/2023 | 2023 | 44 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18249869&pc=QCB | 2023826-2023-05300 | 2023-10-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK; A5: UNK; A6: UNK. CLAIM # (B)(4). | -12.00/+1.5/093 | OD | Physician | 11/09/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18250110&pc=QCB | 2023826-2023-05405 | 2023-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/1.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -12.0/1.0/095 | OD | Physician | 11/28/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18250141&pc=QCB | 2023826-2023-05408 | 2023-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_ 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/1.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -8.0/1.5/094 | OD | Physician | 11/29/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18250152&pc=QCB | 2023826-2023-05407 | 2023-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_ 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/3.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -11.0/3.5/087 | OS | Physician | 11/29/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18251236&pc=QCB | 2023826-2023-05440 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -6.50/2.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON NOTES THERE IS A PLANNED LENS EXCHANGE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -6.50/2.0/090 | OD | Physician | 11/13/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18251614&pc=QCB | 2023826-2023-05404 | 2023-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/2.0/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.0/2.0/073 | OS | Physician | 11/28/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18252143&pc=QCB | 2023826-2023-05441 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -6.50/1.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON NOTES THERE IS A PLANNED LENS EXCHANGE. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -6.50/1.5/089 | OS | Physician | 11/13/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18252502&pc=QCB | 2023826-2023-05430 | 2023-06-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.0/2.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6)2023 LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.0/2.0/093 | OS | Physician | 11/02/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18252947&pc=QCB | 2023826-2023-05431 | 2023-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -14.50/2.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS THAT WAS IMPLANTED IN VERTICAL POSITION AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -14.50/2.0/091 | OS | Physician | 11/02/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18258191&pc=QCB | 2023826-2023-05316 | 2023-11-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6. -5.0/2.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE(OD) ON (B)(6) 2022. SLIGHT RESIDUAL ASTIGMATISM AND LENS ROTATION WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A1 - UNK. A2 - UNK. A3 - UNK. A4 - UNK. A5 - UNK. A6 - UNK. B3 - UNK. D6A - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -5.0/2.0/093 | OD | Physician | 11/16/2023 | 2023 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18258202&pc=QCB | 2023826-2023-05317 | 2022-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6, -7.0/1.0/063 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. REPORTEDLY, ELEVATED IOP CAUSED BY STEROIDS. PATIENT TREATED WITH OPHTHALMIC DROPS. Manufacturer Narrative: CLAIM# (B)(4). | -7.0/1.0/063 | OD | Physician | 11/16/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18258924&pc=QCB | 2023826-2023-05382 | 2023-10-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -18.00/2.0/090 (SPHERE/CYLINDER/AXIS) TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS NOT IMPLANTED. PATIENT CONTACT WAS REPORTED BUT NO PATIENT INJURY. A REPLACEMENT LENS WAS LATER IMPLANTED OF THE SAME MODEL/SIZE AND SIMILAR POWER. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -18.00/2.0/090 | OD | Other Health Care Professional | 11/06/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18261639&pc=QCB | 2023826-2023-05385 |  | Injury | STAAR SURGICAL COMPANY | 2023-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11.00/2.0/072 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS EYE ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS REMOVED AND A REPLACEMENT LENS WAS IMPLANTED OF LONGER LENGTH LENS DUE TO LOW VAULT. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM#: (B)(4). | -11.00/2.0/072 | NA | Other Health Care Professional | 11/08/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18262080&pc=QCB | 2023826-2023-05133 | 2023-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM55_ 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/3.0/000 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "REPLACED THE LENS, ADJUSTED THE VAULT, AND THE RESULT IS GOOD. THERE IS NO CHANGE TO THE LENS SIZE BECAUSE THE VAULT WAS ADJUSTED BY ADJUSTING THE FIXED POSITION OF THE LENS." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.0/3.0/000 | OD | Physician | 11/10/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18263843&pc=QCB | 2023826-2023-05424 | 2023-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+0.5/115 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -12.5/+0.5/115 | OD | Physician | 11/07/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18263844&pc=QCB | 2023826-2023-05420 | 2023-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. ON (B)(6)2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THAT SAME DAY A SHORTER LENGTH LENS WAS IMPLANTED HOWEVER IT IS UNCLEAR IF THIS WAS PERFORMED INTRAOPERATIVELY. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+1.0/092 | OD | Physician | 11/07/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18263845&pc=QCB | 2023826-2023-05418 | 2023-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+1.0/113 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THAT SAME DAY A SHORTER LENGTH LENS WAS IMPLANTED HOWEVER IT IS UNCLEAR IF THIS WAS PERFORMED INTRAOPERATIVELY. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -7.5/+1.0/113 | OD | Physician | 11/07/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18263850&pc=QCB | 2023826-2023-05422 | 2023-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+3.0/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THAT SAME DAY A SHORTER LENGTH LENS WAS IMPLANTED HOWEVER IT IS UNCLEAR IF THIS WAS PERFORMED INTRAOPERATIVELY. T HE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6- DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) AND (ACD < 3.0MM). CLAIM# (B)(4). | -10.0/+3.0/095 | OD | Physician | 11/07/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18263851&pc=QCB | 2023826-2023-05419 | 2023-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.0/106 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THAT SAME DAY A SHORTER LENGTH LENS WAS IMPLANTED HOWEVER IT IS UNCLEAR IF THIS WAS PERFORMED INTRAOPERATIVELY. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -10.0/+2.0/106 | OD | Physician | 11/07/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18268920&pc=QCB | 2023826-2023-05296 | 2023-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.00/+6.0/108 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED SEVERAL TIMES (B)(6) 2022 AND THE PROBLEM WAS RESOLVED. THE LENS ROTATED AGAIN AND WAS REMOVED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.00/+6.0/108 | OS | Physician | 11/10/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18271245&pc=QCB | 2023826-2023-05490 | 2023-10-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+1.0/002 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. AN EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS OF THE SAME LENGTH AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -8.0/+1.0/002 | OS | Physician | 11/10/2023 | 2023 | 45 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18271247&pc=QCB | 2023826-2023-05455 | 2023-05-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+1.5/115 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -12.5/+1.5/115 | OD | Physician | 11/09/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18271249&pc=QCB | 2023826-2023-05491 | 2023-10-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.5/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. AN EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS OF THE SAME LENGTH AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -11.0/+1.5/093 | OD | Physician | 11/10/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18271251&pc=QCB | 2023826-2023-05454 | 2023-05-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+1.0/086 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -12.0/+1.0/086 | OS | Physician | 11/09/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18271256&pc=QCB | 2023826-2023-05489 | 2023-10-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.0/001 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. AN EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS OF THE SAME LENGTH AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH. NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -7.0/+1.0/001 | OD | Physician | 11/10/2023 | 2023 | 45 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18271257&pc=QCB | 2023826-2023-05492 | 2023-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+4.0/080 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) AND (OVER 45YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -10.0/+4.0/080 | OS | Physician | 11/10/2023 | 2023 | 46 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18271673&pc=QCB | 2023826-2023-05504 | 2023-11-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -3.5/2.0/175 WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | -3.5/2.0/175 | OD | Physician | 11/08/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18271724&pc=QCB | 2023826-2023-05507 | 2023-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -9.50/3.5/136 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS REPORTED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.50/3.5/136 | OD | Physician | 11/10/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18271744&pc=QCB | 2023826-2023-05339 | 2023-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/2.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM (B)(4). | -7.0/2.0/086 | OD | Physician | 11/20/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18271745&pc=QCB | 2023826-2023-05458 | 2023-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/3.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -10.5/3.5/095 | OD | Physician | 11/30/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18271747&pc=QCB | 2023826-2023-05465 | 2023-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/1.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -11.0/1.5/086 | OS | Physician | 12/01/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18271749&pc=QCB | 2023826-2023-05459 | 2023-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/2.0/071 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4) . | -8.0/2.0/071 | OS | Physician | 11/30/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18271751&pc=QCB | 2023826-2023-05464 | 2023-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/2.5/168 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -10.5/2.5/168 | OS | Physician | 12/01/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18275561&pc=QCB | 2023826-2023-05474 |  | Injury | STAAR SURGICAL COMPANY | 2023-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -7.0/1.0/105 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). LOW VAULT AND REFRACTIVE MISS WERE REPORTED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: 1 SIMILAR COMPLAINT: CLAIM# (B)(4). (FELLOW EYE). CLAIM# (B)(4). | -7.0/1.0/105 | OD | Physician | 11/20/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18285812&pc=QCB | 2023826-2023-05421 | 2023-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+1.5/109 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B) (6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THAT SAME DAY A SHORTER LENGTH LENS WAS IMPLANTED HOWEVER IT IS UNCLEAR IF THIS WAS PERFORMED INTRAOPERATIVELY. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#(B)(4). | -12.0/+1.5/109 | OD | Physician | 11/07/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18286153&pc=QCB | 2023826-2023-05426 | 2023-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+1.5/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4) | -12.0/+1.5/095 | OD | Physician | 11/07/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18287537&pc=QCB | 2023826-2023-05525 | 2023-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -18.0/4.0/85 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -18.0/4.0/85 | OS | Physician | 12/04/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18287538&pc=QCB | 2023826-2023-05526 | 2023-08-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.0/2.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE SURGEON REPORTS THE LENS HAD ROTATED NOT ASSOCIATED TO LOW VAULTING. THE LENS WAS REPOSITIONED ON (B)(6) 2023 BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT IMPLANTED VERTICALLY AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.0/2.0/091 | OD | Physician | 12/04/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18288020&pc=QCB | 2023826-2023-05450 | 2023-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.00/+2.5/175 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK A5: UNK A6: UNK H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4) | -13.00/+2.5/175 | OD | Physician | 11/13/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18288507&pc=QCB | 2023826-2023-05483 | 2023-10-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A - 8.0/2.0/080 (SPHERE/CYLINDER/AXIS) LENS TORE/BROKE DURING LOADING. ON (B)(6) 2023, A REPLACEMENT LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | - 8.0/2.0/080 | OD | Physician | 11/23/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18288509&pc=QCB | 2023826-2023-05484 | 2023-11-23 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF A - 15.50/1.5/074 (SPHERE/CYLINDER/AXIS) LENS TORE/BROKE DURING LOADING AFTER OPENING THE VIAL. ON (B)(6) 2023, A REPLACEMENT LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | - 15.50/1.5/074 | OS | Physician | 11/24/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18288510&pc=QCB | 2023826-2023-05485 | 2023-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -7.5/1.0/080 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED. SURGICAL ADDITION OF NEW PL WAS PERFORMED. ON (B)(6)2 023, THE LENS WAS REMOVED AND SMALLER LENGTH LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL - AGE<21. CLAIM# (B)(4). | -7.5/1.0/080 | OD | Physician | 11/27/2023 | 2023 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18294796&pc=QCB | 2023826-2023-05496 | 2023-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -18.0/+2.5/072 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -18.0/+2.5/072 | OS | Physician | 11/13/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18301999&pc=QCB | 2023826-2023-05514 | 2023-11-20 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -8.5/1.0/064 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK A5 - UNK A6 - UNK H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. D6A - UNK . (B)(4) | -8.5/1.0/064 | OD | Physician | 11/21/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18302001&pc=QCB | 2023826-2023-05515 | 2023-11-20 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -9.0/1.5/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. D6A - UNK. CLAIM (B)(4). | -9.0/1.5/073 | OS | Physician | 11/21/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18302003&pc=QCB | 2023826-2023-05517 | 2023-11-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.5/0.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -9.5/0.5/093 | OS | Physician | 11/21/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18302015&pc=QCB | 2023826-2023-05516 | 2023-11-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.0/2.0/069 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.0/2.0/069 | OD | Physician | 11/21/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18302028&pc=QCB | 2023826-2023-05520 | 2022-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -6.0/1.0/123 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS DUE TO LENS DAMAGE DURING INSERTION ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: A4 - UNK, A5 - UNK, A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#(B)(4). | -6.0/1.0/123 | OS | Physician | 11/24/2023 | 2023 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18305959&pc=QCB | 2023826-2023-05575 | 2023-11-28 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method; Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -9.5/0.5/089 (SPHERE/CYLINDER/AXIS) WAS INTRAOPERATIVELY IMPLANTED AND REMOVED ON (B)(6) 2023 DUE TO THE LENS BEING IMPLANTED UPSIDE DOWN AND PRODUCT NON-CONFORMITY. REPORTEDLY, "LENS WAS TURNED DURING IMPLANTATION PRECEDURE BECAUSE OF POOR LOADED LENS IN INJECTOR." CAUSE OF THE EVENT IS USER ERROR. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6: USER ERROR. CLAIM# (B)(4). | -9.5/0.5/089 | NA | Physician | 11/30/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18311757&pc=QCB | 2023826-2023-05578 | 2023-11-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.0/2.0/123 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6)2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL ACD<3.0. CLAIM# (B)(4). | -9.0/2.0/123 | OD | Physician | 12/01/2023 | 2023 | 42 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18311770&pc=QCB | 2023826-2023-05547 | 2022-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+3.5/096 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -12.5/+3.5/096 | OD | Physician | 11/14/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18311790&pc=QCB | 2023826-2023-05542 | 2023-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.5/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -11.5/+1.5/094 | OD | Physician | 11/14/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18311801&pc=QCB | 2023826-2023-05543 | 2023-09-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+2.5/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -8.0/+2.5/094 | OD | Physician | 11/14/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18311807&pc=QCB | 2023826-2023-05499 | 2023-11-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.0/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -10.5/+2.0/095 | OD | Physician | 11/15/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18312358&pc=QCB | 2023826-2023-05439 | 2023-11-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -1.00/6.0/018 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LENS ROTATION WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS REPOSITIONED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -1.00/6.0/018 | OD | Physician | 11/14/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18313470&pc=QCB | 2023826-2023-05552 | 2023-11-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.0/090 INTO THE PATIENT'S LEFT EYE (OS) ON 01-NOV-2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -11.0/+1.0/090 | OS | Physician | 11/15/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18313490&pc=QCB | 2023826-2023-05549 | 2023-09-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+2.0/085 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: (B)(4). | -9.0/+2.0/085 | OS | Physician | 11/14/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18315509&pc=QCB | 2023826-2023-05596 | 2023-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/1.0/107 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -9.5/1.0/107 | OD | Physician | 12/05/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18315511&pc=QCB | 2023826-2023-05595 | 2023-10-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/1.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -7.0/1.5/086 | OS | Physician | 12/05/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18315515&pc=QCB | 2023826-2023-05597 | 2023-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/2.5/071 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -7.0/2.5/071 | OS | Physician | 12/05/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18315517&pc=QCB | 2023826-2023-05590 | 2023-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.5/115 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -10.5/1.5/115 | OD | Physician | 12/05/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18315518&pc=QCB | 2023826-2023-05602 | 2023-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/3.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -6.5/3.5/091 | OD | Physician | 12/06/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18315519&pc=QCB | 2023826-2023-05603 | 2023-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/3.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -14.5/3.0/093 | OD | Physician | 12/07/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18315521&pc=QCB | 2023826-2023-05607 | 2023-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/2.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -8.5/2.0/089 | OS | Physician | 12/07/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18315522&pc=QCB | 2023826-2023-05604 | 2023-11-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/1.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -11.5/1.0/088 | OS | Physician | 12/07/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18315523&pc=QCB | 2023826-2023-05600 | 2023-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/1.0/014 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -12.0/1.0/014 | OD | Physician | 12/06/2023 | 2023 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18315524&pc=QCB | 2023826-2023-05608 | 2023-06-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/3.0/066(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -7.0/3.0/066 | OD | Physician | 12/07/2023 | 2023 | 42 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18319699&pc=QCB | 2023826-2023-05579 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-12-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -6.50/1.0/100 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT IS OBSERVED. Manufacturer Narrative: A4-A6:UNK H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | -6.50/1.0/100 | OS | Physician | 12/04/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18319700&pc=QCB | 2023826-2023-05580 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-12-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -6.50/2.50/161 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT IS OBSERVED. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -6.50/2.50/161 | OD | Physician | 12/04/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18321709&pc=QCB | 2023826-2023-05616 | 2023-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.0/094(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -10.5/1.0/094 | OD | Physician | 12/11/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18321713&pc=QCB | 2023826-2023-05613 | 2023-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.0/1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -5.0/1.0/090 | OS | Physician | 12/11/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18323430&pc=QCB | 2023826-2023-05666 | 2023-04-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/1.5/110 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -11.0/1.5/110 | OD | Physician | 12/11/2023 | 2023 | 42 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18323431&pc=QCB | 2023826-2023-05665 | 2023-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.0/167(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -10.0/1.0/167 | OS | Physician | 12/11/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18323432&pc=QCB | 2023826-2023-05663 | 2023-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/1.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -7.0/1.5/095 | OS | Physician | 12/11/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18323433&pc=QCB | 2023826-2023-05662 | 2023-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/1.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -6.5/1.5/092 | OD | Physician | 12/11/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18323434&pc=QCB | 2023826-2023-05664 | 2023-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/0.5/165(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -10.0/0.5/165 | OD | Physician | 12/11/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18330559&pc=QCB | 2023826-2023-05675 | 2023-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.5/2.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS AND THIS RESOLVED THE PROBLEM. THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IT WAS ADDITIONALLY NOTED "THE LENS WAS STUCK IN THE CARTRIDGE WHILE BEING INJECTED INTO THE ANTERIOR CHAMBER CAUSING THE PROXIMAL HAPTIC FOOT PLATE TO RUPTURE. HOWEVER THERE WAS NO LENS AT THAT TIME SO THE REIMPLANTATION WAS (B)(6) 2023". Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.5/2.5/090 | OD | Physician | 12/10/2023 | 2023 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18332893&pc=QCB | 2023826-2023-04653 | 2023-08-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -14.00/4.0/093 (SPHERE/CYLINDER/AXIS) LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS IMPLANTED; REMOVED ON (B)(6) 2023 INTRA-OPERATIVELY. A REPLACEMENT LENS OF THE SAME MODEL AND SIZE WAS LATER IMPLANTED. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). | -14.00/4.0/093 | OS | Physician | 09/20/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18336978&pc=QCB | 2023826-2023-05643 | 2023-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.00/+1.0/135 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THAT SAME DAY A SHORTER LENGTH LENS WAS IMPLANTED HOWEVER IT IS UNCLEAR IF THIS WAS PERFORMED INTRAOPERATIVELY. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.00/+1.0/135 | OD | Physician | 11/20/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18336979&pc=QCB | 2023826-2023-05634 | 2023-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/+3.0/088 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -14.5/+3.0/088 | OS | Physician | 11/17/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18336981&pc=QCB | 2023826-2023-05635 | 2023-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.50/+2.5/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THAT SAME DAY IN A SEPARATE SURGERY, A SHORTER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -5.50/+2.5/090 | OD | Physician | 11/17/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18337565&pc=QCB | 2023826-2023-05531 | 2023-07-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50/+3.0/091 (SPHERE/CYLINDER/AXIS), TORE/BROKE BEFORE LOADING, AFTER OPENING THE VIAL. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -14.50/+3.0/091 | OS | Physician | 11/20/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18341151&pc=QCB | 2023826-2023-05657 | 2023-11-09 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -16.00/2.5/092 (SPHERE/CYLINDER/AXIS) FLIPPED IN PATIENT'S EYE AND WAS EXPLANTED. THE BACKUP LENS WAS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#:(B)(4). | -16.00/2.5/092 | OS | Other Health Care Professional | 11/17/2023 | 2023 | 47 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18341177&pc=QCB | 2023826-2023-05641 | 2023-10-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.5/088 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS TWO SIZES LONGER IN LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -9.5/+1.5/088 | OS | Physician | 11/20/2023 | 2023 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18341178&pc=QCB | 2023826-2023-05636 | 2023-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+2.5/114 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -8.5/+2.5/114 | OS | Physician | 11/17/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18342237&pc=QCB | 2023826-2023-05540 | 2023-10-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00/+2.0/095 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS WELL AND BCVA ACHIEVED. THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -18.00/+2.0/095 | OS | Physician | 11/23/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18342239&pc=QCB | 2023826-2023-05539 | 2023-10-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50/+3.5/136 (SPHERE/CYLINDER/AXIS), WAS TORN/BROKE DURING LOADING, AFTER OPENING VIAL. THE LENS WAS NOT IMPLANTED. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS WELL AND BCVA ACHIEVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.50/+3.5/136 | NA | Physician | 11/23/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18346559&pc=QCB | 2023826-2023-05624 | 2023-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -12.5/2.0/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON 16/MAR/2023. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.5/2.0/073 | OS | Physician | 11/25/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18346560&pc=QCB | 2023826-2023-05627 | 2023-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -8.0/1.0/129 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.0/1.0/129 | OS | Physician | 11/27/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18348332&pc=QCB | 2023826-2023-05591 | 2023-10-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/1.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) (B)(4). | -11.5/1.5/090 | OS | Physician | 12/05/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18353358&pc=QCB | 2023826-2023-05701 | 2023-11-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -1.5/2.5/105 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LOW VAULT WITH ROTATION WAS OBSERVED, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -1.5/2.5/105 | OD | Physician | 11/21/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18353366&pc=QCB | 2023826-2023-05702 | 2023-11-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -1.5/2.5/081 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LOW VAULT WITH ROTATION WAS OBSERVED, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -1.5/2.5/081 | OS | Physician | 11/21/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18353416&pc=QCB | 2023826-2023-05705 | 2023-11-21 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -9.50/1.5/113 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LOW VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.50/1.5/113 | OD | Physician | 11/22/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18353471&pc=QCB | 2023826-2023-05708 | 2023-11-23 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -9.50/2.0/101 (SPHERE/CYLINDER/AXIS) TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS EYE. THE LENS WAS IMPLANTED,REMOVED AND REPLACED INTRAOPERATIVELY WITH NO PATIENT INJURY AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | -9.50/2.0/101 | NA | Physician | 11/24/2023 | 2023 | 43 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18353492&pc=QCB | 2023826-2023-05706 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: 940UM VAULT AFTER ICL IMPLANTATION WAS REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4) | -18.00/1.5/092 | OD | Physician | 11/24/2023 | 2023 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18359603&pc=QCB | 2023826-2023-05732 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT WHILE ATTEMPTING TO LOAD A 13.2MM VTICMO13.2; -16.0/2.0/113 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE CARTRIDGE, IT TORE. A BACK UP LENS WAS USED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -16.0/2.0/113 | NA | Physician | 11/30/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18360945&pc=QCB | 2023826-2023-05554 | 2023-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS INTO THE PATIENTS EYE. THE LENS HAD ROTATED AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: A2: UNK. A3: UNK. A4: UNK. A5: UNK . A6: UNK . B3: UNK. D4: EXPIRATION DATE UNK. D6A: UNK. D6B: UNK. H4: UNK. CLAIM # (B)(4). | -15.00/2.0/110 | OD | Other Health Care Professional | 11/28/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18360958&pc=QCB | 2023826-2023-05556 | 2023-09-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50/+2.0/073 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.50/+2.0/073 | OS | Physician | 11/27/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18360969&pc=QCB | 2023826-2023-05561 | 2023-09-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/110 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. THE PATIENT IS HAPPY WITH NO SYMPTOMS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: D6B: UNK H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.00/+2.0/110 | OD | Physician | 11/29/2023 | 2023 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18360971&pc=QCB | 2023826-2023-05560 | 2023-09-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50/+1.5/066 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON 09-AUG-2022. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. THE PATIENT IS HAPPY WITH NO SYMPTOMS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.50/+1.5/066 | OS | Physician | 11/29/2023 | 2023 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18361085&pc=QCB | 2023826-2023-05645 | 2023-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+1.0/152 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -13.5/+1.0/152 | OD | Physician | 11/21/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18361087&pc=QCB | 2023826-2023-05647 | 2023-09-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+3.5/088 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS TWO SIZES LONGER IN LENGTH DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -13.0/+3.5/088 | OS | Physician | 11/21/2023 | 2023 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18369658&pc=QCB | 2023826-2023-05473 | 2023-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/0.5/055 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -10.0/0.5/055 | OD | Physician | 12/04/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18369659&pc=QCB | 2023826-2023-05472 | 2023-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/1.0/038(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON 03-MAR-2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. (B)(4). | -12.0/1.0/038 | OS | Physician | 12/04/2023 | 2023 | 47 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18369660&pc=QCB | 2023826-2023-05667 | 2023-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/3.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -9.5/3.0/091 | OS | Physician | 12/11/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18369661&pc=QCB | 2023826-2023-05609 | 2023-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/1.5/134 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -14.0/1.5/134 | OS | Physician | 12/07/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18369667&pc=QCB | 2023826-2023-05671 | 2023-12-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/1.5/061(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -7.5/1.5/061 | OS | Physician | 12/12/2023 | 2023 | 45 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18369668&pc=QCB | 2023826-2023-05674 | 2023-11-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/1.0/092(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -6.0/1.0/092 | OS | Physician | 12/12/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18369674&pc=QCB | 2023826-2023-05673 | 2023-11-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/1.0/068(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -5.5/1.0/068 | OD | Physician | 12/12/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18369676&pc=QCB | 2023826-2023-05672 | 2023-12-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/2.5/112 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "LENS OBLIQUE AXIS IMPLANTATION." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -6.5/2.5/112 | OD | Physician | 12/12/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18370057&pc=QCB | 2023826-2023-05685 | 2023-12-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -9.5/1.0/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -9.5/1.0/91 | OD | Physician | 12/12/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18374480&pc=QCB | 2023826-2023-05686 | 2023-12-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.5/1.5/70 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT WAS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -10.5/1.5/70 | OS | Physician | 12/12/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18375245&pc=QCB | 2023826-2023-05738 | 2023-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+2.5/081 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#(B)(4). | -13.0/+2.5/081 | OS | Physician | 11/24/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18375260&pc=QCB | 2023826-2023-05734 | 2023-11-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+2.5/111 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -13.5/+2.5/111 | OD | Physician | 11/24/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18375273&pc=QCB | 2023826-2023-05735 | 2023-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+2.0/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -8.0/+2.0/094 | OD | Physician | 11/24/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18375274&pc=QCB | 2023826-2023-05736 | 2023-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.0/+3.5/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -16.0/+3.5/093 | OS | Physician | 11/24/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18376720&pc=QCB | 2023826-2023-05565 | 2023-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -06.00/+0.5/081 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REMOVED ON (B)(6) 2023 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT WAS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: A4: UNK. A5: UNK. A6: UNK. CLAIM # (B)(4). | -06.00/+0.5/081 | OD | Physician | 12/01/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18376736&pc=QCB | 2023826-2023-05566 | 2023-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -00.50/+6.0/109 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2023 BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -00.50/+6.0/109 | OS | Physician | 12/04/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18376797&pc=QCB | 2023826-2023-05563 | 2023-11-24 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00/+1.0/017 (SPHERE/CYLINDER/AXIS), WAS OBSERVED TO BE DEFORMED. THERE WAS NO PATIENT CONTACT. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: A4: UNK, A5: UNK, A6: UNK, H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.00/+1.0/017 | NA | Physician | 11/30/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18378876&pc=QCB | 2023826-2023-05710 |  | Injury | STAAR SURGICAL COMPANY | 2023-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM513.2, -9.50/1.0/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN OCTOBER 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: D6A: (B)(6) 2023. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -9.50/1.0/111 | OD | Physician | 11/27/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18379196&pc=QCB | 2023826-2023-05719 | 2022-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.00/2.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2023 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | -10.00/2.0/093 | OD | Physician | 11/29/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18379216&pc=QCB | 2023826-2023-05733 | 2023-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+4.0/110 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON 24-OCT-2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.0/+4.0/110 | OD | Physician | 11/24/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18384213&pc=QCB | 2023826-2023-05620 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM TMICL 13.2 IMPLANTABLE COLLAMER LENS OF -15.50/1.5/154 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS EYE. LENS ROTATION WAS OBSERVED. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -15.50/1.5/154 | NA | Other Health Care Professional | 11/29/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18384401&pc=QCB | 2023826-2023-05718 | 2022-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -8.50/2.5/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)2022 LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -8.50/2.5/091 | OS | Physician | 11/29/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18384410&pc=QCB | 2023826-2023-05711 |  | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -15.50/2.5/090 (SPHERE/CYLINDER/AXIS) WAS INSERTED UPSIDE DOWN AND TORN UPON REMOVAL. BACKUP LENS WAS IMPLANTED SUCCESSFULLY. NO PATIENT HARM WAS REPORTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -15.50/2.5/090 | NA | Other Health Care Professional | 11/27/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18385487&pc=QCB | 2023826-2023-05720 | 2023-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -5.50/1.5/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6)2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND IT WAS REPORTED THAT THE PROBLEM RESOLVED. REPORTER STATED " GOOD VAULT OF ABOUT 800 MICRONS, STILL SOMEWHAT HIGH, BUT ANGLES ABOUT 20DEG". THE REPORTER STATED THAT THE CAUSE OF THIS EVENT WAS DUE TO USER ERROR, PRODUCT DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CLAIM#: (B)(4). | -5.50/1.5/091 | OD | Other Health Care Professional | 11/30/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18385725&pc=QCB | 2023826-2023-05721 | 2023-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -5.50/1.5/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6)2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND IT WAS REPORTED THAT THE PROBLEM RESOLVED. REPORTER STATED " GOOD VAULT OF ABOUT 800 MICRONS, STILL SOMEWHAT HIGH, BUT ANGLES ABOUT 20DEG". THE REPORTER STATED THAT THE CAUSE OF THIS EVENT WAS DUE TO USER ERROR, PRODUCT DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: CLAIM#: (B)(4). | -5.50/1.5/090 | OS | Other Health Care Professional | 11/30/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18387534&pc=QCB | 2023826-2023-05728 | 2023-11-09 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.00/+2.5/112 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. THE LENS HAD A LOW VAULT AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINED IMPLANTED. CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: DEVICE CODE: 1494 THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.00/+2.5/112 | OD | Physician | 12/12/2023 | 2023 | 44 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18387546&pc=QCB | 2023826-2023-05688 |  | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF A -14.5/1.0/098 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -14.5/1.0/098 | OD | Physician | 12/13/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18387553&pc=QCB | 2023826-2023-05567 | 2023-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -03.00/+5.0/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2023. THE LENS WAS EXPLANTED ON (B)(6) 2023 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2023 BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4: UNK A5: UNK A6: UNK H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -03.00/+5.0/086 | OD | Physician | 12/04/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18387656&pc=QCB | 2023826-2023-05713 | 2023-11-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.00/1.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE/BROKE DURING LOADING. LENS WAS IMPLANTED,REMOVED AND REPLACED INTRAOPERATIVELY. CAUSE OF EVENT UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -15.00/1.5/090 | NA | Physician | 11/28/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18387739&pc=QCB | 2023826-2023-05772 | 2023-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+3.0/095 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -12.0/+3.0/095 | OS | Physician | 11/27/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18387740&pc=QCB | 2023826-2023-05768 | 2023-11-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.0/085 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -10.5/+1.0/085 | OS | Physician | 11/27/2023 | 2023 | 44 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18387741&pc=QCB | 2023826-2023-05769 | 2023-11-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.0/075 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -10.0/+1.0/075 | OS | Physician | 11/27/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18387742&pc=QCB | 2023826-2023-05770 | 2023-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+1.0/067 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -13.5/+1.0/067 | OD | Physician | 11/27/2023 | 2023 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18387787&pc=QCB | 2023826-2023-05774 | 2023-11-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+0.5/141 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4) | -8.0/+0.5/141 | OD | Physician | 11/28/2023 | 2023 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18389261&pc=QCB | 2023826-2023-05767 | 2023-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.5/053 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4) | -9.5/+1.5/053 | OS | Physician | 11/27/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18393705&pc=QCB | 2023826-2023-05693 | 2023-12-01 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -14.0/2.5/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT(LENS TOUCHED THE EYE). NO PATIENT INJURY. THE LENS WAS IMPLANTED, AND EXPLANTED INTRAOPERATIVELY. PROBLEM RESOLVED. REPORTEDLY A NEW LENS IS TO BE IMPLANTED AT A LATER DATE. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -14.0/2.5/095 | OD | Physician | 12/04/2023 | 2023 | 48 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18393709&pc=QCB | 2023826-2023-05696 | 2023-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.0/2.5/135 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2023. ON (B)(6) 2023 PATIENTS 9 WEEK POST REPOSITIONING, LENS DID NOT ROTATE AGAIN. PROBLEM RESOLVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR AND DEVICE: LENGTH NOT APPROPRIATE TO WHITE TO WHITE MEASUREMENT. Manufacturer Narrative: A4 - UNK, A5 - UNK, A6 - UNK, H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.0/2.5/135 | OD | Physician | 12/05/2023 | 2023 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18393710&pc=QCB | 2023826-2023-05697 | 2023-12-01 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -14.0/1.5/122 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY PATIENT IS HAPPY. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -14.0/1.5/122 | OD | Physician | 12/05/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18394871&pc=QCB | 2023826-2023-05794 | 2023-11-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -4.0/1.0/067(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "THE SECOND TICL IS IN THE VERTICAL POSITION." Manufacturer Narrative: B5: LENS WAS REPOSITIONED ON (B)(6) 2023. THE REPOSITIONING DID NOT RESOLVED THE ISSUE. (B)(4) Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -4.0/1.0/067 | OD | Physician | 12/17/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18394873&pc=QCB | 2023826-2023-05796 | 2023-10-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -10.0/1.5/099 (SPHERE/CYILNDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. EXCESSIVE VAULT IS OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: 1 SIMILAR COMPLAINT WAS FOUND: CLAIM# (B)(4). CLAIM# (B)(4). | -10.0/1.5/099 | OD | Physician | 12/18/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18394878&pc=QCB | 2023826-2023-05797 | 2023-10-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -9.50/1.0/075 (SPHERE/CYILNDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. EXCESSIVE VAULT AND REFRACTIVE SUPRISE ARE OBSERVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE FOUND. CLAIM# (B)(4). | -9.50/1.0/075 | OS | Physician | 12/18/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18395488&pc=QCB | 2023826-2023-05798 | 2023-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.0/2.0/073 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6)2023, THE LENS WAS EXCHANGED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "THE LEFT EYE WAS IMPLANTED IN THE VERTICAL DIRECTION." Manufacturer Narrative: A4-A6:UNK. H6: OFF-LABEL -AGE <45. CLAIM# (B)(4). | -11.0/2.0/073 | OS | Physician | 12/18/2023 | 2023 | 46 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18401515&pc=QCB | 2023826-2023-05741 | 2023-07-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/3.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -8.5/3.5/094 | OS | Physician | 12/14/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18401516&pc=QCB | 2023826-2023-05744 | 2023-08-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/1.0/069 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A DIFFERENT MODEL/ SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -12.0/1.0/069 | OS | Physician | 12/14/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18401517&pc=QCB | 2023826-2023-05740 | 2023-07-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/4.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -9.5/4.0/095 | OD | Physician | 12/14/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18401518&pc=QCB | 2023826-2023-05742 | 2023-06-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/1.0/046 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING WITHOUT ROTATION. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK CLAIM # (B)(4). | -8.5/1.0/046 | OS | Physician | 12/14/2023 | 2023 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18401519&pc=QCB | 2023826-2023-05743 | 2023-08-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/1.0/102 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -12.0/1.0/102 | OD | Physician | 12/14/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18401522&pc=QCB | 2023826-2023-05746 | 2023-06-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.5/079 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING WITHOUT ROTATION. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -10.5/1.5/079 | OS | Physician | 12/14/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18401523&pc=QCB | 2023826-2023-05747 | 2023-08-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.0/038 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING WITHOUT ROTATION. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -10.0/1.0/038 | OD | Physician | 12/14/2023 | 2023 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18401528&pc=QCB | 2023826-2023-05745 | 2023-06-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/1.5/099 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING WITHOUT ROTATION. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -11.0/1.5/099 | OD | Physician | 12/14/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18407299&pc=QCB | 2023826-2023-05801 | 2023-12-19 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF A -8.5/4.0/076 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2023. LENS ROTATON WAS OBSERVED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -8.5/4.0/076 | OS | Physician | 12/19/2023 | 2023 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18411465&pc=QCB | 2023826-2023-05470 | 2023-08-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -2.50/6.0/013 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -2.50/6.0/013 | OD | Physician | 12/04/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18411929&pc=QCB | 2023826-2023-05815 | 2023-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+3.5/087 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/+3.5/087 | OS | Physician | 11/29/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18411930&pc=QCB | 2023826-2023-05812 | 2023-10-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.0/108 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS REPORTEDLY NOT RESOLVED, AND THE REPORTER STATED THAT THEY WILL "KEEP WATCHING". Manufacturer Narrative: CLAIM# (B)(4). | -10.0/+1.0/108 | OD | Physician | 11/29/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18411934&pc=QCB | 2023826-2023-05817 | 2023-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.0/072 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) . CLAIM# (B)(4). | -7.0/+1.0/072 | OS | Physician | 11/29/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18411937&pc=QCB | 2023826-2023-05818 | 2022-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+2.0/176 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/+2.0/176 | OD | Physician | 11/29/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18411970&pc=QCB | 2023826-2023-05814 | 2023-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+5.5/086 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -10.0/+5.5/086 | OD | Physician | 11/29/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18413877&pc=QCB | 2023826-2023-05831 | 2023-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -13.50/3.0/102 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 LENS REPOSITIONING WAS PERFORMED DUE TO LOW VAULT WITH ROTATION BUT THE PROBLEM DID NOT RESOLVE. ON (B)(6) 2023 THE LENS WAS REMOVED AND A REPLACEMENT LENS OF LONGER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. IN THE REPORTER OPINION THE CAUSE OF THE EVENT WAS THE DEVICE "ICL TOO SHORT DESPITE SIZING RECOMMENDATION". Manufacturer Narrative: CLAIM# (B)(4) | -13.50/3.0/102 | OS | Physician | 11/28/2023 | 2023 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18414576&pc=QCB | 2023826-2023-05836 | 2023-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.0/087 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -7.0/+1.0/087 | OS | Physician | 12/03/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18414872&pc=QCB | 2023826-2023-05755 | 2023-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/1.0/087(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -8.5/1.0/087 | OS | Physician | 12/18/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18414873&pc=QCB | 2023826-2023-05754 | 2023-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO LOW VAULT WITHOUT ROTATION. ON (B)(6) 2023 A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED. REPORTEDLY, "THE SECOND LENS IS PLACED IN AN OBLIQUE POSITION." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -10.5/1.5/093 | OD | Physician | 12/18/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18414874&pc=QCB | 2023826-2023-05752 | 2023-12-15 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/1.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING WITHOUT ROTATION. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -14.0/1.5/074 | OD | Physician | 12/18/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18414876&pc=QCB | 2023826-2023-05751 | 2023-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/2.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -11.0/2.5/089 | OS | Physician | 12/14/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18415614&pc=QCB | 2023826-2023-05750 | 2023-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/3.0/084 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | -13.0/3.0/084 | OS | Physician | 12/14/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18415867&pc=QCB | 2023826-2023-05804 | 2023-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -7.5/2.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED AND A LARGER LENGTH LENS WAS IMPLANTED DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -7.5/2.0/092 | OD | Physician | 12/20/2023 | 2023 | 43 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18416180&pc=QCB | 2023826-2023-05753 | 2023-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/2.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK . CLAIM # (B)(4). | -9.0/2.0/094 | OS | Physician | 12/18/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18418970&pc=QCB | 2023826-2023-05830 | 2023-08-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -8.50/1.0/129 DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. FOR STATUS OF THE EYE THE REPORTER STATED THAT THE EYE WAS FIRST DAY POSTOPERATIVE WITH CORNEAL EDEMA AND MIOSIS FROM SURGERY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -8.50/1.0/129 | OD | Physician | 12/05/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18421045&pc=QCB | 2023826-2023-05811 | 2023-12-11 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF A -16.5/2.5/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). REPORTEDLY, "THERE WAS A LINT-LIKE SUBSTANCE ON SFC-45." THERE WAS NO PROBLEM WITH THE EYE. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS. CLAIM# (B)(4). | -16.5/2.5/086 | OS | Physician | 12/22/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18422406&pc=QCB | 2023826-2023-05757 | 2023-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/1.5/075 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -7.0/1.5/075 | OD | Physician | 12/17/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18422407&pc=QCB | 2023826-2023-05756 | 2023-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/1.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -7.5/1.5/094 | OS | Physician | 12/17/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18422410&pc=QCB | 2023826-2023-05761 | 2023-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.0/4.0/107 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -17.0/4.0/107 | OD | Physician | 12/18/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18422479&pc=QCB | 2023826-2023-05843 | 2023-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.5/+1.0/091 | OD | Physician | 12/04/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18422481&pc=QCB | 2023826-2023-05846 | 2023-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.5/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -10.5/+1.5/090 | OS | Physician | 12/04/2023 | 2023 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18422482&pc=QCB | 2023826-2023-05841 | 2023-11-25 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+1.5/073 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -6.5/+1.5/073 | OS | Physician | 12/04/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18422483&pc=QCB | 2023826-2023-05840 | 2023-11-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.0/085 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -9.0/+1.0/085 | OS | Physician | 12/04/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18422484&pc=QCB | 2023826-2023-05844 | 2023-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.5/+1.0/091 | OS | Physician | 12/04/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18422486&pc=QCB | 2023826-2023-05845 | 2023-10-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+3.0/095 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -7.5/+3.0/095 | OS | Physician | 12/04/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18422488&pc=QCB | 2023826-2023-05847 | 2023-10-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+1.5/172 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -12.0/+1.5/172 | OD | Physician | 12/04/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18427318&pc=QCB | 2023826-2024-00003 | 2023-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -5.5/3.5/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LENS IMPLANTED UPSIDE DOWN. PATIENT CONTACT(LENS TOUCHED THE EYE). NO PATIENT INJURY. ON (B)(6) 2023 A SAME MODEL/LENGTH LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK; A5 - UNK; A6 - UNK; H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -5.5/3.5/099 | OS | Physician | 12/11/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18427319&pc=QCB | 2023826-2024-00001 | 2023-05-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.5/2.0/069 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK; A5 - UNK; A6 - UNK; CLAIM# (B)(4). | -10.5/2.0/069 | OS | Physician | 12/11/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18427320&pc=QCB | 2023826-2024-00002 | 2023-05-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.0/2.0/132 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION ON (B)(6)2023. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: A4 - UNK; A5 - UNK; A6 - UNK; CLAIM# (B)(4). | -15.0/2.0/132 | OD | Physician | 12/11/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18427325&pc=QCB | 2023826-2024-00007 | 2023-12-05 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event; Ejection Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -13.0/1.5/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE LENS STUCK IN CARTRIDGE/INJECTION SYSTEM DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE; LENS STUCK IN THE INJECTOR. Manufacturer Narrative: A4 - UNK. A5 - UNK. A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.0/1.5/070 | OS | Physician | 12/13/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18427327&pc=QCB | 2023826-2024-00005 | 2023-11-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED VERTICALLY A 13.2MM VTICMO13.2. -18.0/1.0/070 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT(LENS TOUCHED THE EYE). NO PATIENT INJURY. THE LENS WAS REPLACED WITH A SAME LENGTH LENS ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT: UNKNOWN. BACK UP LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY. Manufacturer Narrative: H6 - DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -18.0/1.0/070 | OD | Physician | 12/11/2023 | 2023 | 44 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18427328&pc=QCB | 2023826-2024-00004 | 2023-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2. -11.5/1.5/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT(LENS TOUCHED THE EYE). NO PATIENT INJURY. THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: A4 - UNK; A5 - UNK; A6 - UNK; H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.5/1.5/087 | OD | Physician | 12/11/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18427573&pc=QCB | 2023826-2024-00045 | 2023-11-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2 VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -8.00/2.0/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). ON 16-NOV-2023 THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY DUE TO THE LENS BEING IMPLANTED UPSIDE DOWN. NO PATIENT INJURY REPORTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -8.00/2.0/095 | OD | Physician | 12/05/2023 | 2023 | 46 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18428693&pc=QCB | 2023826-2024-00061 | 2023-12-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT WHILE ATTEMPTING TO IMPLANT A 12.6MM VTICM5_12.6; -6.50/3.0/096 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS), IT WAS INDICATED THAT THE LENS HAD FLIPPED UPSIDE DOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -6.50/3.0/096 | OS | Physician | 12/11/2023 | 2023 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18432115&pc=QCB | 2023826-2024-00020 | 2023-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+2.5/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON 15-OCT-2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.5/+2.5/094 | OS | Physician | 12/05/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18432116&pc=QCB | 2023826-2024-00017 | 2023-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+3.0/94 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -13.5/+3.0/94 | OD | Physician | 12/05/2023 | 2023 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18432126&pc=QCB | 2023826-2024-00018 | 2023-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+2.0/096 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -7.0/+2.0/096 | OS | Physician | 12/05/2023 | 2023 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18432127&pc=QCB | 2023826-2024-00019 | 2023-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0/+1.5/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON 07-OCT-2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: (B)(4). | -15.0/+1.5/090 | OS | Physician | 12/05/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18432771&pc=QCB | 2023826-2024-00027 |  | Injury | STAAR SURGICAL COMPANY | 2024-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF A -5.0/0.5/68 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, AND PERIPHERAL ANTERIOR SYNECHIAE. IN A SUBSEQUENT SURGERY LATER THAT DAY, A SMALLER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -5.0/0.5/68 | OS | Physician | 12/26/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18432773&pc=QCB | 2023826-2024-00026 |  | Injury | STAAR SURGICAL COMPANY | 2024-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF A -5.0/0.5/155 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES, AND PERIPHERAL ANTERIOR SYNECHIAE. IN A SUBSEQUENT SURGERY LATER THAT DAY, A SMALLER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -5.0/0.5/155 | OD | Physician | 12/26/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18437844&pc=QCB | 2023826-2024-00031 | 2023-10-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/091 (SPHERE/CYLINDER/AXIS) LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED, AND REPLACED ON (B)(6) 2023. THE PROBLEM WAS RESOLVED. THERE WAS NO PATIENT INJURY. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -9.0/1.0/091 | NA | Physician | 12/26/2023 | 2023 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18440773&pc=QCB | 2023826-2024-00022 | 2023-10-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+1.5/173 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -7.5/+1.5/173 | OD | Physician | 12/06/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18441205&pc=QCB | 2023826-2024-00081 | 2023-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/2.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -5.5/2.0/091 | OS | Physician | 12/27/2023 | 2023 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18441227&pc=QCB | 2023826-2024-00064 | 2023-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.5/074(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. REPORTEDLY, "ICL LENS IN VERTICAL POSITION." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -10.0/1.5/074 | OS | Physician | 12/22/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18441228&pc=QCB | 2023826-2024-00068 | 2023-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.0/4.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -16.0/4.5/090 | OD | Physician | 12/26/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18441229&pc=QCB | 2023826-2024-00067 | 2023-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/3.5/172 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -12.5/3.5/172 | OD | Physician | 12/26/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18441231&pc=QCB | 2023826-2024-00063 | 2023-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/1.0/092(SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. REPORTEDLY, "ICL LENS IN VERTICAL POSITION." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK CLAIM #(B)(4). | -11.0/1.0/092 | OD | Physician | 12/22/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18441235&pc=QCB | 2023826-2024-00073 | 2023-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/2.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -7.0/2.0/090 | OD | Physician | 12/24/2023 | 2023 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18441236&pc=QCB | 2023826-2024-00076 | 2023-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.0/1.5/109 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) CLAIM #(B)(4) | -5.0/1.5/109 | OD | Physician | 12/26/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18441237&pc=QCB | 2023826-2024-00071 | 2023-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0/2.5/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, "THE SURGEON OPTED FOR VERTICAL IMPLANTATION." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -15.0/2.5/082 | OS | Physician | 12/25/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18441238&pc=QCB | 2023826-2024-00077 | 2023-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/1.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -12.0/1.5/090 | OS | Physician | 12/26/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18441239&pc=QCB | 2023826-2024-00072 | 2023-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.0/177 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -10.0/1.0/177 | OD | Physician | 12/26/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18441240&pc=QCB | 2023826-2024-00074 | 2023-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/2.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). | -12.0/2.0/091 | OS | Physician | 12/24/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18441244&pc=QCB | 2023826-2024-00080 | 2023-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/1.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -9.0/1.5/095 | OD | Physician | 12/27/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18441245&pc=QCB | 2023826-2024-00078 | 2023-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/1.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -12.5/1.5/095 | OD | Physician | 12/26/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18446417&pc=QCB | 2023826-2024-00051 | 2023-10-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -9.00/2.0/090 (SPHERE/CYLINDER/AXIS) LENS HAD A TEAR/BREAK DURING INJECTION/DELIVERY INTO THE PATIENTS RIGHT EYE (OD). ON (B)(6) 2023 THE LENS WAS IMPLANTED; REMOVED INTRA-OPERATIVELY. A REPLACEMENT LENS WAS LATER IMPLANTED OF THE SAME MODEL;SIZE AND POWER. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) CLAIM# (B)(4). | -9.00/2.0/090 | OD | Physician | 12/11/2023 | 2023 | 43 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18448887&pc=QCB | 2023826-2024-00021 | 2023-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+2.5/113 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON 01-NOV-2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -11.5/+2.5/113 | OD | Physician | 12/06/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18449691&pc=QCB | 2023826-2024-00088 | 2023-12-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) AS A REPLACEMENT LENS TO CORRECT LOW VAULT WITHOUT ROTATION ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -10.5/1.5/092 | OS | Physician | 12/28/2023 | 2023 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18450512&pc=QCB | 2023826-2024-00100 | 2023-07-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.0/+2.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -3.0/+2.0/093 | OD | Physician | 12/08/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18450634&pc=QCB | 2023826-2024-00140 | 2023-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -12.00/+2.0/069 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE PATIENT EXPERIENCED LENS ROTATION. ON (B)(6) 2023 THE LENS WAS REPOSITIONED AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -12.00/+2.0/069 | OS | Physician | 12/13/2023 | 2023 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18451022&pc=QCB | 2023826-2024-00086 | 2023-12-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/2.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -10.5/2.5/089 | OD | Physician | 12/28/2023 | 2023 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18451037&pc=QCB | 2023826-2024-00105 | 2023-09-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+1.5/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -7.5/+1.5/091 | OD | Physician | 12/08/2023 | 2023 | 46 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18451124&pc=QCB | 2023826-2024-00087 | 2023-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/1.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "REPLACE THE CRYSTAL TO ACHIEVE THE DESIRED ARCH HEIGHT." THE PROBLEM WAS NOT RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.5/1.5/092 | OS | Physician | 12/28/2023 | 2023 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18451220&pc=QCB | 2023826-2024-00101 | 2023-07-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.5/+2.0/082 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -4.5/+2.0/082 | OS | Physician | 12/08/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18451345&pc=QCB | 2023826-2024-00084 | 2023-05-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.5/1.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM #(B)(4). | -3.5/1.5/095 | OS | Physician | 12/28/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18451406&pc=QCB | 2023826-2024-00083 | 2023-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/1.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -6.0/1.0/094 | OD | Physician | 12/27/2023 | 2023 | 49 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18452189&pc=QCB | 2023826-2024-00082 | 2023-11-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.5/1.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON 05-NOV-2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -3.5/1.0/086 | OS | Physician | 12/27/2023 | 2023 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18456608&pc=QCB | 2023826-2024-00037 |  | Injury | STAAR SURGICAL COMPANY | 2024-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -5.0/4.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE. A POSSIBLE LENS EXPLANT IS SCHEDULED. Manufacturer Narrative: A4-A6:UNK. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -5.0/4.0/091 | NA | Physician | 12/28/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18457201&pc=QCB | 2023826-2024-00110 |  | Malfunction | STAAR SURGICAL COMPANY | 2024-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -7.00/2.0/071 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). A PLANNED LENS EXCHANGE DUE TO HIGH VAULT WAS REPORTED . LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -7.00/2.0/071 | OS | Other Health Care Professional | 12/11/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18460692&pc=QCB | 2023826-2024-00056 | 2023-12-05 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -5.00/4.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). ON (B)(6) 2023 THE LENS TORE/ BROKE DURING INJECTION/DELIVERY INTO THE EYE AND ALSO WAS IMPLANTED UPSIDE DOWN. THE LENS WAS IMPLANTED; REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SAME MODEL SIZE AND POWER LENS. THE REPLACEMENT DID NOT RESOLVE THE PROBLEM. SEE MFR# 2023826-2024-00057 FOR REPLACEMENT LENS. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- ONE SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -5.00/4.0/090 | OS | Other Health Care Professional | 12/08/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18460744&pc=QCB | 2023826-2024-00057 | 2023-12-05 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -5.00/4.0/091 (SPHERE/CYLINDER/AXIS) LENS WAS BEING IMPLANTED AS A REPLACEMENT LENS AND THE LENS TORE. ON (B)(6) 2023 THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY. NO NEW LENS WAS IMPLANTED. SEE MFR# 2023826-2024-00056 FOR INITIAL LENS. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- ONE SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -5.00/4.0/091 | NA | Other Health Care Professional | 12/08/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18472164&pc=QCB | 2023826-2024-00114 | 2023-11-30 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON 30-NOV-2023 A 13.2MM, VTICM5_13.2, -9.5/2.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE LEFT EYE (OS). THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. THERE WAS NO PATIENT INJURY. ON (B)(6) 2023 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -9.5/2.0/083 | OS | Physician | 12/14/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18472191&pc=QCB | 2023826-2024-00115 | 2023-11-22 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2023 A 13.2MM, VTICM5_13.2, -5.5/1.5/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS NOTED TO BE BROKEN AFTER INJECTION/DELIVERY INTO THE LEFT EYE (OS). THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. THERE WAS NO PATIENT INJURY. ON (B)(6) 2023 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -5.5/1.5/077 | OS | Physician | 12/14/2023 | 2023 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18472297&pc=QCB | 2023826-2024-00158 | 2023-07-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -9.00/1.50/88 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.00/1.50/88 | OS | Physician | 12/14/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18472299&pc=QCB | 2023826-2024-00159 | 2023-08-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -13.0/2.50/88 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.0/2.50/88 | OD | Physician | 12/14/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18472302&pc=QCB | 2023826-2024-00160 | 2023-09-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -7.50/2.50/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -7.50/2.50/89 | OD | Physician | 12/14/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18472316&pc=QCB | 2023826-2024-00156 | 2023-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -8.0/2.0/090 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: A4-A6:UNK. CLAIM# (B)(4). | -8.0/2.0/090 | OD | Physician | 12/14/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18472381&pc=QCB | 2023826-2024-00146 | 2023-11-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0/+2.0/088 INTO THE PATIENT'S RIGHT EYE (OD) ON 02-NOV-2023. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -15.0/+2.0/088 | OD | Physician | 12/11/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18472383&pc=QCB | 2023826-2024-00145 | 2023-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.5/+5.0/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS TWO SIZES SHORTER IN LENGTH DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -15.5/+5.0/090 | OD | Physician | 12/11/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18472385&pc=QCB | 2023826-2024-00143 | 2023-12-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+1.0/110 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -6.5/+1.0/110 | OS | Physician | 12/11/2023 | 2023 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18472386&pc=QCB | 2023826-2024-00144 | 2023-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+0.5/068 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -14.0/+0.5/068 | OD | Physician | 12/11/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18473793&pc=QCB | 2023826-2024-00123 |  | Injury | STAAR SURGICAL COMPANY | 2024-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -13.50/3.5/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.50/3.5/112 | NA | Physician | 12/18/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18473797&pc=QCB | 2023826-2024-00129 | 2021-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.5/4.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021 AND (B)(6) 2022, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS REPLACED FOR A SAME LENGTH SPHERICAL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: A4 - UNK, A5 - UNK, A6 - UNK. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.5/4.0/092 | OS | Physician | 12/19/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18473798&pc=QCB | 2023826-2024-00128 | 2023-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.0/3.0/093 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT(LENS TOUCHED THE EYE). NO PATIENT INJURY. THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.0/3.0/093 | OD | Physician | 12/19/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18473800&pc=QCB | 2023826-2024-00125 | 2023-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.5/4.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS EXPLANTED ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.5/4.0/091 | OD | Physician | 12/19/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18473802&pc=QCB | 2023826-2024-00126 | 2023-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.0/4.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2023. THIS DID NOT RESOLVE THE PROBLEM. THE LENS WAS EXPLANTED ON (B)(6) 2023. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: A4 - UNK A5 - UNK A6 - UNK H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.0/4.0/094 | OS | Physician | 12/19/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18479215&pc=QCB | 2023826-2024-00055 | 2023-11-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF -8.00/3.0/072 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS EYE. EXCESSIVE VAULT WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -8.00/3.0/072 | NA | Other Health Care Professional | 12/13/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18479518&pc=QCB | 2023826-2024-00058 | 2023-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -12.50/2.5/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2023. LOW VAULT WITH LENS ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2023. THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -12.50/2.5/095 | OS | Physician | 12/14/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18479816&pc=QCB | 2023826-2024-00179 | 2023-11-22 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 OF -0.50/4.5/072 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE/BROKE DURING LOADING ON (B)(6) 2023. THE LENS WAS NOT IMPLANTED INTO THE PATIENTS RIGHT EYE (OD). SURGEON USED THE STANDBY LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -0.50/4.5/072 | OD | Physician | 12/14/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18480550&pc=QCB | 2023826-2024-00091 | 2023-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_ 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/0.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK. CLAIM # (B)(4). | -12.0/0.5/087 | OS | Physician | 12/28/2023 | 2023 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18480551&pc=QCB | 2023826-2024-00092 | 2023-11-27 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.0/2.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK . CLAIM # (B)(4). | -4.0/2.0/089 | OD | Physician | 12/28/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18480552&pc=QCB | 2023826-2024-00096 | 2023-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/4.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -10.0/4.0/087 | OD | Physician | 12/30/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18480556&pc=QCB | 2023826-2024-00093 | 2023-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_ 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.5/2.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -4.5/2.0/089 | OS | Physician | 12/28/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18487248&pc=QCB | 2023826-2024-00075 | 2023-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/1.5/081 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK CLAIM #(B)(4). | -12.0/1.5/081 | OS | Physician | 12/24/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18498395&pc=QCB | 2023826-2024-00119 | 2023-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -11.5/3.5/071 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE DID NOT RESOLVE THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -11.5/3.5/071 | OS | Physician | 12/18/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18498405&pc=QCB | 2023826-2024-00212 |  | Malfunction | STAAR SURGICAL COMPANY | 2024-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5/3.5/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023 AS A REPLACEMENT LENS TO ADDRESS LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. PER UPDATED INFORMATION THE PATIENTS VISUAL CONDITION IS GOOD AND MEDICAL OR SURGICAL INTERVENTION IS NOT NECESSARY. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR # 2023826-2024-00119 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -11.5/3.5/087 | OS | Physician | 12/18/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18499297&pc=QCB | 2023826-2024-00198 | 2023-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+1.5/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS TWO SIZES SHORTER IN LENGTH DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.5/+1.5/094 | OD | Physician | 12/17/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18499298&pc=QCB | 2023826-2024-00197 | 2023-10-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. AN EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS OF THE SAME LENGTH AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.0/+2.0/091 | OD | Physician | 12/17/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18506914&pc=QCB | 2023826-2024-00214 | 2023-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -7.5/2.0/055 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO LOW VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -7.5/2.0/055 | OS | Physician | 12/18/2023 | 2023 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18509114&pc=QCB | 2023826-2024-00226 |  | Malfunction | STAAR SURGICAL COMPANY | 2024-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM513.2, -10.50/1.0/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2023. LOW VAULT WAS REPORTED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.50/1.0/099 | OD | Physician | 12/20/2023 | 2023 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18509225&pc=QCB | 2023826-2024-00228 | 2023-12-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.00/1.5/087 (SPHERE/CYLINDER/AXIS) TORE DURING INSERTION PROCESS, USED BACKUP LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -11.00/1.5/087 | NA | Other Health Care Professional | 12/20/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18509883&pc=QCB | 2023826-2024-00203 | 2023-12-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.0/+1.0/074 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4) | -17.0/+1.0/074 | OS | Physician | 12/18/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18509885&pc=QCB | 2023826-2024-00202 | 2023-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+3.0/068 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -14.0/+3.0/068 | OS | Physician | 12/18/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18509886&pc=QCB | 2023826-2024-00219 | 2023-10-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.0/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6)2023 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4) | -11.0/+2.0/094 | OS | Physician | 12/19/2023 | 2023 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18509888&pc=QCB | 2023826-2024-00222 | 2023-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+2.0/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4) | -11.5/+2.0/095 | OD | Physician | 12/21/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18509889&pc=QCB | 2023826-2024-00220 | 2023-12-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.5/011 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON 18-DEC-2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.0/+1.5/011 | OS | Physician | 12/20/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18509890&pc=QCB | 2023826-2024-00223 | 2023-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.0/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4) | -11.0/+2.0/090 | OS | Physician | 12/21/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18509891&pc=QCB | 2023826-2024-00201 | 2023-11-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/+1.0/163 INTO THE PATIENT'S RIGHT EYE (OD) ON 14-NOV-2023. AN EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS OF THE SAME LENGTH AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -5.5/+1.0/163 | OD | Physician | 12/18/2023 | 2023 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18509892&pc=QCB | 2023826-2024-00218 | 2023-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+3.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -7.5/+3.0/093 | OD | Physician | 12/19/2023 | 2023 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18525762&pc=QCB | 2023826-2024-00164 | 2023-12-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -11.0/1.5/001 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED WITH A SAME LENGTH LENS INTRAOPERATIVELY ON (B)(6) 2023 FROM PATIENT'S LEFT EYE (OS) DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. PATIENT CONTACT BUT NO PATIENT INJURY. . THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.0/1.5/001 | OS | Physician | 12/26/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18525776&pc=QCB | 2023826-2024-00172 |  | Malfunction | STAAR SURGICAL COMPANY | 2024-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -11.50/4.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM##(B)(4). | -11.50/4.0/090 | OD | Physician | 12/27/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18525777&pc=QCB | 2023826-2024-00173 |  | Malfunction | STAAR SURGICAL COMPANY | 2024-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -10.5/3.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.5/3.0/090 | OS | Physician | 12/27/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18540475&pc=QCB | 2023826-2024-00289 | 2023-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+3.5/010 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.0/+3.5/010 | OS | Physician | 12/24/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18540476&pc=QCB | 2023826-2024-00288 | 2023-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+2.0/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/+2.0/094 | OS | Physician | 12/24/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18548442&pc=QCB | 2023826-2024-00291 | 2022-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.5/087 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) . CLAIM# (B)(4). | -7.0/+1.5/087 | OD | Physician | 12/24/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18548444&pc=QCB | 2023826-2024-00296 | 2023-12-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+1.5/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.5/+1.5/092 | OS | Physician | 12/25/2023 | 2023 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18558967&pc=QCB | 2023826-2024-00315 | 2023-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+4.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -13.5/+4.0/093 | OD | Physician | 12/26/2023 | 2023 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18558970&pc=QCB | 2023826-2024-00314 | 2023-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+4.5/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. ON (B)(6)2023, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: (B)(4). | -13.0/+4.5/089 | OS | Physician | 12/26/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18558971&pc=QCB | 2023826-2024-00319 | 2023-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+3.0/169 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS OF THE SAME LENGTH AND IMPLANTED VERTICALLY DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -7.0/+3.0/169 | OD | Physician | 12/26/2023 | 2023 | 49 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18558980&pc=QCB | 2023826-2024-00312 | 2023-12-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.5/+1.5/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -3.5/+1.5/090 | OD | Physician | 12/25/2023 | 2023 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18558981&pc=QCB | 2023826-2024-00316 | 2023-10-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.5/110 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -11.0/+1.5/110 | OD | Physician | 12/26/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18558982&pc=QCB | 2023826-2024-00318 | 2023-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+3.5/080 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM#: (B)(4). | -13.0/+3.5/080 | OS | Physician | 12/26/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18559938&pc=QCB | 2023826-2024-00321 | 2023-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+1.5/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON 30-NOV-2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.0/+1.5/091 | OD | Physician | 12/27/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18560037&pc=QCB | 2023826-2024-00094 |  | Malfunction | STAAR SURGICAL COMPANY | 2024-01-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING WITH ROTATION. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4-A6: UNK (B)(4). | -10.0/1.5/091 | OD | Physician | 12/29/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18560094&pc=QCB | 2023826-2024-00234 | 2023-12-07 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1 -11.5/1.0/081 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE/BROKE BEFORE LOADING. IT WAS REPORTED THAT THE SURGEON BELIEVES THAT THE LENS WAS DAMAGED DURING REMOVAL FROM LENS VIAL. THERE WAS NO PATIENT CONTACT. ON (B)(6) 2023 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.5/1.0/081 | NA | Physician | 12/26/2023 | 2023 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18560135&pc=QCB | 2023826-2024-00235 | 2023-11-11 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6 -11.5/3.5/080 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE/BROKE DURING LOADING. THERE WAS NO PATIENT CONTACT. ON (B)(6) 2023 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.5/3.5/080 | NA | Physician | 12/26/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18560331&pc=QCB | 2023826-2024-00362 | 2023-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -14.00/4.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) IN VERTICAL POSITION ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS THAT WAS IMPLANTED IN HORIZONTAL POSITION DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -14.00/4.0/094 | OD | Physician | 12/26/2023 | 2023 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18560345&pc=QCB | 2023826-2024-00363 | 2023-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.5/1.5/073 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) IN VERTICAL POSITION ON (B)(6) 2023. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS THAT WAS IMPLANTED IN HORIZONTAL POSITION DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -11.5/1.5/073 | OS | Physician | 12/26/2023 | 2023 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18560360&pc=QCB | 2023826-2024-00364 | 2023-11-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.50/1.0/073 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023.THE LENS WAS INTRAOPERATIVELY REMOVED DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -9.50/1.0/073 | OD | Physician | 12/26/2023 | 2023 | 49 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18565306&pc=QCB | 2023826-2024-00330 | 2023-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+2.0/075 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+2.0/075 | OD | Physician | 12/28/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18565307&pc=QCB | 2023826-2024-00329 | 2023-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.0/095 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.5/+1.0/095 | OS | Physician | 12/28/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18565308&pc=QCB | 2023826-2024-00328 | 2023-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/075 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.0/+1.5/075 | OS | Physician | 12/28/2023 | 2023 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18565309&pc=QCB | 2023826-2024-00331 | 2023-11-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.5/+0.5/113 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) (B)(4) | -4.5/+0.5/113 | OS | Physician | 12/28/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18565320&pc=QCB | 2023826-2024-00324 | 2023-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+2.0/087 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -14.0/+2.0/087 | OS | Physician | 12/27/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18567447&pc=QCB | 2023826-2024-00398 | 2023-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/087 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -9.5/+1.0/087 | OD | Physician | 12/29/2023 | 2023 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18567453&pc=QCB | 2023826-2024-00400 | 2023-12-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+2.0/068 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. AN EXCESSIVE VAULT WAS NOTED. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LENS OF THE SAME LENGTH AND IT WAS IMPLANTED VERTICALLY. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -7.5/+2.0/068 | OS | Physician | 12/29/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18567454&pc=QCB | 2023826-2024-00397 | 2023-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+1.5/105 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -7.5/+1.5/105 | OS | Physician | 12/29/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18567456&pc=QCB | 2023826-2024-00399 | 2023-12-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2024-01-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.0/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE LENS WAS EXCHANGED ON (B)(6) 2023 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) (B)(4) | -10.5/+1.0/094 | OD | Physician | 12/29/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18567908&pc=QCB | 2023826-2024-00373 |  | Injury | STAAR SURGICAL COMPANY | 2024-01-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -10.5/1.5/088 WAS EXPLANTED FROM THE PATIENT'S EYE DUE TO WRONG SIZE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -10.5/1.5/088 | NA | Other Health Care Professional | 12/28/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18567918&pc=QCB | 2023826-2024-00374 |  | Malfunction | STAAR SURGICAL COMPANY | 2024-01-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -9.50/1.0/091 WAS THE WRONG SIZE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: D6A:UNK. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -9.50/1.0/091 | NA | Other Health Care Professional | 12/28/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18567997&pc=QCB | 2023826-2024-00375 | 2023-11-21 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-01-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -6.00/2.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED UPSIDE DOWN INTO THE PATIENTS RIGHT EYE (OD). THE LENS WAS INTRAOPERATIVELY REMOVED AND REPLACED AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -6.00/2.5/093 | OD | Other Health Care Professional | 12/29/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18804187&pc=QCB | 2023826-2024-00925 |  | Malfunction | STAAR SURGICAL COMPANY | 2024-02-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -10.0/1.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023 AS A REPLACEMENT LENS TO ADDRESS LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. SEE MFR # 2023826-2023-05812 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | -10.0/1.0/091 | OD | Physician | 11/28/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18821187&pc=QCB | 2023826-2024-00959 | 2023-10-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2024-03-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/2.5/077 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE PATIENT EXPERIENCED LOW VAULTING WITHOUT ROTATION. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: A4: A6: UNK. H6: DEVICE CODE: 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #733907. | -12.5/2.5/077 | OD | Physician | 10/21/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=18947108&pc=QCB | 2023826-2023-05384 | 2023-09-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2024-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -11.50/2.0/022 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. LENS ROTATION WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6-DEVICE CODE 1494: OFF-LABEL USE : PRE -EXISTING: PROGRESSIVE MYOPIA. CLAIM# (B)(4). | -11.50/2.0/022 | OD | Physician | 11/06/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10511686&pc=QCB | 2023826-2020-01991 | 2017-10-23 04:00 | Injury | STAAR SURGICAL COMPANY AG | 2020-09-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -18.50/+3.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS RESIDUE ON LENS. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.50/+3.0/088 | OD |  | 08/14/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11087582&pc=QCB | 2023826-2020-03098 | 2008-11-07 05:00 | Injury | STAAR SURGICAL COMPANY AG | 2020-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A TICM120V4, -11.0/6.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2008. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B-5: CORRECTED TO ADD "THE LENS WAS REPOSITIONED FOUR TIMES, BUT THE PROBLEM WAS NOT RESOLVED." CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/6.0/096 | OD |  | 12/09/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11405570&pc=QCB | 2023826-2021-00563 | 2021-02-02 05:00 | Injury | STAAR SURGICAL COMPANY AG | 2021-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | No Clinical Signs, Symptoms or Conditions | NA#S-P#NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -15.50/+1.0/156, (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2010. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: B2 IS NOT APPLICABLE IN INITIAL MDR CLAIM # (B)(4). Manufacturer Narrative: B1: ADVERSE EVENT B5 - THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS; -15.50/+1.0/156 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2010. THE PATIENT EXPERIENCED REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED ON 04/01/2022 FOR A DIFFERENT MODEL/ LONGER LENGTH LENS OF A DIFFERENT DIOPTER. THIS RESOLVED THE PROBLEM. H10: SERIOUS INJURY CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN (B)(6) BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -15.50/+1.0/156 | OD | Physician | 02/10/2021 | 2021 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13502421&pc=QCB | 2023826-2022-00409 | 2021-07-14 04:00 | Injury | STAAR SURGICAL COMPANY AG | 2022-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -22.50/+4.0/109 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -22.50/+4.0/109 | OD | Physician | 01/11/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=4479976&pc=QCB | 2023826-2015-00122 | 2014-12-23 | Injury | STAAR SURGICAL COMPANY | 2015-02-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2014. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS NOTED ON (B)(6) 2014. THE LENS WAS REPOSTIONED ON (B)(6) 2014. ON PATIENT'S LAST VISIT TO THE DOCTOR'S OFFICE ON (B)(6) 2014, UCVA: 20/80. BCVA: 20/30. TWO OR THREE DAYS AFTER REPOSTIONING THE LENS, IT ROTATED AGAIN. THE ICL REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). (OTHER): LENS REMAINS IMPLANTED. EVALUATION CODES: METHOD - (OTHER): WORK ORDER SEARCH RESULTS - (OTHER): A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINTS WERE FOUND WITHIN THE WORK ORDER. CONCLUSIONS (NO CONCLUSION CAN BE DRAWN): BASED ON THE COMPLAINT HISTORY AND WORK ORDER SEARCH, A SPECIFIC ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED. (B)(4). Manufacturer Narrative: THIS IS A RESUBMISSION TO CORRECT MFR # FOR SUPP# 1 IN WHICH THE MFR NUMBER WAS REPORTED INCORRECT. CLAIM# (B)(4). Manufacturer Narrative: UPDATED INFO: ON (B)(4) 2015 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS AND THE PROBLEM IS RESOLVED. (B)(4). | -15.5/+2.5/064 | OS | Physician | 01/06/2015 | 2015 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6205829&pc=QCB | 2023826-2016-01509 | 2016-06-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2016-12-27 05:00 | QCB | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -15.0/+2.0/85 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WEIGHT: UNK. (B)(4). WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). THIS IS A RESUBMISSION OF THE INITIAL MDR PER FDA REQUEST. | -15.0/+2.0/85 | OS | Physician | 09/27/2016 | 2016 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6323476&pc=QCB | 2023826-2016-01551 | 2016-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2017-02-10 05:00 | QCB | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -17.0/+3.0/070 (SPHERE/CYLINDER/AXIS) DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S UNCORRECTED VISUAL ACUITY WAS 20/32. Manufacturer Narrative: PMA/510(K): NA THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). THIS IS A RESUBMISSION OF THE INITIAL MDR PER FDA REQUEST. | -17.0/+3.0/070 | OS | Physician | 10/13/2016 | 2016 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6339288&pc=QCB | 2023826-2016-01631 | 2016-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2017-02-17 05:00 | QCB | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -6.0/+4.0/+150 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PMA/510(K): NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). THIS IS A RESUBMISSION OF THE INITIAL MDR PER FDA REQUEST. | -6.0/+4.0/+150 | OS | Physician | 10/26/2016 | 2016 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6339374&pc=QCB | 2023826-2016-01665 | 2016-10-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2017-02-17 05:00 | QCB | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.0/120 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). THIS IS A RESUBMISSION OF THE INITIAL MDR PER FDA REQUEST. | -11.00/+1.0/120 | OD | Physician | 10/31/2016 | 2016 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6343294&pc=QCB | 2023826-2016-01279 | 2015-11-12 | Injury | STAAR SURGICAL COMPANY | 2017-02-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7 MM VTICMO 13.7 IMPLANTABLE COLLAMER LENS -11.0/+2.0/108 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2015 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION - THE LENS WAS RETURNED DRY IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE NO VISIBLE DAMAGE TO LENS. CORRECTED DATA: (B)(4). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). REPORT RESUBMITTED PER FDA REQUEST. | -11.0/+2.0/108 | OS | Physician | 08/23/2016 | 2016 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6343302&pc=QCB | 2023826-2016-01293 | 2016-03-03 | Injury | STAAR SURGICAL COMPANY | 2017-02-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -10.5/+1.5/075 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION - THE LENS WAS RETURNED IN LIQUID IN CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS TO HAVE FOREIGN MATERIAL ON LENS SURFACE (BLOOD). CONCLUSION: THE ANTERIOR ENDOTHELIUM CHAMBER DEPTH (ACD) IS BELOW 3.0MM AND PER DFU FOR THIS LENS MODEL, THIS LENS MODEL IS CONTRAINDICATED FOR ACD BELOW 3.0MM. CORRECTED DATA: (B)(4). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). REPORT RESUBMITTED PER FDA REQUEST. | -10.5/+1.5/075 | OS | Physician | 08/24/2016 | 2016 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6344271&pc=QCB | 2023826-2016-01295 | 2016-03-24 | Injury | STAAR SURGICAL COMPANY | 2017-02-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -10.5/1.0/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION - THE LENS WAS RETURNED IN LIQUID IN CASE/VIAL. CORRECTED DATA: (B)(4).. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/1.0/082 | OD | Physician | 08/24/2016 | 2016 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6349340&pc=QCB | 2023826-2016-01505 | 2016-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2017-02-22 05:00 | QCB | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS - 10.0/+2.0/96 DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: 709548. THIS IS A RESUBMISSION OF THE INITIAL MDR PER FDA REQUEST. | - 10.0/+2.0/96 | OS | Physician | 09/27/2016 | 2016 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6373202&pc=QCB | 2023826-2016-01291 |  | Injury | STAAR SURGICAL COMPANY | 2017-03-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -13.5/+2.5/094 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: LENS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION PERFORMED AND OBSERVED NO VISABLE DAMAGED TO LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: (B)(4).. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). REPORT RESUBMITTED PER FDA REQUEST. | -13.5/+2.5/094 | OS | Physician | 08/09/2016 | 2016
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6439519&pc=QCB | 2023826-2016-01572 |  | Injury | STAAR SURGICAL COMPANY | 2017-03-28 04:00 | QCB | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -13.5/+1.5/141 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO A SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/25. Manufacturer Narrative: NA- THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). THIS IS A RESUBMISSION OF THE INITIAL MDR PER FDA REQUEST. | -13.5/+1.5/141 | OD | Physician | 10/21/2016 | 2016 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6453248&pc=QCB | 2023826-2016-01554 | 2015-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2017-04-03 04:00 | QCB | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -7.50/+1.5/059 (SPHERE/CYLINDER/AXIS) DIOPTER, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO EXCESSIVE VAULTING AND MYDRIATIC PUPIL. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP UNCORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: PMA/510(K): NA.THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). THIS IS A RESUBMISSION OF THE INITIAL MDR PER FDA REQUEST. | -7.50/+1.5/059 | OS | Physician | 10/13/2016 | 2016 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6454037&pc=QCB | 2023826-2016-01630 | 2016-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2017-04-03 04:00 | QCB | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -7.00/+5.5/+016 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PMA/510(K): NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). THIS IS A RESUBMISSION OF THE INITIAL MDR PER FDA REQUEST. | -7.00/+5.5/+016 | OD | Physician | 10/26/2016 | 2016 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=6519186&pc=QCB | 2023826-2016-01555 | 2016-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2017-04-25 04:00 | QCB | ICL (IMPLANTABLE COLLAMER LENS) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -10.00/+1.0/092 DIOPTER, IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2016 DUE TO LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT'S POST-OP BEST CORRECTED VISUAL ACUITY WAS 20/20. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). THIS IS A RESUBMISSION OF THE INITIAL MDR PER FDA REQUEST. | -10.00/+1.0/092 | OD | Physician | 10/14/2016 | 2016 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=7906913&pc=QCB | 2023826-2018-01438 | 2018-06-21 | Injury | STAAR SURGICAL COMPANY | 2018-09-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.5/+2.0/ 83 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WAS OBSERVED, THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENS ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.5/+2.0/ 83 | OD | Physician | 09/18/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8090050&pc=QCB | 2023826-2018-01694 | 2018-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -6.5/+1.0/105 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO OVER VAULTING. THE PROBLEM WAS RESOLVED. "NO PATIENT INJURY" IS REPORTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN THE LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). | -6.5/+1.0/105 | OS |  | 10/25/2018 | 2018 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8090056&pc=QCB | 2023826-2018-01693 | 2018-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2018-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM TMICL13.2 LENS, -8.5/+2.0/111 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO OVER VAULTING. THE PROBLEM WAS RESOLVED. "NO PATIENT INJURY" IS REPORTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND TEARS IN THE HAPTIC. (B)(4). | -8.5/+2.0/111 | OD |  | 10/25/2018 | 2018 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8215231&pc=QCB | 2023826-2019-00022 | 2018-11-13 | Injury | STAAR SURGICAL COMPANY | 2019-01-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+2.0/085 | OD | Physician | 12/10/2018 | 2018 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8250946&pc=QCB | 2023826-2019-00116 | 2018-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICMO12.6, -14.0/3.0/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMR LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGED RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.0/3.0/086 | OD |  | 12/19/2018 | 2018 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8254840&pc=QCB | 2023826-2019-00094 | 2018-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/+6.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REPOSITIONED, THE SURGEON NOTING LOW VAULT AND LENS ROTATION. REPORTEDLY, THE LENS RE-ROTATED ON (B)(6) 2018. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -16.5/6.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. LOW VAULT WITH LENS ROTATION WAS NOTED ON (B)(6) 2018. THE LENS WAS REPOSITIONED ON (B)(6) 2018 BUT THE ROTATION REEMERGED ON (B)(6) 2018. THE LENS REMAINS IMPLANTED. CLAIM#:(B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | -16.5/+6.0/087 | OD |  | 12/20/2018 | 2018 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8256522&pc=QCB | 2023826-2019-00041 | 2018-04-08 | Injury | STAAR SURGICAL COMPANY | 2019-01-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.00/1.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: DATE SHOULD BE CORRECTED TO 02-JAN-2019 IN THE ORIGINAL MDR. COMMON DEVICE NAME SHOULD BE CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/1.5/097 | OD | Physician | 01/02/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8256529&pc=QCB | 2023826-2019-00042 | 2018-05-20 | Injury | STAAR SURGICAL COMPANY | 2019-01-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -11.00/1.0/099 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: DATE CORRECTED TO 02-JAN-2019 IN THE ORIGINAL MDR. COMMON DEVICE NAME CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/1.0/099 | OD | Physician | 01/02/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8256524&pc=QCB | 2023826-2019-00067 | 2018-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/2.0/095 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/2.0/095 | OS |  | 01/02/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8270742&pc=QCB | 2023826-2019-00176 | 2018-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -17.50/3.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. REPOSITIONING AND PROBLEM RESOLVED WAS REPORTED. IT WAS ALSO REPORTED THAT THE LENS WAS REMOVED AND EXCHANGED FOR A NON TORIC LENS WITHIN THE SAME SURGERY. THIS EXCHANGE RESOLVED THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICL RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THAT A 13.2MM VTICM13.2, -17.50/3.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. REPOSITIONING AND PROBLEM RESOLVED WAS REPORTED. ON (B)(6) 2018 THE LENS WAS REMOVED, AND ON (B)(6) 2018 THE LENS WAS REPLACED WITH A NON TORIC LENS AND THE PROBLEM WAS RESOLVED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. (EXPLANTED DATE): UNK TO (B)(6) 2018 CLAIM# (B)(4). Manufacturer Narrative: PROCODE SHOULD BE CORRECTED TO "PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB" IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -17.50/3.0/081 | OD |  | 12/24/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8271129&pc=QCB | 2023826-2019-00172 | 2017-11-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICMO12.6, -15.50/3.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO OBSERVATION OF LOW VAULT WITH ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: CORRECTED TO: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.50/3.0/086 | OS |  | 12/26/2018 | 2018 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8272685&pc=QCB | 2023826-2019-00167 | 2018-05-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/118 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PATIENT EXPERIENCED MENTAL SICKNESS. ANOTHER ICL LENS WAS NOT IMPLANTED. THE REPORTER INDICATED THE EVENT WAS DUE TO PATIENT RELATED FACTOR AND WAS NOT PRODUCT RELATED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+1.5/118 | OD |  | 12/26/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8272700&pc=QCB | 2023826-2019-00154 | 2018-12-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/098 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CORRECTED DATA: DATE CORRECTED TO (B)(6) 2018 IN ORIGINAL MDR. COMMON DEVICE NAME CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+1.5/098 | OD |  | 12/28/2018 | 2018 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8274015&pc=QCB | 2023826-2019-00173 | 2018-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM, VTICM5_12.1, -12.00/2.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO OBSERVATION OF LOW VAULT WITH ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: CORRECTED TO: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/2.0/082 | OS |  | 12/26/2018 | 2018 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8274412&pc=QCB | 2023826-2019-00174 | 2018-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.00/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CORRECTED DATA: COMMON DEVICE NAME CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/2.0/089 | OD |  | 12/26/2018 | 2018 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8276360&pc=QCB | 2023826-2019-00168 | 2018-05-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+1.0/075 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO PATIENT EXPERIENCED MENTAL SICKNESS. ANOTHER ICL LENS WAS NOT IMPLANTED. THE REPORTER INDICATED THE EVENT WAS DUE TO PATIENT RELATED FACTOR AND WAS NOT PRODUCT RELATED. Manufacturer Narrative: CORRECTED DATA: METHOD CODE 4110 ADDED RESULT CODE 213 ADDED RESULT CODE 114: PREVIOUS CODE NOT APPLICABLE CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOSITURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+1.0/075 | OS |  | 12/26/2018 | 2018 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8276370&pc=QCB | 2023826-2019-00155 | 2018-10-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.0/+3.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.0/+3.0/091 | OS |  | 12/26/2018 | 2018 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8282060&pc=QCB | 2023826-2019-00164 | 2018-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -10.00/2.5/105 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS.(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/2.5/105 | OD |  | 01/02/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8282073&pc=QCB | 2023826-2019-00184 | 2018-09-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.00/2.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -18.00/2.5/095 | OD |  | 01/10/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8287268&pc=QCB | 2023826-2019-00185 | 2018-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-01-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICM5_13.2, -9.00/1.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/1.0/089 | OD |  | 08/08/2018 | 2018 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8296783&pc=QCB | 2023826-2019-00134 | 2018-11-28 | Injury | STAAR SURGICAL COMPANY | 2019-02-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/1.0/070 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/1.0/070 | OS | Physician | 01/10/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8296785&pc=QCB | 2023826-2019-00137 | 2018-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/1.5/095 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS. COMMON DEVICE NAME SHOULD BE CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS IN ORIGINAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.50/1.5/095 | OD |  | 01/10/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8307833&pc=QCB | 2023826-2019-00223 | 2018-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED; THE STATUS OF THE EYE IS REPORTED AS "QUIET" AND "STABLE." Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CORRECTED DATA: COMMON DEVICE NAME CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.5/089 | OD |  | 01/10/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8322856&pc=QCB | 2023826-2019-00204 | 2018-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+2.0/106 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. COMMON DEVICE NAME SHOULD BE CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS IN ORIGINAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.0/+2.0/106 | OD |  | 01/10/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8335870&pc=QCB | 2023826-2019-00234 | 2018-12-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.50/2.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/2.0/094 | OS |  | 02/06/2019 | 2019 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8335871&pc=QCB | 2023826-2019-00233 | 2018-12-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.00/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/2.0/090 | OD |  | 02/06/2019 | 2019 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8337068&pc=QCB | 2023826-2019-00243 | 2019-01-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+3.5/102 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+3.5/102 | OD |  | 01/14/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8337592&pc=QCB | 2023826-2019-00251 |  | Injury | STAAR SURGICAL COMPANY | 2019-02-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.5/+2.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -8.5/2.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD). LENS ROTATION NOT ASSOCIATED TO LOW VAULT WAS REPORTED. THE LENS WAS REPOSITIONED BUT IT DID NOT RESOLVE THE ISSUE. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER PROVIDED CONFLICTING INCIDENT, IMPLANT, AND EXPLANT DATES. FOLLOW-UPS TO CLARIFY DID NOT YIELD ANY ADDITIONAL INFORMATION. COMMON DEVICE NAME SHOULD BE CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS IN ORIGINAL MDR. DATE IN ORIGINAL MDR SHOULD BE CORRECTED TO UNK. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+2.0/091 | OD |  | 01/14/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8355980&pc=QCB | 2023826-2019-00283 | 2018-08-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, - 8.5/+1.5/078 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS QUIET. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND DEFORMED HAPTICS. CORRECTED DATA: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | - 8.5/+1.5/078 | OS |  | 12/21/2018 | 2018 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8365952&pc=QCB | 2023826-2019-00269 | 2019-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.00/4.0/050 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019 DUE TO EXCESSIVE VAULT. ON (B)(6) 2019 A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.00/4.0/050 | OS |  | 02/13/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372291&pc=QCB | 2023826-2019-00304 | 2018-05-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -13.5/+3.0/062 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS BUT DIFFERENT AXIS POWER AND THE PROBLEM WAS RESOLVED. THE EXCHANGED LENS WAS STILL SLIGHTLY OFF-AXIS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.5/+3.0/062 | OS |  | 01/28/2019 | 2019 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372521&pc=QCB | 2023826-2019-00321 | 2018-02-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION WHICH DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED AND RESIDUE ON THE LENS. CORRECTED DATA: COMMON DEVICE NAME SHOULD BE CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS IN ORIGINAL MDR. PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS REPOSITIONING, LENS EXCHANGE). CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -12.0/+2.5/090 | OS |  | 02/01/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8372533&pc=QCB | 2023826-2019-00323 | 2019-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.5/+3.0/094 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS. CORRECTED DATA: COMMON DEVICE NAME SHOULD BE CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+3.0/094 | OD |  | 02/01/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8382860&pc=QCB | 2023826-2019-00343 | 2018-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -13.5/1.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATED THAT THE VAULT BECAME NORMAL. AC IS DEEP AND PATIENT IS DOING GOOD. Manufacturer Narrative: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.5/1.0/090 | OD |  | 01/30/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8409305&pc=QCB | 2023826-2019-00382 | 2019-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -16.0/+3.0/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE STATUS OF THE EYE IS "GOOD" AFTER EXPLANT. THE CAUSE OF THE EVENT IS REPORTED AS "SIZING." Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -16.0/+3.0/083 | OS |  | 02/14/2019 | 2019 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8409320&pc=QCB | 2023826-2019-00384 | 2019-02-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -17.5/+1.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO LOW VAULT AND LENS ROTATION. THE CAUSE OF THE EVENT IS REPORTED AS ZONULAR DEHISCIENCE. AFTER EXPLANT, THE "EYE IS NORMAL." ADDITIONALLY, IT'S REPORTED THAT "TORIC ICL WAS DECENTERED DUE TO ZONULAR DEHISCIENCE. REMOVED DE-CENTERED LENS AND LEFT THE PATIENT WITHOUT DOING ANY FURTHER PROCEDURE." Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -17.5/+1.5/090 | OD |  | 02/14/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8422932&pc=QCB | 2023826-2019-00403 | 2018-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+2.5/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+2.5/074 | OS |  | 02/12/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8422940&pc=QCB | 2023826-2019-00404 | 2018-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION- THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+1.5/082 | OD |  | 02/12/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8422952&pc=QCB | 2023826-2019-00406 | 2018-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.5/088 | OD |  | 02/12/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8422961&pc=QCB | 2023826-2019-00405 | 2018-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+1.5/083 | OD |  | 02/12/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8426409&pc=QCB | 2023826-2019-00377 | 2019-01-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICMO12.6, -18.0/4.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2018. ON (B)(6) 2019, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CORRECTED DATA: COMMON DEVICE NAME SHOULD BE CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS IN ORIGINAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -18.0/4.5/082 | OS |  | 02/14/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8434585&pc=QCB | 2023826-2019-00433 | 2019-01-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. A PORTION OF THE LENS IS MISSING. CORRECTED DATA: COMMON DEVICE NAME SHOULD BE CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS IN ORIGINAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.5/084 | OD |  | 02/17/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8437599&pc=QCB | 2023826-2019-00444 | 2019-02-17 | Injury | STAAR SURGICAL COMPANY | 2019-03-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM513.2, -9.50/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) /2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.5/089 | OS | Physician | 02/19/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8437899&pc=QCB | 2023826-2019-00450 | 2019-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/081 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON THE SAME DATE, IN A 2ND SURGERY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+2.0/081 | OD |  | 02/22/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8437906&pc=QCB | 2023826-2019-00449 | 2019-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -13.0/+1.0/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON THE SAME DATE, IN A 2ND SURGERY THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -13.0/+1.0/085 | OD |  | 02/22/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8437912&pc=QCB | 2023826-2019-00453 | 2019-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM, VTICMO_12.6 IMPLANTABLE COLLAMER LENS, -14.5/+2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS DUE TO LENS ROTATION. THE PROBLEM IS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -14.5/+2.0/092 | OD |  | 02/22/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8438494&pc=QCB | 2023826-2019-00434 | 2019-01-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 11.6MM VTICH11.6 IMPLANTABLE COLLAMER LENS, -7.50/+5.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.50/+5.5/092 | OD |  | 02/18/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8442012&pc=QCB | 2023826-2019-00437 | 2019-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, --15.0/+4.0/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. A SHORTER LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | --15.0/+4.0/084 | OS |  | 02/19/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8452667&pc=QCB | 2023826-2019-00496 | 2018-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -13.50/3.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGE WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.50/3.5/090 | OS |  | 02/27/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8452962&pc=QCB | 2023826-2019-00497 | 2018-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1, -11.5/+3.0/81 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL DATA DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND PRESENCE OF RESIDUE ON LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/+3.0/81 | OS |  | 02/27/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8456190&pc=QCB | 2023826-2019-00502 | 2019-01-24 | Injury | STAAR SURGICAL COMPANY | 2019-03-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM TMICL13.7, -10.0/3.5/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. . Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. RETURNED INSIDE A ZIPLOC BAG. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). | -10.0/3.5/93 | OD | Other Health Care Professional | 02/27/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8456719&pc=QCB | 2023826-2019-00498 | 2019-01-24 | Injury | STAAR SURGICAL COMPANY | 2019-03-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM TMICL13.7, -10.5/4.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. RETURNED INSIDE A ZIPLOC BAG. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). | -10.5/4.0/081 | OS | Other Health Care Professional | 02/27/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8457838&pc=QCB | 2023826-2019-00474 | 2019-02-21 | Injury | STAAR SURGICAL COMPANY | 2019-03-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM TMICL13.7 IMPLANTABLE COLLAMER LENS, -14.0/+1.5/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED DUE TO SIZE. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: UNK. DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -14.0/+1.5/087 | OD | Other Health Care Professional | 02/25/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8461320&pc=QCB | 2023826-2019-00461 | 2019-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -3.50/3.0/034 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WAS OBSERVED, THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR AND DEVICE. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -3.50/3.0/034 | OS |  | 03/13/2019 | 2019 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8461401&pc=QCB | 2023826-2019-00468 | 2019-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -12.50/2.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON 13/MAR/2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION- THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY- UNK. RACE- UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.50/2.5/094 | OD |  | 03/19/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8466193&pc=QCB | 2023826-2019-00515 | 2019-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -8.0/+6.0/081 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+6.0/081 | OD |  | 03/01/2019 | 2019 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8469942&pc=QCB | 2023826-2019-00459 | 2019-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -6.00/2.0/122 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN RESIDUE ON LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.00/2.0/122 | OD |  | 03/05/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8475387&pc=QCB | 2023826-2019-00522 | 2019-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+3.0/113 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: SURGEON NAME: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.5/+3.0/113 | OD |  | 03/08/2019 | 2019 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8475407&pc=QCB | 2023826-2019-00521 | 2018-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/062 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 WAS EXCHANGED WITH A LONGER LENS DUE TO LENS ROTATION. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS VIAL WITH CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.0/+1.5/062 | OS |  | 03/07/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8475408&pc=QCB | 2023826-2019-00526 | 2019-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+2.5/084 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.0/+2.5/084 | OS |  | 03/08/2019 | 2019 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8479552&pc=QCB | 2023826-2019-00535 | 2017-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICH13.2, -16.00/1.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGE FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.00/1.0/075 | OD |  | 03/06/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8480698&pc=QCB | 2023826-2019-00547 | 2018-11-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -11.0/+0.5/120 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+0.5/120 | OD |  | 03/11/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8480709&pc=QCB | 2023826-2019-00548 | 2018-10-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -4.5/+1.0/090 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -4.5/+1.0/090 | OS |  | 03/11/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8480751&pc=QCB | 2023826-2019-00540 | 2019-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.00/3.5/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2016. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS RETURNED DRY WITH RESIDUE IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND: NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/3.5/087 | OS |  | 03/08/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8485060&pc=QCB | 2023826-2019-00549 | 2018-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.0/+1.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TEAR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.0/+1.5/089 | OS |  | 03/11/2019 | 2019 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8485082&pc=QCB | 2023826-2019-00550 | 2018-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.0/+1.0/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY WITH CLEAR SURGICAL RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN, BENT AND DEFORMED WITH FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.0/+1.0/083 | OD |  | 03/11/2019 | 2019 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8489331&pc=QCB | 2023826-2019-00553 | 2018-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/3.0/093 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS) ON (B)(6)2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.0/3.0/093 | OS |  | 03/11/2019 | 2019 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8489352&pc=QCB | 2023826-2019-00556 | 2018-07-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/1.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/1.5/088 | OS |  | 03/11/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8489361&pc=QCB | 2023826-2019-00554 | 2018-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/1.5/088 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/1.5/088 | OD |  | 03/11/2019 | 2019 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8492908&pc=QCB | 2023826-2019-00557 | 2018-07-25 | Injury | STAAR SURGICAL COMPANY | 2019-04-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/1.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/1.0/089 | OD | Physician | 03/11/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8492651&pc=QCB | 2023826-2019-00545 | 2018-09-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-08 04:00 | QCB | IMPLANTABLE COLLAME LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.00/1.5/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.00/1.5/095 | OS |  | 03/11/2019 | 2019 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8492672&pc=QCB | 2023826-2019-00544 | 2018-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -4.50/1.5/084 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT:UNK. ETHNICITY:UNK. RACE:UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -4.50/1.5/084 | OD |  | 03/11/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8492677&pc=QCB | 2023826-2019-00546 | 2018-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.00/ 0.5/ 106 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.00/ 0.5/ 106 | OD |  | 03/11/2019 | 2019 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8492805&pc=QCB | 2023826-2019-00543 | 2018-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -4.00/1.0/114 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULT. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -4.00/1.0/114 | OS |  | 03/11/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8497513&pc=QCB | 2023826-2019-00579 | 2019-01-09 | Injury | STAAR SURGICAL COMPANY | 2019-04-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.0/+2.0/069 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.0/+2.0/069 | OS | Physician | 03/15/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8497481&pc=QCB | 2023826-2019-00582 | 2019-03-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Failure to Unfold or Unwrap; Adverse Event Without Identified Device or Use Problem | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -13.5/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WOULD NOT OPEN/UNFOLD PROPERLY DURING INJECTION INTO THE RIGHT EYE (OD). THE INCISION WAS WIDENED AND SUTURES WERE USED. THE LENS WAS IMPLANTED, REMOVED AND EXCHANGE WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE REPORTER STATED "NO PROBLEMS" FOR PATIENTS CURRENT CONDITION. Manufacturer Narrative: DEVICE CODE 2993 IN PREVIOUS MDR IS NOT APPLICABLE. (B)(4). Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE EVALUATION: LENS WAS RETURNED STUCK IN CARTRIDGE. THE CARTRIDGE WAS RETURNED IN LENS BOX. THERE WAS RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE DEVICE AND RESIDUE ON CARTRIDGE. CLAIM # (B)(4). | -13.5/2.0/093 | OD |  | 03/13/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8497494&pc=QCB | 2023826-2019-00575 | 2019-02-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -12.0/+1.5/091 | OS |  | 03/15/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8497499&pc=QCB | 2023826-2019-00570 | 2018-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -11.0/+1.5/087 | OS |  | 03/15/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8497500&pc=QCB | 2023826-2019-00572 |  | Injury | STAAR SURGICAL COMPANY | 2019-04-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -11.5/+1.5/068 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS DUE TO LOW VAULT AND LENS ROTATION. IT IS NOT REPORTED WHETHER THE PROBLEM WAS RESOLVED HOWEVER, OPERATIVE REPORT STATES THE PATIENT WAS TRANSFERRED TO THE RECOVERY ROOM "IN GOOD POST-OPERATIVE CONDITION." Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN THE LENS VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.5/+1.5/068 | OS |  | 03/14/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8516309&pc=QCB | 2023826-2019-00624 | 2018-06-21 | Injury | STAAR SURGICAL COMPANY | 2019-04-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM,VTICMO13.2, -8.5/1.0/078 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/1.0/078 | OS | Physician | 03/21/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8515442&pc=QCB | 2023826-2019-00620 | 2018-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -9.00/1.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGE FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/1.5/083 | OS |  | 03/20/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8515952&pc=QCB | 2023826-2019-00621 | 2018-12-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM,VTICMO13.2, -10.5/2.0/092 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/2.0/092 | OD |  | 03/20/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8516055&pc=QCB | 2023826-2019-00622 | 2018-09-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM,VTICMO13.2, -16.5/1.0/088 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.5/1.0/088 | OD |  | 03/21/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8516164&pc=QCB | 2023826-2019-00623 | 2018-09-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM,VTICMO12.6, -17.0/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -17.0/2.0/095 | OS |  | 03/21/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8518157&pc=QCB | 2023826-2019-00560 | 2019-03-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.50/2.5/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -13.50/2.5/073 | OS |  | 04/02/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8518182&pc=QCB | 2023826-2019-00566 | 2019-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.00/3.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -13.00/3.0/090 | OD |  | 04/02/2019 | 2019 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8518993&pc=QCB | 2023826-2019-00614 | 2019-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.0/+2.0/091 | OD |  | 03/22/2019 | 2019 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8518996&pc=QCB | 2023826-2019-00616 | 2018-10-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE SURGEON REPORTS THE LENS ROTATED TWICE TO 90 DEGREES ROTATION. ON (B)(6) 2018 THE LENS HAD BEEN REPOSITIONED, WHICH DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SAME SIZE, DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, HOWEVER THE SURGEON STATES, "COULD BE ZONULAR WEAKNESS." Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.0/+1.5/085 | OD |  | 03/22/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8525040&pc=QCB | 2023826-2019-00627 | 2019-01-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.5+3.0/170 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO LOW VAULT WITH ROTATION. A REPLACEMENT LENS OF LONGER LENGTH WAS LATER IMPLANTED IN A SECOND SURGERY. STATUS OF THE EYE WAS REPORTED AS "OK". IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS USER ERROR/OTHER: SMALL LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5+3.0/170 | OS |  | 03/20/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8525169&pc=QCB | 2023826-2019-00604 | 2018-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+1.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.0/+1.5/090 | OS |  | 03/22/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8525173&pc=QCB | 2023826-2019-00603 | 2018-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/098 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.5/+2.0/098 | OD |  | 03/22/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8525193&pc=QCB | 2023826-2019-00605 | 2019-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+1.0/103 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. A SHORTER LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+1.0/103 | OD |  | 03/22/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8525205&pc=QCB | 2023826-2019-00606 | 2019-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.5/+3.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2019 BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.5/+3.5/089 | OS |  | 03/22/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8545623&pc=QCB | 2023826-2019-00671 | 2019-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT 13.2MM, VTICM5_13.2, -11.50/1.0/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID; IN A LENS VIAL. VISUAL INSPECTION FOUND THE HAPTIC HAS A TEAR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.50/1.0/100 | OD |  | 03/27/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8552444&pc=QCB | 2023826-2019-00655 | 2018-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.0/+2.0/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS EXPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.0/+2.0/086 | OS |  | 03/29/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8552446&pc=QCB | 2023826-2019-00656 | 2019-01-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO_12.6 IMPLANTABLE COLLAMER LENS, -15.0/+5.5/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT AXIS LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -15.0/+5.5/097 | OD |  | 03/29/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8552492&pc=QCB | 2023826-2019-00658 | 2019-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+2.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LENS ROTATION; THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/+2.5/090 | OS |  | 03/29/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8552973&pc=QCB | 2023826-2019-00665 | 2019-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1 MM, TMICL12.1, -9.0/3.5/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. IT WAS REPORTED THAT THIS EXCHAGE RESOLVED THE PROBLEM. REPORTEDLY, REPOSITIONING WAS NOT PERFORMED. HOWEVER, IT WAS REPORTED THAT THE REPOSITIONING RESOLVED THE PROBLEM. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.0/3.5/070 | OS |  | 03/27/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8556850&pc=QCB | 2023826-2019-00669 | 2019-01-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -8.5/1.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXCHANGED ON (B)(6) 2019 DUE TO LOW VAULT. IT WAS REPORTED THAT THIS EXCHANGE DID NOT RESOLVE THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED FOR THIS CASE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4). Manufacturer Narrative: "EXCHANGE DID NOT RESOLVE THE PROBLEM" IN ORIGINAL MDR IS NOT APPLICABLE. THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -8.5/1.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018 . THE LENS WAS EXCHANGED ON (B)(6) 2019 DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.5/1.5/089 | OS |  | 03/28/2019 | 2019 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8564610&pc=QCB | 2023826-2019-00678 | 2018-09-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/2.0/076 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PRODUCT EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM #: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/2.0/076 | OS |  | 04/17/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8564614&pc=QCB | 2023826-2019-00676 | 2019-03-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.50/2.5/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -14.50/2.5/077 | OS |  | 04/17/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8564616&pc=QCB | 2023826-2019-00679 | 2018-11-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.00/2.5/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -9.00/2.5/101 | OD |  | 04/17/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8564622&pc=QCB | 2023826-2019-00680 | 2018-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.00/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC WAS DEFORMED AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY - UNK. RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/1.0/090 | OD |  | 04/17/2018 | 2018 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8564642&pc=QCB | 2023826-2019-00685 | 2018-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.00/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE- UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -12.00/1.5/092 | OD |  | 04/17/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8564656&pc=QCB | 2023826-2019-00692 | 2018-12-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-04-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.50/3.0/122 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: H3 - CORRECTED TO YES IN SUPPLEMENTAL MDR #1. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -7.50/3.0/122 | OD |  | 04/17/2019 | 2019 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8570514&pc=QCB | 2023826-2019-00726 | 2018-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -9.0/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT DIOPTER NON TORIC LENS. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY WITH RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND; NO VISIBLE DAMAGE TO LENS. 1494 OFF-LABEL USE: PATIENT AGE <21. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.0/2.0/090 | OD |  | 04/04/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8571508&pc=QCB | 2023826-2019-00696 | 2019-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.50/+0.5/033 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS. THE LENS WAS DISCARDED. Manufacturer Narrative: DATE OF EVENT: (B)(6) 2019. THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.50/+0.5/033 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON 31-JAN-2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE ICL SIZE WAS TOO BIG. THE PATIENT IS HAPPY AND NO NEED FOR GLASSES. THE LENS WAS DISCARDED. IMPLANT DATE (B)(6) 2018. EXPLANT DATE: (B)(6) 2019. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.50/+0.5/033 | OD |  | 04/01/2019 | 2019 | 49
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8571527&pc=QCB | 2023826-2019-00697 | 2019-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -5.50/+1.5/169 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO ICL SIZE WAS TOO BIG. Manufacturer Narrative: A YAG WAS PERFORMED FOR ADDITION OF NEW PERIPHERAL IRIDECTOMY. THE PATIENT IS HAPPY AND NO NEED FOR GLASSES. DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.50/+1.5/169 | OS |  | 04/01/2019 | 2019 | 49
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8575880&pc=QCB | 2023826-2019-00716 | 2018-09-28 | Injury | STAAR SURGICAL COMPANY | 2019-05-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/1.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). | -10.50/1.0/095 | OD | Physician | 03/11/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8575926&pc=QCB | 2023826-2019-00701 | 2019-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND CLEAR SURGICAL RESIDUE ON THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/+1.0/090 | OD |  | 04/04/2019 | 2019 | 55
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8575953&pc=QCB | 2023826-2019-00703 | 2019-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50/+1.0/070 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.50/+1.0/070 | OD |  | 04/05/2019 | 2019 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8575974&pc=QCB | 2023826-2019-00702 | 2019-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.0/+1.0/088 | OS |  | 04/04/2019 | 2019 | 55
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8577482&pc=QCB | 2023826-2019-00727 | 2018-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -9.00/2.0/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018 AND THE LENS HAS ROTATED OFF AXIS TWICE BY 20 DEGREES. IT WAS REPORTED THAT REPOSITIONING OF THE LENS WAS PERFORMED ON (B)(6) 2018 AND IT DID NOT RESOLVE THE PROBLEM. THE REPORTER STATED " DAY 1 SITTING CORRECTLY, BY WEEK 1 LENS HAD ROTATED TO 20 DEGREES AND IS STILL SITTING AT 15-20 DEGREES CURRENTLY. LENS NEEDS TO BE EXCHANGED AS POSITION IS NOT STABLE". LENS REMAINS IMPLANTED, IN THE REPORTER OPINION THE CAUSE OF THIS EVENT IS DUE TO A PATIENT RELATED FACTOR. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS EXPLANTED ON (B)(6) 2019 AND LATER REPLACED WITH A LONGER LENGTH LENS. THIS REPLACEMENT LENS RESOLVED THE PROBLEM. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXPLANT. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: STATEMENT ON FOLLOW UP #1 SHOULD BE CORRECTED FROM "THE LENS WAS EXPLANTED ON (B)(6) 2019, AND LATER REPLACED WITH A LONGER LENS" TO "ON (B)(6) 2019 A LONGER LENGTH REPLACEMENT LENS WAS IMPLANTED WITHIN THE SAME SURGICAL PROCEDURE AS THE REMOVAL OF THE INITIAL LENS." (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.00/2.0/098 | OS |  | 04/03/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8586199&pc=QCB | 2023826-2019-00733 | 2019-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, - 8.50/+1.5/070 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LENS ROTATION.THE LENS WAS EXCHANGED FOR A LONGER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE LENS EXCHANGE RESOLVED THE PROBLEM. THE PATIENT IS DOING VERY WELL WITH NO ROTATION, EXCELLENT VISION, NORMAL VAULT AND PRESSURE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.50/+1.5/070 | OS |  | 04/08/2019 | 2019 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8588886&pc=QCB | 2023826-2019-00736 | 2018-12-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.5/+3.0/094 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+3.0/094 | OD |  | 04/08/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8593083&pc=QCB | 2023826-2019-00747 | 2019-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+2.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+2.5/093 | OS |  | 04/08/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8593130&pc=QCB | 2023826-2019-00744 | 2019-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+2.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.5/+2.5/094 | OD |  | 04/08/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8593204&pc=QCB | 2023826-2019-00759 | 2018-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/077 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS HAD LOW VAULTING AND LENS ROTATION. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED. THE LENS IS PLANNED TO BE EXPLANTED. Manufacturer Narrative: THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2019 FOR A DIFFERENT MODEL BUT SAME LENGTH LENS. THE EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS UNKNOWN. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGED. (B)(4). Manufacturer Narrative: EVENT DESCRIPTION: "(B)(6) 2018" IN INITIAL MDR SHOULD BE "(B)(6) 2018"; "THE LENS WAS EXPLANTED AND EXCHANGED" SHOULD BE CHANGED TO "THE LENS WAS EXHANGED" IN FOLLOW-UP MDR #1. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.0/+2.0/077 | OS |  | 04/08/2019 | 2019 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8593221&pc=QCB | 2023826-2019-00752 |  | Injury | STAAR SURGICAL COMPANY | 2019-05-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/102 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED WITH MOISTURE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+2.0/102 | OD |  | 04/08/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8595418&pc=QCB | 2023826-2019-00766 | 2019-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND CLEAR SURGICAL RESIDUE ON THE LENS AND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -12.0/+2.0/092 | OD |  | 04/12/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8595419&pc=QCB | 2023826-2019-00767 | 2019-04-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AND THE LENS HAPTIC HAS A TEAR. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -12.0/+2.5/090 | OS |  | 04/12/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8595443&pc=QCB | 2023826-2019-00773 | 2019-03-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.0/+1.5/065 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THAT THE LENS OPTIC HAS A TEAR. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -6.0/+1.5/065 | OD |  | 04/12/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8600627&pc=QCB | 2023826-2019-00807 | 2019-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-7.00/1.5/094 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.00/1.5/094 | OD |  | 04/17/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8600744&pc=QCB | 2023826-2019-00808 | 2018-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-5.50/4.5/088 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.50/4.5/088 | OD |  | 04/17/2019 | 2019 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8601511&pc=QCB | 2023826-2019-00809 | 2018-11-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-11.00/1.5/093 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/1.5/093 | OD |  | 04/17/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8601767&pc=QCB | 2023826-2019-00813 | 2018-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-10.50/1.0/104 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO OBSERVATION OF LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC WAS TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/1.0/104 | OD |  | 04/17/2019 | 2019 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8609179&pc=QCB | 2023826-2019-00835 | 2019-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.5/+2.0/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.5/+2.0/095 | OD |  | 04/18/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8609180&pc=QCB | 2023826-2019-00833 | 2019-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.0/+3.5/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LENS ROTATION AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A LENS VIAL WITH LIQUID. VISUAL INSPECTION FOUND THERE WAS NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.0/+3.5/080 | OS |  | 04/17/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8609193&pc=QCB | 2023826-2019-00831 | 2019-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+2.5/068 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN LIQUID WITH RESIDUE IN A LENS VIAL. VISUAL INSPECTION FOUND THERE WAS NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/+2.5/068 | OS |  | 04/17/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8609207&pc=QCB | 2023826-2019-00834 | 2019-04-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -7.5/+1.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS "OCOS". Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -7.5/+1.5/086 | OD |  | 04/18/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8609208&pc=QCB | 2023826-2019-00837 | 2019-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.5/+2.5/007 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC IS BENT AND HAS A TEAR. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+2.5/007 | OD |  | 04/18/2019 | 2019 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8609214&pc=QCB | 2023826-2019-00836 | 2019-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -5.5/+1.0/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE AND RESIDUE IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.5/+1.0/092 | OS |  | 04/18/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8613511&pc=QCB | 2023826-2019-00812 | 2018-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-12.0/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISULA INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.0/1.5/095 | OS |  | 04/17/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8613683&pc=QCB | 2023826-2019-00799 | 2019-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/041 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND SHALLOW ANTERIOR CHAMBER DEPTH. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.5/041 | OD |  | 04/15/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8617574&pc=QCB | 2023826-2019-00788 | 2019-01-10 | Injury | STAAR SURGICAL COMPANY | 2019-05-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.50/2.0/098 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -15.50/2.0/098 | OD | Physician | 04/29/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8621857&pc=QCB | 2023826-2019-00801 | 2019-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, - 12.0/+4.0/091 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | - 12.0/+4.0/091 | OD |  | 04/19/2019 | 2019 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8628180&pc=QCB | 2023826-2019-00888 | 2019-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICTED THAT A 12.1MM, VTICM5_12.1, -11.0/3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2019 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A NON TORIC LENS AND IT WAS REPORTED THAT THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY WITH RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND; NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -11.0/3.0/093 | OS |  | 04/22/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8628565&pc=QCB | 2023826-2019-00890 | 2018-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -14.0/2.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC HAD TEARS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.0/2.5/085 | OS |  | 04/22/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8631803&pc=QCB | 2023826-2019-00843 | 2019-05-02 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -13.50/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT WAS OBSERVED, THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019, THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: (B)(4). | -13.50/2.5/092 | OS | Other Health Care Professional | 04/30/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8631508&pc=QCB | 2023826-2019-00846 | 2019-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/1.5/093 | OD |  | 05/06/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8631519&pc=QCB | 2023826-2019-00850 | 2019-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/3.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.50/3.5/089 | OS |  | 05/06/2019 | 2019 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8631522&pc=QCB | 2023826-2019-00849 | 2019-04-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -11.00/2.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEFORMED WITH FOREIGN RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/2.5/090 | OS |  | 05/06/2019 | 2019 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8631524&pc=QCB | 2023826-2019-00853 | 2019-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.50/1.5/126 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.50/1.5/126 | OS |  | 05/06/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8631527&pc=QCB | 2023826-2019-00852 | 2019-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -4.50/1.0/089 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC HAD A TEAR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -4.50/1.0/089 | OD |  | 05/06/2019 | 2019 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8631528&pc=QCB | 2023826-2019-00851 | 2019-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/1.5/089 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.00/1.5/089 | OS |  | 05/06/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8631530&pc=QCB | 2023826-2019-00854 | 2019-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -6.50/2.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.50/2.5/083 | OS |  | 05/06/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8631539&pc=QCB | 2023826-2019-00859 | 2018-12-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.50/3.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -9.50/3.0/089 | OS |  | 05/06/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8631801&pc=QCB | 2023826-2019-00790 | 2019-04-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.50/2.0/115 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SAME LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. SERIAL NUMBER: "(B)(4)" SHOULD BE CORRECTED TO "(B)(4)" IN THE INITIAL MDR. EXPIRATION DATE: "31-MAY-2021" SHOULD BE CORRECTED TO "30-JUN-2021" IN THE INITIAL MDR. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/2.0/115 | OS |  | 04/30/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8633249&pc=QCB | 2023826-2019-00876 | 2019-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO LOW VAULT AND LENS ROTATION. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. CURRENTLY, THERE IS NO REPLACEMENT LENS IN THE PATIENT'S EYE. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. DEVICE EVALUATED BY MFR: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). | -18.0/+1.5/097 | OS |  | 04/25/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8633253&pc=QCB | 2023826-2019-00875 | 2019-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.0/+2.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO LOW VAULT AND LENS ROTATION. THE SURGEON ALSO REPORTS THAT THE LENS TORE DURING INJECTION INTO THE EYE. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR FOR THE LENS TEAR AND PATIENT-RELATED FACTOR FOR THE LOW VAULT. CURRENTLY, THERE IS NO REPLACEMENT LENS IN THE PATIENT'S EYE. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BENT. (B)(4). | -18.0/+2.5/089 | OD |  | 04/25/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8634351&pc=QCB | 2023826-2019-00805 | 2019-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS ICL SIZING. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED IN A LENS VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+2.0/095 | OD |  | 04/25/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8634396&pc=QCB | 2023826-2019-00806 | 2019-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.0/+4.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS ICL SIZING. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A LENS VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+4.0/094 | OS |  | 04/25/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8638575&pc=QCB | 2023826-2019-00906 | 2019-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+3.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATED BY MFR: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+3.0/089 | OS |  | 04/27/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8645457&pc=QCB | 2023826-2019-00896 | 2019-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.00/2.5/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/2.5/079 | OS |  | 05/06/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8645459&pc=QCB | 2023826-2019-00897 | 2019-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.00/1.5/088 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON 31/JAN/2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEFORMED. THE HAPTIC WAS CAUGHT (SMASHED) WITH THE VIAL CAP.(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.00/1.5/088 | OD |  | 05/06/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8645461&pc=QCB | 2023826-2019-00898 | 2019-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.00/1.0/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: EXPLANTED DATE: (B)(6) 2019 TO (B)(6) 2019. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/1.0/099 | OD |  | 05/06/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8645465&pc=QCB | 2023826-2019-00900 | 2010-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -6.50/1.0/088 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.50/1.0/088 | OS |  | 05/06/2019 | 2019 | 16
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8645475&pc=QCB | 2023826-2019-00863 | 2019-04-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Unintended Movement | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.00/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE PATIENT EXPERIENCED LOW VAULT DUE TO THE LENS WAS IMPLANTED UPSIDE DOWN. THE DOCTOR DID NOT NOTICE THAT IT WAS UPSIDE DOWN AT THE TIME OF IMPLANTATION. AT A LATER EXAMINATION THE DOCTOR FOUND THE ISSUE AND EXCHANGED THE INITIAL LENS WITH THE SAME SIZE LENS AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/1.5/091 | OS |  | 05/07/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8651113&pc=QCB | 2023826-2019-00971 | 2019-03-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -9.5/2.0/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. THE REPORTER STATED THAT AFTER EXPLANTING THE LENS THE PROBLEM RESOLVED, PATIENT IS WEARING CONTACT LENSES AS BEFORE AND IS IN GOOD CONDITION. AT LAST VISIT BCVA WAS 20/20. THE SURGEON DOES NOT PLAN TO IMPLANT ANOTHER LENS. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY, IN LENS CASE/VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMGE TO THE LENS. DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: CONCLUSION CODE: 4316- CODE NOT AVAILABLE IS NOT APPLICABLE AND SHOULD BE CORRECTED TO CONCLUSION CODE: 4315- CAUSE NOT ESTABLISHED IN INITIAL MDR. RESULT CODE: 114- OPERATIONAL PROBEM IDENTIFIED IS NOT APPLICABLE AND SHOULD BE CORRECTED TO RESULT CODE: 213- NO DEVICE PROBLEM FOUND IN PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.5/2.0/94 | OD |  | 04/30/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8651875&pc=QCB | 2023826-2019-00972 | 2019-03-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -9.0/2.0/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS ) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. THE REPORTER STATED THAT AFTER EXPLANTING THE LENS THE PROBLEM RESOLVED, PATIENT IS WEARING CONTACT LENSES AS BEFORE AND IS IN GOOD CONDITION. AT LAST VISIT BCVA WAS 20/20. THE SURGEON DOES NOT PLAN TO IMPLANT ANOTHER LENS. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN LENS CASE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS.DIMENSIONAL INSPECTION FOUND LENS WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.0/2.0/94 | OS |  | 04/30/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8651987&pc=QCB | 2023826-2019-00973 | 2019-03-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -12.5/3.0/85 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.5/3.0/85 | OS |  | 04/30/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652039&pc=QCB | 2023826-2019-00974 | 2019-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.0/3.5/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS ) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: EVENT DESCRIPTION: "-9.0/3.5/91" SHOULD BE CORRECTED TO "-9.5/3.5/91" IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.0/3.5/91 | OS |  | 04/30/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652057&pc=QCB | 2023826-2019-00931 | 2019-03-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/118 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS "NO 12.9MM AVAILABLE". Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+1.5/118 | OD |  | 05/06/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652059&pc=QCB | 2023826-2019-00927 | 2018-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/124 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS LISTED AS UNKNOWN. THE PROBLEM IS NOTED AS NOT RESOLVED. REFERENCE MFR REPORT# 2023826-2019-00928 FOR CASE CONTINUATION. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BROKEN. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -8.5/+1.5/124 | OD |  | 05/06/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652066&pc=QCB | 2023826-2019-00924 | 2018-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS LISTED AS UNKNOWN. Manufacturer Narrative: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -11.5/+1.5/091 | OD |  | 05/06/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652101&pc=QCB | 2023826-2019-00932 | 2019-03-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+1.0/065 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED AND THE STATUS OF THE EYE IS "CALM." THE CAUSE OF THE EVENT IS REPORTED AS "NO 12.9MM AVAILABLE." Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+1.0/065 | OS |  | 05/06/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652103&pc=QCB | 2023826-2019-00930 | 2019-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/100 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.0/100 | OS |  | 05/06/2019 | 2019 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652104&pc=QCB | 2023826-2019-00929 | 2018-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.5/102 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: PRODUCT EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE WITH DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -11.0/+2.5/102 | OD |  | 05/06/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652107&pc=QCB | 2023826-2019-00925 | 2018-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+3.5/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS LISTED AS UNKNOWN. Manufacturer Narrative: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -14.0/+3.5/097 | OS |  | 05/06/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652109&pc=QCB | 2023826-2019-00926 | 2019-02-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS LISTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -13.0/+2.5/088 | OD |  | 05/06/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8652118&pc=QCB | 2023826-2019-00923 | 2019-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.0/+1.0/071 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -16.0/+1.0/071 | OS |  | 05/06/2019 | 2019 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8656511&pc=QCB | 2023826-2019-00964 | 2019-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY WITH RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC WAS BENT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.0/+1.5/092 | OD |  | 05/09/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8656517&pc=QCB | 2023826-2019-00965 | 2019-04-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+2.5/089 | OD |  | 05/09/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8660540&pc=QCB | 2023826-2019-00984 | 2019-01-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -18.0/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE CODE: 1494 OFF-LABEL USE. DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -18.0/2.0/093 | OD |  | 05/06/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8660541&pc=QCB | 2023826-2019-00983 | 2019-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.0/2.5/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.0/2.5/107 | OD |  | 05/06/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8660596&pc=QCB | 2023826-2019-00985 | 2019-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -6.0/2.5/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.0/2.5/070 | OD |  | 05/06/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8660769&pc=QCB | 2023826-2019-00968 |  | Injury | STAAR SURGICAL COMPANY | 2019-05-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -8.0/+0.5/098 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+0.5/098 | OS |  | 05/06/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8660775&pc=QCB | 2023826-2019-00952 | 2019-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.0/+1.0/125 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+1.0/125 | OD |  | 05/06/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8660783&pc=QCB | 2023826-2019-00954 | 2019-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+2.0/072 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS NOT EXCHANGED FOR ANOTHER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.5/+2.0/072 | OS |  | 05/06/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8660798&pc=QCB | 2023826-2019-00959 | 2018-10-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+2.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+2.5/084 | OD |  | 05/06/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8660802&pc=QCB | 2023826-2019-00960 | 2019-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+3.0/078 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+3.0/078 | OD |  | 05/06/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8660839&pc=QCB | 2023826-2019-00963 | 2019-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+2.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.5/+2.0/089 | OD |  | 05/07/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8660841&pc=QCB | 2023826-2019-00962 | 2019-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/081 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+2.0/081 | OD |  | 05/07/2019 | 2019 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8664329&pc=QCB | 2023826-2019-00866 | 2019-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -8.50/3.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.50/3.0/090 | OD |  | 05/15/2019 | 2019 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8664331&pc=QCB | 2023826-2019-00868 | 2018-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -10.50/2.0/088 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/2.0/088 | OD |  | 05/16/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8664800&pc=QCB | 2023826-2019-00992 | 2019-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.5/4.5/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO OBSERVATION OF LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.5/4.5/096 | OS |  | 05/06/2019 | 2019 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8666880&pc=QCB | 2023826-2019-00995 | 2018-12-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1 MM, VTICMO 12.1, -9.0/1.5/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DEVICE CODE 1494: ACD<3.0MM. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.0/1.5/078 | OS |  | 05/06/2019 | 2019 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8666982&pc=QCB | 2023826-2019-00998 | 2019-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -6.5/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.5/1.0/093 | OD |  | 05/06/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8667128&pc=QCB | 2023826-2019-01001 | 2019-03-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.5/1.5/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/1.5/100 | OD |  | 05/06/2019 | 2019 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8667617&pc=QCB | 2023826-2019-01002 | 2019-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.5/1.0/076 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.5/1.0/076 | OS |  | 05/06/2019 | 2019 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8668383&pc=QCB | 2023826-2019-01004 | 2018-08-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -8.5/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC WAS DEFORMED AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/1.0/091 | OD |  | 05/07/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8668522&pc=QCB | 2023826-2019-01005 | 2018-08-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.5/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: VTICMO LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.5/1.0/085 | OS |  | 05/07/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8668934&pc=QCB | 2023826-2019-01044 | 2019-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+1.0/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.50/+1.0/080 | OD |  | 05/27/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8671664&pc=QCB | 2023826-2019-01012 | 2019-05-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL13.2 LENS, -12.5/+3.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: "TMICL13.2" SHOULD BE CORRECTED TO "TMICL12.6" IN INITIAL MEDWATCH REPORT. "PATIENT'S RIGHT (OD) EYE" SHOULD BE CORRECTED TO "PATIENT'S LEFT EYE (OS)" "THE LENS WAS EXCHANGED" SHOULD BE CORRECTED TO "THE LENS WAS EXPLANTED AND REPLACED ". "PATIENT CODE 3191 : NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE)" SHOULD BE CORRECTED TO "PATIENT CODE 3191 : NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT)". (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.5/+3.0/092 | OS |  | 05/07/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8672625&pc=QCB | 2023826-2019-01035 | 2019-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+3.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND DEBRIS ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+3.5/088 | OS |  | 05/16/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8675666&pc=QCB | 2023826-2019-01010 |  | Injury | STAAR SURGICAL COMPANY | 2019-06-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM, TMICL12.1, -16.00/4.0/093 (SPHERE/ CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO VAULT/SIZING. IT WAS REPORTED THAT A LONGER LENGTH LENS WAS LATER IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -16.00/4.0/093 | OD | Physician | 05/08/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8676928&pc=QCB | 2023826-2019-01041 | 2018-06-21 | Injury | STAAR SURGICAL COMPANY | 2019-06-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+1.5/082 | OD | Physician | 05/17/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8676796&pc=QCB | 2023826-2019-01038 | 2019-04-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.5/090 | OD |  | 05/17/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8676867&pc=QCB | 2023826-2019-01036 | 2018-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/179 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.0/179 | OS |  | 05/16/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8676902&pc=QCB | 2023826-2019-01039 | 2019-04-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/081 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+1.5/081 | OS |  | 05/17/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8676907&pc=QCB | 2023826-2019-01046 | 2019-01-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -6.0/+2.0/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.0/+2.0/085 | OS |  | 05/31/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8676915&pc=QCB | 2023826-2019-01042 | 2018-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.0/+1.5/063 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.0/+1.5/063 | OS |  | 05/17/2019 | 2019 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8676922&pc=QCB | 2023826-2019-01048 | 2019-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/099 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATIONS: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON IT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+1.5/099 | OS |  | 05/31/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8685162&pc=QCB | 2023826-2019-01075 |  | Injury | STAAR SURGICAL COMPANY | 2019-06-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM,VTICM5_13.7, -6.00/0.5/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.00/0.5/101 | OD | Physician | 05/17/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8683515&pc=QCB | 2023826-2019-01009 | 2019-04-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM,VTICMO13.2, -4.5/6.0/103 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED. IT WAS REPORTED THAT THE REPLACEMENT RESOLVED THE PROBLEM, VAULT BECAME NORMAL AND PATIENT IS DOING VERY WELL. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE AVAILABLE FOR EVALUATION; YES. RETURNED TO MANUFACTURER ON 17/05/2019(. B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -4.5/6.0/103 | OD |  | 05/13/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8683769&pc=QCB | 2023826-2019-01071 | 2019-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM,VTICMO13.7, -18.0/3.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM, PATIENT IS HAPPY. Manufacturer Narrative: DEVICE AVAILABLE FOR EVALUATION; YES. RETURNED TO MANUFACTURER ON 17/05/2019. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -18.0/3.5/083 | OD |  | 05/13/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8684971&pc=QCB | 2023826-2019-01074 | 2019-04-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM,VTICM5_13.2, -14.00/2.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.00/2.0/092 | OD |  | 05/14/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8686764&pc=QCB | 2023826-2019-00936 | 2018-12-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.50/3.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2018 AND (B)(6) 2019. ON (B)(6) 2019 THE LENS EXCHANGED AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL OBSERVATION FOUND HAPIC TORN/BROKEN/BENT/DEFORMED AND RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/3.0/087 | OS |  | 05/16/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8692879&pc=QCB | 2023826-2019-01066 | 2018-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K #: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+2.0/088 | OS |  | 05/16/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8692889&pc=QCB | 2023826-2019-01064 | 2018-12-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/112 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DATE OF EVENT CORRECTED FROM (B)(6). CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.0/+1.5/112 | OD |  | 05/16/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8695703&pc=QCB | 2023826-2019-00941 | 2019-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6 MM, VTICMO12.6, -17.50/3.5/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2018. ON (B)(6) 2019 THE LENS EXCHANGED AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT UNKNOWN, ETHNICITY UNKNOWN, RACE UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -17.50/3.5/077 | OS |  | 05/18/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8695708&pc=QCB | 2023826-2019-00940 | 2019-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -18.00/2.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: "VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS." IN PREVIOUS MDR SHOULD BE CHANGED TO "VISUAL INSPECTION FOUND THE HAPTIC TORN." (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -18.00/2.5/094 | OD |  | 05/18/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8705830&pc=QCB | 2023826-2019-01094 | 2019-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.5/1.5/167 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/1.5/167 | OD |  | 05/23/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8705899&pc=QCB | 2023826-2019-01095 | 2019-05-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.5/1.5/017 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/1.5/017 | OS |  | 05/23/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8709922&pc=QCB | 2023826-2019-01093 | 2019-05-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -17/3.0/88 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. HOWEVER, REPOSITIONING DID NOT RESOLVE THE PROBLEM AND THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2019. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: RETURNED TO MANUFACTURER: 06-APR-2019 SHOULD BE CORRECTED TO 04-JUN-2019. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -17/3.0/88 | OD |  | 05/21/2019 | 2019 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8716008&pc=QCB | 2023826-2019-01055 | 2018-12-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -12.50/3.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -12.50/3.5/089 | OD |  | 06/04/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8716010&pc=QCB | 2023826-2019-01056 | 2018-12-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.00/4.5/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON 26/DEC/2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US . (B)(4). | -9.00/4.5/102 | OS |  | 06/04/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8716015&pc=QCB | 2023826-2019-01058 | 2018-11-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.00/1.5/094 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/1.5/094 | OS |  | 06/04/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8716016&pc=QCB | 2023826-2019-01057 | 2018-11-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.50/1.5/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.5/100 | OD |  | 06/04/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8716017&pc=QCB | 2023826-2019-01060 | 2019-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.50/1.0/169 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PRODUCT EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH DEBRIS ON THE LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.50/1.0/169 | OD |  | 06/04/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8716022&pc=QCB | 2023826-2019-01081 | 2019-03-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.00/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/2.5/092 | OS |  | 06/05/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8716026&pc=QCB | 2023826-2019-01085 | 2019-05-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.00/1.0/056 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: "LOW VAULT" SHOULD BE CHANGED TO "LOW VAULT WITH ROTATION" IN INITIAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/1.0/056 | OD |  | 06/07/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8716027&pc=QCB | 2023826-2019-01086 | 2019-05-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -12.00/1.0/152 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: "LOW VAULT" SHOULD BE CHANGED TO "LOW VAULT WITH ROTATION" IN INITIAL MDR. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/1.0/152 | OS |  | 06/07/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8728233&pc=QCB | 2023826-2019-01111 | 2019-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.00/2.00/94 (SPHERE/AXIS/CYLINDER), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PRODUCT EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/2.00/94 | OS |  | 05/30/2019 | 2019 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8736580&pc=QCB | 2023826-2019-01123 | 2018-06-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/085 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND BENT WITH RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+2.0/085 | OD |  | 06/04/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8742031&pc=QCB | 2023826-2019-01153 | 2019-03-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -7.50/+4.5/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING, LENS ROTATION AND LENS DISLOCATION/SUBLUXATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR, THERE WAS A MISTAKE IN CALCULATION. Manufacturer Narrative: CONCLUSION CODE 61: USER REPORTED MISTAKE IN CALCULATION - PREVIOUS CONCLUSION CODE, 4315 NO LONGER APPLIABLE, SUPPLEMENTAL MDR SHOULD HAVE INCLUDED "RESIDUE ON LENS" FOR THE DEVICE EVALUATION. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISUAL DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.50/+4.5/087 | OS |  | 05/29/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8746015&pc=QCB | 2023826-2019-01142 | 2017-05-25 | Injury | STAAR SURGICAL COMPANY | 2019-06-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -15.0/1.5/166 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2017 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO OBSERVATION OF LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.0/1.5/166 | OD | Physician | 06/04/2019 | 2019 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8745460&pc=QCB | 2023826-2019-01138 | 2019-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -3.0/3.0/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM, PATIENT IS HAPPY AND ALL IS FINE. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US, (B)(4). | -3.0/3.0/93 | OD |  | 06/03/2019 | 2019 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8745932&pc=QCB | 2023826-2019-01139 | 2018-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -8.0/1.5/086 (SPHERE/CYLINDER/AXIS), , IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC DEFORMED AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K #: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.0/1.5/086 | OD |  | 06/04/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8751289&pc=QCB | 2023826-2019-01167 | 2019-05-09 | Injury | STAAR SURGICAL COMPANY | 2019-07-01 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/110 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO THE LENS WAS "TOO LARGE." THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+1.0/110 | NA | Physician | 06/10/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8751286&pc=QCB | 2023826-2019-01168 | 2019-03-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -11.0/+1.5/089 | OD |  | 06/05/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8754599&pc=QCB | 2023826-2019-01145 | 2019-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -15.00/+3.5/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE RETURNED TO MANUFACTURER DATE IN SUPPLEMENTAL MDR IS NOT APPLICABLE; THE CORRECT DATE RETURNED IS 26-JUN-2019. CLAIM#: (B)(4). Manufacturer Narrative: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON LENS. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US, (B)(4). | -15.00/+3.5/100 | OD |  | 06/05/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8755914&pc=QCB | 2023826-2019-01143 | 2019-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICM5_13.2, -11.00/2.0/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE PATIENT WAS DISSATISFIED WITH THE QUALITY OF VISION AFTER ICL IMPLANT. THEREFORE, ON (B)(6) 2019 THE LENS WAS EXPLANTED. IT WAS REPORTED THAT THE PATIENTS CURRENT CONDITION IS GOOD, THE PROBLEM RESOLVED AFTER EXPLANT. THE PATIENTS CONDITION RETURNED TO THE PREOPERATIVE CONDITION. Manufacturer Narrative: RESULT CODE: 144 SHOULD BE CORRECTED TO 114 IN SUPPLEMENTAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.00/2.0/94 | OD |  | 06/04/2019 | 2019 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8755960&pc=QCB | 2023826-2019-01161 | 2018-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.5/+2.5/077 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+2.5/077 | OD |  | 06/04/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8755961&pc=QCB | 2023826-2019-01160 | 2018-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+2.5/099 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: IT WAS REPORTED THAT 'THERE ARE CYSTS IN CILIARY SULCUS, SO BE CAREFUL OF THE LENS ROTATION'. (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.5/+2.5/099 | OD |  | 06/04/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8755964&pc=QCB | 2023826-2019-01162 | 2018-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+2.0/095 | OS |  | 06/04/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8756019&pc=QCB | 2023826-2019-01144 | 2019-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICM5_13.2, -12/2.0/2 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE PATIENT WAS DISSATISFIED WITH THE QUALITY OF VISION AFTER ICL IMPLANT. THEREFORE, ON (B)(6) 2019 THE LENS WAS EXPLANTED. IT WAS REPORTED THAT THE PATIENTS CURRENT CONDITION IS GOOD, THE PROBLEM RESOLVED AFTER EXPLANT. THE PATIENTS CONDITION RETURNED TO THE PREOPERATIVE CONDITION. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY AND WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT WITH RESIDENT ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12/2.0/2 | OS |  | 06/04/2019 | 2019 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8759061&pc=QCB | 2023826-2019-01146 | 2019-06-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2, -14.5/2.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT, PIGMENT DISPERSION AND UNREACTIVE (FIXED) PUPIL. A LENS OF SHORTER LENGTH WAS LATER IMPLANTED. PATIENTS CURRENT CONDITION WAS REPORTED AS BETTER, PUPIL STILL SLIGHTLY DILATED. LENS EXCHANGE DID NOT COMPLETELY RESOLVE THE PROBLEM. Manufacturer Narrative: PRODUCT EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH RESIDUE ON THE LENS. DATE OF REPORT: 20-JUN-2019 SHOULD BE CORRECTED TO 04-JUN-2019 IN INITIAL MDR. DATE RECEIVED BY MANUFACTURER: 20-JUN-2019 SHOULD BE CORRECTED TO 04-JUN-2019 IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: AWARE DATE RECORRECTED TO THE DATE REPORTED IN THE INITIAL MDR. 20-JUN-2019 IS THE CORRECT AWARE DATE. AWARE DATE REPORTED IN FOLLOW-UP#1 IS NO LONGER APPLICABLE. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -14.5/2.0/094 | OS |  | 06/20/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8764082&pc=QCB | 2023826-2019-01104 | 2019-05-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.00/2.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED, THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE; PATIENT AGE < 21 YEARS OLD. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.00/2.5/095 | OS |  | 06/19/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8764092&pc=QCB | 2023826-2019-01129 | 2019-05-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -12.50/4.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.50/4.0/095 | OD |  | 06/20/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8764904&pc=QCB | 2023826-2019-01148 | 2019-01-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -10.00/2.00/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGE FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND RESIDUE ON THE LENS WITH NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K #: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/2.00/89 | OS |  | 06/13/2019 | 2019 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8764905&pc=QCB | 2023826-2019-01183 | 2019-05-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A TICM120V4, -15.00/1.00/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2012. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. A REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON LENS. VISUAL INSPECTION FOUND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: IMPLANTED DATE 25-AUG-2012 SHOULD BE CORRECTED TO (B)(6) 2012. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K #: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.00/1.00/89 | OD |  | 06/13/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8769358&pc=QCB | 2023826-2019-01185 | 2019-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -11.0/5.5/80 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE ON (B)(6) 2019. ON (B)(6)2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM, PATIENT IS HAPPY. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/5.5/80 | OS |  | 06/12/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8772264&pc=QCB | 2023826-2019-01134 | 2018-11-24 | Injury | STAAR SURGICAL COMPANY | 2019-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -13.50/2.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.50/2.5/083 | OS | Physician | 06/24/2019 | 2019 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8772146&pc=QCB | 2023826-2019-01151 | 2019-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICM5_13.2, -8.5/1.5/74 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION : LENS WAS RETURNED IN A MICRO- CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/1.5/74 | OS |  | 06/19/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8772260&pc=QCB | 2023826-2019-01133 | 2018-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -15.50/1.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.50/1.5/097 | OD |  | 06/24/2019 | 2019 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8772273&pc=QCB | 2023826-2019-01174 | 2019-06-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.00/1.5/136 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/1.5/136 | OS |  | 06/28/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8774157&pc=QCB | 2023826-2019-01189 | 2019-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+3.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -11.5/+3.0/091 | OD |  | 06/13/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8774164&pc=QCB | 2023826-2019-01197 | 2019-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/157 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REPOSITIONED AND THEN EXCHANGED FOR A LONGER, SPHERICAL LENS DUE TO LENS ROTATION (REPOSITIONING DID NOT RESOLVE THE PROBLEM). REPORTEDLY, LENS REPOSITIONING AND EXCHANGE OCCURRED WITHIN THE SAME SURGERY. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.5/+1.5/157 | OD |  | 06/18/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8791001&pc=QCB | 2023826-2019-01228 | 2019-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM, VTICMO12.6, -9.5/2.5/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS EXCHANGED AND THE PROBLEM RESOLVED.. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC DEFORMED AND RESIDUE ON LENS. METHOD CODE 10- FOR PRODUCT EVALUATION, RESULT CODE 114- OPERATIONAL PROBLEM IDENTIFIED. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K #: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.5/2.5/084 | OS |  | 06/20/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8796303&pc=QCB | 2023826-2019-01265 | 2019-06-13 | Injury | STAAR SURGICAL COMPANY | 2019-07-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -9.50/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO THE LENS BEING TOO LONG (DIDNT FIT), NO PATIENT INJURY NOTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THE LENS TORN IN HALF. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.50/1.5/095 | OD | Physician | 06/17/2019 | 2019 | 69
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8793820&pc=QCB | 2023826-2019-01211 | 2019-05-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -14.50/3.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LENS ROTATION AND LOW VAULT WERE OBSERVED. THE LENS WAS REMOVED AND EXCHANGED WITH A LONGER LENGHT LENS ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K #: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.50/3.5/081 | OS |  | 07/04/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8795451&pc=QCB | 2023826-2019-01233 |  | Injury | STAAR SURGICAL COMPANY | 2019-07-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -11.0/+2.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT AND LENS ROTATION. THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH, SEX, WIGHT, ETHNICITY-UNK. DATE OF EVENT: UNK. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.0/+2.0/089 | OD |  | 06/24/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8795472&pc=QCB | 2023826-2019-01241 | 2019-05-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+1.5/118 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED AND IN A 2ND SURGERY ON THE SAME DATE, A SHORTER LENS WAS IMPLANTED DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS "LENS SIZING NOMOGRAM FROM STAAR PREDICTS TOO LARGE A LENS NEEDED." Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A LENS VIAL IN LIQUID. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNKNOWN. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.5/+1.5/118 | OD |  | 06/21/2019 | 2019 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8798429&pc=QCB | 2023826-2019-01103 | 2019-05-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6 MM, VTICM5_12.6, -11.50/3.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2019, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED FOR A SPHERICAL LENS AND THE PROBLEM WAS RESOLVED. IN THE REPORTER'S OPINION, THE CAUSE OF THE EVENT IS ATTRIBUTED TO THE DEVICE. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE; AGE < 21 CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.50/3.0/090 | OD |  | 06/19/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8799639&pc=QCB | 2023826-2019-01277 |  | Injury | STAAR SURGICAL COMPANY | 2019-07-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -11.0/2.0/173 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED, AXIS WAS OFF BY 90/ CALCULATE WRONG. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: RETURNED TO MANUFACTURER DATE IN ORIGINAL MDR IS NOT APPLICABLE; CORRECT DATE IS 08-JUL-2019. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/2.0/173 | OS |  | 06/20/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8806863&pc=QCB | 2023826-2019-01208 | 2019-02-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6 MM, VTICMO12.6, -13.00/3.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(6). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.00/3.0/091 | OS |  | 07/02/2019 | 2019 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8813867&pc=QCB | 2023826-2019-01286 | 2019-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/094 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TEAR. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+2.5/094 | OD |  | 07/02/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8819679&pc=QCB | 2023826-2019-01299 | 2019-05-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.5/+2.5/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+2.5/083 | OS |  | 06/26/2019 | 2019 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8826259&pc=QCB | 2023826-2019-01212 |  | Injury | STAAR SURGICAL COMPANY | 2019-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -11.50/2.5/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2014. LENS ROTATION AND LOW VAULT WERE OBSERVED. THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: METHOD CODE 4110: WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH RED RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION:LENS RETURNED DRY IN LENS CASE/VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.50/2.5/074 | OD |  | 07/09/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8832979&pc=QCB | 2023826-2019-01306 | 2019-06-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -15.00/2.00/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019 AS A REPLACEMENT LENS DUE TO EXCESSIVE VAULT, PIGMENT DISPERSION AND UNREACTIVE (FIXED) PUPIL BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE REPORTER STATED THAT THE PATIENT IS DOING BETTER, LESS LIGHT SENSITIVITY. HOWEVER, THE PUPIL IS STILL MILDLY DILATED WITH SOME IRIS TRAUMA AND PIGMENTATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: OUTCOMES ATTRIBUTED TO ADVERSE EVENT: OTHER SERIOUS SHOULD HAVE BEEN INCLUDED IN INTIAL MDR. DATE OF EVENT: CORRECTED TO UNK IN INITAL MDR. "LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019" CORRECTED TO "LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019". CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -15.00/2.00/088 | OS |  | 06/28/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8834086&pc=QCB | 2023826-2019-01334 | 2019-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+1.0/097 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND PERIPHERAL CORNEAL-IRIS TOUCH. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THE EXCHANGED LENS STILL HAD A HIGH VAULT AND DID NOT RESOLVE THE PROBLEM. SEE MFR REPORT # 2023826-2019-01543. THE CAUSE OF THE EVENT WAS DUE TO PATIENT SULCUS WAS TOO SMALL AND THE VAULT WAS THEREFORE TOO HIGH. THE PATIENT NEEDED A SMALLER LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO ROTATION OF CILIARY BODY SMALLER THAN MEASURED EYE. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS TRAY BOX. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.0/+1.0/097 | OD |  | 06/27/2019 | 2019 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8836106&pc=QCB | 2023826-2019-01250 | 2019-06-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.50/3.0/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN/BENT AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.50/3.0/077 | OS |  | 07/11/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8836275&pc=QCB | 2023826-2019-01346 | 2019-01-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2, -12./1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12./1.0/093 | OD |  | 07/02/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8837408&pc=QCB | 2023826-2019-01347 | 2018-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM, VTICMO12.6, -18.0/1.5/130 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -18.0/1.5/130 | OS |  | 07/02/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8837704&pc=QCB | 2023826-2019-01343 | 2019-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM, VTICM5_13.2, -6.0/6.0/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO LOW VAULT WITH ROTATION. A LONGER LENGTH REPLACEMENT LENS WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: 'DEVICE EVALUATION ANTICIPATED, BUT NOT YET BEGUN' MARKED IN ERROR. THIS STATEMENT IS NOT APPLICABLE IN INITIAL MDR. WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS FOUND WITHIN ASSOCIATED LOTS. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND A HAPTIC TEAR AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.0/6.0/099 | OD |  | 07/05/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8843590&pc=QCB | 2023826-2019-01340 | 2019-03-26 | Injury | STAAR SURGICAL COMPANY | 2019-07-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+2.5/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON 1(B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+2.5/088 | OD | Physician | 07/09/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8843592&pc=QCB | 2023826-2019-01338 | 2019-06-03 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/138 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON NOTES EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). REPORTEDLY THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/+2.0/138 | OS | Physician | 07/04/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8843594&pc=QCB | 2023826-2019-01341 | 2019-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-07-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.0/+3.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SAME SIZE AND MODEL, SIMILAR POWER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.0/+3.5/089 | OS |  | 07/09/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8851486&pc=QCB | 2023826-2019-01352 | 2019-07-03 | Injury | STAAR SURGICAL COMPANY | 2019-08-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.00/3.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/3.0/094 | OS | Physician | 07/19/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8851458&pc=QCB | 2023826-2019-01269 | 2018-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, - 12.00/2.5/102 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PRODUCT EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | - 12.00/2.5/102 | OD |  | 07/16/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8851461&pc=QCB | 2023826-2019-01270 | 2019-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/1.5/088 SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.50/1.5/088 | OS |  | 07/16/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8851470&pc=QCB | 2023826-2019-01271 | 2019-01-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -14.00/3.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.00/3.0/094 | OD |  | 07/16/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8852162&pc=QCB | 2023826-2019-01381 | 2019-07-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+2.5/089 | OS |  | 07/07/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8871968&pc=QCB | 2023826-2019-01398 | 2019-04-29 | Injury | STAAR SURGICAL COMPANY | 2019-08-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+3.5/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR ANOTHER MANUFACTURER LENS. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR AND WAS NOT RELATED TO THE PRODUCT. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+3.5/087 | OS | Physician | 07/09/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8870916&pc=QCB | 2023826-2019-01414 | 2018-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/077 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6)2019 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/ VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#:(B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -9.5/+1.5/077 | OD |  | 07/16/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8870918&pc=QCB | 2023826-2019-01412 | 2018-10-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/100 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -9.5/+1.0/100 | OD |  | 07/16/2019 | 2019 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8870919&pc=QCB | 2023826-2019-01410 | 2019-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+2.0/099 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIUDE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+2.0/099 | OD |  | 07/16/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8870920&pc=QCB | 2023826-2019-01409 | 2019-02-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/100 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.5/100 | OD |  | 07/16/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8870921&pc=QCB | 2023826-2019-01407 | 2019-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+3.0/082 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+3.0/082 | OS |  | 07/14/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8870922&pc=QCB | 2023826-2019-01406 | 2019-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.0/+5.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -15.0/+5.5/092 | OS |  | 07/11/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8870923&pc=QCB | 2023826-2019-01411 | 2018-10-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.0/+1.0/090 | OS |  | 07/16/2019 | 2019 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8871973&pc=QCB | 2023826-2019-01397 | 2019-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A DIFFERENT MANUFACTURER LENS. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR AND WAS NOT RELATED TO THE PRODUCT. Manufacturer Narrative: ADDITIONAL DEVICE EVALUATION WAS PERFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/+2.5/089 | OD |  | 07/09/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8883521&pc=QCB | 2023826-2019-01355 | 2019-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -5.50/1.5/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.50/1.5/104 | OS |  | 07/22/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8883526&pc=QCB | 2023826-2019-01356 | 2019-04-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.00/2.0/171 SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/2.0/171 | OD |  | 07/22/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8883531&pc=QCB | 2023826-2019-01359 | 2019-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.50/1.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(4) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(4) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE, CLEAR SURGICAL RESIDUE AND DEBRIS ON LENS. VISUAL INSPECTION FOUND DEBRIS AND RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.0/075 | OD |  | 07/22/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8883536&pc=QCB | 2023826-2019-01415 | 2019-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -9.50/1.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.0/095 | OD |  | 07/22/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8887958&pc=QCB | 2023826-2019-01360 | 2019-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.50/1.5/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2019 THE LENS WAS REMOVED AND A NEW LENS IMPLANTED ON (B)(6) 2019, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTER NOTED "AFTER THE PATIENT IMPLANTED ICL, THE LENS ROTATED IN THE EYE. AFTER 3 DAYS, THE LENS ROTATED 30°. THE LENS ROTATED INTO A VERTICAL POSITION ONE WEEK AFTER SURGERY. SURGEON DECIDED TO IMPLANT LENS TO ASSIST THE LISAK SURGERY, PATIENT ASTIGMATISM." Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE/DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PATIENT CODE 3191: "SECONDARY SURGICAL INTERVENTION- EXCHANGE" SHOULD BE CORRECTED TO "SECONDARY SURGICAL INTERVENTION- EXPLANT" IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.50/1.5/078 | OS |  | 07/22/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8889729&pc=QCB | 2023826-2019-01361 | 2019-04-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/2.0/092 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATED BY MFR: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/2.0/092 | OD |  | 07/24/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8889735&pc=QCB | 2023826-2019-01362 | 2019-04-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.00/1.5/080 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -09.00/1.5/080 | OS |  | 07/24/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8894311&pc=QCB | 2023826-2019-01436 | 2019-01-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+0.5/098 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+0.5/098 | OS |  | 07/22/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8894314&pc=QCB | 2023826-2019-01438 | 2018-11-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+2.0/093 | OS |  | 07/22/2019 | 2019 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8894319&pc=QCB | 2023826-2019-01434 | 2019-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -11.0/+2.0/088 | OS |  | 07/17/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8894321&pc=QCB | 2023826-2019-01437 | 2018-11-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.0/+2.5/083 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.0/+2.5/083 | OD |  | 07/22/2019 | 2019 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8895320&pc=QCB | 2023826-2019-01455 | 2019-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.0/1.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DEVICE CODE 1494: OFF-LABEL USE: AGE OF PATIENT < 21 YEARS AT THE TIME OF IMPLANTATION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.0/1.0/084 | OD |  | 07/16/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8899055&pc=QCB | 2023826-2019-01425 | 2019-07-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.50/+5.0/103 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.50/+5.0/103 | OD |  | 07/16/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8899062&pc=QCB | 2023826-2019-01426 | 2019-06-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.0/+2.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+2.5/095 | OD |  | 07/17/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8899067&pc=QCB | 2023826-2019-01427 | 2019-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, - 9.50/+2.0/169 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | - 9.50/+2.0/169 | OD |  | 07/19/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8899076&pc=QCB | 2023826-2019-01428 | 2019-06-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, - 9.50/+3.0/018 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | - 9.50/+3.0/018 | OS |  | 07/19/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8902525&pc=QCB | 2023826-2019-01476 | 2019-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -7.00/4.0/018 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. EXCESSIVE VAULT AND FOCAL IRIDOCORNEAL CONTACT WAS REPORTED. REPOSITIONING OF THE LENS WAS PERFORMED BUT IT DID NOT RESOLVE THE PROBLEM. THE REPORTER STATED THAT THE SURGEON PLANS TO EXCHANGE THE LENS, PATIENT IS BEING MONITOR CLOSELY. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: REPORTER STATED "THE SMALLER TORIC ICL FIXED THE OVERCROWDING ISSUE WITHIN THE PATIENTS EYE. DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT) EXPLANTED DATE IN PREVIOUS MDR CORRECTED TO (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A ZIPLOC BIOHAZARD BAG. VISUAL INSPECTON FOUND NO VISIBLE DAMAGE. CLAIM#: (B)(4). Manufacturer Narrative: "THE REPORTER STATED THAT A THE SURGEON PLANS TO EXCHANGE THE LENS, PATIENT IS BEING MONITORED CLOSELY." SHOULD BE CORRECTED TO "ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM." IN THE INITIAL MDR. EXPLANT DATE: (B)(6) 2019. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXCHANGE. (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.00/4.0/018 | OS |  | 07/18/2019 | 2019 | 51
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8910144&pc=QCB | 2023826-2019-01448 | 2019-07-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.50/4.5/091 SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: LENS WAS IMPLANTED ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019. (B)(4). Manufacturer Narrative: (B)(4). | -17.50/4.5/091 | OD |  | 07/31/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8912793&pc=QCB | 2023826-2019-01521 | 2019-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO_13.2, -9.00/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/1.0/091 | OD |  | 07/22/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8912814&pc=QCB | 2023826-2019-01511 | 2019-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.5/+2.5/105 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2019. ON (B)(6) 2019, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). ON (B)(6) 2019 A SHORTER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: THE OPTIC AND HAPTIC ARE TORN. (B)(4). Manufacturer Narrative: PATIENT CODE 3191: NO CODE AVAILABLE "(SECONDARY SURGERY, LENS EXCHANGE)" SHOULD BE CORRECTED TO "LENS EXPLANT" INTHE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+2.5/105 | OD |  | 07/29/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8912817&pc=QCB | 2023826-2019-01515 | 2019-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.0/+2.0/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -15.0/+2.0/083 | OS |  | 07/30/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8912818&pc=QCB | 2023826-2019-01506 | 2019-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -11.5/+1.5/089 | OD |  | 07/24/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8912821&pc=QCB | 2023826-2019-01505 | 2019-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.5/+2.0/090 | OD |  | 07/24/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8913028&pc=QCB | 2023826-2019-01522 | 2019-01-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO_12.6, -9.5/3.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CONCOMMITANT MEDICAL PRODUCTS: INJECTOR MODEL-MSI-PF,LOT#: 1395302, CARTRIDGE MODEL-SFC-45; LOT#: 1404251, FOAM TIP PLUNGER MODEL-FTP; LOT#: 1407639. DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.5/3.0/082 | OD |  | 07/24/2019 | 2019 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8913100&pc=QCB | 2023826-2019-01523 | 2019-01-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO_12.6, -9.5/3.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: DEVICE EVALUATED BY MFR: DEVICE EVALUATION: VISUAL INSPECTION ALSO FOUND A TEAR ON THE LENS OPTIC. SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL #1 CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.5/3.5/091 | OS |  | 07/24/2019 | 2019 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8913314&pc=QCB | 2023826-2019-01543 | 2019-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.0/+1.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO ANTERIOR NOTATION OF CILIARY BODY SMALLER THAN MEASURED EYE. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.0/+1.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO ANTERIOR ROTATION OF CILIARY BODY SMALLER THAN MEASURED EYE. THE LENS WAS DESTROYED AND WILL NOT BE RETURNED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+1.0/094 | OD |  | 07/25/2019 | 2019 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8913318&pc=QCB | 2023826-2019-01484 | 2019-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -5.5/+5.5/019 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS EYE, ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE AND DEBRIS ON LENS. VISUAL INSPECTION FOUND THE HAPTICS BENT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.5/+5.5/019 | NA |  | 07/26/2019 | 2019 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8918813&pc=QCB | 2023826-2019-01493 | 2019-04-17 | Injury | STAAR SURGICAL COMPANY | 2019-08-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+3.5/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+3.5/096 | OD | Physician | 07/22/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8917675&pc=QCB | 2023826-2019-01524 | 2019-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO_12.6, -14.0/3.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN, HAPTIC TORN AND DEFORMED, AND RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.0/3.5/085 | OS |  | 07/24/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8917802&pc=QCB | 2023826-2019-01525 | 2019-02-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.5/1.5/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR LONGER LENGTH LENS DUE TO LOW VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/1.5/102 | OD |  | 07/24/2019 | 2019 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8918791&pc=QCB | 2023826-2019-01485 | 2019-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.5/+3.5/029 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.5/+3.5/029 | OD |  | 07/26/2019 | 2019 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8918805&pc=QCB | 2023826-2019-01492 | 2018-12-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+4.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+4.5/095 | OD |  | 07/22/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8918817&pc=QCB | 2023826-2019-01495 | 2019-05-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/+2.0/092 | OS |  | 07/23/2019 | 2019 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8918818&pc=QCB | 2023826-2019-01490 | 2019-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+0.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+0.5/088 | OD |  | 07/22/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8930920&pc=QCB | 2023826-2019-01463 | 2019-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -14.00/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED FOR A SPHERICAL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -14.00/1.5/095 | OD |  | 08/05/2019 | 2019 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8930927&pc=QCB | 2023826-2019-01464 | 2019-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.50/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED FOR A SPHERICAL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.50/1.5/090 | OS |  | 08/05/2019 | 2019 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941303&pc=QCB | 2023826-2019-01566 | 2019-07-17 | Injury | STAAR SURGICAL COMPANY | 2019-08-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -14.0/+2.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED LIQUID, IN A LENS VIAL AND ZIPLOC BAG. VISUAL INSPECTION FOUND NO VISUAL DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.0/+2.5/093 | OD | Physician | 07/31/2019 | 2019 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941311&pc=QCB | 2023826-2019-01574 | 2019-07-24 | Injury | STAAR SURGICAL COMPANY | 2019-08-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -9.0/+3.0/101 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+3.0/101 | OD | Physician | 07/31/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941314&pc=QCB | 2023826-2019-01567 | 2019-07-17 | Injury | STAAR SURGICAL COMPANY | 2019-08-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -12.0/+1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THE LENS WAS RETURNED IN A LENS VIAL AND (B)(6) BAG. VISUAL INSPECTION FOUND NO VISUAL DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.0/+1.5/091 | OS | Physician | 07/31/2019 | 2019 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8939420&pc=QCB | 2023826-2019-01531 | 2019-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.00/2.5/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WAS OBSERVED. THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL DRY WITH CLEAR SURGICAL RESIDUE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.00/2.5/099 | OD |  | 08/07/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941307&pc=QCB | 2023826-2019-01582 | 2019-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FOREIGN RESIDUE ON LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -12.0/+1.5/094 | OS |  | 08/05/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8941315&pc=QCB | 2023826-2019-01581 | 2019-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE. VISUAL INSPECTION FOUND HAPTIC TORN AND FOREIGN RESIDUE ON LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -12.0/+1.0/089 | OD |  | 08/05/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8951569&pc=QCB | 2023826-2019-01618 | 2019-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -17.5/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -17.5/2.0/090 | OS |  | 08/19/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8952253&pc=QCB | 2023826-2019-01619 | 2019-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.0/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM, PATIENT IS HAPPY. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND A BENT HAPTIC. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.0/2.0/093 | OD |  | 08/19/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8957683&pc=QCB | 2023826-2019-01539 | 2019-05-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -13.50/1.5/127 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED DURING A SECOND SURGERY ON (B)(6) 2019 FOR A SMALLER AXIS LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION - DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.50/1.5/127 | OS |  | 08/08/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8962357&pc=QCB | 2023826-2019-01615 | 2019-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -13.00/+2.00/70 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT BUT IT DID NOT RESOLVE THE PROBLEM. Manufacturer Narrative: DEVICE CODE 1494: AGE < 21 YEARS OLD SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. REFERENCE MFR# 2023826-2019-01665 FOR EXCHANGE LENS. DEVICE EVALUATION: DIMENSIONAL AND FUNCTIONAL INSPECTION DID NOT FIND THE LENS TO BE WITHIN SPECIFICATIONS. INVESTIGATION FOUND THAT IT IS LIKELY THAT THE CUSTOMERS SWITCHED THE INITIAL AND EXCHANGE LENSES WHEN PACKAGING THEM FOR RETURN. CLAIM#: (B)(4). Manufacturer Narrative: H6 - METHOD CODE 3331: DEVICE HISTORY RECORD REVIEW FOUND ASSOCIATED SOURCE LOTS WERE MANUFACTURED WITHIN ESTABLISHED PARAMETERS AND LIMITS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.00/+2.00/70 | OD |  | 08/16/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8962452&pc=QCB | 2023826-2019-01589 | 2019-05-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.00/+1.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE LENS HAS A CURVED SHAPE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.00/+1.0/093 | OS |  | 08/05/2019 | 2019 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8962480&pc=QCB | 2023826-2019-01591 | 2019-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+2.0/086 | OD |  | 08/06/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8962493&pc=QCB | 2023826-2019-01592 | 2019-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+2.5/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2019 BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.5/+2.5/087 | OD |  | 08/06/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8962498&pc=QCB | 2023826-2019-01593 | 2019-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/003 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. UDI CORRECTED TO N/A IN INITIAL MDR. DEVICE CODE 1494 (OFF-LABEL USE, ACD < 3.0MM) SHOULD HAVE BEEN INCLUDED IN INITIAL MDR, METHOD CODE 4111 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/+1.5/003 | OD |  | 08/06/2019 | 2019 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8962917&pc=QCB | 2023826-2019-01616 | 2019-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT LENS A 13.2MM, VTICMO13.2, -15.00/+1.00/114 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: PREVIOUS MDR IS NOT APPLICABLE FOR THIS CLAIM. THE MFR REPORT NUMBER WAS INPUT INCORRECTLY. THE PREVIOUS MDR WAS SUPPOSED TO BE SUBMITTED FOR CLAIM# (B)(4) BUT WAS ACCIDENTALLY SUBMITTED FOR CLAIM# (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). 16-JUN-2019 SHOULD BE CORRECTED TO (B)(6) 2019 IN INITIAL MEDWATCH REPORT. DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. DEVICE EVALUATION: LENS WAS RETURNED DRY INSIDE A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.00/+1.00/114 | OS |  | 08/16/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8963063&pc=QCB | 2023826-2019-01620 | 2019-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -17.0/1.5/131 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO OBSERVATION OF EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND A TORN HAPTIC. CLAIM#: (B)(4). Manufacturer Narrative: DATE OF EVENT: CORRECTED TO (B)(6) 2019 IN INITIAL MDR. DATE OF REPORT: CORRECTED TO 19-AUG-2019 IN INITIAL MDR. DEVICE RETURNED TO MANUFACTURER: CORRECTED TO 23-AUG-2019 IN INTIAL MDR. CONCOMITANT MEDICAL PRODUCTS: INJ SYSTEM FTP, LOT# UNK SHOULD HAVE BEEN INCLUDED IN INTIAL MDR RESULT CODE 213 CORRECTED TO 114 IN INTIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -17.0/1.5/131 | OD |  | 08/19/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966502&pc=QCB | 2023826-2019-01641 | 2019-01-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/087 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.0/087 | OS |  | 08/08/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966507&pc=QCB | 2023826-2019-01651 | 2019-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+0.5/008 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: B5 - PER UPDATED INFORMATION, THE 'PATIENT IS HAPPY AND ALL IS FINE.' THE SURGEON PLANS TO REPLACE THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.0/+0.5/008 | OS |  | 08/14/2019 | 2019 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966516&pc=QCB | 2023826-2019-01639 | 2019-05-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+1.5/088 | OS |  | 08/08/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966522&pc=QCB | 2023826-2019-01634 | 2019-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+2.5/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+2.5/089 | OS |  | 08/08/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966524&pc=QCB | 2023826-2019-01633 | 2019-06-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.0/+3.0/069 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.0/+3.0/069 | OS |  | 08/08/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8970007&pc=QCB | 2023826-2019-01542 |  | Injury | STAAR SURGICAL COMPANY | 2019-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.50/2.5/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED. THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS "SIZING ISSUE." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.50/2.5/079 | OS |  | 08/16/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8971769&pc=QCB | 2023826-2019-01596 | 2019-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.5/+1.5/098 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -17.5/+1.5/098 | OS |  | 08/08/2019 | 2019 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8971773&pc=QCB | 2023826-2019-01600 | 2019-06-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/079 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+2.0/079 | OS |  | 08/08/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8971786&pc=QCB | 2023826-2019-01602 | 2019-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -4.0/+1.5/097 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -4.0/+1.5/097 | OS |  | 08/08/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8971799&pc=QCB | 2023826-2019-01606 | 2019-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+1.5/080 | OD |  | 08/08/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8971807&pc=QCB | 2023826-2019-01607 | 2019-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS EXCHANGED FOR ANOTHER ICL LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. UDI CORRECTED TO N/A IN INITIAL MDR. DEVICE CODE 1494 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. METHOD CODE 4111 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. | -13.0/+1.5/082 | OS |  | 08/08/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8983240&pc=QCB | 2023826-2019-01691 | 2019-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.5/+2.0/095 | OD |  | 08/14/2019 | 2019 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8983242&pc=QCB | 2023826-2019-01693 | 2019-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.5/+0.5/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -9.5/+0.5/085 | OD |  | 08/15/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8983243&pc=QCB | 2023826-2019-01690 | 2019-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/167 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: DEVICE CODE 1494: OFF-LABEL USE (PREVIOUS LASER REFRACTIVE EYE SURGERY). (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -9.0/+1.5/167 | OS |  | 08/14/2019 | 2019 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8997595&pc=QCB | 2023826-2019-01679 | 2019-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00/+1.50/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.00/+1.50/087 | OD |  | 08/16/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8997608&pc=QCB | 2023826-2019-01680 | 2019-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50/+3.00/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.50/+3.00/092 | OS |  | 08/16/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9012373&pc=QCB | 2023826-2019-01622 | 2019-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.5/2.0/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO LOW VAULT. A REPLACEMENT LENS OF LONGER LENGTH WAS GOING TO BE IMPLANTED DURING THE SAME SURGERY AS REMOVAL BUT THE LENS GOT DAMAGE WHILE INJECTING INTO THE EYE. THE EYE WAS LEFT WITHOUT ANY PKAKIC LENS. Manufacturer Narrative: DATE OF EVENT: CORRECTED TO (B)(6) 2019 IN INTIAL MDR. CONCOMITANT MEDICAL PRODUCTS: INJECTOR SYSTEM MODEL: FTP, LOT# UNK SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/2.0/099 | OD |  | 08/20/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9038168&pc=QCB | 2023826-2019-01657 | 2019-07-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.5/+2.5/118 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -11.5/+2.5/118 | OS |  | 08/26/2019 | 2019 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9056501&pc=QCB | 2023826-2019-01681 | 2019-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, - 11.00/+1.0/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#:(B)(4). Manufacturer Narrative: (B)(4). | - 11.00/+1.0/086 | OD |  | 08/19/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9070254&pc=QCB | 2023826-2019-01706 |  | Injury | STAAR SURGICAL COMPANY | 2019-09-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL 13.2, -9.00/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: INITIAL MDR SHOULD HAVE INCLUDED: VERY NARROW ANGLES DUE TO IRIS BOWING WERE ALSO REPORTED. PATIENT CODE 3191: VERY NARROW ANGLES DUE TO IRIS BOWING SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: LENS WAS RETURNED IN A ZIPLOC BAG, IN A VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). | -9.00/1.5/089 | NA |  | 08/27/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9076355&pc=QCB | 2023826-2019-01719 | 2019-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+4.5/069 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+4.5/069 | OS |  | 08/26/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9076357&pc=QCB | 2023826-2019-01721 | 2019-01-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL DRY WITH CLEAR SURGICAL RESIDUE. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -13.0/+1.5/085 | OS |  | 08/27/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9076358&pc=QCB | 2023826-2019-01720 | 2019-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+3.5/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+3.5/091 | OS |  | 08/27/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9076361&pc=QCB | 2023826-2019-01716 | 2019-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -16.0/+1.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED BUT NOT MARKETED IN THE U.S. (B)(4). | -16.0/+1.0/091 | OS |  | 08/21/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9076364&pc=QCB | 2023826-2019-01714 |  | Injury | STAAR SURGICAL COMPANY | 2019-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | No Code Available | P030016 | Event Description: THE REPORTER STATES THAT A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -9.0/+3.0/072 (SPHERE/CYLINDER/AXIS) WAS EXPLANTED FROM THE PATIENT'S LEFT (OS) EYE DUE TO IT BEING THE WRONG POWER LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION ADDED TO PREVIOUS DEVICE EVALUATION: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. METHOD CODE 3331, DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: A1: PATIENT ID: (B)(6). A2: DOB: (B)(6) 1996. D6- IMPLANT DATE: (B)(6) 2019. CLAIM# (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN A BIOHAZARD ZIPLOC BAG WRAPPED IN GAUZE AND WITH LIQUID IN THE BAG. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.0/+3.0/072 | OS |  | 08/21/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9079006&pc=QCB | 2023826-2019-01655 | 2019-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.0/+2.0/109 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT WITH LENS ROTATION. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -7.0/+2.0/109 | OS |  | 08/22/2019 | 2019 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9087432&pc=QCB | 2023826-2019-01725 | 2019-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Insufficient Information | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -10.50/2.5/055 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019 AND EXPLANTED ON (B)(6) 2019. REASON FOR EXPLANT IS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS ROTATION WAS REPORTED. THE LENS WAS EXPLANTED AND REPLACED WITH A LONGER LENGTH/ DIFFERENT POWER LENS ON (B)(6) 2019, BUT THE PROBLEM WAS NOT RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: B5 - 'THE STATUS OF THE EYE IS "IRRITATED AND TEARING A LOT". POST-OP UCVA IS 20/20.' SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MDR#1 CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.50/2.5/055 | OS |  | 08/21/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9087939&pc=QCB | 2023826-2019-01726 |  | Injury | STAAR SURGICAL COMPANY | 2019-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Insufficient Information | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019 AND EXPLANTED ON (B)(6) 2019. REASON FOR EXPLANT IS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EVENT: UPDATED TO STATE: A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.00/1.5/134 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH NOT APPLICABLE IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.00/1.5/134 | OD |  | 08/21/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9088394&pc=QCB | 2023826-2019-01710 | 2019-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/0.5/120 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND OPTIC TORN/HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/0.5/120 | OD |  | 08/30/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9098128&pc=QCB | 2023826-2019-01727 | 2019-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | UNK | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.00/3.0/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019 AS A REPLACEMENT LENS TO CORRECT LOW VAULT. ON (B)(6) 2019 A NEAR FLAT ICL OVER THE CRYSTALLINE LENS WITH NEARLY NO VAULT WAS OBSERVED. DUE TO FEAR OF ICL TOUCHING THE CRYSTALLINE LENS THE SUREGEON DECIDE TO EXPLANT THE LENS WITHOUT ANY FURTHER EXCHANGE. IT WAS REPORTED THAT THE LENS WAS EXPLANTED ON (B)(6) 2019, PATIENTS POST OP BCVA WAS REPORTED TO BE 6/9 PARTIAL AND IOP 14MMHG. Manufacturer Narrative: REFERENCE MFR# 2023826-2019-01728, DEVICE EVALUATION: DIMENSIONAL INSPECTION DID NOT FIND THE LENS TO BE WITHIN SPECIFICATIONS. INVESTIGATION FOUND THAT IT IS LIKELY THAT THE CUSTOMERS SWITCHED THE INITIAL AND EXCHANGE LENSES WHEN PACKAGING THEM FOR RETURN. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: DATE RECEIVED BY MANUFACTURER IN PREVIOUS MDR 24-NOV-2019 IS CORRECTED TO 24-DEC-2019. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CORRECTED DATA: REFERENCE MFR#: 2023826-2019-01728 FOR INITIAL LENS. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TEAR. METHOD CODE 4110: WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/3.0/89 | OD |  | 08/23/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9098129&pc=QCB | 2023826-2019-01728 | 2019-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | UNK | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.00/3.0/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND ROTATION BUT IT DID NOT RESOLVE THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: DATE RECEIVED BY MANUFACTURER IN PREVIOUS MDR 24-NOV-2019 IS CORRECTED TO 24-DEC-2019 (B)(4). Manufacturer Narrative: CORRECTED DATA DEVICE EVALUATION: LENS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC TORN. REFERENCE MFR# 2023826-2019-01727 FOR EXCHANGE LENS. ADDITIONAL INFORMATION: METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. METHOD CODE 4110: LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: REFERENCE MFR# 2023826-2019-01727. DEVICE EVALUATION: DIMENSIONAL INSPECTION DID NOT FIND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION WAS PERFORMED AND FOUND LENS TO BE WITHIN SPECIFICATION. INVESTIGATION FOUND THAT IT IS LIKELY THAT THE CUSTOMERS SWITCHED THE INITIAL AND EXCHANGE LENSES WHEN PACKAGING THEM FOR RETURN. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-DEVICE CODE: 2923 LENS ROTATION. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/3.0/90 | OD |  | 08/23/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9107362&pc=QCB | 2023826-2019-01623 | 2019-08-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.5/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.5/1.5/093 | OS |  | 08/25/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9107447&pc=QCB | 2023826-2019-01624 | 2019-08-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -13.0/3.0/140 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.0/3.0/140 | OS |  | 08/25/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9112372&pc=QCB | 2023826-2019-01625 | 2019-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -5.0/1.0/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FUND NO VISIBLE DAMAGE TO THE LENS. DEVICE CODE 1494: OFF-LABEL USE - ACD < 3.00MM SHOULD BE ADDED TO THE INITIAL MDR. DEVICE CODE 1494: OFF-LABEL USE - AGE < 21 YEARS OLD SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.0/1.0/073 | OS |  | 08/26/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9115520&pc=QCB | 2023826-2019-01209 | 2019-07-06 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/3.0/090 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.00/3.0/090 | OD | Physician | 09/03/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9117726&pc=QCB | 2023826-2019-01781 | 2019-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.0/1.0/161 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.0/1.0/161 | OD |  | 08/30/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9119033&pc=QCB | 2023826-2019-01783 | 2019-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.5/2.0/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019 AS A REPLACEMENT LENS TO CORRECT LOW VAULT WITH ROTATION BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED WITH A NON TORIC LENS OF SAME LENGTH, DIFFERENT DIOPTER AND THE PROBLEM RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CORRECTED DATA: INITIAL MDR SHOULD HAVE STATED: 'THE PATIENT EXPERIENCED LOW VAULT WITH ROTATION AFTER THE INITIAL LENS EXCHANGE.' INITIAL MDR SHOULD HAVE INCLUDED DEVICE CODE 2923 CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.5/2.0/91 | OD |  | 08/30/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9120430&pc=QCB | 2023826-2019-01740 | 2019-08-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+2.5/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -14.5/+2.5/085 | OS |  | 09/01/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9120431&pc=QCB | 2023826-2019-01739 | 2019-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Patient-Device Incompatibility; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.5/+2.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER, DIFFERENT MODEL LENS. REPORTEDLY, THE LENS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THE LENS WAS EXCHANGED DUE TO LENS ROTATION. THE REPORTER STATES THAT THE CAUSE OF "PATIENT-RELATED FACTOR" MEANS THE SURGEON BELIEVES THIS EVENT WAS NOT A SURGICALLY-INDUCED PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.5/+2.0/095 | OS |  | 09/01/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9120487&pc=QCB | 2023826-2019-01787 | 2019-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICTED THAT A 12.6MM, VTICMO12.6, -14.50/3.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2019 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A NON TORIC LENS AND IT WAS REPORTED THAT THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.50/3.0/094 | OD |  | 08/30/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9120890&pc=QCB | 2023826-2019-01750 | 2019-08-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.0/+1.0/101 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CORRECTED DATA: DEVICE CODE 1494: OFF-LABEL USE- ACD < 3.00MM SHOULD BE ADDED TO THE INITIAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.0/+1.0/101 | OS |  | 08/26/2019 | 2019 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9120901&pc=QCB | 2023826-2019-01779 | 2019-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.50/+3.5/088 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING.THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.50/+3.5/088 | OS |  | 06/25/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9120909&pc=QCB | 2023826-2019-01780 | 2019-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50/+3.5/107 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE LENS WAS DISCARDED. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.50/+3.5/107 | OD |  | 06/25/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9125515&pc=QCB | 2023826-2019-01769 | 2019-06-04 | Injury | STAAR SURGICAL COMPANY | 2019-09-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/2.5/067 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON 18/JUL/2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.50/2.5/067 | OS | Physician | 09/10/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9125490&pc=QCB | 2023826-2019-01766 | 2019-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -4.00/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -4.00/1.0/085 | OS |  | 09/10/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9125506&pc=QCB | 2023826-2019-01768 | 2019-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/2.0/068 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/2.0/068 | OD |  | 09/10/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9125517&pc=QCB | 2023826-2019-01771 | 2019-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -8.50/1.0/087 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN, HAPTIC TORN AND RESIDUE ONTHE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.50/1.0/087 | OD |  | 09/10/2019 | 2019 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9125524&pc=QCB | 2023826-2019-01773 | 2019-04-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.00/3.0/109 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE- UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -18.00/3.0/109 | OS |  | 09/10/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9126490&pc=QCB | 2023826-2019-01756 | 2018-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.50/+2.5/101 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REPOSITIONED ON (B)(6) 2018 AND THE PROBLEM WAS NOT RESOLVED. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING AND LENS ROTATION. ANOTHER LENS WAS NOT IMPLANTED. Manufacturer Narrative: WORK ORDER SEARCH FOUND ONE OTHER SIMILAR COMPLAINT TYPE WITHIN ASSOCIATED LOTS. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.50/+2.5/101 | OD |  | 08/27/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9126494&pc=QCB | 2023826-2019-01755 | 2018-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.50/+2.5/071 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REPOSITIONED ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING AND LENS ROTATION. ANOTHER LENS WAS NOT IMPLANTED. Manufacturer Narrative: CORRECTED DATA: DATE OF EVENT:(B)(6) 2018 IN INITIAL MDR IS CORRECTED TO (B)(6) 2018. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.50/+2.5/071 | OS |  | 08/27/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9136934&pc=QCB | 2023826-2019-01805 | 2019-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/078 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.5/+1.0/078 | OS |  | 09/06/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9136935&pc=QCB | 2023826-2019-01814 | 2019-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -6.5/+1.5/100 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS THE LENS WAS EXPLANTED DUE TO LENS ROTATION. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION:LENS RETURNED IN LIQUID IN LENS CASE/VIAL. VISUAL INSPECTION FOUND OPTIC SPLIT AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.5/+1.5/100 | OD |  | 09/10/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9136937&pc=QCB | 2023826-2019-01804 | 2019-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/126 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTED LOW VAULT AND LENS ROTATION. ON (B)(6) 2019 AND (B)(6) 2019 THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SAME SIZE, DIFFERENT AXIS LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID WITH CLEAR SURGICAL RESIDUE/DEBRIS IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND SCRATCH TRACES ON THE LENS. DIMENSIONAL INSPECTION FOUND LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.0/+2.0/126 | OS |  | 09/04/2019 | 2019 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9136938&pc=QCB | 2023826-2019-01812 | 2019-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/082 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEFORMED. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+2.5/082 | OD |  | 09/10/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9136939&pc=QCB | 2023826-2019-01813 | 2019-05-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/170 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.0/170 | OD |  | 09/10/2019 | 2019 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9136942&pc=QCB | 2023826-2019-01811 | 2019-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/096 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+2.0/096 | OS |  | 09/10/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9136943&pc=QCB | 2023826-2019-01810 | 2019-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+0.5/082 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: DEVICE CODE 1494: OFF-LABEL USE- ACD < 3.00MM SHOULD BE ADDED TO THE INITIAL MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+0.5/082 | OD |  | 09/09/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9136945&pc=QCB | 2023826-2019-01808 | 2018-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -13.0/+1.5/094 | OS |  | 09/09/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9139953&pc=QCB | 2023826-2019-01819 | 2019-08-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -14.5/1.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATED BY MFR: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.5/1.0/075 | OS |  | 09/03/2019 | 2019 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9141015&pc=QCB | 2023826-2019-01823 | 2019-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.00/4.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/4.0/090 | OS |  | 09/04/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9141612&pc=QCB | 2023826-2019-01827 | 2019-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1 -13.00/2.00/87 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND A HAPTIC TEAR AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.00/2.00/87 | OS |  | 09/06/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9146049&pc=QCB | 2023826-2019-01824 | 2019-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.5/2.5/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PATIENT CODE 1494: OFF LABEL-USE (PATIENT'S AGE <21) SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. DEVICE CODE 2923: LENS ROTATION SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/2.5/90 | OS |  | 09/04/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9146264&pc=QCB | 2023826-2019-01849 | 2019-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO_12.1, -13.0/2.5/67 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE CODE:2923 FOR LENS ROTATION. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.0/2.5/67 | OS |  | 09/03/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9147829&pc=QCB | 2023826-2019-01795 | 2019-01-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -3.00/6.0/058 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: CORRECTED DATA: EVENT: THE LENS WAS EXPLANTED ON (B)(6) 2019 AND IN A SEPARATE SURGERY LATER THAT DAY A SHORTER LENGTH REPLACEMENT LENS WAS IMPLANTED. THE REPLACEMENT LENS RESOLVED THE PROBLEM. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; LENS EXPLANT. ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT;DEFORMED. METHOD CODE 4110: LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -3.00/6.0/058 | OD |  | 09/16/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9157455&pc=QCB | 2023826-2019-01829 |  | Injury | STAAR SURGICAL COMPANY | 2019-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.50/3.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. LOW VAULT WITH ROTATION WAS OBSERVED. ON (B)(6) 2019 THE LENS WAS REPOSITIONED, BUT THE LENS ROTATED AGAIN VERTICALLY. FOR PATIENTS CURRENT CONDITION THE REPORTER REPORTED, HIGH MIXED ASTIGMATISM AND BCVA WAS REPORTED TO BE 6/12. THE REPORTED INDICATED THAT THE LENS NEEDS TO BE REPLACED WITH A LONGER LENGTH LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE CODE 2923 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. METHOD CODE 4117 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.50/3.5/105 | OD |  | 09/03/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9158765&pc=QCB | 2023826-2019-01837 | 2019-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+2.0/081 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN/AND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+2.0/081 | OS |  | 09/10/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9158774&pc=QCB | 2023826-2019-01839 | 2019-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+1.5/095 | OD |  | 09/10/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9158778&pc=QCB | 2023826-2019-01838 | 2019-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN/AND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+2.0/090 | OD |  | 09/10/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9164176&pc=QCB | 2023826-2019-01876 | 2019-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00/5.0/106 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: EVENT: THE LENS WAS EXCHANGED ON AN UNKNOWN DATE. CORRECTED TO UNK IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.00/5.0/106 | OD |  | 09/10/2019 | 2019 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9164218&pc=QCB | 2023826-2019-01877 | 2019-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.50/3.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.50/3.0/093 | OS |  | 09/10/2019 | 2019 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9164285&pc=QCB | 2023826-2019-01880 | 2019-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50/1.0/081 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATED BY MFR: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.50/1.0/081 | OS |  | 09/10/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9164361&pc=QCB | 2023826-2019-01869 | 2019-08-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+2.0/115 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.5/+2.0/115 | OS |  | 09/12/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9164364&pc=QCB | 2023826-2019-01872 | 2019-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.5/+2.0/077 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT AND LENS ROTATION. THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR AND DEVICE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN THE LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -12.5/+2.0/077 | OS |  | 09/16/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9164366&pc=QCB | 2023826-2019-01871 | 2019-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.5/+4.5/099 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON REPORTS LOW VAULT AND LENS ROTATION. THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR AND DEVICE. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN THE LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -11.5/+4.5/099 | OD |  | 09/16/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9164367&pc=QCB | 2023826-2019-01867 | 2019-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: EVENT: PLEASE DISREGARD PREVIOUS EVENT STATEMENT. DEVICE EVALUATED BY MFR: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.0/+1.0/092 | OD |  | 09/11/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9164416&pc=QCB | 2023826-2019-01882 | 2019-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50/2.0/066 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6)2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.50/2.0/066 | OS |  | 09/10/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9167923&pc=QCB | 2023826-2019-01883 | 2019-05-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.0/1.0/176 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -12.0/1.0/176 | OD |  | 09/10/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9169934&pc=QCB | 2023826-2019-01898 | 2019-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.50/1.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6)2012. SIGNIFFICANT ENDOTHELIAL CELL LOSS WAS OBSERVED, LENS REMAINS IMPLANTED. CAUSE OF EVENT IS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). LENS WAS EXPLANTED ON (B)(6) 2019 AND A POST OP REPORT STATED "NO ENDOTHELIUM ISSUES IN THE FAMILY REPORTED (ECC). CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - METHOD CODE 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.50/1.5/097 | OS |  | 09/09/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9169953&pc=QCB | 2023826-2019-01899 | 2019-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.50/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2012. SIGNIFFICANT ENDOTHELIAL CELL LOSS WAS OBSERVED, LENS REMAINS IMPLANTED. CAUSE OF EVENT IS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: LENS WAS EXPLANTED ON (B)(6) 2019 AND IT WAS REPORTED POST OP THAT "THERE ARE NO ISSUES WITH ENDOTHELIUM ECC REPORTED". PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXPLANT). CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - METHOD CODE 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.50/2.0/090 | OD |  | 09/09/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9173558&pc=QCB | 2023826-2019-01887 | 2017-04-12 | Injury | STAAR SURGICAL COMPANY | 2019-10-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/1.5/104 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.0/1.5/104 | OD | Physician | 09/10/2019 | 2019 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9173575&pc=QCB | 2023826-2019-01875 | 2019-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/1.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: DEVICE CODE 1494 (OFF LABEL USE: AGE < 21 YEARS OLD) SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. ADDITIONAL INFORMATION: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/1.0/093 | OS |  | 09/10/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9174706&pc=QCB | 2023826-2019-01884 | 2019-04-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.0/1.0/081 (SPHERE/CYLINDER/AXIS) , INTO THE PATIENTS LEFT EYE (OS) ON (B)(6)2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT, ACD <3.0MM) SHOULD HAVE BEEN INCLUDED. DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.0/1.0/081 | OS |  | 09/10/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9175114&pc=QCB | 2023826-2019-01890 | 2019-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.0/2.0/92 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL DEBRIS IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.0/2.0/92 | OD |  | 09/12/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9175297&pc=QCB | 2023826-2019-01891 | 2019-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -5.0/2.5/179 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTIVE DATA: EVENT: "5.0/2.5/179" SHOULD BE CORRECTED TO "5.0/2.5/079" IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.0/2.5/179 | OS |  | 09/12/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9175891&pc=QCB | 2023826-2019-01830 | 2019-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.00/3.5/077 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/3.5/077 | OS |  | 09/17/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9176365&pc=QCB | 2023826-2019-01894 | 2019-07-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6 MM VTICM5_12.6, -5.50/5.00/80 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. ON PATIENTS LAST VISIT BCVA WAS 20/25 AND VAULT VALUE WAS 367. Manufacturer Narrative: DEVICE EVALUATED BY MFR: LENS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BENT AND DEFORMED. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.50/5.00/80 | OS |  | 09/10/2019 | 2019 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9176657&pc=QCB | 2023826-2019-01896 | 2019-09-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM TMICL13.2, -8.0/3.0/067 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR NON TORIC LENS AND THE PROBLEM RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE IN A CONTACT LENS CONTAINER. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.0/3.0/067 | OS |  | 09/11/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9179688&pc=QCB | 2023826-2019-01940 | 2019-09-03 | Injury | STAAR SURGICAL COMPANY | 2019-10-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.50/1.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH RESIDUE AND DEBRIS ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.50/1.5/093 | OS | Physician | 09/19/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9179279&pc=QCB | 2023826-2019-01937 | 2019-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50/3.5/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.50/3.5/086 | OD |  | 09/17/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9179424&pc=QCB | 2023826-2019-01939 | 2019-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/1.0/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# : (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.00/1.0/092 | OD |  | 09/17/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9180350&pc=QCB | 2023826-2019-01943 | 2019-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.50/1.50/81 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DATE OF EVENT: (B)(6) 2019. DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.50/1.50/81 | OD |  | 09/19/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9186482&pc=QCB | 2023826-2019-01859 | 2018-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.00/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 FOR A SPHERICAL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: RESULT CODE: 213 SHOULD BE CORRECTED TO 114. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED MOISTURE ON THE LENS, IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/2.0/090 | OS |  | 09/19/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9188677&pc=QCB | 2023826-2019-01949 | 2018-12-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0/+2.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+2.0/090 | OS |  | 09/23/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9188692&pc=QCB | 2023826-2019-01946 | 2019-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.0/+2.5/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A SIZING ISSUE. Manufacturer Narrative: DEVICE EVALUATION : THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THE EVENT DESCRIPTION INITIALLY STATED THAT SURGERY OCCURRED IN THE RIGHT (OD) EYE. CORRECTED INFORMATION: SURGERY OCCURRED IN THE LEFT (OS) EYE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.0/+2.5/089 | OS |  | 09/19/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9193930&pc=QCB | 2023826-2019-01912 | 2019-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS, -9.00/2.5/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE MANUFACTURE DATE: 04-AUG-2019. DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.00/2.5/080 | OD |  | 09/24/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9193968&pc=QCB | 2023826-2019-01913 | 2019-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS, -8.0/1.5/100 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE MANUFACTURE DATE: 04-AUG-2019. DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/1.5/100 | OS |  | 09/24/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9199772&pc=QCB | 2023826-2019-01944 | 2019-08-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT 12.6MM VTICM5_12.6, -10.5/1.5/004 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER AND DIFFERENT DIOPTER LENS DUE TO THE LENS POWER NOT FITTING THE PATIENTS REFRACTION. THE REPORTER STATED THAT THE CAUSE OF THIS EVENT WAS USER ERROR, THE REPORTER STATED " WRONG OCOS CALCULATION OF THE SURGEON, HE USED +CYLINDER INSTEAD OF - CYLINDER". Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: EVENT: "LEFT EYE (OS)" SHOULD BE ADDED INTO THE INITIAL MDR. MFR NARRATIVE: "FOREIGN RESIDUE AND DEBRIS" SHOULD BE ADDED TO THE PREVIOUS MDR. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/1.5/004 | OS |  | 09/19/2019 | 2019 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9199780&pc=QCB | 2023826-2019-01921 | 2018-09-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0/+3.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+3.5/093 | OD |  | 09/17/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9200280&pc=QCB | 2023826-2019-01956 | 2019-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM , TMICL12.6, -13/2.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO RISK OF ANGLE CLOSURE. THIS EXCHANGE RESOLVED THE PROBLEM. PATIENT RELATED FACTOR WAS REPORTED FOR CAUSE OF THIS EVENT. Manufacturer Narrative: EVENT DESCRIPTION: THE LENS WAS EXCHANGED ON (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. (B)(4). DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND A TEAR ON A HAPTIC. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13/2.5/094 | OS |  | 09/18/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9203009&pc=QCB | 2023826-2019-01914 | 2018-01-23 | Injury | STAAR SURGICAL COMPANY | 2019-10-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.50/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE REPORTER INDICATED THE LENS WAS REMOVED ON (B)(6) 2018 DUE TO CEPHALGIA PSYCHOLOGICAL REASONS. REPORTEDLY PATIENT BELIEVES "THAT THE COLLAMER MATERIAL OF THE LENS CAUSES CEPHALGIA/TRIGEMINAL NEURALGIA. TREATMENT FOR "REMAINING CEPHALGIA, NEUROLOGICAL CONSULTATION, TRIGEMINAL NEURALGIA. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.50/1.0/091 | OD | Physician | 09/23/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9203002&pc=QCB | 2023826-2019-01917 | 2019-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -8.00/1.0/094 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.00/1.0/094 | OD |  | 09/25/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9203007&pc=QCB | 2023826-2019-01915 | 2018-05-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.50/2.0/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE REPORTER INDICATED THE LENS WAS REMOVED ON (B)(6) 2018 DUE TO CEPHALGIA PSYCHOLOGICAL REASONS. REPORTEDLY PATIENT BELIEVES "THAT THE COLLAMER MATERIAL OF THE LENS CAUSES CEPHALGIA/TRIGEMINAL NEURALGIA. TREATMENT FOR "REMAINING CEPHALGIA, NEUROLOGICAL CONSULTATION, TRIGEMINAL NEURALGIA. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: CORRECTED DATA: METHOD CODE: 4110 SHOULD BE REMOVED AS IT IS NOT APPLICABLE. RESULT CODE: 213 SHOULD BE CORRECTED TO 114. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/2.0/070 | OS |  | 09/23/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9203011&pc=QCB | 2023826-2019-01916 | 2018-03-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.50/1.5/090 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.50/1.5/090 | OD |  | 09/23/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9203017&pc=QCB | 2023826-2019-01918 | 2019-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -9.00/1.0/088 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN, THE HAPTIC BENT, RED DEBRIS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/1.0/088 | OS |  | 09/25/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9205126&pc=QCB | 2023826-2019-01935 | 2019-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -08.00/2.5/059 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -08.00/2.5/059 | OD |  | 09/30/2019 | 2019 | 53
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9215902&pc=QCB | 2023826-2019-01930 | 2019-09-12 | Injury | STAAR SURGICAL COMPANY | 2019-10-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.50/+2.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO THE LENS WAS IMPLANTED UPSIDE DOWN. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH/DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.50/+2.0/093 | OS | Physician | 09/20/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9215014&pc=QCB | 2023826-2019-01982 | 2019-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM, VTICMO12.1, -14.0/2.0/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. LOW VAULT WITH ROTATION WAS OBSERVED, ON (B)(6) 2019 THE LENS WAS REPOSITIONED BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS EXPLANTED. IT WAS REPORTED THAT ON (B)(6) 2019 A LONGER LENGTH LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.0/2.0/087 | OD |  | 09/22/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9216117&pc=QCB | 2023826-2019-01983 | 2019-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICM5_13.2, -8.00/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). (B)(4). | -8.00/1.5/090 | OD |  | 09/23/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9216264&pc=QCB | 2023826-2019-01984 | 2019-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICM5_12.6, -3.5/1.5//87 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). (B)(4). | -3.5/1.5//87 | OD |  | 09/25/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9221389&pc=QCB | 2023826-2019-01969 | 2019-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.0/+4.5/055 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS "OCOS." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -6.0/+4.5/055 | OD |  | 09/25/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9227389&pc=QCB | 2023826-2019-02007 | 2019-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.50/+2.0/155 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.50/+2.0/155 | OD |  | 09/26/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9227394&pc=QCB | 2023826-2019-02006 | 2019-02-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.0/082 | OD |  | 09/25/2019 | 2019 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9241885&pc=QCB | 2023826-2019-01958 | 2018-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Break | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/036 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION:LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.5/+1.5/036 | OS |  | 09/28/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9245636&pc=QCB | 2023826-2019-01909 | 2019-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS,-10.50/1.0/126 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, AND ANGLE CLOSURE. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED. Manufacturer Narrative: DEVICE EVALUATION : TWO LENSES WERE RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND A SCRATCH ON A HAPTIC AND RESIDUE ON THE LENS SURFACE. REFER TO MDR 2019-01910 FOR FELLOW LENS, CLAIM#: (B)(4). Manufacturer Narrative: EVENT: THE LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD). DEVICE EVALUATION SHOULD HAVE BEEN REPORTED IN SUPPLEMENTAL MDR#1. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. HEALTH PROFESSIONAL, OCCUPATION: UNK. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.50/1.0/126 | OD |  | 10/03/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9245708&pc=QCB | 2023826-2019-01910 | 2019-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS,-12.0/1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT, AND ANGLE CLOSURE. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED. Manufacturer Narrative: EVENT: THE SURGEON IMPLANTED A 13.2MM TMICL13.2 LENS, -10.50/1.5/048 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS). IF FOLLOW-UP, WHAT TYPE: DEVICE EVALUATION SHOULD HAVE BEEN REPORTED IN SUPPLEMENTAL MDR#1. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION : TWO LENSES WERE RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. REFER TO MDR 2019-01909 FOR FELLOW LENS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. HEALTH PROFESSIONAL, OCCUPATION: UNK. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.0/1.5/091 | OS |  | 10/03/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9252137&pc=QCB | 2023826-2019-02034 | 2019-06-07 | Injury | STAAR SURGICAL COMPANY | 2019-10-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO 13.2, -17.0/4.0/85 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2015. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -17.0/4.0/85 | OS | Physician | 10/06/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9252668&pc=QCB | 2023826-2019-02031 |  | Injury | STAAR SURGICAL COMPANY | 2019-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, - 10.50/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO SIZING ISSUE. THE LENS WAS NOT EXCHANGED FOR ANOTHER ICL LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, WRAPPED IN GAUZE IN A BIOHAZARD BAG. VISUAL INSPECTION FOUND A HAPTIC TORN AND BENT. THERE WAS RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. AGE: UNK. DATE OF EVENT: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | - 10.50/+2.0/090 | OD |  | 10/02/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9271543&pc=QCB | 2023826-2019-02039 | 2019-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -12.00/1.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED WITH ADDITIONAL SUTURE ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. LENS "WAS TOO SMALL FOR THE PATIENT EYE, SO THE LENS TURNED AND ENDED UP AT THE WRONG AXIS. I THOUGHT IT WAS PATIENT RELATED, NOT RELATED TO THE ICL ITSELF." PER PATIENT'S LAST VISIT (B)(6) 2019; "CORNEAL OEDEMA ONLY AT THE INCISION SITE PLUS 1 SUTURE AT THE INCISION SITE. THE SUTURE WILL BE REMOVED IN ONE MONTH." REPORTEDLY LENS EXCHANGE WAS TO GET HIGHER VAULT AND PREVENT FUTURE ROTATION. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH RED RESIDUE ON THE LENS. PATIENT CODE 3191: NO CODE AVAILABLE (SUTURE ADDED). CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - AT A FOLLOW UP VISIT IT WAS REPORTED THAT THE CORNEAL EDEMA RESOLVED, THE SUTURE WAS REMOVED ON (B)(6) 2019 AND CORNEAL EDEMA AT THE INCISION SITE WAS OF NORMAL POST-OP CONDITIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/1.5/105 | OS |  | 10/08/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9279659&pc=QCB | 2023826-2019-02068 | 2019-05-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+1.5/090 | OS |  | 10/14/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9284946&pc=QCB | 2023826-2019-02045 | 2019-06-12 | Injury | STAAR SURGICAL COMPANY | 2019-11-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2, -10.0/1.5/134 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.0/1.5/134 | OD | Physician | 10/18/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9284944&pc=QCB | 2023826-2019-02048 | 2019-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2, -13.0/2.0/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.0/2.0/89 | OD |  | 10/18/2019 | 2019 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9284945&pc=QCB | 2023826-2019-02047 | 2019-04-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2, -10.00/1.0/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/1.0/097 | OD |  | 10/18/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9284949&pc=QCB | 2023826-2019-02046 | 2019-01-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2, -16.5/2.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.5/2.0/087 | OD |  | 10/18/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9284954&pc=QCB | 2023826-2019-02055 | 2019-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM, VTICMO13.2, -11.0/1.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON 1(B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIA WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -11.0/1.0/075 | OD |  | 10/21/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9284960&pc=QCB | 2023826-2019-02088 | 2019-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.00/5.5/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT IDENTIFIER - (B)(4). EVENT - POST OP THE REPORTER INDICATED THAT THE EXCHANGED LENS (13.7MM) RESULTED IN EXCESS VAULT. SO THE SURGEON ATTEMPTED TO EXCHANGE THE LENS BACK TO A SHORTER LENGTH LENS (13.2MM) BUT WITH ANOTHER POSITION OF THE LENS AXIS. REPORTER HAD STATED "DURING EXPLANTATION OF THE LENS (13.7MM) THERE WAS AN IRIS PROLAPSE AND BECAUSE OF THIS ADVERSE EVENT THE SURGEON HAD DECIDED TO CANCEL THE INTRAOCULAR SURGERY AND WILL DO A PRK FOR REFRACTIVE CORRECTION AFTERWARDS". SEE CLAIM XXXXXX FOR THE REPLACEMENT LENS COMPLAINT. (B)(4). Manufacturer Narrative: EVENT- REPORTER STATED "PATIENT IS FINE". CLAIM XXXXXX SHOULD BE CORRECTED TO CLAIM# (B)(4). (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL/CASE IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN AND THE OPTIC STRAINED. DIAMETER MEASUREMENTS NOT POSSIBLE DUE TO BROKEN HAPTIC. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: G4 - DATE RECEIVED BY MANUFACTURER: 09-DEC-2019 SHOULD BE CORRECTED TO 05-DEC-2019 IN SUPPLEMENTAL MEDWATCH REPORT #1. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.00/5.5/104 | OS |  | 10/14/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9292254&pc=QCB | 2023826-2019-01976 | 2019-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -08.50/1.5/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A LONGER LENGTH LENS ON (B)(6) 2019 DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CORRECTED DATA: D11- MEDICAL PRODUCTS: INJECTOR MODEL MSI-PF, LOT # 1405501; CARTRIDGE MODEL SFC-45, LOT # 1415538; FOAMTIPPLUNGER MODEL, LOT # 1406581. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -08.50/1.5/079 | OS |  | 10/16/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9292272&pc=QCB | 2023826-2019-02087 | 2018-06-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -5.00/6.0/091 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.00/6.0/091 | OD |  | 10/18/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9292288&pc=QCB | 2023826-2019-01980 | 2019-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -5.00/1.0/091 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND BLACK GROWTH ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H3- DEVICE EVALUATION: "VISUAL INSPECTION FOUND THE HAPTIC TORN" SHOULD BE ADDED TO THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -5.00/1.0/091 | OD |  | 10/18/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9292289&pc=QCB | 2023826-2019-01981 | 2019-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -5.50/1.0/079 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.50/1.0/079 | OS |  | 10/18/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9300094&pc=QCB | 2023826-2019-02103 | 2019-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -08.50/+2.0/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. AFTER THE LENS EXCHANGE, THE ARCH HEIGHT IS NORMAL RANGE, THE INTRAOCULAR PRESSURE IS NORMAL AND VISION IS GOOD. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON LENS WITH CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -08.50/+2.0/083 | OS |  | 10/20/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9305822&pc=QCB | 2023826-2019-02158 | 2019-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM, VTICMO12.6, -12.00/1.5/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE PROBLEM OF LOW VAULT RESOLVED AFTER EXPLANTING THE LENS. THE PATIENTS UCVA IS 20/2000 AFTER EXPLANTING THE LENS, SAME AS BEFORE SURGERY. THE SURGEON DOES NOT PLAN ON IMPLANTING ANOTHER LENS BECAUSE OF THE STATUS OF THE PATIENTS EYE. THE REPORTER STATED " AFTER THE SURGERY THE SURGEON HAD FOUND A PART OF SUSPENSORY LIGAMENT OF THE PATIENT HAD BEEN RUPTURE AND THE VAULT WAS LOW. AFTER RULING OUT THE RUPTURE OF THE SUSPENSORY LIGAMENT CAUSED BY THE SURGERY, THE SURGEON BELIEVED THAT IT WAS RELATED TO THE PATIENTS HIGH MYOPIA AND HE DIDNT NOTICED BEFORE THE SURGERY." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL, DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -12.00/1.5/111 | OD |  | 10/13/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9306584&pc=QCB | 2023826-2019-02097 | 2019-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -14.00/4.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CORRECTED DATA: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.00/4.5/093 | OS |  | 10/17/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9306626&pc=QCB | 2023826-2019-02098 | 2019-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.50/2.50/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A NON TORIC LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.50/2.50/94 | OS |  | 10/17/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9306665&pc=QCB | 2023826-2019-02099 | 2019-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6 MM, VTICM5_12.6, -12.00/2.00/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A NON TORIC LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT:UNK. ETHNICITY:UNK. RACE:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/2.00/90 | OD |  | 10/17/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9306692&pc=QCB | 2023826-2019-01999 | 2019-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -8.00/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.00/1.5/095 | OS |  | 10/21/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9306707&pc=QCB | 2023826-2019-02004 | 2019-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -6.50/1.0/089 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.50/1.0/089 | OD |  | 10/21/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9312745&pc=QCB | 2023826-2019-02081 | 2019-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+2.0/091 | OS |  | 10/14/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9312757&pc=QCB | 2023826-2019-02084 | 2019-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#(B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.5/092 | OD |  | 10/16/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9312764&pc=QCB | 2023826-2019-02085 | 2019-08-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.0/+2.5/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED AND VISUAL INVESTIGATION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#(B)(4). Manufacturer Narrative: CORRECTED DATA: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+2.5/083 | OS |  | 10/16/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9316386&pc=QCB | 2023826-2019-02096 | 2019-05-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICTED THAT A 12.6MM, VTICM5_12.6, -5.00/3.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2019 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -5.00/3.0/096 | OS |  | 10/17/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9317870&pc=QCB | 2023826-2019-02136 | 2018-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/051 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.5/051 | OS |  | 10/21/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9317876&pc=QCB | 2023826-2019-02144 | 2019-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1 MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -7.5/+1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE SURGEON REPORTED "VON HENRICK 2/3 (TEMPORAL) AND 1/4 (NASAL), ANGLE OPEN, PIGMENTS ON THE ICL SURFACE 3+." THE PROBLEM WAS RESOLVED WITH THE EXCHANGE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -7.5/+1.0/091 | OS |  | 10/21/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9317877&pc=QCB | 2023826-2019-02131 | 2019-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+3.5/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#:(B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -15.5/+3.5/085 | OS |  | 10/18/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9317879&pc=QCB | 2023826-2019-02142 | 2019-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.5/+2.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -8.5/+2.5/087 | OD |  | 10/21/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9317881&pc=QCB | 2023826-2019-02132 | 2019-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.5/091 | OD |  | 10/18/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9317887&pc=QCB | 2023826-2019-02134 | 2019-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/079 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -9.5/+2.0/079 | OS |  | 10/18/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9317890&pc=QCB | 2023826-2019-02128 | 2019-01-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -8.5/+2.0/085 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+2.0/085 | OS |  | 10/16/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9317908&pc=QCB | 2023826-2019-02126 | 2019-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -13.5/+1.5/093 | OD |  | 10/16/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9317910&pc=QCB | 2023826-2019-02127 | 2019-01-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7 MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/105 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON 2(B)(6)2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN WITH FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.0/105 | OD |  | 10/16/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9317911&pc=QCB | 2023826-2019-02143 | 2019-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -6.0/+0.5/098 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE SURGEON REPORTED "VON HENRICK 2/3 (TEMPORAL) AND 1/4 (NASAL), ANGLE OPEN, PIGMENTS ON THE ICL SURFACE 3+." THE PROBLEM WAS RESOLVED WITH THE EXCHANGE. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.0/+0.5/098 | OD |  | 10/21/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9319110&pc=QCB | 2023826-2019-02086 | 2019-07-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+1.0/089 | OD |  | 10/16/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9319142&pc=QCB | 2023826-2019-02104 | 2019-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.0/+3.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH RESIDUE AND MOISTURE ON THE LENS. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+3.0/091 | OS |  | 10/16/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9319156&pc=QCB | 2023826-2019-02107 | 2019-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.5/+3.0/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+3.0/083 | OS |  | 10/18/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9319159&pc=QCB | 2023826-2019-02108 | 2019-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.5/+3.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN, BENT AND DEFORMED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+3.0/093 | OD |  | 10/18/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9331843&pc=QCB | 2023826-2019-02110 | 2018-12-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE AVAILABLE FOR EVALUATION: "NO" SHOULD BE REMOVED FROM SUPPLEMENTAL MEDWATCH REPORT #1 AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.0/090 | OD |  | 10/20/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9331845&pc=QCB | 2023826-2019-02112 | 2019-07-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND FOREIGN DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.5/094 | OD |  | 10/20/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9331852&pc=QCB | 2023826-2019-02113 | 2019-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/069 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE AVAILABLE FOR EVALUATION IN PREVIOUS MDR REPORTED AS NO IS NOT APPLICABLE. CORRECTED TO YES. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+2.0/069 | OD |  | 10/20/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9331880&pc=QCB | 2023826-2019-02114 | 2019-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: DEVICE CODE 1494: OFF-LABEL USE; ACD< 3.00MM SHOULD BE ADDED IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: CONCOMITANT MEDICAL PRODUCTS: "NO" IS NOT APPLICABLE IN SUPPLEMENTAL MDR #1. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+1.0/100 | OS |  | 10/20/2019 | 2019 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9357257&pc=QCB | 2023826-2019-02193 | 2019-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/069 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. ANOTHER LENS WAS NOT IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE/DEBRIS ON THE LENS. CLAIM#: 714962. Manufacturer Narrative: H3 - DEVICE EVALUATION: DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.5/+1.5/069 | OS |  | 10/31/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9357282&pc=QCB | 2023826-2019-02192 | 2019-08-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.50/+3.0/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. ANOTHER LENS WAS NOT IMPLANTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS CASE/ VIAL WITH CLEAR SURGICAL RESIDUE/ DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.50/+3.0/083 | OD |  | 10/30/2019 | 2019 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9357322&pc=QCB | 2023826-2019-02194 | 2019-09-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+3.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN WITH FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/+3.5/094 | OS |  | 10/31/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9357363&pc=QCB | 2023826-2019-02195 | 2019-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/110 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. ANOTHER LENS WAS NOT IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE/DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.5/+1.0/110 | OD |  | 11/01/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9357397&pc=QCB | 2023826-2019-02216 | 2019-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.0/+0.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.0/+0.5/093 | OS |  | 10/28/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9362222&pc=QCB | 2023826-2019-02182 | 2019-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT HE ATTEMPTED TO INSERT A13.2MM, TMICL13.2, -13.00/2.00/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) TWO TIMES AND THEN SWITCHED TO THE BACKUP LENS. THE REPORTER STATED " WHEN I INITIALLY TRIED TO INSERT THE ICL, IT WAS NOT GLIDING THROUGH THE INCISION AND UNFOLDING PROPERLY. I THOUGHT THAT MAY BE DUE TO MY INCISION BEING TOO SMALL SO I THEN I ENLARGED MY INCISION SLIGHTLY. WHEN I TRIED TO INSERT THE ICL AGAIN THE PLUNGER STUCK TO THE ICL, I DIDN'T KNOW IF IT WAS AN ISSUE WITH THE ICL, PLUNGER, CARTRIDGE, OR INJECTOR SO I SWITCHED EVERYTHING OUT IN ORDER TO ENSURE I HAD AN ICL WITH PROPER INTEGRITY FOR THE PATIENT." THE ICL TOUCHED THE PATIENTS EYE BUT WAS NEVER FULLY INSERTED INTO THE PATIENTS EYE. IT WAS ONLY PARTIALLY INSERTED. THE BACKUP LENS WAS USED WITH NO COMPLICATIONS. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.00/2.00/092 | OD |  | 10/25/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9364593&pc=QCB | 2023826-2019-02156 | 2019-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/1.0/085 | OD |  | 11/05/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9372932&pc=QCB | 2023826-2019-02214 |  | Injury | STAAR SURGICAL COMPANY | 2019-11-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, - 10.50/+2.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO SIZING ISSUE. THE LENS WAS REPLACED WITH ANOTHER TORIC LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | - 10.50/+2.0/093 | OD | Other Health Care Professional | 11/11/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9372951&pc=QCB | 2023826-2019-02238 | 2019-10-23 | Injury | STAAR SURGICAL COMPANY | 2019-11-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, - 14.00/+2.5/100 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD). THE LENS WAS EXPLANTED DUE TO BEING TOO LARGE. THE LENS WAS REPLACED WITH ANOTHER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | - 14.00/+2.5/100 | OD | Other Health Care Professional | 11/07/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9373322&pc=QCB | 2023826-2019-02218 | 2019-10-22 | Injury | STAAR SURGICAL COMPANY | 2019-11-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -16.0/+3.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. REPORTEDLY THE LENS FLIPPED AFTER IMPLANT. DURING EXPLANT THE LENS WAS DAMAGED. A BACKUP LENS WAS IMPLANTED. IT IS UNKNOWN IF THE LENS WAS REMOVED AND/OR REPLACED INTRAOPERATIVELY. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DEVICE EVALUATION: THE LENS WAS RETURNED DRY WRAPPED IN A PEEL POUCH. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.0/+3.5/088 | OD | Other Health Care Professional | 11/04/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9373327&pc=QCB | 2023826-2019-02220 | 2019-07-17 | Injury | STAAR SURGICAL COMPANY | 2019-11-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+3.5/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+3.5/092 | OD | Physician | 10/30/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9369703&pc=QCB | 2023826-2019-02252 | 2019-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.00/6.00/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: D7 SHOULD HAVE BEEN INCLUDED IN THE INTIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE: ACD< 3.00MM SHOULD BE ADDED IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED WITH FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.00/6.00/102 | OS |  | 10/30/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9372984&pc=QCB | 2023826-2019-02247 | 2019-10-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+4.5/107 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED SO THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: D10 - DEVICE AVAILABLE FOR EVALUATION: "NO" SHOULD BE REMOVED FROM SUPPLEMENTAL MEDWATCH REPORT #1, AS LENS WAS AVAILABLE. B5 - ON (B)(6) 2019 THE LENS WAS REPOSITIONED BUT DID NOT RESOLVE THE PROBLEM. THE LENS WAS EXCHANGED ON (B)(6) 2019, INTRAOPERATIVELY DURING INTIAL LENS REMOVAL. H6 - DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.5/+4.5/107 | OD |  | 11/06/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9373317&pc=QCB | 2023826-2019-02235 | 2017-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+3.5/097 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2019, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. REPORTEDLY, THE SURGEON HAS NO PLANS TO IMPLANT A REPLACEMENT LENS. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. VISUAL INSPECTION FOUND THERE WAS CLEAR SURGICAL RESIDUE/DEBRIS AND BLOOD-LIKE RESIDUE ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6-METHOD CODE 4110: LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CORRECTED DATA: H6-DEVICE CODE: 1494 OFF LABEL IN INITIAL MDR IS NOT APPLICABLE. CLAIM# (B)(4). Manufacturer Narrative: B4 - CORRECTED TO 06-NOV-2019 IN INITIAL MDR G4 - CORRECTED TO 06-NOV-2019 IN INITIAL MDR H3 - 'VISUAL INSPECTION ALSO FOUND NO VISIBLE DAMAGE TO THE LENS.' SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MDR#1 CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+3.5/097 | OS |  | 11/06/2019 | 2019 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9373318&pc=QCB | 2023826-2019-02233 | 2019-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.0/+3.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE DEVICE WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.0/+3.0/093 | OD |  | 11/04/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9373319&pc=QCB | 2023826-2019-02219 | 2019-05-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.5/+2.5/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION, WHICH DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/+2.5/080 | OS |  | 10/30/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9373325&pc=QCB | 2023826-2019-02230 | 2019-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.5/+3.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B4 - CORRECTED TO (B)(6)2019 IN INITIAL MDR. G4 - CORRECTED TO (B)(6)2019 IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -12.5/+3.0/087 | OD |  | 11/04/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9373329&pc=QCB | 2023826-2019-02222 | 2019-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/083 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+1.5/083 | OS |  | 10/31/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9373330&pc=QCB | 2023826-2019-02232 | 2019-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+4.0/072 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE DEVICE WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+4.0/072 | OS |  | 11/04/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9373332&pc=QCB | 2023826-2019-02234 | 2017-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+3.0/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2019, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. REPORTEDLY, THE SURGEON HAS NO PLANS TO IMPLANT A REPLACEMENT LENS. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN THE LENS CASE/VIAL. THERE WAS CLEAR SURGICAL DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: METHOD CODE 4110: LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. DEVICE CODE: 1494 OFF LABEL IN INITIAL MDR IS NOT APPLICABLE. CLAIM# (B)(4). Manufacturer Narrative: B4 - CORRECTED TO 06-NOV-2019 IN INITIAL MDR G4 - CORRECTED TO 06-NOV-2019 IN INITIAL MDR CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+3.0/080 | OD |  | 11/06/2019 | 2019 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9386161&pc=QCB | 2023826-2019-02215 | 2019-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.0/+1.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.0/+1.5/093 | OS |  | 11/04/2019 | 2019 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9386162&pc=QCB | 2023826-2019-02239 | 2019-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.0/+2.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.0/+2.0/093 | OD |  | 11/04/2019 | 2019 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9386176&pc=QCB | 2023826-2019-02241 | 2019-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-11-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.0/+5.5/101 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. AGE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.0/+5.5/101 | OD |  | 11/04/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9396906&pc=QCB | 2023826-2019-02256 | 2019-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -7.5/1.5/165 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND CLEAR RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: B4 - CORRECTED TO 04-NOV-2019 IN INITIAL MDR. G4 - CORRECTED TO 04-NOV-2019 IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.5/1.5/165 | OD |  | 11/04/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9398697&pc=QCB | 2023826-2019-02259 | 2019-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.0/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: G4 - DATE RECEIVED BY MANUFACTURER: 03-NOV-2019 SHOULD BE CORRECTED TO 04-NOV-2019 IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#:(B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.0/1.0/090 | OD |  | 11/04/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9399004&pc=QCB | 2023826-2019-02301 | 2019-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -15.0/3.5/167 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2019 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON 02 OCT 2019 THE LENS WAS EXCHANGED FOR THE SAME SIZE DIFFERENT AXIS LENS. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: EVENT DATE: "(B)(6) 2019" SHOULD BE CORRECTED TO "(B)(6) 2018" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.0/3.5/167 | OS |  | 11/03/2019 | 2019 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9399096&pc=QCB | 2023826-2019-02300 | 2019-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -6.5/1.0/030 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED. IT WAS REPORTED THAT THE PATIENT REQUESTED REMOVAL. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT WAS PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED. CLAIM# (B)(4). Manufacturer Narrative: DATE OF REPORT: "(B)(6) 2019" SHOULD BE CORRECTED TO "(B)(6) 2019" IN THE INITIAL MDR. DATE RECEIVED BY MANUFACTURER: "(B)(6) 2019" SHOULD BE CORRECTED TO "(B)(6) 2019" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.5/1.0/030 | OS |  | 11/04/2019 | 2019 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9402944&pc=QCB | 2023826-2019-02260 | 2019-09-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -17.0/2.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -17.0/2.5/080 | OD |  | 11/05/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9411436&pc=QCB | 2023826-2019-02306 | 2019-06-18 | Injury | STAAR SURGICAL COMPANY | 2019-12-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.5/1.5/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). | -8.5/1.5/93 | OS | Physician | 11/06/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9412429&pc=QCB | 2023826-2019-02290 | 2019-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5/+1.0/097 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.5/+1.0/097 | OS |  | 11/11/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9412432&pc=QCB | 2023826-2019-02283 | 2019-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+2.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.0/+2.5/092 | OD |  | 11/11/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9412433&pc=QCB | 2023826-2019-02282 | 2019-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/077 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -12.0/+2.5/077 | OS |  | 11/11/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9412435&pc=QCB | 2023826-2019-02277 | 2019-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+1.0/081 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+1.0/081 | OS |  | 11/11/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9412443&pc=QCB | 2023826-2019-02279 | 2019-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -12.5/+1.0/103 | OS |  | 11/11/2019 | 2019 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9412446&pc=QCB | 2023826-2019-02286 | 2019-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+0.5/013 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -12.0/+0.5/013 | OS |  | 11/11/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9412448&pc=QCB | 2023826-2019-02291 | 2019-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/081 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+1.5/081 | OD |  | 11/11/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9412450&pc=QCB | 2023826-2019-02287 | 2019-08-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/095 (SPHERE /CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC DEFORMED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.5/095 | OS |  | 11/11/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9412454&pc=QCB | 2023826-2019-02276 | 2019-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+4.0/070 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.0/+4.0/070 | OS |  | 11/11/2019 | 2019 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9416956&pc=QCB | 2023826-2019-02198 | 2019-06-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.00/1.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.00/1.0/088 | OS |  | 11/11/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9416968&pc=QCB | 2023826-2019-02199 | 2019-08-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.00/2.5/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: B5: "13.2MM VTICMO13.2" SHOULD BE CORRECTED TO "12.6MM VTICMO12.6" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.00/2.5/102 | OD |  | 11/11/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9417013&pc=QCB | 2023826-2019-02203 | 2019-08-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.50/1.5/091 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). | -9.50/1.5/091 | OD |  | 11/11/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9417031&pc=QCB | 2023826-2019-02207 | 2019-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -3.50/1.0/086 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC DEFORMED. CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -3.50/1.0/086 | OS |  | 11/11/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9417040&pc=QCB | 2023826-2019-02208 | 2019-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.50/1.5/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.50/1.5/099 | OD |  | 11/11/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9426994&pc=QCB | 2023826-2019-02264 | 2019-10-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -3.00/5.5/122 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -3.00/5.5/122 | OD |  | 11/14/2019 | 2019 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9427136&pc=QCB | 2023826-2019-02265 | 2019-09-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.00/2.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2019, "REMOVING LENS RESOLVED PROBLEM, THE PATIENT'S BCVA IS 20/33 THE SAME AS PRE-OP, NO PLAN TO IMPLANT ANOTHER LENS." Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS FOREIGN MATERIAL ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -17.00/2.5/089 | OS |  | 11/18/2019 | 2019 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9427140&pc=QCB | 2023826-2019-02269 | 2019-11-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -6.00/0.5/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND A HAPTIC TEAR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.00/0.5/072 | OD |  | 11/18/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9427277&pc=QCB | 2023826-2019-02266 | 2019-09-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.50/2.5/085 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER REPORTER ON (B)(6) 2019, "REMOVING LENS SOLVED PROBLEM, PATIENT'S BCVA IS 20/33 THE SAME AS PRE-OP, NO PLAN TO IMPLANT ANOTHER LENS." Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.50/2.5/085 | OD |  | 11/18/2019 | 2019 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9432237&pc=QCB | 2023826-2019-02322 | 2019-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/2.0/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.5/2.0/073 | OS |  | 11/11/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9432279&pc=QCB | 2023826-2019-02327 | 2019-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/2.0/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE, RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AS WELL AS RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -11.5/2.0/082 | OD |  | 11/11/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9434019&pc=QCB | 2023826-2019-02329 | 2019-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/2.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -11.5/2.0/093 | OS |  | 11/11/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9434133&pc=QCB | 2023826-2019-02330 | 2019-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/2.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -8.5/2.5/092 | OD |  | 11/11/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9441076&pc=QCB | 2023826-2019-02340 | 2019-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6 MM VTICM5_12.6, -11.5/1.5/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM RESOLVED. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -11.5/1.5/91 | OD |  | 11/14/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9441172&pc=QCB | 2023826-2019-02356 | 2019-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.5/+2.0/101 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS THE LENS WAS EXCHANGED WITH A LONGER LENS WHICH RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION (LENS REPOSITIONING) SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: LENS WAS RETURNED WITH MOISTURE AND FOREIGN RESIDUE IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -14.5/+2.0/101 | OD |  | 11/18/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9441176&pc=QCB | 2023826-2019-02348 | 2019-10-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/149 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE ON (B)(6) 2019. ON THE SAME DAY THE LENS WAS EXCHANGED WITH AN ALTERNATE LENS BECAUSE THE LENS SIZE WAS "TOO BIG." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.5/+1.0/149 | NA |  | 11/15/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9441179&pc=QCB | 2023826-2019-02357 | 2019-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+3.0/071 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER AND DIFFERENT MODEL LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -11.0/+3.0/071 | OS |  | 11/18/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9441183&pc=QCB | 2023826-2019-02355 | 2019-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+6.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.0/+6.0/090 | OD |  | 11/18/2019 | 2019 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9441186&pc=QCB | 2023826-2019-02351 | 2019-01-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.0/+0.5/074 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.0/+0.5/074 | OS |  | 11/13/2019 | 2019 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9441187&pc=QCB | 2023826-2019-02350 | 2019-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+3.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+3.0/089 | OS |  | 11/13/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9448284&pc=QCB | 2023826-2019-02338 |  | Injury | STAAR SURGICAL COMPANY | 2019-12-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM, TMICL13.7, -9.5/4.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS LEFT EYE (OS) ON (B)(6)2019. IT WAS REPORTED THAT THE LENS WAS LATER EXPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANTED: UNK. (B)(4). WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.5/4.0/093 | OS | Other Health Care Professional | 11/11/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9448440&pc=QCB | 2023826-2019-02339 |  | Injury | STAAR SURGICAL COMPANY | 2019-12-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM, TMICL13.7, -9.5/4.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. IT WAS REPORTED THAT THE LENS WAS LATER EXPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANTED:UNK. (B)(4). WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.5/4.0/091 | OD | Other Health Care Professional | 11/11/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9446581&pc=QCB | 2023826-2019-02341 | 2019-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.0MM, TICM120V4, -11.0/4.0/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2012. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PROBLEM RESOLVED. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT A TICM120V4 12.0MM; -11.0/4.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2012. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A VTICM5 13.2MM; -8.50/3.0/091 (SPHERE/CYLINDER/AXIS) AND THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE /DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/4.0/89 | OS |  | 11/14/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9447599&pc=QCB | 2023826-2019-02336 | 2019-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.50/4.5/138 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. IT WAS ALSO REPORTED THAT REPOSITIONING OF THE LENS WAS PERFORMED AND THE PROBLEM RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -9.50/4.5/138 | OS |  | 11/11/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9447766&pc=QCB | 2023826-2019-02337 | 2019-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.50/5.5/22 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019 AND EXCESSIVE VAULT WAS OBSERVED. LENS REPOSITIONING RESOLVED THE PROBLEM. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FOREIGN PURPLE RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -9.50/5.5/22 | OD |  | 11/11/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9454406&pc=QCB | 2023826-2019-02369 |  | Injury | STAAR SURGICAL COMPANY | 2019-12-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H3- "HAPTIC TORN" SHOULD BE CORRECTED TO "OPTIC TORN" IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.0/+1.5/096 | OS |  | 11/11/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9454418&pc=QCB | 2023826-2019-02213 | 2019-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+2.0/088 | OS |  | 11/11/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9454432&pc=QCB | 2023826-2019-02370 | 2019-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # 715023 | -8.0/+1.5/085 | OS |  | 11/11/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9454433&pc=QCB | 2023826-2019-02365 | 2019-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -13.0/+2.0/090 | OS |  | 11/11/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9454441&pc=QCB | 2023826-2019-02373 | 2019-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.5/087 | OD |  | 11/11/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9454450&pc=QCB | 2023826-2019-02374 | 2019-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.0/+1.5/077 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -16.0/+1.5/077 | OD |  | 11/11/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9454473&pc=QCB | 2023826-2019-02378 | 2019-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5/+1.0/078 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6- DEVICE CODE 1494- OFF-LABEL USE: AGE < 21 YEARS OLD SHOULD BE ADDED IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEFORMED. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # 715031 | -6.5/+1.0/078 | OD |  | 11/11/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9454485&pc=QCB | 2023826-2019-02368 | 2019-01-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+2.0/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.0/+2.0/085 | OD |  | 11/11/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9461815&pc=QCB | 2023826-2019-02381 | 2019-03-06 | Injury | STAAR SURGICAL COMPANY | 2019-12-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS WAS DISCARDED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.0/+2.5/084 | OS | Physician | 11/13/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9461818&pc=QCB | 2023826-2019-02382 | 2019-03-06 | Injury | STAAR SURGICAL COMPANY | 2019-12-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+3.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS WAS DISCARDED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -14.0/+3.0/090 | OD | Physician | 11/13/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9477066&pc=QCB | 2023826-2019-02414 | 2019-10-01 | Injury | STAAR SURGICAL COMPANY | 2019-12-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -9.0/1.0/135 (SPHERE/CYLINDER\AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -9.0/1.0/135 | OD | Physician | 11/18/2019 | 2019 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9476085&pc=QCB | 2023826-2019-02412 | 2019-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/1.5/085 (SPHERE/CYLINDER\AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -11.0/1.5/085 | OS |  | 11/18/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9476330&pc=QCB | 2023826-2019-02413 | 2019-08-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/2.5/094 (SPHERE/CYLINDER\AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -12.0/2.5/094 | OD |  | 11/18/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9490525&pc=QCB | 2023826-2019-02420 | 2019-10-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -6.50/+6.00/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019, EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS OBSERVED. THE REPORTER STATED THAT THE SURGEON HAS DECIDED TO MONITOR THE PATIENT FOR NOW, IF NECESSARY AN EXCHANGE WILL BE PERFORMED. PATIENTS CURRENT CONDITION WAS REPORTED AS " STABLE VISION, SHALLOW ANTERIOR CHAMBER WITH NARROW ANGLES AND NORMAL IOP". BCVA WAS REPORTED TO BE 20/30. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT IS DUE TO THE SIZE OF THE ICL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: D11 - CONCOMITANT MEDICAL PRODUCTS AND THERAPY DATES: INJECTOR MODEL-MSI-TF; LOT#-UNK SHOULD BE CORRECTED TO INJECTOR MODEL-MSI-PF; LOT#-UNK. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (NARROW ANGLES, SHALLOW AC). H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.50/+6.00/90 | OS |  | 11/22/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9492297&pc=QCB | 2023826-2019-02402 | 2019-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.0/091 | OD |  | 11/18/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9512603&pc=QCB | 2023826-2019-02427 | 2019-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.5/+2.5/084 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER, DIFFERENT MODEL LENS DUE TO LENS TILTING. THE PROBLEM WAS NOT RESOLVED. REFERENCE MFR REPORT 2023826-2019-02429 FOR REPLACEMENT LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.5/+2.5/084 | OS |  | 11/28/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9512610&pc=QCB | 2023826-2019-02426 | 2019-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -7.0/+5.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.0/+5.5/091 | OD |  | 11/28/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9512614&pc=QCB | 2023826-2019-02428 | 2019-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -7.0/+4.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.0/+4.0/091 | OS |  | 11/28/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9512616&pc=QCB | 2023826-2019-02442 | 2019-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/098 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM #: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -9.5/+1.0/098 | OS |  | 12/03/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9512737&pc=QCB | 2023826-2019-02480 | 2019-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICM5_13.2, -6.5/1.5/79 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.5/1.5/79 | OD |  | 12/01/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9515902&pc=QCB | 2023826-2019-02313 | 2019-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.00/2.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM#(B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4) | -18.00/2.0/092 | OS |  | 12/03/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9515907&pc=QCB | 2023826-2019-02317 | 2019-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -11.50/1.5/063 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.50/1.5/063 | OD |  | 12/03/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9515912&pc=QCB | 2023826-2019-02316 | 2019-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -11.00/1.5/094 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/1.5/094 | OS |  | 12/03/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9516570&pc=QCB | 2023826-2019-02319 | 2019-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/0.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON 05/JUL/2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/0.5/094 | OD |  | 12/04/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9516580&pc=QCB | 2023826-2019-02452 | 2019-11-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -18.00/2.5/105 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA) SHOULD ALSO HAVE BEEN REPORTED IN INITIAL MEDWATCH REPORT (MDR). F3, F4, F5-INFORMATION WAS ENTERED IN ERROR IN INITIAL MDR. H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS CLEAR SURGICAL RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. H6-PATIENT CODE 3191: NO CODE AVAILABLE (SIGNIFICANT REDUCTION OF ICA) CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -18.00/2.5/105 | OD |  | 12/05/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9530979&pc=QCB | 2023826-2019-02487 | 2019-03-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/1.5/086 (SPHERE/CYLINDER\AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO A PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS . VISUAL INSPECTION FOUND THE HAPTIC TORN AS WELL AS RESIDUE AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/1.5/086 | OD |  | 12/04/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9531933&pc=QCB | 2023826-2019-02488 | 2019-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/1.5/091 (SPHERE/CYLINDER\AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE, AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.0/1.5/091 | OD |  | 12/04/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9532303&pc=QCB | 2023826-2019-02490 | 2019-11-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/1.5/073 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.5/1.5/073 | OS |  | 12/04/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9532309&pc=QCB | 2023826-2019-02462 | 2019-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.0/+2.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO LENS ROTATION. ON (B)(6) 2019 A LONGER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS "FORMULA FOR LENGTH CALCULATION ERROR." Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AND THE LENS CURVED IN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT,ETHNICITY,RACE: UNK. PMA/510K: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -7.0/+2.5/086 | OD |  | 12/04/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9532310&pc=QCB | 2023826-2019-02463 | 2019-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+4.0/079 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT,ETHNICITY,RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+4.0/079 | OS |  | 12/04/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9532312&pc=QCB | 2023826-2019-02464 | 2019-11-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -2.5/+0.5/078 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND BENT WITH DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -2.5/+0.5/078 | OS |  | 12/04/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9532313&pc=QCB | 2023826-2019-02465 | 2019-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+3.5/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT,ETHNICITY,RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) | -11.5/+3.5/089 | OD |  | 12/08/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9532322&pc=QCB | 2023826-2019-02466 | 2019-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.0/+4.0/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT,ETHNICITY,RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+4.0/091 | OS |  | 12/08/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9532328&pc=QCB | 2023826-2019-02469 | 2019-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -13.0/+0.5/081 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. AN IRIS-FIXATED LENS FROM A DIFFERENT MANUFACTURER WAS THEN IMPLANTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL, DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+0.5/081 | OS |  | 12/08/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9532332&pc=QCB | 2023826-2019-02467 | 2019-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/084 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS.+ CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.5/+1.5/084 | OS |  | 12/08/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9532554&pc=QCB | 2023826-2019-02514 | 2019-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/2.0/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS DUE TO A PATIENT RELATED FACTOR. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/2.0/088 | OS |  | 12/04/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9532664&pc=QCB | 2023826-2019-02515 | 2019-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/1.0/085 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS DUE TO A PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/1.0/085 | OD |  | 12/04/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9533416&pc=QCB | 2023826-2019-02484 | 2019-10-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.00/3.50/92 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.00/3.50/92 | OD |  | 12/05/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9533576&pc=QCB | 2023826-2019-02491 | 2019-11-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2019-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICTED THAT A 12.1MM, VTICM5_12.1, -12.00/1.50/82 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A NON TORIC LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE AND DEBRIS ON THE LENS. DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/1.50/82 | OS |  | 12/05/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9538839&pc=QCB | 2023826-2020-00002 | 2019-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.50/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2019, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED FOR LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5 - LENS WAS REPOSITIONED ON (B)(6) 2019 SHOULD BE CORRECTED TO (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.50/1.5/091 | OS |  | 12/08/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9538852&pc=QCB | 2023826-2020-00004 | 2018-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -5.000/2.0/079 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.000/2.0/079 | OD |  | 12/08/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9541065&pc=QCB | 2023826-2020-00021 |  | Injury | STAAR SURGICAL COMPANY | 2020-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.50/3.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO LOW VAULT. A REPLACEMENT LENS OF LONGER LENGTH WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/3.5/090 | OD |  | 12/05/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9541351&pc=QCB | 2023826-2020-00020 |  | Injury | STAAR SURGICAL COMPANY | 2020-01-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -7.50/2.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO LOW VAULT. A REPLACEMENT LENS OF LONGER LENGTH WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.50/2.0/094 | OS |  | 12/05/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9559178&pc=QCB | 2023826-2020-00029 | 2019-11-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/2.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/2.0/090 | OS |  | 12/08/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9559342&pc=QCB | 2023826-2020-00041 | 2019-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.5/+1.0/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT,ETHNICITY,RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.5/+1.0/092 | OD |  | 12/15/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9559365&pc=QCB | 2023826-2020-00038 | 2019-10-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+4.0/112 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY,RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+4.0/112 | OD |  | 12/12/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9559366&pc=QCB | 2023826-2020-00042 | 2019-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+2.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: : (B)(4). Manufacturer Narrative: WEIGHT,ETHNICITY,RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+2.0/090 | OS |  | 12/15/2019 | 2019 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9559871&pc=QCB | 2023826-2020-00030 | 2019-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.0/4.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM# : (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.0/4.5/089 | OD |  | 12/08/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9560036&pc=QCB | 2023826-2020-00027 | 2019-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.5/2.0/88 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION PATIENT RELATED FACTOR WAS THE CAUSE OF THIS EVENT. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM #: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/2.0/88 | OS |  | 12/08/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9560168&pc=QCB | 2023826-2020-00031 | 2019-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.5/2.0/87 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION PATIENT RELATED FACTOR WAS THE CAUSE OF THIS EVENT. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM #: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/2.0/87 | OD |  | 12/08/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9560415&pc=QCB | 2023826-2020-00032 | 2019-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/2.5/111 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS LATER EXPLANTED DUE TO EXCESSIVE VAULT. ON (B)(6) 2019 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B4-DATE OF THIS REPORT: (B)(6) 2019 IN INITIAL MDR IS CORRECTED TO (B)(6) 2019. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -17.5/2.5/111 | OD |  | 12/10/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9560507&pc=QCB | 2023826-2020-00033 | 2019-11-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0/2.5/074 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ON (B)(6) 2019 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: B4 - DATE OF THIS REPORT: (B)(6) 2019. G4 - DATE RECEIVED BY MANUFACTURER: (B)(6) 2019 SHOULD BE CORRECTED TO (B)(6) 2019 IN INITIAL MEDWATCH REPORT. H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.0/2.5/074 | OS |  | 12/09/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9560605&pc=QCB | 2023826-2020-00063 | 2019-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/2.5/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.5/2.5/095 | OD |  | 12/11/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9560628&pc=QCB | 2023826-2020-00026 | 2019-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.5/1.0/084 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.5/1.0/084 | OS |  | 12/12/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9560716&pc=QCB | 2023826-2020-00057 | 2019-11-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -7.5/3.5/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.5/3.5/092 | OD |  | 12/11/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9565442&pc=QCB | 2023826-2020-00007 |  | Injury | STAAR SURGICAL COMPANY | 2020-01-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -10.50/2.5/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REMOVED AND EXCHANGED FOR A SPHERICAL LENS ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. (B)(4). | -10.50/2.5/070 | OS | Administrator/Supervisor | 12/10/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9565596&pc=QCB | 2023826-2020-00060 | 2019-12-03 | Injury | STAAR SURGICAL COMPANY | 2020-01-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -12.0/3.0/006 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. IT WAS REPORTED THAT AN IOL LENS WAS IMPLANTED ON THE SAME DATE THE ICL WAS REMOVED DUE TO EARLY CATARACT THAT WAS NOT CAUSED BY ICL. PATIENT IS IN GOOD CONDITION. Manufacturer Narrative: DEVICE EVALUATION : LENS WAS RETURNED IN A EYE DROP BOTTLE WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND DEBRIS ON LENS. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.0/3.0/006 | OS | Other Health Care Professional | 12/11/2019 | 2019 | 55
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9565039&pc=QCB | 2023826-2020-00068 | 2019-11-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -13.0/1.0/091 | OD |  | 12/15/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9565710&pc=QCB | 2023826-2020-00006 | 2019-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM513.7, -6.50/5.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. REPORTEDLY, "PATIENT WILL HAVE PRK TREATMENT" DUE TO "IRIS PROLAPSE DURING REMOVAL OF LENS. PLANNED LENS EXCHANGE WAS CANCELED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION- LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: B4- "(B)(6) 2019" SHOULD BE CORRECTED TO "(B)(6) 2019" IN THE INITIAL MDR. B5- "PROBLEM WAS RESOLVED." SHOULD BE CORRECTED TO "REPORTER STATED ALL IS FINE." IN THE INITIAL MDR. H6- PATIENT CODE 3191- IRIS PROLAPSE SHOULD BE ADDED IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.50/5.5/086 | OS |  | 12/05/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9565711&pc=QCB | 2023826-2020-00010 | 2019-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -10.50/3.0/069 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/3.0/069 | OS |  | 12/11/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9565732&pc=QCB | 2023826-2020-00011 | 2019-10-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.50/1.5/153 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.50/1.5/153 | OS |  | 12/11/2019 | 2019 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9565846&pc=QCB | 2023826-2020-00012 | 2019-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.00/2.5/092 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/ 510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.00/2.5/092 | OD |  | 12/11/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9565847&pc=QCB | 2023826-2020-00013 |  | Injury | STAAR SURGICAL COMPANY | 2020-01-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.00/1.0/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WAS OBSERVED, THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/1.0/92 | OD |  | 12/13/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9566305&pc=QCB | 2023826-2020-00052 | 2019-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+2.0/095 | OD |  | 12/11/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9566333&pc=QCB | 2023826-2020-00053 | 2019-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/072 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.5/072 | OS |  | 12/11/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9566337&pc=QCB | 2023826-2020-00054 | 2019-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+3.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+3.0/091 | OD |  | 12/11/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9566355&pc=QCB | 2023826-2020-00055 | 2019-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+2.0/070 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+2.0/070 | OS |  | 12/11/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9578488&pc=QCB | 2023826-2020-00085 | 2019-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.50/+3.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.50/+3.0/093 | OS |  | 12/16/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9578494&pc=QCB | 2023826-2020-00087 | 2019-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.50/+4.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS RE-POSITIONED BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN AND THE LENS WAS RETURNED IN THREE PIECES. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.50/+4.0/094 | OS |  | 12/16/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9578495&pc=QCB | 2023826-2020-00088 | 2019-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+5.0/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: H3- PRODUCT EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. ADDITIONAL INFORMATION: B5: PREVIOUS STATEMENT: THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CORRECTED STATEMENT: AFTER REPOSITIONING EXCESSIVE VAULT WAS OBSERVED. THE LENS WAS EXCHANGED FOR A SHORTER LENS ON 30-NOV- 2019 AND THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6- DEVICE CODE 1583 SHOULD BE ADDED TO THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.00/+5.0/084 | OS |  | 12/16/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9583392&pc=QCB | 2023826-2020-00070 | 2019-12-10 | Injury | STAAR SURGICAL COMPANY | 2020-01-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.0/3.5/099 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -12.0/3.5/099 | OS | Physician | 12/16/2019 | 2019 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9586885&pc=QCB | 2023826-2020-00066 | 2019-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -18.00/3.5/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2019, BUT THIS DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN AND ICL ROTATION. REPORTER NOTED, "ICL ROTATION." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.00/3.5/074 | OS |  | 12/16/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9587396&pc=QCB | 2023826-2020-00110 | 2019-01-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -18.0/+2.5/103 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS NOT RESOLVED. REFERENCE MFR REPORT# 2023826-2019-01312 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H6-METHOD CODE 4110: LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION UNABLE TO BE PERFORMED DUE TO SIGNIFICANT DAMAGE OF THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/+2.5/103 | OD |  | 12/18/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9587408&pc=QCB | 2023826-2020-00108 | 2019-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT,ETHNICITY,RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.0/091 | OD |  | 12/18/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9587409&pc=QCB | 2023826-2020-00109 | 2019-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/129 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENGHT, DIFFERENT MODEL LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+1.0/129 | OD |  | 12/18/2019 | 2019 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9587447&pc=QCB | 2023826-2020-00069 | 2019-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.00/1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019 . ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL THAT WAS SHATTERED, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -9.00/1.5/091 | OS |  | 12/16/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9589371&pc=QCB | 2023826-2020-00101 | 2019-11-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CORRECTED DATA: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+1.5/090 | OS |  | 12/18/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9595576&pc=QCB | 2023826-2020-00079 | 2019-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.50/4.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2019, THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED FOR A SPHERICAL LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -13.50/4.0/090 | OS |  | 12/18/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9597940&pc=QCB | 2023826-2020-00124 | 2019-06-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -10.5/1.0/130 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/1.0/130 | OD |  | 12/22/2019 | 2019 | 49
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9597948&pc=QCB | 2023826-2020-00126 | 2019-10-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -17.00/2.0/125 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -17.00/2.0/125 | OS |  | 12/24/2019 | 2019 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9597954&pc=QCB | 2023826-2020-00127 | 2019-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/1.5/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 -11.0/1.5/104 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXCHANGED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/1.5/104 | OD |  | 12/24/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9604390&pc=QCB | 2023826-2020-00136 | 2019-11-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.5/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+2.0/090 | OS |  | 12/22/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9604394&pc=QCB | 2023826-2020-00137 | 2019-11-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.5/+2.0/073 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH DEBRIS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+2.0/073 | OD |  | 12/22/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9608012&pc=QCB | 2023826-2020-00169 | 2019-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/1.0/163 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -10.5/1.0/163 | OS |  | 12/24/2019 | 2019 | 50
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9608200&pc=QCB | 2023826-2020-00162 |  | Injury | STAAR SURGICAL COMPANY | 2020-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.00/2.0/172 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2019, THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED FOR A "LENS WITH 90 DEGREE ORIENTATION." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: D11-INFORMATION SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A LENS VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/2.0/172 | OD |  | 12/26/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9608693&pc=QCB | 2023826-2020-00173 | 2019-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2019 A 13.2MM, VTICM5_13.2, -12.0./2.0/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE WHILE THE DOCTOR WAS IMPLANTING/INSERTING THE LENS INTO THE PATIENTS LEFT EYE (OS). ON (B)(6) 2019 THE LENS WAS REMOVED AND A REPLACEMENT LENS WAS IMPLANTED WITHIN THE SAME SURGERY. PROBLEM RESOLVED. IN THE REPORTER OPINION THE CAUSE OF THE EVENT WAS THE DEVICE, LENS STUCK IN CARTRIDGE. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE AND CLEAR SURGICAL RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.0./2.0/92 | OS |  | 12/24/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9612920&pc=QCB | 2023826-2020-00153 | 2019-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -6.0/+2.5/079 (SPHERE/CYLINDAR/AXIS) INTO THE PATIENT'S EYE ON (B)(6) 2008. THE SURGEON REPORTS LOW VAULT AND RESIDUAL REFRACTION. REPORTEDLY, THE LENS REMAINS IMPLANTED.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -5.5/+5.5/114 (SPHERE/CYLINDAR/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2008. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND RESIDUAL REFRACTION. THE PROBLEM IS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AND THE LENS CURVED IN. CLAIM#: (B)(4). Manufacturer Narrative: B5 - PLEASE DISREGARD PREVIOUS B5 STATEMENT. H6-ADDITIONAL PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.0/+2.5/079 | OD |  | 12/30/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9612923&pc=QCB | 2023826-2020-00150 | 2019-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -13.0/+3.5/098 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON THE SAME DATE IN A SECONDARY SURGERY, THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -13.0/+3.5/098 | OD |  | 12/26/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9612928&pc=QCB | 2023826-2020-00152 | 2019-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+2.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -14.0/+2.5/094 | OS |  | 12/30/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9624214&pc=QCB | 2023826-2020-00146 | 2019-11-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/101 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+1.5/101 | OS |  | 12/26/2019 | 2019 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9624215&pc=QCB | 2023826-2020-00140 | 2019-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+1.0/115 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.0/+1.0/115 | OD |  | 12/24/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9624219&pc=QCB | 2023826-2020-00145 | 2019-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/056 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE LENS TO NOT BE WITHIN SPECIFICATION. H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS INCLUDING THE SUSPECTED DEVICE HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H10 - SUPP 1 SUBMITTED FOR 2023826-2020-00145 IS NOT ASSOCIATED WITH CLAIM# (B)(4) AND SHOULD BE DISREGARDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.0/056 | OS |  | 12/24/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9624227&pc=QCB | 2023826-2020-00148 | 2019-10-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -13.50/+2.0/099 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL/SAME LENGTH/DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.50/+2.0/099 | OD |  | 12/26/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9624868&pc=QCB | 2023826-2020-00176 | 2019-08-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/1.0/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE REPORTER STATED THAT THE EXPLANT RESOLVED THE PROBLEM AND THAT THE DOCTOR PLANS TO IMPLANT ANOTHER LENS BUT THERE IS NO EXACT TIME. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL WITH CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND FIBERS ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -10.5/1.0/082 | OD |  | 12/25/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9630050&pc=QCB | 2023826-2020-00193 | 2019-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+1.0/115 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+1.0/115 | OS |  | 12/29/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9633579&pc=QCB | 2023826-2020-00154 | 2019-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.50/2.0/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/2.0/103 | OD |  | 01/06/2020 | 2020 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9633582&pc=QCB | 2023826-2020-00155 | 2020-01-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -12.50/2.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.50/2.5/086 | OS |  | 01/06/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9633725&pc=QCB | 2023826-2020-00157 |  | Injury | STAAR SURGICAL COMPANY | 2020-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.50/1.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED. THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN AND WTW MEASUREMENT. Manufacturer Narrative: ADDITIONAL DATA: H3 - DEVICE EVALUATION: LENS WAS RETURNED DRY, IN MICRO-CENTRIFUGE VIAL, CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND DEBRIS ON THE LENS AND LENS CURVED IN CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.50/1.0/083 | OS |  | 01/08/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9633730&pc=QCB | 2023826-2020-00158 | 2019-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.00/2.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WAS OBSERVED, THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "VAULT IS ABSOLUTELY NORMAL WITH THE EXCHANGE AND AC IS DEEP. PATIENT IS DOING WELL." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/2.5/095 | OD |  | 01/09/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9633736&pc=QCB | 2023826-2020-00159 | 2019-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -8.50/3.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. PER REPORTER, "PATIENT IS HAPPY." Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN DEBRIS ON THE LENS. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.50/3.0/092 | OD |  | 01/10/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9635040&pc=QCB | 2023826-2020-00197 | 2019-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+1.5/084 | OS |  | 12/30/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9639291&pc=QCB | 2023826-2020-00191 | 2019-12-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-01-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/070 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON 3(B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -13.0/+2.0/070 | OD |  | 01/07/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9639455&pc=QCB | 2023826-2020-00187 |  | Injury | STAAR SURGICAL COMPANY | 2020-01-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM TMICL12.1 IMPLANTABLE COLLAMER LENS, (-7.5/+1.5/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2019. REPORTEDLY, APPROXIMATELY 1-MONTH POSTOP THE PATIENT HAS MORE CYLINDER THAN SHE STARTED WITH. IT WAS DETERMINED THAT THE ICL IS SITTING VERTICALLY IN THE EYE. ON (B)(6) 2020 THE LENS WAS REPOSITIONED. 1-WEEK POST-REPOSITION, THE UNCORRECTED VISION IS 20/40. Manufacturer Narrative: H6-METHOD CODE INFORMATION SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT: A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -7.5/+1.5/090 | OD |  | 01/06/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9656180&pc=QCB | 2023826-2020-00213 | 2020-01-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.5/109 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: ON (B)(6) 2020 THE REPORTER INDICATED ORIGINAL LENS REPORTED WAS THE INCORRECT SERIAL#, CORRECT SERIAL# SHOULD BE (B)(6). B5 - THE LENS (SPHERE/CYLINDER/AXIS) SHOULD BE CORRECTED TO -11.0/+2.5/110. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+2.5/109 | OD |  | 01/06/2020 | 2020 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9663138&pc=QCB | 2023826-2020-00218 | 2019-12-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, - 17.0/+4.0/077 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS IS WELL IN PLACE AND PATIENT IS HAPPY. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | - 17.0/+4.0/077 | OS |  | 01/08/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9663192&pc=QCB | 2023826-2020-00214 | 2020-01-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.5/+2.0/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: ON 06-MAR-2020 REPORTER INDICATED ORIGINAL LENS REPORTED WAS THE INCORRECT SERIAL #, CORRECT SERIAL NUMBER IS #(B)(6). B5 - THE LENS (SPHERE/CYLINDER/AXIS) SHOULD BE CORRECTED TO -8.5/+2.0/084. CORRECTED DATA: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H10 - CORRECTION TO "ADDITIONAL MANUFACTURER NARRATIVE" IN SUPPLEMENTAL #1, CORRECT LENS SERIAL IS #T385696. B5 - THE LENS (SPHERE/CYLINDER/AXIS) SHOULD BE CORRECTED TO -8.50/+2.0/088. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+2.0/084 | OS |  | 01/06/2020 | 2020 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9666972&pc=QCB | 2023826-2020-00181 | 2019-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -4.50/0.5/024 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED. THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -4.50/0.5/024 | OS |  | 01/17/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9666978&pc=QCB | 2023826-2020-00210 | 2020-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -9.00/1.0/135 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH SPHERICAL LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, PATIENT RELATED FACTOR. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A ZIPLOC BAG WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT ALSO THE CAUSE OF EVENT WAS " SMALLER THAN ANTICIPATED ANATOMY BASED ON WTW + UBM". CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -9.00/1.0/135 | OD |  | 01/14/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9666990&pc=QCB | 2023826-2020-00182 | 2019-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -5.00/0.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.00/0.5/097 | OD |  | 01/17/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9668713&pc=QCB | 2023826-2020-00217 | 2019-07-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+2.0/088 | OD |  | 01/07/2020 | 2020 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9668725&pc=QCB | 2023826-2020-00222 | 2019-05-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, - 18.0/+4.0/088 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS REPOSITIONED ON (B)(6)2019 BUT THE PROBLEM WAS NOT RESOLVED. H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. H6 - PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS REPOSITIONED. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B3-DATE OF EVENT IN PREVIOUS MDR IS CORRECTED TO (B)(6) 2019. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). TWO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.0/+4.0/088 | OD |  | 01/13/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9668832&pc=QCB | 2023826-2020-00239 | 2019-12-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -6.5/+1.0/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SAME SIZE AND MODEL, IMPLANTED AT A DIFFERENT AXIS (NOT HORIZONTALLY) DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A LENS CASE/VIAL. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND THE HAPTIC BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.5/+1.0/091 | OD |  | 01/12/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9668838&pc=QCB | 2023826-2020-00240 | 2019-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5/+5.0/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.5/+5.0/095 | OS |  | 01/14/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9684588&pc=QCB | 2023826-2020-00258 | 2019-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -3.50/+1.5/041 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS REPOSITIONED 2 TIMES BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -3.50/+1.5/041 | OS |  | 01/13/2020 | 2020 | 69
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9689283&pc=QCB | 2023826-2020-00285 | 2020-01-20 | Injury | STAAR SURGICAL COMPANY | 2020-02-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -13.00/1.5/170 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. REPORTEDLY, "PATIENT IS HAPPY." Manufacturer Narrative: WEIGHT, ETHNICITY, & RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.00/1.5/170 | OS | Physician | 01/21/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9689310&pc=QCB | 2023826-2020-00284 | 2020-01-20 | Injury | STAAR SURGICAL COMPANY | 2020-02-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -13.50/2.5/175 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.50/2.5/175 | OD | Physician | 01/21/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9690062&pc=QCB | 2023826-2020-00273 |  | Injury | STAAR SURGICAL COMPANY | 2020-02-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM TMICL12.1 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/085 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2019. REPORTEDLY, APPROXIMATELY 1-MONTH POSTOP THE PATIENT HAS MORE CYLINDER THAN SHE STARTED WITH. IT WAS DETERMINED THAT THE ICL IS SITTING VERTICALLY IN THE EYE. ON (B)(6) 2020 THE LENS WAS REPOSITIONED. 1-WEEK POST-REPOSITION, THE UNCORRECTED VISION IS 20/20 PLANO. Manufacturer Narrative: (B)(4). | -8.5/+1.5/085 | OS | Other Health Care Professional | 01/17/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9690221&pc=QCB | 2023826-2020-00277 | 2019-10-01 | Injury | STAAR SURGICAL COMPANY | 2020-02-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/112 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND HAPTIC DEFORMED AND FOREIGN RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+1.5/112 | OD | Physician | 01/17/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9689474&pc=QCB | 2023826-2020-00276 | 2019-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/087 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC DEFORMED AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.5/087 | OS |  | 01/17/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9690134&pc=QCB | 2023826-2020-00279 | 2019-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.0/+5.0/076 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. ON THE SAME DATE IN A SEPARATE SURGERY, A SHORTER LENGTH LENS WAS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE, UNKNOWN, AND "CALCULATION SAID 13.2MM BUT VAULT IS EXCESSIVE." Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED WITH FOREIGN RESIDUE AND DEBRIS ON THE LENS. THE LENS HAS A CURVED SHAPE. CLAIM# (B)(4). Manufacturer Narrative: AGE, WEIGHT, ETHNICITY, AND RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.0/+5.0/076 | OS |  | 01/20/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9690137&pc=QCB | 2023826-2020-00275 | 2019-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). ON (B)(6) 2019 A SHORTER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CNETRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.0/092 | OD |  | 01/17/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9690372&pc=QCB | 2023826-2020-00280 | 2018-06-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.0/+3.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS REPOSITIONED DUE TO LOW VAULT AND LENS ROTATION. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -12.0/+3.0/091 | OS |  | 01/21/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9703253&pc=QCB | 2023826-2020-00261 | 2019-11-14 | Injury | STAAR SURGICAL COMPANY | 2020-02-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+2.0/087 | OD | Physician | 01/16/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9703263&pc=QCB | 2023826-2020-00262 | 2019-11-14 | Injury | STAAR SURGICAL COMPANY | 2020-02-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/110 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+1.0/110 | OS | Physician | 01/16/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9700364&pc=QCB | 2023826-2020-00298 | 2020-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -10.00/4.0/109 (SPHERE/CYLINDER/AXIS), IMLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE REPORTER STATED THA THE LENS CONTINUES TO ROTATE OFF AXIS. SECONDARY SURGICAL INTERVENTION (REPOSITIONING/ROTATION) WAS PERFORMED ON (B)(6) 2019 BUT THE PROBLEM WAS NOT RESOLVED. LENS HAS ROTATED OFF AXIS AGAIN AND LOW VAULT WAS OBSERVED. EXCHANGE SURGERY IS PLANNED FOR (B)(6) 2020. IF ADDITIONAL INFORMATION IS RECIEVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: H3 - LENS WAS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION THE LENS HAPTIC TORN. ADDITIONAL INFORMATION: B5 - IT WAS REPORTED THAT ON (B)(6) 2020, THE LENS WAS EXPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.00/4.0/109 | OD |  | 01/12/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9700537&pc=QCB | 2023826-2020-00299 | 2020-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -11.50/3.5/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE REPORTER STATED THA THE LENS CONTINUES TO ROTATE OFF AXIS. SECONDARY SURGICAL INTERVENTION (REPOSITIONING/ROTATION) WAS PERFORMED ON (B)(6) 2019 BUT THE PROBLEM WAS NOT RESOLVED. LENS HAS ROTATED OFF AXIS AGAIN AND LOW VAULT WAS OBSERVED. EXCHANGE SURGERY IS PLANNED FOR (B)(6) 2020. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: H3 - LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. ADDITIONAL INFORMATION: B5 - IT WAS REPORTED THAT ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED AS "ICL SIZE TOO SMALL, ROTATION". CLAIM#:(B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.50/3.5/070 | OS |  | 01/12/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9703290&pc=QCB | 2023826-2020-00270 | 2019-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.0/+2.0/073 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D11- CARTRIDGE MODEL SFC-45 - LOT# "1445291" SHOULD BE CORRECTED TO "1451713" IN THE INITIAL MDR. ADDITIONAL INFORMATION: B3- "(B)(6) 2019" SHOULD BE IN THE INITIAL MDR. B5- "THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2019. THE EXCHANGE RESOLVED THE PROBLEM." SHOULD BE ADDED TO THE INITIAL MDR. D7- "(B)(6) 2019" SHOULD BE IN THE INITIAL MDR. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.0/+2.0/073 | OS |  | 01/15/2020 | 2020 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9709029&pc=QCB | 2023826-2020-00271 | 2019-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+5.5/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE/DEBRIS IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE WITH FOREIGN RESIDUE AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "THE LENS WAS EXCHANGED FOR A SHORTER LENS ON (B)(6) 2019. THE EXCHANGE RESOLVED THE PROBLEM." SHOULD BE ADDED TO THE INITIAL MDR. D7- "(B)(6) 2019" SHOULD BE ADDED TO THE INITIAL MDR. D11- "FOAM TIP PLUNGER MODEL FTP- LOT# 1454829" SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B3- "(B)(6) 2019" SHOULD BE REMOVED FROM SUPPLEMENTAL MDR #2. B5- STATEMENT SHOULD BE REMOVED FROM SUPPLEMENTAL MDR #2. D7- "(B)(6) 2019" SHOULD BE REMOVED FROM SUPPLEMENTAL MDR #2. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.0/+5.5/100 | OD |  | 01/15/2020 | 2020 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9711468&pc=QCB | 2023826-2020-00252 | 2019-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A VTICMO13.7, -15.50/4.5/148 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B4- "28-JAN-2020" SHOUDL BE CORRECTED TO "26-JAN-2020" IN THE INITIAL MDR. G4- "28-JAN-2020" SHOUDL BE CORRECTED TO "26-JAN-2020" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN BLUE RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.50/4.5/148 | OD |  | 01/26/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9711469&pc=QCB | 2023826-2020-00251 | 2020-01-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -8.50/1.0/068 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS FLIPPED DURING INSERTION INTO PATIENT EYE. THE LENS WAS REMOVED AND EXCHANGED FOR THE SAME SIZE, SIMILAR (SPHERE/CYLINDER/AXIS) LENS ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -8.50/1.0/068 | OS |  | 01/27/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9746626&pc=QCB | 2023826-2020-00381 | 2018-06-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -14.00/+4.00/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND A TEAR ON THE OPTIC AND HAPTIC. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS REPORTED THAT ROTATION WAS ALSO EXPERIENCED WITH THE LOW VAULTING. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.00/+4.00/92 | OS |  | 01/29/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9753266&pc=QCB | 2023826-2020-00371 | 2019-12-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/087 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+2.0/087 | OS |  | 02/03/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9753272&pc=QCB | 2023826-2020-00368 | 2019-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+2.0/088 | OD |  | 01/30/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9753273&pc=QCB | 2023826-2020-00372 | 2019-12-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/100 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+2.0/100 | OD |  | 02/03/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9753274&pc=QCB | 2023826-2020-00369 | 2019-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+1.5/091 | OD |  | 01/30/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9759599&pc=QCB | 2023826-2020-00326 | 2020-01-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-02-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/073 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+1.5/073 | OD |  | 02/03/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9777366&pc=QCB | 2023826-2020-00426 | 2020-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.00/1.00/94, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND A SHALLOW ANTERIOR CHAMBER. THE SURGEON STATED THAT HE WORRIED ABOUT POSSIBLE INCREASE IN IOP AND PRESCRIBED PRESSURE LOWERING EYEDROPS PRIOR TO ICL EXCHANGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.00/1.00/94, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT,SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND A SHALLOW ANTERIOR CHAMBER. THIS EXCHANGE RESOLVED THE PROBLEM. THE SURGEON STATED THAT HE WORRIED ABOUT POSSIBLE INCREASE IN IOP AND PRESCRIBED PRESSURE LOWERING EYEDROPS PRIOR TO ICL EXCHANGE. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE: 3191: "ANGLE CLOSURE" - SHOULD BE REMOVED AS IS IT NOT APPLICABLE. H6 - PATIENT CODE: 3191: SHALLOW ANTERIOR CHAMBER, SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/1.00/94 | OD |  | 02/04/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9777367&pc=QCB | 2023826-2020-00427 | 2020-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.50/+1.00/94, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND A SHALLOW ANTERIOR CHAMBER. THE SURGEON STATED THAT HE WORRIED ABOUT POSSILBLE INCREASE IN IOP AND PRESCRIBED PRESSURE LOWERING EYEDROPS PRIOR TO ICL EXCHANGE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.50/+1.00/94, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(4) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND A SHALLOW ANTERIOR CHAMBER. THIS EXCHANGE RESOLVED THE PROBLEM. THE SURGEON STATED THAT HE WORRIED ABOUT POSSIBLE INCREASE IN IOP AND PRESCRIBED PRESSURE LOWERING EYEDROPS PRIOR TO ICL EXCHANGE. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - PATIENT CODE: 3191: "ANGLE CLOSURE" - SHOULD BE REMOVED AS IS IT NOT APPLICABLE. H6 - PATIENT CODE: 3191: SHALLOW ANTERIOR CHAMBER, SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND DEBRIS ON THE LENS. CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.50/+1.00/94 | OS |  | 02/04/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781538&pc=QCB | 2023826-2020-00343 | 2020-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -4.00/1.0/089 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. REPORTEDLY, "SURGEON AND PATIENT ARE HAPPY." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -4.00/1.0/089 | OS |  | 02/13/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781693&pc=QCB | 2023826-2020-00399 | 2019-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/+1.5/095 | OS |  | 02/12/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781695&pc=QCB | 2023826-2020-00400 | 2019-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.0/+1.0/089 | OS |  | 02/12/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781697&pc=QCB | 2023826-2020-00396 | 2019-11-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -19.5/+4.0/118 (SPHERE/CYLINDAR/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2009. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -19.5/+4.0/118 | OD |  | 02/11/2020 | 2020 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781710&pc=QCB | 2023826-2020-00421 | 2019-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+2.0/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENGTH AND DIFFERENT MODEL LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS WITH CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -9.0/+2.0/080 | OD |  | 02/17/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781718&pc=QCB | 2023826-2020-00404 | 2019-06-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETUNRED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+2.0/091 | OS |  | 02/17/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781729&pc=QCB | 2023826-2020-00394 | 2019-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.0/+6.0/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS, AND CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.0/+6.0/085 | OD |  | 02/10/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9781751&pc=QCB | 2023826-2020-00397 | 2020-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.5/+5.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS LOW VAULT AND REFRACTIVE ERROR. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -8.5/+5.5/091 | OD |  | 02/11/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9789011&pc=QCB | 2023826-2020-00431 | 2020-01-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -14.50/2.00/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2017. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.50/2.00/90 | OD |  | 02/10/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9789152&pc=QCB | 2023826-2020-00436 | 2019-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -11.0/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION- LENS RETURNED DRY IN A MICO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/BENT AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/1.5/093 | OD |  | 02/17/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9789221&pc=QCB | 2023826-2020-00439 | 2019-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.0/1.5/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.0/1.5/074 | OS |  | 02/17/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9789259&pc=QCB | 2023826-2020-00440 | 2019-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -7.0/1.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.0/1.0/095 | OD |  | 02/17/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9789266&pc=QCB | 2023826-2020-00411 | 2020-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS BUT THE LENS WAS REPORTED AS HAVING A LOW VAULT - SEE MFR REPORT # 2023826-2020-00412. THE CAUSE OF THE EVENT WAS UNKNOWN/SPECIAL EYE. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.0/+2.0/089 | OS |  | 02/09/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9814258&pc=QCB | 2023826-2020-00474 | 2019-12-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, - 15.50/+4.5/153 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING WELL AND IS SATISFIED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.50/+4.5/153 | OD |  | 02/20/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9819153&pc=QCB | 2023826-2020-00461 | 2019-10-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/094 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.0/094 | OD |  | 02/18/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9819154&pc=QCB | 2023826-2020-00467 | 2019-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/130 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENGTH AND DIFFERENT MODEL LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.5/+1.5/130 | OS |  | 02/20/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9819158&pc=QCB | 2023826-2020-00466 | 2019-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+3.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -15.0/+3.5/087 | OS |  | 02/20/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9819159&pc=QCB | 2023826-2020-00463 | 2019-03-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+6.0/096 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -15.0/+6.0/096 | OD |  | 02/18/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9819160&pc=QCB | 2023826-2020-00459 | 2019-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.5/+1.0/098 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.5/+1.0/098 | OD |  | 02/18/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9819161&pc=QCB | 2023826-2020-00465 | 2019-11-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.5/+1.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -6.5/+1.5/093 | OS |  | 02/19/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9819163&pc=QCB | 2023826-2020-00457 | 2019-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+5.0/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: H10: "LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE" SHOULD BE CORRECTED TO "LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL." CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H10 STATEMENT IS CORRECTED TO H3- DEVICE EVALUATION: LENS WAS NOT RETURNED DRY. LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH "MOISTURE ON THE LENS." CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.5/+5.0/083 | OS |  | 02/18/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9819164&pc=QCB | 2023826-2020-00468 | 2019-10-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/087 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+1.5/087 | OD |  | 02/20/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9819165&pc=QCB | 2023826-2020-00460 | 2019-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/075 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+1.5/075 | OS |  | 02/18/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9819166&pc=QCB | 2023826-2020-00458 | 2019-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/074 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/+2.0/074 | OS |  | 02/18/2020 | 2020 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9819170&pc=QCB | 2023826-2020-00462 | 2019-03-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.0/+6.0/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -16.0/+6.0/083 | OS |  | 02/18/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9819172&pc=QCB | 2023826-2020-00451 | 2019-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION:LENS RETURNED DRY IN LENS CASE/VIAL. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC BROKEN. DIMENSIONAL INSPECTION WAS UNABLE TO BE PERFORMED DUE TO MISING HAPTIC ON BOTH SIDES. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/+1.5/089 | OD |  | 02/17/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824556&pc=QCB | 2023826-2020-00484 | 2019-11-15 | Injury | STAAR SURGICAL COMPANY | 2020-03-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.50/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -09.50/1.5/092 | OD | Physician | 02/19/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824383&pc=QCB | 2023826-2020-00345 | 2019-06-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.50/3.0/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.50/3.0/074 | OS |  | 02/17/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824462&pc=QCB | 2023826-2020-00346 | 2019-06-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.00/3.0/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: G4- " (B)(6) 2020" SHOULD BE CORRECTED TO "(B)(6) 2020" IN THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.00/3.0/109 | OD |  | 02/17/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824506&pc=QCB | 2023826-2020-00350 | 2019-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.00/1.5/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC DEFORMED. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/1.5/102 | OS |  | 02/17/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824510&pc=QCB | 2023826-2020-00351 | 2019-11-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/3.5/084 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO SIGNIFICATE REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: DEVICE EVALUATION DONE UNDER MFR# - 2023826-2020-01079 SUPP2. ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.00/3.5/084 | OS |  | 02/17/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824512&pc=QCB | 2023826-2020-00352 | 2019-11-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -5.00/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.00/1.0/091 | OS |  | 02/17/2020 | 2020 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824520&pc=QCB | 2023826-2020-00353 | 2019-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.00/1.0/087 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND OPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/1.0/087 | OD |  | 02/17/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824528&pc=QCB | 2023826-2020-00356 | 2019-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.50/2.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.50/2.5/081 | OD |  | 02/18/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824545&pc=QCB | 2023826-2020-00361 | 2019-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.50/1.0/069 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.50/1.0/069 | OS |  | 02/18/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824569&pc=QCB | 2023826-2020-00485 | 2019-12-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.00/1.0/059 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.00/1.0/059 | OS |  | 02/20/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824577&pc=QCB | 2023826-2020-00488 | 2019-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.00/1.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -7.00/1.5/088 | OS |  | 02/20/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9824578&pc=QCB | 2023826-2020-00487 | 2019-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.50/2.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MIRCRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/2.5/090 | OD |  | 02/20/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9826513&pc=QCB | 2023826-2020-00479 | 2019-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.5/+2.5/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.5/+2.5/083 | OS |  | 02/17/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9826517&pc=QCB | 2023826-2020-00480 | 2019-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.5/+3.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+3.0/095 | OD |  | 02/17/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9826535&pc=QCB | 2023826-2020-00481 | 2019-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+3.5/075 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+3.5/075 | OS |  | 02/17/2020 | 2020 | 17
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9835633&pc=QCB | 2023826-2020-00441 | 2019-01-10 | Injury | STAAR SURGICAL COMPANY | 2020-03-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -11.5/1.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/1.0/095 | OD | Physician | 02/17/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9837832&pc=QCB | 2023826-2020-00527 | 2019-08-23 | Injury | STAAR SURGICAL COMPANY | 2020-03-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/+3.0/081 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.00/+3.0/081 | OD | Physician | 02/18/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9835956&pc=QCB | 2023826-2020-00434 | 2018-12-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -9.5/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.5/1.5/089 | OD |  | 02/17/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9836264&pc=QCB | 2023826-2020-00435 | 2019-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.0/2.0/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.0/2.0/072 | OD |  | 02/17/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9836392&pc=QCB | 2023826-2020-00538 | 2019-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.5/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADIITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/1.5/091 | OS |  | 02/18/2020 | 2020 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9836894&pc=QCB | 2023826-2020-00539 | 2019-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -7.5/1.0/060 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FORIEGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.5/1.0/060 | OS |  | 02/18/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9837048&pc=QCB | 2023826-2020-00542 | 2019-06-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -7.5/1.0/068 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.5/1.0/068 | OS |  | 02/18/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9837245&pc=QCB | 2023826-2020-00543 | 2019-06-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -8.5/1.5/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/1.5/107 | OD |  | 02/18/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9837767&pc=QCB | 2023826-2020-00558 | 2019-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.5/2.0/130 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.5/2.0/130 | OD |  | 02/20/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9837777&pc=QCB | 2023826-2020-00482 | 2019-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.5/100 | OS |  | 02/17/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9837789&pc=QCB | 2023826-2020-00525 | 2019-07-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+2.0/164 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.5/+2.0/164 | OD |  | 02/18/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9837790&pc=QCB | 2023826-2020-00531 | 2019-11-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/098 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.5/098 | OD |  | 02/20/2020 | 2020 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9837803&pc=QCB | 2023826-2020-00483 | 2019-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -16.0/+3.0/079 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -16.0/+3.0/079 | OS |  | 02/18/2020 | 2020 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9837823&pc=QCB | 2023826-2020-00526 | 2019-01-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/072 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B7-PRE-EXISTING CONDITION: CILIARY BODY CYST. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.5/072 | OD |  | 02/18/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9837887&pc=QCB | 2023826-2020-00555 | 2019-12-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.5/1.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.5/1.5/088 | OS |  | 02/20/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9837982&pc=QCB | 2023826-2020-00557 | 2019-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.0/3.0/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.0/3.0/105 | OS |  | 02/20/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9838171&pc=QCB | 2023826-2020-00530 | 2019-11-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.00/+1.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.00/+1.0/093 | OD |  | 02/20/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9838182&pc=QCB | 2023826-2020-00532 | 2019-09-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.00/+0.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.00/+0.5/090 | OD |  | 02/20/2020 | 2020 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9841880&pc=QCB | 2023826-2020-00552 | 2019-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -5.5/1.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.5/1.0/086 | OD |  | 02/20/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9842307&pc=QCB | 2023826-2020-00556 | 2019-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.5/5.0/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.5/5.0/109 | OD |  | 02/20/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9842724&pc=QCB | 2023826-2020-00521 | 2020-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.5/+1.5/015 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- RESULT CODE 213 SHOULD BE ADDED TO THE PREVIOUS MDR. H6- CONCLUSION CODE 4310 SHOULD BE ADDED TO THE PREVIOUS MDR. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: CODES WERE ACCIDENTALLY SUBMITTED AND ARE NOT APPLICABLE FROM PREVIOUS MDR SUPPLEMENTAL #2. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -15.5/+1.5/015 | OD |  | 02/25/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9842728&pc=QCB | 2023826-2020-00512 | 2019-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.0/096 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICLA DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+2.0/096 | OS |  | 02/20/2020 | 2020 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9842747&pc=QCB | 2023826-2020-00510 | 2019-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.5/+1.5/085 | OD |  | 02/20/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9850050&pc=QCB | 2023826-2020-00562 | 2019-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/075 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+1.5/075 | OS |  | 02/20/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9857162&pc=QCB | 2023826-2020-00579 | 2019-08-29 | Injury | STAAR SURGICAL COMPANY | 2020-03-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.50/+1.50/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO SIZING ISSUE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.50/+1.50/084 | OS | Physician | 02/27/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9855373&pc=QCB | 2023826-2020-00500 | 2019-10-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -14.00/2.0/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS IN MICRO-CENTRIFUGE VIAL, CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.00/2.0/107 | OS |  | 02/28/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9857159&pc=QCB | 2023826-2020-00577 | 2019-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0/+2.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -16.0/+2.5/095 | OD |  | 02/25/2020 | 2020 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9869304&pc=QCB | 2023826-2020-00628 | 2019-02-11 | Injury | STAAR SURGICAL COMPANY | 2020-03-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12/+2.5/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12/+2.5/94 | OD | Physician | 02/25/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9868296&pc=QCB | 2023826-2020-00632 | 2019-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -7.0/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.0/1.5/095 | OS |  | 02/28/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9868443&pc=QCB | 2023826-2020-00631 | 2019-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -13.5/3.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.5/3.5/090 | OD |  | 02/28/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9869215&pc=QCB | 2023826-2020-00629 | 2019-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.0/5.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.0/5.0/087 | OD |  | 02/25/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9869323&pc=QCB | 2023826-2020-00622 | 2019-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/099 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+1.0/099 | OD |  | 03/03/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9869331&pc=QCB | 2023826-2020-00616 | 2018-09-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+3.0/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -18.0/+3.0/103 | OD |  | 03/02/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9869332&pc=QCB | 2023826-2020-00609 | 2020-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+1.5/086 | OS |  | 02/26/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9869333&pc=QCB | 2023826-2020-00611 | 2019-12-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+1.0/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL AND SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/+1.0/080 | OD |  | 02/28/2020 | 2020 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9869334&pc=QCB | 2023826-2020-00617 | 2020-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -6.5/+4.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO THE PATIENT WAS "DISSATISFIED WITH THE QUALITY OF HIS VISION." THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.5/+4.0/094 | OS |  | 03/03/2020 | 2020 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9880951&pc=QCB | 2023826-2020-00664 | 2019-07-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+3.0/080 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+3.0/080 | NA |  | 03/03/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9880987&pc=QCB | 2023826-2020-00667 | 2019-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.5/+1.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.5/+1.0/089 | OD |  | 03/04/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9880988&pc=QCB | 2023826-2020-00662 | 2019-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3-LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+1.5/092 | OD |  | 03/03/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9881006&pc=QCB | 2023826-2020-00670 | 2020-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50/+1.5/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, WITH DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.50/+1.5/091 | OS |  | 03/06/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9881015&pc=QCB | 2023826-2020-00655 | 2018-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.0/+2.5/078 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+2.5/078 | OS |  | 03/01/2020 | 2020 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9881018&pc=QCB | 2023826-2020-00657 | 2018-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+3.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+3.0/091 | OD |  | 03/03/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9881029&pc=QCB | 2023826-2020-00661 | 2019-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/079 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+1.0/079 | OD |  | 03/03/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9881032&pc=QCB | 2023826-2020-00660 | 2019-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.5/+4.0/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.5/+4.0/083 | OS |  | 03/03/2020 | 2020 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9881037&pc=QCB | 2023826-2020-00671 | 2020-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00/+1.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, WITH DEBRIS ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC BENT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.00/+1.0/089 | OD |  | 03/06/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9889364&pc=QCB | 2023826-2020-00582 | 2020-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.5, -13.00/4.5/093 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LENS ROTATION AND LOW VAULT. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B4-DATE OF THIS REPORT IN INITIAL MDR IS CORRECTED TO (B)(6) 2020 B5- "THE EXCHANGE RESOLVED THE PROBLEM." NEEDS TO BE INCULDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: DATE IN THE INITIAL MDR IS CORRECTED TO (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.00/4.5/093 | OS |  | 03/01/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9889366&pc=QCB | 2023826-2020-00587 | 2019-09-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -18.00/3.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -18.00/3.5/080 | OS |  | 03/03/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9889368&pc=QCB | 2023826-2020-00588 | 2019-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.00/2.5/059 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. EVENT DATE: (B)(6) 2020. CORRECT DATE IN INITIAL MEDWATCH REPORT SHOULD BE CORRECTED TO (B)(6) 2020. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: D10: INFORMATION ADDED WHICH SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MEDWATCH REPORT #1. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -18.00/2.5/059 | OS |  | 03/02/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9889374&pc=QCB | 2023826-2020-00590 | 2019-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.00/2.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA : B4- DATE OF THIS REPORT IN INITIAL MDR IS CORRECTED TO 02-MAR-2020 G4-DATE RECEIVED BY MANUFACTURER IN INITIAL MDR IS CORRECTED TO 02-MAR-2020 CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.00/2.5/088 | OD |  | 03/02/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9889379&pc=QCB | 2023826-2020-00583 | 2020-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.5, -13.00/5.0/100 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LENS ROTATION AND LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B4-DATE OF THIS REPORT IN INITIAL MDR IS CORRECTED TO 01-MAR-2020. CLAIM# : (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: DATE IN THE INITIAL MDR IS CORRECTED TO (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.00/5.0/100 | OD |  | 03/01/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9889389&pc=QCB | 2023826-2020-00586 | 2019-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/0.5/163 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH SPHERICAL LENS ON (B)(6) 2019 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM.. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND CLEAR SURGICAL RESIDUE ON THE LENS AND A TORN HAPTIC. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA : D7- EXPLANT DATE IN INITIAL MDR IS CORRECTED TO (B)(6) 2019. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/0.5/163 | OD |  | 03/03/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9889391&pc=QCB | 2023826-2020-00591 | 2019-11-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -8.00/1.0/133 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND RESIDUE ON THE LENS. CORRECTED DATA" B3 - 02-MAR-2020. G4 - CORRECTED DATE IN INITIAL MEDWATCH REPORT SHOULD BE, 02-MAR-2020. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.00/1.0/133 | OD |  | 03/02/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9889394&pc=QCB | 2023826-2020-00589 | 2019-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -18.00/3.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON 26/FEB/2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND OTIC TORN AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA : B4- DATE OF THIS REPORT IN INITIAL MDR IS CORRECTED TO 02-MAR-2020. G4-DATE RECEIVED BY MANUFACTURER IN INITIAL MDR IS CORRECTED TO 02-MAR-2020. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -18.00/3.0/083 | OS |  | 03/02/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9889398&pc=QCB | 2023826-2020-00594 | 2019-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.00/1.5/083 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS ON (B)(6) 2019 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# : (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.00/1.5/083 | OS |  | 03/03/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9889404&pc=QCB | 2023826-2020-00600 | 2019-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.50/1.0/132 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON 20/APR/2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.50/1.0/132 | OS |  | 03/05/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9889429&pc=QCB | 2023826-2020-00596 | 2019-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -5.50/2.0/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.50/2.0/104 | OD |  | 03/02/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9901532&pc=QCB | 2023826-2020-00720 | 2020-01-27 | Injury | STAAR SURGICAL COMPANY | 2020-03-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -4.50/5.00/003 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -4.50/5.00/003 | OS | Physician | 03/06/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9897127&pc=QCB | 2023826-2020-00642 | 2019-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.0/1.0/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.0/1.0/087 | OD |  | 03/02/2020 | 2020 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9897197&pc=QCB | 2023826-2020-00641 | 2019-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.5/1.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.5/1.0/092 | OS |  | 03/02/2020 | 2020 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9897284&pc=QCB | 2023826-2020-00645 | 2019-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.5/1.0/077 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.5/1.0/077 | OD |  | 03/02/2020 | 2020 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9898226&pc=QCB | 2023826-2020-00646 | 2019-11-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.0/1.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.0/1.0/075 | OS |  | 03/03/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9898333&pc=QCB | 2023826-2020-00647 | 2019-11-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -9.0/1.5/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.0/1.5/075 | OD |  | 03/03/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9898473&pc=QCB | 2023826-2020-00653 | 2019-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.5/1.5/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/1.5/072 | OS |  | 03/05/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9899215&pc=QCB | 2023826-2020-00652 | 2019-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.0/1.0/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/1.0/098 | OD |  | 03/05/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9900192&pc=QCB | 2023826-2020-00651 | 2019-12-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -16.0/2.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: DATE SHOULD BE CORRECTED TO (B)(6) 2020 IN INITIAL MEDWATCH REPORT. DATE SHOULD BE CORRECTED TO (B)(6) 2020 IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.0/2.5/083 | OS |  | 03/02/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9900469&pc=QCB | 2023826-2020-00648 | 2019-12-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -16.0/2.5/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B4 - DATE SHOULD BE CORRECTED TO 02-MAR-2020 IN INITIAL MEDWATCH REPORT. G4 - DATE SHOULD BE CORRECTED TO 02-MAR-2020 IN INITIAL MEDWATCH REPORT. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.0/2.5/104 | OD |  | 03/02/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9910952&pc=QCB | 2023826-2020-00681 | 2019-12-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/098 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: PMA/510K: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -12.5/+1.0/098 | OS |  | 03/06/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9910958&pc=QCB | 2023826-2020-00683 | 2020-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS WITH CLEAR SURGICAL RESIDUE DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+1.0/088 | OD |  | 03/06/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9910959&pc=QCB | 2023826-2020-00691 | 2019-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/082 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+2.5/082 | OD |  | 03/09/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9910960&pc=QCB | 2023826-2020-00677 | 2019-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -7.5/+0.5/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.5/+0.5/080 | OD |  | 03/04/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9910967&pc=QCB | 2023826-2020-00682 | 2020-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+1.5/090 | OS |  | 03/06/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9910969&pc=QCB | 2023826-2020-00690 | 2019-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+3.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+3.0/089 | OS |  | 03/09/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9910971&pc=QCB | 2023826-2020-00676 | 2019-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -6.5/+0.5/104 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: PREVIOUS SUPPLEMENTAL MDR WAS SUBMITTED IN ERROR. INFORMATION INCLUDED IN SUPPLEMENTAL MDR#1 IS NOT APPLICABLE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.5/+0.5/104 | OS |  | 03/04/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9910974&pc=QCB | 2023826-2020-00686 | 2019-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+1.5/069 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN AND DEFORMED. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+1.5/069 | OD |  | 03/09/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9913465&pc=QCB | 2023826-2020-00725 | 2019-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED LATER THAT SAME DAY ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+2.0/083 | OS |  | 03/08/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9913483&pc=QCB | 2023826-2020-00724 | 2019-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.5/080 | OS |  | 03/08/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9913503&pc=QCB | 2023826-2020-00730 | 2019-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+2.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISIBLE INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.5/+2.0/094 | OD |  | 03/09/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9913528&pc=QCB | 2023826-2020-00726 | 2019-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+3.0/070 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISIBLE INSPECTION FOUND THE NO VISIBLE DAMAGE TO THE LENS. (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.0/+3.0/070 | OS |  | 03/08/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9916620&pc=QCB | 2023826-2020-00734 | 2019-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/076 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.5/076 | OS |  | 03/11/2020 | 2020 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9916623&pc=QCB | 2023826-2020-00731 | 2019-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+2.0/088 | OS |  | 03/11/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9921817&pc=QCB | 2023826-2020-00824 | 2019-10-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.5/1.5/126 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.5/1.5/126 | OD |  | 03/09/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9921922&pc=QCB | 2023826-2020-00825 | 2019-10-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.5/1.5/068 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.5/1.5/068 | OS |  | 03/09/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9923218&pc=QCB | 2023826-2020-00829 | 2019-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -16.5/2.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.5/2.5/083 | OS |  | 03/09/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9923718&pc=QCB | 2023826-2020-00831 | 2019-05-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.0/1.0/111 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION; H3- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/1.0/111 | OD |  | 03/09/2020 | 2020 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9924220&pc=QCB | 2023826-2020-00833 | 2020-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -14.0/2.0/94 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. H6 - DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.0/2.0/94 | OS |  | 03/10/2020 | 2020 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9924291&pc=QCB | 2023826-2020-00832 | 2019-12-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -6.5/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: B4- "(B)(6) 2019" SHOULD BE CORRECTED TO "(B)(6) 2020" IN THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.5/2.0/093 | OD |  | 03/10/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9928341&pc=QCB | 2023826-2020-00702 | 2020-03-06 | Injury | STAAR SURGICAL COMPANY | 2020-04-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.50/1.5/055 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -17.50/1.5/055 | OS | Physician | 03/09/2020 | 2020 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9928328&pc=QCB | 2023826-2020-00706 | 2020-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM, VTICMO13.2, -12.50/3.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN, DEFORMED AND STRETCHED OUT WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.50/3.5/083 | OS |  | 03/09/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9928342&pc=QCB | 2023826-2020-00703 | 2019-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -6.50/2.5/170 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS "TOO LONG DIAMETER." Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.50/2.5/170 | OD |  | 03/09/2020 | 2020 | 47
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9928344&pc=QCB | 2023826-2020-00700 | 2019-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -8.50/2.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON 01/OCT/2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.50/2.0/085 | OD |  | 03/09/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9928345&pc=QCB | 2023826-2020-00699 | 2019-12-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.50/0.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS ON (B)(6) 2019 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM.. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.50/0.5/095 | OD |  | 03/09/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9928346&pc=QCB | 2023826-2020-00701 | 2019-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.50/2.0/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/2.0/074 | OS |  | 03/09/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9928363&pc=QCB | 2023826-2020-00793 | 2020-03-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -10.50/3.0/002 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. EXCESSIVE VAULT WAS OBSERVED AND THE LENS WAS REMOVED ON (B)(6) 2020. A REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/3.0/002 | OS |  | 03/12/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9928367&pc=QCB | 2023826-2020-00714 | 2019-06-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -17.00/1.5/059 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -17.00/1.5/059 | OS |  | 03/13/2020 | 2020 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9936251&pc=QCB | 2023826-2020-00750 | 2019-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+2.0/084 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+2.0/084 | OS |  | 03/11/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9936256&pc=QCB | 2023826-2020-00771 | 2019-10-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN/BROKEN/BENT/DEFORMED. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION CORRECTED TO: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+2.0/080 | OD |  | 03/17/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9936270&pc=QCB | 2023826-2020-00749 | 2019-11-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.5/+1.0/150 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.5/+1.0/150 | OD |  | 03/11/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9936287&pc=QCB | 2023826-2020-00772 | 2019-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.5/090 | OS |  | 03/17/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9936290&pc=QCB | 2023826-2020-00765 | 2019-05-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/069 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -8.5/+1.5/069 | OS |  | 03/16/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9936299&pc=QCB | 2023826-2020-00751 | 2019-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Adverse Event Without Identified Device or Use Problem | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+3.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REMOVED AND EXCHANGED WITH AN ALTERNATE LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, NO PATIENT INJURY OCCURRED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.0/+3.5/087 | OD |  | 03/11/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9936305&pc=QCB | 2023826-2020-00764 | 2019-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+2.0/086 | OS |  | 03/17/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9936307&pc=QCB | 2023826-2020-00760 | 2019-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3- DEVICE EVALUATION: LENS RETURNED ON A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND OPTIC TORN/ AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+2.5/089 | OD |  | 03/16/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9938066&pc=QCB | 2023826-2020-00708 | 2019-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.00/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "THE PATIENT IS IN URGENT NEED OF SURGERY, AFTER FULL COMMUNICATION, CHOOSE ICL." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS, AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -18.00/1.5/091 | OD |  | 03/11/2020 | 2020 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9938116&pc=QCB | 2023826-2020-00717 | 2018-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.50/1.5/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DEFORMED WITH FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.50/1.5/084 | OD |  | 03/16/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9938167&pc=QCB | 2023826-2020-00713 | 2019-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -8.50/1.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.50/1.0/086 | OS |  | 03/13/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9938193&pc=QCB | 2023826-2020-00715 | 2019-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -8.50/3.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.50/3.5/085 | OS |  | 03/13/2020 | 2020 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9940065&pc=QCB | 2023826-2020-00842 | 2019-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.0/3.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.0/3.0/092 | OS |  | 03/13/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9944282&pc=QCB | 2023826-2020-00811 | 2019-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3- LENS RETURNED DRY WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+2.0/093 | OS |  | 03/17/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9944285&pc=QCB | 2023826-2020-00820 | 2019-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+2.0/089 | OS |  | 03/18/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9944296&pc=QCB | 2023826-2020-00821 | 2019-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.5/083 | OD |  | 03/19/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9944305&pc=QCB | 2023826-2020-00822 | 2019-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+1.5/080 | OD |  | 03/19/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9944313&pc=QCB | 2023826-2020-00818 | 2019-10-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+3.0/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+3.0/085 | OD |  | 03/18/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9944316&pc=QCB | 2023826-2020-00803 | 2020-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5/+2.0/079 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.5/+2.0/079 | OD |  | 03/17/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9944320&pc=QCB | 2023826-2020-00804 | 2020-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5/+2.0/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.5/+2.0/096 | OS |  | 03/17/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9944324&pc=QCB | 2023826-2020-00819 | 2019-10-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+3.5/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+3.5/083 | OS |  | 03/18/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9953454&pc=QCB | 2023826-2020-00873 | 2019-11-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-7.5/2.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.5/2.0/082 | OD |  | 03/16/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9953738&pc=QCB | 2023826-2020-00876 | 2019-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-7.5/3.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.5/3.0/096 | OD |  | 03/16/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9954004&pc=QCB | 2023826-2020-00885 | 2019-12-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.0/2.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE AND MOISTURE ON THE LENS. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6- DEVICE EVALUATION STATEMENT IN PREVIOUS MDR IS NOT APPLICABLE. CORRECT STATEMENT- "THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS." CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.0/2.0/092 | OD |  | 03/17/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9954013&pc=QCB | 2023826-2020-00877 | 2019-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-6.0/3.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3-DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.0/3.0/087 | OS |  | 03/16/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9954066&pc=QCB | 2023826-2020-00878 | 2019-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1,-15.0/4.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.0/4.0/084 | OD |  | 03/16/2020 | 2020 | 47
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9954341&pc=QCB | 2023826-2020-00879 | 2019-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1,-10.0/3.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.0/3.5/088 | OS |  | 03/16/2020 | 2020 | 47
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9954463&pc=QCB | 2023826-2020-00880 | 2020-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-14.5/4.5/052 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION:LENS RETURNED IN A MICRO-CENRIFUGE VIAL WITH MOISTURE ON LENS WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND RESIDUE ON LENS. EVICE CODE 1494-OFF-LABEL; KERATOCONUS. KERATOCONUS NEEDS TO BE INCLUDED IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H3- THE STATEMENT IN PREVIOUS MDR SHOULD INCLUDE"VISUAL OBSERVATION FOUND NO VISBLE DAMAGE AND RESIDUE ON LENS." CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.5/4.5/052 | OD |  | 03/16/2020 | 2020 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9954528&pc=QCB | 2023826-2020-00890 | 2019-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -12.0/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. RETURNED PRODUCT EVALUATION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.0/2.0/088 | OD |  | 03/18/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9954554&pc=QCB | 2023826-2020-00881 | 2020-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-15.0/2.00/10 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION:LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS. DEVICE CODE 1494 OFF-LABEL; KERATOCONUS. KERATOCONUS TO BE INCLUDED IN INITIAL MDR . CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.0/2.00/10 | OS |  | 03/16/2020 | 2020 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9954850&pc=QCB | 2023826-2020-00883 | 2020-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1,-18.0/2.0/84 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 2.8MM) CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -18.0/2.0/84 | OD |  | 03/18/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9955121&pc=QCB | 2023826-2020-00884 | 2020-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1,-18.0/2.5/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFTT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD < 2.8MM) SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -18.0/2.5/91 | OS |  | 03/18/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9957594&pc=QCB | 2023826-2020-00780 | 2019-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -10.50/2.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/2.5/087 | OD |  | 03/17/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9957629&pc=QCB | 2023826-2020-00788 | 2019-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.50/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.50/1.0/091 | OD |  | 03/18/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9957630&pc=QCB | 2023826-2020-00789 | 2019-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.50/3.5/169 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.50/3.5/169 | OD |  | 03/18/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9957631&pc=QCB | 2023826-2020-00796 | 2019-10-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/1.0/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED ON (B)(6) 2019 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2020, "NOW THE PATIENT IS IN GOOD CONDITION, AND HIS VISUAL ACUITY HAS RECOVERED TO PRE-OPERATIVE UCVA AND BCVA IN THE LAST VISIT." Manufacturer Narrative: CORRECTED DATA: D11 - LOT NUMBERS UPDATED FROM "UNK" TO THE CORRECT REPORTED NUMBERS. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/1.0/099 | OD |  | 03/18/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9960148&pc=QCB | 2023826-2020-00861 | 2019-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/104 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -13.5/+2.0/104 | OS |  | 03/22/2020 | 2020 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9960152&pc=QCB | 2023826-2020-00855 | 2019-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/068 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+1.0/068 | OS |  | 03/16/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9960155&pc=QCB | 2023826-2020-00868 | 2019-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/079 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -11.0/+2.0/079 | OS |  | 03/24/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9960157&pc=QCB | 2023826-2020-00858 | 2020-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.0/+2.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020, THE LENS WAS EXCHANGED WITH A SAME SIZE AND MODEL, DIFFERENT SPHERE AND AXIS LENS DUE TO THE PATIENT WAS DISSATISFIED WITH THE POSTOPERATIVE VISUAL ACUITY. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# : (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -9.0/+2.5/092 | OD |  | 03/19/2020 | 2020 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9960158&pc=QCB | 2023826-2020-00864 | 2019-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+3.0/079 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+3.0/079 | OS |  | 03/24/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9960161&pc=QCB | 2023826-2020-00859 | 2020-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.0/+2.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020, THE LENS WAS EXCHANGED WITH A SAME SIZE AND MODEL, DIFFERENT SPHERE AND AXIS LENS DUE TO THE PATIENT WAS DISSATISFIED WITH THE POSTOPERATIVE VISUAL ACUITY. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED LIQUID IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# : (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -9.0/+2.0/091 | OS |  | 03/19/2020 | 2020 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9960162&pc=QCB | 2023826-2020-00867 | 2019-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.5/+1.5/087 | OD |  | 03/24/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9960164&pc=QCB | 2023826-2020-00865 | 2019-11-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K:THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -11.0/+1.0/085 | OS |  | 03/24/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9960166&pc=QCB | 2023826-2020-00863 | 2019-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+2.5/102 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.0/+2.5/102 | OD |  | 03/24/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9960168&pc=QCB | 2023826-2020-00862 | 2019-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+2.0/174 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -8.5/+2.0/174 | OD |  | 03/22/2020 | 2020 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9965168&pc=QCB | 2023826-2020-00899 | 2019-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.00/+2.0/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC DEFORMED AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.00/+2.0/085 | OD |  | 03/24/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9965179&pc=QCB | 2023826-2020-00900 | 2019-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.50/+2.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS SLIGHTLY CURVED AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H10 - "CLAIM#: (B)(4) SHOULD BE CORRECTED TO CLAIM# (B)(4) IN SUPPLEMENTAL MEDWATCH REPORT #1. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.50/+2.5/093 | OD |  | 03/24/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9967621&pc=QCB | 2023826-2020-00906 | 2020-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.00/1.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. "THE PATIENT WAS DISATISFIED WITH TOO MUCH VISUAL ACUITY." THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2020 WITH A "LOWER POWER LENS." THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION-LENS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.00/1.5/081 | OD |  | 03/23/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9967622&pc=QCB | 2023826-2020-00907 | 2020-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.50/1.5/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. "THE PATIENT WAS DISSATISFIED WITH TOO MUCH VISUAL ACUITY." THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2020 WITH A SAME SIZE, "LOWER POWER LENS." THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND RESIDUE ON LENS AND LENS HAS A CURVED SHAPE. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -9.50/1.5/079 | OS |  | 03/23/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9967623&pc=QCB | 2023826-2020-00910 | 2018-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.00/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND DEFORMED. CORRECTED DATA: D11 - INJECTOR MODEL# LIOLO-24, LOT#UNK SHOULD BE REMOVED FROM THE INITIAL MEDWATCH REPORT, AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/1.0/087 | OD |  | 03/24/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9967624&pc=QCB | 2023826-2020-00914 | 2020-03-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -17.50/3.0/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -17.50/3.0/109 | OD |  | 03/26/2020 | 2020 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9967627&pc=QCB | 2023826-2020-00909 | 2018-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -6.50/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.50/2.0/093 | OD |  | 03/24/2020 | 2020 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9967628&pc=QCB | 2023826-2020-00912 | 2019-10-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.00/0.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -9.00/0.5/105 | OS |  | 03/24/2020 | 2020 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9977662&pc=QCB | 2023826-2020-00918 | 2020-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/1.0/076 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -11.00/1.0/076 | OD |  | 03/29/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9977664&pc=QCB | 2023826-2020-00917 | 2020-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -11.00/1.0/093 | OS |  | 03/29/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9977916&pc=QCB | 2023826-2020-00919 | 2020-03-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.00/2.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -8.00/2.0/083 | OD |  | 03/29/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9979178&pc=QCB | 2023826-2020-00923 | 2019-12-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.00/1.5/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/1.5/074 | OD |  | 04/03/2020 | 2020 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9979181&pc=QCB | 2023826-2020-00922 | 2020-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -7.00/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED AND THE LENS WAS REPOSITIONED ON (B)(6) 2019. THIS DID NOT RESOLVE THE PROBLEM. THE LENS WAS THEN REMOVED AND EXCHANGED WITH A SAME LENGHT LENS ON (B)(6) 2020. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.00/2.0/095 | OS |  | 04/03/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9987714&pc=QCB | 2023826-2020-00939 | 2020-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6,-13.5/3.00/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM, LENS IS IN POSITION AND PATIENT IS HAPPY. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.5/3.00/95 | OD |  | 03/23/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9987806&pc=QCB | 2023826-2020-00938 | 2020-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6,-13.5/3.00/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM, LENS IS IN POSITION AND PATIENT IS HAPPY. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.5/3.00/92 | OS |  | 03/23/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9988702&pc=QCB | 2023826-2020-00987 | 2020-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-7.0/2.5/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.0/2.5/094 | OS |  | 03/29/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9988739&pc=QCB | 2023826-2020-00988 | 2020-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-9.5/3.5/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.5/3.5/093 | OD |  | 03/29/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9991662&pc=QCB | 2023826-2020-00946 | 2019-04-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/ +1.5/ 088 (SPHERE/ CYLINDER/ AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/ +1.5/ 088 | OS |  | 03/29/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9991665&pc=QCB | 2023826-2020-00957 | 2019-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+1.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.0/+1.5/088 | OS |  | 04/07/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10001301&pc=QCB | 2023826-2020-00992 | 2020-01-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-7.5/1.5/093 (SPHERE/ CYLINDER/ AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.5/1.5/093 | OS |  | 04/03/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10002747&pc=QCB | 2023826-2020-00993 |  | Injury | STAAR SURGICAL COMPANY | 2020-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM TMICL13.2, -9.5/2.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED AND A SHORTER LENGTH LENS WAS IMPLANTED DUE TO OVER VAULT. IT WAS REPORTED THAT THE PROBLEM RESOLVED. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- "LEFT EYE (OD)" SHOULD BE CORRECTED TO "LEFT EYE (OS)" IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -9.5/2.5/086 | OS |  | 03/30/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10002859&pc=QCB | 2023826-2020-00982 | 2020-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/100 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.5/+1.5/100 | OD |  | 04/02/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10002863&pc=QCB | 2023826-2020-00980 | 2020-03-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/089 (SPHERE/ CYLINDER/ AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SAME LENGTH, DIFFERENT MODEL LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -13.0/+2.0/089 | OS |  | 04/02/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10002879&pc=QCB | 2023826-2020-00981 | 2020-03-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SAME SIZE, DIFFERENT MODEL LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.0/+1.0/082 | OD |  | 04/02/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10013318&pc=QCB | 2023826-2020-00961 | 2018-11-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.00/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/1.5/095 | OD |  | 04/13/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10013319&pc=QCB | 2023826-2020-00962 | 2018-11-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.00/3.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/3.5/093 | OS |  | 04/13/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10013320&pc=QCB | 2023826-2020-00959 | 2020-01-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/1.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/1.0/079 | OD |  | 04/13/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10013327&pc=QCB | 2023826-2020-00971 | 2020-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.50/2.5/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2020, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED FOR A SAME LENGTH SPHERICAL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.50/2.5/084 | OS |  | 04/20/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10013331&pc=QCB | 2023826-2020-00965 | 2019-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.50/1.5/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.5/100 | OD |  | 04/14/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10013332&pc=QCB | 2023826-2020-00968 | 2019-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.00/1.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.00/1.5/080 | OS |  | 04/14/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10013333&pc=QCB | 2023826-2020-00969 | 2020-04-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: HE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/2.0/088 | OS |  | 04/15/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10013337&pc=QCB | 2023826-2020-00967 | 2019-09-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.00/1.0/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/1.0/099 | OS |  | 04/14/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10013338&pc=QCB | 2023826-2020-00966 | 2019-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.50/1.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.5/083 | OS |  | 04/14/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10013340&pc=QCB | 2023826-2020-00972 | 2020-01-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -14.00/3.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2020, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.00/3.0/079 | OS |  | 04/20/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10013342&pc=QCB | 2023826-2020-00970 | 2020-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-04-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.50/3.0/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2020, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED FOR A SAME LENGTH SPHERICAL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.50/3.0/098 | OD |  | 04/20/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10025650&pc=QCB | 2023826-2020-01001 | 2018-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/094 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.0/094 | OD |  | 04/13/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10025654&pc=QCB | 2023826-2020-01003 | 2020-03-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/065 (SPHERE/ CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). | -11.5/+1.0/065 | OS |  | 04/13/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10025897&pc=QCB | 2023826-2020-01014 | 2019-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -7.00/+2.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.00/+2.0/094 | OS |  | 04/13/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10025909&pc=QCB | 2023826-2020-01011 | 2019-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50/+1.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -09.50/+1.5/094 | OD |  | 04/12/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10025934&pc=QCB | 2023826-2020-01012 | 2020-01-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.00/+2.5/077 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.00/+2.5/077 | OS |  | 04/12/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10025943&pc=QCB | 2023826-2020-01013 | 2020-04-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.00/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL. CONDITION WAS DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC DEFORMED AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.00/+2.0/090 | OD |  | 04/12/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10025950&pc=QCB | 2023826-2020-01019 | 2020-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED LATER THE SAME DAY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+1.0/096 | OS |  | 04/16/2020 | 2020 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10033108&pc=QCB | 2023826-2020-01028 | 2019-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.12, -16.50/1.0/089 (SPHERE/ CYLINDER/ AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: A1 - WJ. CLAIM#: (B)(4). Manufacturer Narrative: "VTICMO12.12" SHOULD BE CORRECTED TO "VTICMO12.1" IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.50/1.0/089 | OD |  | 04/21/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10033111&pc=QCB | 2023826-2020-01026 | 2020-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/2.0/098 (SPHERE/ CYLINDER/ AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.00/2.0/098 | OD |  | 04/20/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10043597&pc=QCB | 2023826-2020-01032 | 2019-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.50/2.5/111 (SPHERE/ CYLINDER/ AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.50/2.5/111 | OS |  | 04/26/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10043633&pc=QCB | 2023826-2020-01051 | 2018-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-13.00/2.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC DEFORMED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.00/2.5/095 | OD |  | 04/14/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10043634&pc=QCB | 2023826-2020-01050 | 2019-09-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-16.5/2.0/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: B5 - IT WAS ALSO REPORTED THAT THE PATIENT EXPERIENCED UNREACTIVE (FIXED) PUPIL. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (UNREACTIVE FIXED PUPIL). H6 - DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - PATIENT ALSO EXPERIENCED UNREACTIVE (FIXED) PUPIL SHOULD BE REMOVED FROM SUPPLEMENTAL MEDWATCH REPORT #1, AS IT IS NOT APPLICABLE. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (UNREACTIVE FIXED PUPIL) SHOULD BE REMOVED FROM SUPPLEMENTAL MEDWATCH REPORT #1, AS IT IS NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.5/2.0/100 | OD |  | 04/13/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10045074&pc=QCB | 2023826-2020-01056 | 2019-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-11.0/1.0/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/1.0/102 | OD |  | 04/14/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10045103&pc=QCB | 2023826-2020-01057 | 2019-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1,-7.5/1.0/053 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.5/1.0/053 | OS |  | 04/14/2020 | 2020 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10045306&pc=QCB | 2023826-2020-01048 | 2019-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-13.5/2.0/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.5/2.0/100 | OS |  | 04/13/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10045344&pc=QCB | 2023826-2020-01049 | 2019-11-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-5.0/1.5/177 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.0/1.5/177 | OD |  | 04/13/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10049181&pc=QCB | 2023826-2020-01061 | 2020-03-08 | Injury | STAAR SURGICAL COMPANY | 2020-05-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS WAS LOST IN THE OR AND WILL NOT BE RETURNED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.50/+1.0/091 | OD | Physician | 04/20/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10048950&pc=QCB | 2023826-2020-01039 | 2019-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN, BENT AND DEFORMED WITH RESIDUE ON THE LENS. CLAIM#:(B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -11.5/+2.0/093 | OS |  | 04/20/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10048962&pc=QCB | 2023826-2020-01035 | 2019-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.5/+4.5/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE LENS, IMPLANTED IN A VERTICAL POSITION DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.5/+4.5/088 | OS |  | 04/15/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10049200&pc=QCB | 2023826-2020-01062 | 2020-03-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50/+2.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -09.50/+2.0/095 | OS |  | 04/21/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10049209&pc=QCB | 2023826-2020-01063 | 2020-03-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00/+1.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -09.00/+1.0/088 | OD |  | 04/21/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10057297&pc=QCB | 2023826-2020-01068 | 2020-04-22 | Injury | STAAR SURGICAL COMPANY | 2020-05-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -6.00/1.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. REPORTEDLY, "SURGEON IMPLANTED LENS INTO THE WRONG EYE" OF PATIENT. PER THE REPORTER ON (B)(6) 2020, "LENS WAS REPOSITIONED ON (B)(6) 2020. PATIENT IS HAPPY." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.00/1.0/081 | OD | Physician | 04/23/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10057310&pc=QCB | 2023826-2020-01069 | 2020-04-22 | Injury | STAAR SURGICAL COMPANY | 2020-05-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -6.00/1.0/127 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. REPORTEDLY, "SURGEON IMPLANTED LENS INTO THE WRONG EYE" OF PATIENT. PER THE REPORTER ON (B)(6) 2020, "LENS WAS REPOSITIONED ON (B)(6) 2020. PATIENT IS HAPPY." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.00/1.0/127 | OS | Physician | 04/23/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10061487&pc=QCB | 2023826-2020-01100 | 2019-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -16.0/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.0/2.0/088 | OS |  | 04/19/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10061560&pc=QCB | 2023826-2020-01101 | 2019-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -16.5/1.5/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.5/1.5/077 | OD |  | 04/19/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10061812&pc=QCB | 2023826-2020-01104 | 2012-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.0MM, TICM120V4, -16.5/3.50/108 (SPHERE/ CYLINDER/ AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2012. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.5/3.50/108 | OD |  | 04/22/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10061894&pc=QCB | 2023826-2020-01105 | 2018-07-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.0MM, TICM120V4, -17.5/3.50/82 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2012. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -17.5/3.50/82 | OS |  | 04/22/2020 | 2020 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10062022&pc=QCB | 2023826-2020-01106 | 2020-01-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -12.00/3.50/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/3.50/91 | OD |  | 04/22/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10075051&pc=QCB | 2023826-2020-01091 | 2020-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.50/+3.0/090 (SPHERE/ CYLINDER/ AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.50/+3.0/090 | OD |  | 04/26/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10075061&pc=QCB | 2023826-2020-01092 | 2020-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.00/+3.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.00/+3.0/090 | OS |  | 04/26/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10075961&pc=QCB | 2023826-2020-01083 | 2020-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -4.5/+2.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS LENS ROTATION. ON (B)(6) 2020 THE LENS WAS REPOSITIONED WHICH DID NOT RESOLVED THE PROBLEM. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B1- ADVERSE EVENT B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE B5-THE REPORTER STATES THAT THE SURGEON IMPLANTED A 13.7MM VTICM5 13.7 OF -4.50/2.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE SURGEON REPORTS LENS ROTATION. ON (B)(6) 2020 THE LENS WAS REPOSITIONED, THIS DID NOT RESOLVE THE PROBLEM. THE LENS WAS REMOVED ON (B)(6) 2020. FOUR DAYS LATER ON (B)(6)2020 AN EXCHANGE LENS OF SAME MODEL AND LENGTH BUT DIFFERENT POWER WAS IMPLANTED AND POSITIONED TO 158 DEGREES. THIS RESOLVED THE PROBLEM. D7- EXPLANT DATE: (B)(6) 2020 H6- PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION; EXPLANT CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -4.5/+2.0/095 | OD |  | 04/23/2020 | 2020 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10076815&pc=QCB | 2023826-2020-01107 | 2019-12-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -11.50/2.00/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR THE SAME MODEL AND LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. G4 - G4 DATE IN INITIAL MEDWATCH REPORT OF "20-APR-2022" SHOULD BE CORRECTED TO "22-APR-2020". CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.50/2.00/90 | OS |  | 04/22/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10092931&pc=QCB | 2023826-2020-01145 | 2020-01-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-15.50/3.00/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: LENS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.50/3.00/90 | OD |  | 04/28/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10093051&pc=QCB | 2023826-2020-01143 | 2020-04-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6,-9.00/1.50/99 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/1.50/99 | OS |  | 04/28/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10093190&pc=QCB | 2023826-2020-01147 | 2020-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-14.50/5.5/85 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. PATIENT IS DOING WELL AND THE PROBLEM WAS RESOLVED AFTER EXPLANTING THE LENS. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY, IN LENS CASE/VIAL. VISUAL INSPECTION FOUND A HAPTIC TORN. DIMENSIONAL INSPECTION WAS PERFORMED, OVERALL LENGTH VERIFICATION FOUND THE LENS TO BE WITHIN SPEC. HOWEVER, WIDTH MEASUREMENT WAS NOT PERFORMED DUE TO MISSING/DAMAGED HAPTIC. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.50/5.5/85 | OS |  | 04/28/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10093686&pc=QCB | 2023826-2020-01075 | 2019-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.00/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/2.5/093 | OS |  | 05/06/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10093994&pc=QCB | 2023826-2020-01119 | 2020-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/069 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.50/+1.5/069 | OS |  | 05/05/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10093997&pc=QCB | 2023826-2020-01122 | 2020-01-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.50/+1.5/095 | OD |  | 05/11/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10094016&pc=QCB | 2023826-2020-01130 | 2020-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/102 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.0/102 | OD |  | 04/29/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10094018&pc=QCB | 2023826-2020-01133 | 2020-03-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+1.5/088 | OD |  | 05/05/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10094037&pc=QCB | 2023826-2020-01132 | 2020-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/085 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.5/085 | OD |  | 04/29/2020 | 2020 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10094039&pc=QCB | 2023826-2020-01131 | 2018-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.5/+1.0/081 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS DUE TO LENS ROTATION. REPORTEDLY, THE LENS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.5/+1.0/081 | OS |  | 04/29/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10105252&pc=QCB | 2023826-2020-01078 | 2018-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.00/2.5/086 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CORRECTED DATA: B5: DIOPTER SHOULD BE CORRECTED TO (SPHERE/CYLINDER/AXIS). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/2.5/086 | OS |  | 05/11/2020 | 2020 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10105264&pc=QCB | 2023826-2020-01112 | 2020-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -3.50/0.5/169 (SPHERE/ CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT WITH FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -3.50/0.5/169 | OD |  | 05/11/2020 | 2020 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10105971&pc=QCB | 2023826-2020-01076 | 2020-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.50/2.0/095 (SPHERE / CYLINDER / AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.50/2.0/095 | OS |  | 05/10/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10112134&pc=QCB | 2023826-2020-01153 | 2020-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+2.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.50/+2.0/092 | OD |  | 05/14/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10112199&pc=QCB | 2023826-2020-01154 | 2020-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+2.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.50/+2.0/089 | OS |  | 05/14/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10114935&pc=QCB | 2023826-2020-01116 | 2020-01-18 | Injury | STAAR SURGICAL COMPANY | 2020-06-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -11.50/4.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.50/4.0/084 | OS | Physician | 05/13/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10114921&pc=QCB | 2023826-2020-01114 | 2020-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -8.50/1.0/002 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. REPORTEDLY PATIENT'S NEXT VISIT IS IN (B)(6). "LITTLE CORNEAL EDEMA. INFORMATION ABOUT VAULT ALSO IN (B)(6)." ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.50/1.0/002 | OS |  | 05/15/2020 | 2020 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10116240&pc=QCB | 2023826-2020-01159 | 2019-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+3.5/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.5/+3.5/080 | OS |  | 05/11/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10135953&pc=QCB | 2023826-2020-01173 | 2019-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+4.0/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.00/+4.0/082 | OS |  | 05/18/2020 | 2020 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140629&pc=QCB | 2023826-2020-01212 | 2020-04-24 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -3.5/+1.5/099 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES (ICA). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -3.5/+1.5/099 | OS | Physician | 05/19/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10137636&pc=QCB | 2023826-2020-01219 | 2018-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: HE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.50/1.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2018 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.50/1.0/088 | OD |  | 05/14/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138560&pc=QCB | 2023826-2020-01117 | 2020-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.00/2.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REMOVED AND A LONGER LENGTH LENS WAS LATER IMPLANTED DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.00/2.5/089 | OD |  | 05/19/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138573&pc=QCB | 2023826-2020-01115 | 2020-04-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -8.00/1.0/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.00/1.0/100 | OD |  | 05/18/2020 | 2020 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138872&pc=QCB | 2023826-2020-01226 | 2018-08-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -6.5/1.0/172 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.5/1.0/172 | OD |  | 05/11/2020 | 2020 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138895&pc=QCB | 2023826-2020-01224 | 2019-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.1MM, VTICMO13.2, -9.0/3.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION : H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.0/3.0/090 | OD |  | 05/11/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10138982&pc=QCB | 2023826-2020-01227 | 2019-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -9.5/1.5/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.5/1.5/075 | OD |  | 05/12/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10139453&pc=QCB | 2023826-2020-01228 | 2019-01-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -14.0/3.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC DEFORMED AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.0/3.0/095 | OD |  | 05/12/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140106&pc=QCB | 2023826-2020-01232 | 2020-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -11.50/4.00/84 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.50/4.00/84 | OS |  | 05/13/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140246&pc=QCB | 2023826-2020-01245 | 2020-04-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -13.00/2.00/81 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ON (B)(6) 2020 A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH DEBRIS AND MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE 1494: OFF-LABEL USE (UNDER 21 YRS OF AGE AT DATE OF IMPLANTATION) ADDED TO INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: B5 - "ON (B)(6) 2020 A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED AND PROBLEM RESOLVED." CORRECTED TO "ON (B)(6) 2020 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED." IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.00/2.00/81 | OS |  | 05/13/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140591&pc=QCB | 2023826-2020-01203 | 2020-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.5/+3.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO SUSPICION OF LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: TWO DIFFERENT SAME SIZE, TORIC MODEL LENSES WERE RETURNED IN A VIAL, IN LIQUID. VISUAL INSPECTION FOUND BOTH LENSES NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -13.5/+3.5/092 | OS |  | 05/18/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140602&pc=QCB | 2023826-2020-01202 | 2020-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+3.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO SUSPICION OF LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: TWO DIFFERENT MODELS, SAME LENGTH TORIC LENSES WERE RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND BOTH LENSES WITH NO VISIBLE DAMAGE. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -15.5/+3.5/092 | OD |  | 05/18/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140608&pc=QCB | 2023826-2020-01199 | 2019-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.0/+1.0/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY, WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND STRETCHED OUT WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.0/+1.0/088 | OD |  | 05/14/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140611&pc=QCB | 2023826-2020-01208 | 2019-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/106 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -11.0/+2.0/106 | OD |  | 05/18/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10140615&pc=QCB | 2023826-2020-01207 | 2018-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.5/+1.0/107 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS DUE TO LENS ROTATION. REPORTEDLY, THE LENS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.5/+1.0/107 | OS |  | 05/18/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10158160&pc=QCB | 2023826-2020-01243 | 2020-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -09.00/+1.5/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE ICL ROTATED 60 DEGREES COUNTERCLOCKWISE 5 MONTHS AFTER SURGERY AND THE LENS WAS REPOSITIONED. THE ICL ROTATED AGAIN TO THE SAME POSITION. THE SURGEON PLANS TO EXPLANT/EXCHANGE THE LENS IN (B)(6) 2020. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS REPORTED THAT ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THE PROBLEM RESOLVED. H6 - PATIENT CODE 3191: NO CODE AVAILABLE (SECONDARY SURGERY, LENS EXCHANGE) CLAIM#: (B)(4), Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -09.00/+1.5/096 | OD |  | 05/28/2020 | 2020 | 50
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10158208&pc=QCB | 2023826-2020-01240 | 2019-01-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50/+0.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: B4 - CORRECTED TO (B)(6) 2019, IN INITIAL MDR. G4 - CORRECTED TO (B)(6) 2019, IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -06.50/+0.5/088 | OD |  | 04/25/2019 | 2019 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10163011&pc=QCB | 2023826-2020-01274 | 2020-03-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 11.6MM, VTICH11.6, -4.00/4.5/014 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2019 AND (B)(6) 2019, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -4.00/4.5/014 | OD |  | 05/25/2020 | 2020 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10163016&pc=QCB | 2023826-2020-01250 | 2020-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.00/1.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH RESIDUE AND MOISTURE ON THE LENS. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/1.5/087 | OS |  | 05/28/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10163017&pc=QCB | 2023826-2020-01253 | 2019-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.00/3.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH RESIDUE AND MOISTURE ON THE LENS. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/3.0/092 | OD |  | 05/28/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10163018&pc=QCB | 2023826-2020-01255 | 2019-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICM5_12.6, -14.00/3.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH RESIDUE AND MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE- UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -14.00/3.0/090 | OD |  | 05/28/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10163022&pc=QCB | 2023826-2020-01252 | 2020-05-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -13.50/2.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION- LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.50/2.0/085 | OD |  | 05/28/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10163043&pc=QCB | 2023826-2020-01263 | 2020-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.50/1.5/095 (SPHERE/ CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). (B)(6). | -7.50/1.5/095 | OS |  | 06/01/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10163044&pc=QCB | 2023826-2020-01259 | 2019-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -4.00/3.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -4.00/3.5/080 | OS |  | 06/01/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10163046&pc=QCB | 2023826-2020-01260 | 2020-03-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -11.00/1.0/150 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE- UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -11.00/1.0/150 | OS |  | 06/01/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10163052&pc=QCB | 2023826-2020-01266 | 2020-03-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -16.50/4.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE- UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -16.50/4.0/081 | OS |  | 06/01/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10163054&pc=QCB | 2023826-2020-01264 | 2019-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1, -5.50/1.5/115 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#(B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE- UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -5.50/1.5/115 | OD |  | 06/01/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10163157&pc=QCB | 2023826-2020-01297 | 2020-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2 MM, VTICMO13.2, -11.0/2.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: H6 - DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT) SHOULD BE CORRECTED TO DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DATE OF EVENT WAS CORRECTED TO (B)(6) 2020. DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE- UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -11.0/2.0/092 | OD |  | 05/19/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10163259&pc=QCB | 2023826-2020-01298 | 2020-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2 MM, VTICMO13.2, -12.5/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE- UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -12.5/2.5/093 | OS |  | 05/19/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10167935&pc=QCB | 2023826-2020-01302 | 2020-05-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.50/1.50/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH RESIDUE, DEBRIS AND MOISTURE ON THE LENS. VISUAL INSPECTION FOUND RESIDUE AND DEBRIS ON THE LENS AS WELL AS NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/1.50/90 | OD |  | 05/20/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10167981&pc=QCB | 2023826-2020-01303 | 2020-05-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.50/1.5/85 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND RESIDUE AND DEBRIS ON THE LENS AS WELL AS NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/1.5/85 | OS |  | 05/20/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10168119&pc=QCB | 2023826-2020-01300 | 2020-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -7.50/2.0/87 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND CLEAR SURGICAL RESIDUE AND DEBRIS ON THE LENS AND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.50/2.0/87 | OS |  | 05/20/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10168618&pc=QCB | 2023826-2020-01305 | 2020-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -14.5/2.0/068 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.5/2.0/068 | OS |  | 05/20/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10183530&pc=QCB | 2023826-2020-01342 | 2020-05-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.0/1.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC AND OPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.0/1.0/096 | OD |  | 05/28/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10183642&pc=QCB | 2023826-2020-01343 | 2020-05-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.0/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/1.0/085 | OS |  | 05/28/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10184789&pc=QCB | 2023826-2020-01314 | 2019-12-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.00/6.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B4 - (B)(6) 2020. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.00/6.0/094 | OD |  | 06/05/2020 | 2020 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10184800&pc=QCB | 2023826-2020-01319 | 2020-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.50/1.5/093 | OD |  | 06/09/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10184803&pc=QCB | 2023826-2020-01315 | 2019-12-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -4.00/2.5/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: B4 - CORRECTED TO (B)(6) 2020 IN INITIAL MDR. G3 - CORRECTED TO (B)(6) 2020 IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -4.00/2.5/072 | OD |  | 06/05/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10184820&pc=QCB | 2023826-2020-01313 | 2019-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -12.00/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DAT: B4-DATE OF THIS REPORT IN INITIAL MDR SHOULD BE CORRECTED TO (B)(6) 2020 G4- DATE RECEIVED BY MANUFACTURER IN INITIAL MDR SHOULD BE CORRECTED TO (B)(6) 2020 CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/1.0/087 | OD |  | 06/05/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10184824&pc=QCB | 2023826-2020-01311 | 2020-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -12.50/2.0/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THE PROBLEM WAS NOT RESOLVED AND WAS REPLACED WITH A SPHERICAL LENS AND THAT RESOLVED THE PROBLEM. SEE MFR FOR CLAIM# (B)(4) FOR REPLACEMENT LENS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.50/2.0/107 | OD |  | 06/02/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10184832&pc=QCB | 2023826-2020-01317 | 2019-12-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.50/4.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND DEFORMED. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B4 - (B)(6) 2020. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.50/4.0/079 | OS |  | 06/05/2020 | 2020 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10184839&pc=QCB | 2023826-2020-01318 | 2020-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.50/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: B4 CORRECTED TO 05-JUN 2020, IN INTIAL MDR. G4 CORRECTED TO 05-JUN 2020, IN INTIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.50/2.5/092 | OD |  | 06/05/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10184853&pc=QCB | 2023826-2020-01312 | 2020-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -12.00/2.5/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. B5 - IT WAS LATER REPORTED THAT THE PROBLEM WAS NOT RESOLVED AND REPLACED WITH ANOTHER SPHERICAL LONGER LENGTH LENS WHICH RESOLVED THE PROBLEM. SEE MFR FOR CLAIM# (B)(4) FOR REPLACEMENT LENS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/2.5/077 | OS |  | 06/02/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10184877&pc=QCB | 2023826-2020-01320 | 2020-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.00/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CORRECTED DATA: B5 - VTICMO13.2; 12.00/1.0/087 SHOULD BE CORRECTED TO VTICMO13.2; -12.0/1.5/087. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/1.0/087 | OS |  | 06/09/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10188281&pc=QCB | 2023826-2020-01361 | 2019-03-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.00/2.0/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND DEFORMED WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.00/2.0/112 | OD |  | 06/01/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10188361&pc=QCB | 2023826-2020-01365 | 2020-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.1/1.0/73 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.1/1.0/73 | OD |  | 06/02/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10188459&pc=QCB | 2023826-2020-01364 | 2019-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1,-13.50/2.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND FOREIGN RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.50/2.0/094 | OS |  | 06/01/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10188596&pc=QCB | 2023826-2020-01372 | 2020-03-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.0/1.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.0/1.0/086 | OD |  | 06/01/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10188631&pc=QCB | 2023826-2020-01370 | 2020-04-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-10.5/2.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/2.0/087 | OD |  | 06/01/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10188673&pc=QCB | 2023826-2020-01368 | 2020-04-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -11.0/1.5/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/1.5/104 | OD |  | 06/01/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10188708&pc=QCB | 2023826-2020-01375 | 2019-12-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-6.5/2.0/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: B4- CORRECTED TO (B)(6) 2020, IN INITIAL MDR. G4- CORRECTED TO (B)(6) 2020, IN INITIAL MDR. CLAIM# 7(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.5/2.0/109 | OD |  | 06/05/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10188804&pc=QCB | 2023826-2020-01380 | 2019-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-9.5/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DAT: B4-DATE OF THIS REPORT IN INITIAL MDR SHOULD BE CORRECTED TO 05-JUN-2020. G4- DATE RECEIVED BY MANUFACTURER IN INITIAL MDR SHOULD BE CORRECTED TO 05-JUN-2020. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.5/1.5/095 | OD |  | 06/05/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10188872&pc=QCB | 2023826-2020-01374 | 2019-12-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.0/1.0/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: B4 - CORRECTED TO 05-JUN-2020 IN INITIAL MDR. G4 - CORRECTED TO 05-JUN-2020 IN INITIAL MDR. (B)(4) Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/1.0/072 | OS |  | 06/05/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10188943&pc=QCB | 2023826-2020-01377 | 2020-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-10.0/2.0/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.0/2.0/105 | OD |  | 06/05/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10189271&pc=QCB | 2023826-2020-01373 | 2020-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-11.0/2.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND IT WAS REPORTED THAT THE PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION H3- DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND HAPTIC TORN/BENT AND DEFORMED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/2.0/092 | OD |  | 06/01/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10189305&pc=QCB | 2023826-2020-01371 | 2020-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -4.5/1.0/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. IT WAS REPORTED THAT LENS EXPLANT RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -4.5/1.0/072 | OD |  | 06/01/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10189461&pc=QCB | 2023826-2020-01288 | 2020-03-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50/+2.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -09.50/+2.5/092 | OD |  | 06/01/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10189464&pc=QCB | 2023826-2020-01291 | 2020-04-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+3.0/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND OPTIC TORN/HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.00/+3.0/100 | OS |  | 06/01/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10189503&pc=QCB | 2023826-2020-01282 | 2019-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+5.0/079 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.00/+5.0/079 | OS |  | 06/01/2020 | 2020 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10197348&pc=QCB | 2023826-2020-01334 | 2019-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -9.0/+1.5/085 | OS |  | 06/01/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10197349&pc=QCB | 2023826-2020-01332 | 2020-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/179 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT AXIS LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# : (B)(4). Manufacturer Narrative: H3 - UPDATED DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FIBER ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CORRECTED DATA: B5: IT WAS LATER REPORTED THAT "DUE TO LOW VAULT, DOCTOR EXCHANGED THE LENS WITH SAME LENGTH 90 DEGREE DIFFERENT, AND IMPLANTED HORIZONTALLY. THAT ACTUALLY BASED ON SULCUS TO SULCUS VERTICAL LARGER THAN HORIZONTAL". SHOULD HAVE BEEN INCLUDED IN PREVIOUS MEDWATCH REPORT. H6: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -8.0/+1.5/179 | OD |  | 06/01/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10197351&pc=QCB | 2023826-2020-01340 | 2020-04-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/081 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+1.0/081 | OS |  | 06/02/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10197352&pc=QCB | 2023826-2020-01338 | 2019-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/098 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+2.0/098 | OD |  | 06/01/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10197355&pc=QCB | 2023826-2020-01339 | 2020-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.5/+2.0/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -14.5/+2.0/103 | OD |  | 06/02/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10197356&pc=QCB | 2023826-2020-01335 | 2020-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/113 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.5/113 | OD |  | 06/01/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10197359&pc=QCB | 2023826-2020-01326 | 2019-08-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Adverse Event Without Identified Device or Use Problem | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/098 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON THE SAME DATE THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: CORRECTED DATA: H6 - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. CLAIM# (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.0/098 | OD |  | 06/01/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10197360&pc=QCB | 2023826-2020-01327 | 2020-05-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/083 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.5/083 | OD |  | 06/01/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10197375&pc=QCB | 2023826-2020-01337 | 2020-04-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+1.0/091 | OD |  | 06/01/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10214799&pc=QCB | 2023826-2020-01360 | 2020-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.00/+2.5/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.00/+2.5/091 | OS |  | 06/05/2020 | 2020 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10214819&pc=QCB | 2023826-2020-01359 | 2020-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.00/+2.0/099 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.00/+2.0/099 | OD |  | 06/05/2020 | 2020 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10214845&pc=QCB | 2023826-2020-01358 | 2020-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50/+1.0/102 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CORRECTED DATA: A3 - FEMALE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -08.50/+1.0/102 | OD |  | 06/05/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10214880&pc=QCB | 2023826-2020-01355 | 2020-01-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+3.0/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.00/+3.0/100 | OD |  | 06/05/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10214909&pc=QCB | 2023826-2020-01353 | 2020-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-06-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.50/+2.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION. H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.50/+2.5/090 | OS |  | 06/05/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10222171&pc=QCB | 2023826-2020-01415 | 2020-03-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION : H3-DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.5/+1.5/083 | OS |  | 06/05/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10222174&pc=QCB | 2023826-2020-01414 | 2020-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0/+3.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -8.0/+3.5/089 | OS |  | 06/05/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10222186&pc=QCB | 2023826-2020-01417 | 2020-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#(B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.0/+1.5/095 | OS |  | 06/07/2020 | 2020 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10224726&pc=QCB | 2023826-2020-01395 | 2020-05-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -14.50/3.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#(B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B3 - (B)(6) 2020. CORRECTED DATA: H6 - DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.50/3.0/096 | OD |  | 06/12/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10234849&pc=QCB | 2023826-2020-01387 | 2020-06-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.0/3.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.0/3.5/083 | OS |  | 06/12/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10235894&pc=QCB | 2023826-2020-01438 | 2020-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.5/2.5/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/2.5/110 | OS |  | 06/12/2020 | 2020 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10242453&pc=QCB | 2023826-2020-01424 | 2020-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.50/+2.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -08.50/+2.0/092 | OS |  | 06/08/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10242466&pc=QCB | 2023826-2020-01422 | 2020-04-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.00/+4.0/067 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS RETURNED IN A MICRO CENTRIFUGE VIAL WITH RESIDUE, DEBRIS AND MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -08.00/+4.0/067 | OS |  | 06/08/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10242478&pc=QCB | 2023826-2020-01426 | 2019-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.50/+3.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS REPOSITIONED ON (B)(6) 2019 BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.50/+3.0/088 | OS |  | 06/08/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10242500&pc=QCB | 2023826-2020-01434 | 2020-04-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+3.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM # : (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.00/+3.0/087 | OS |  | 06/19/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10242528&pc=QCB | 2023826-2020-01431 | 2020-06-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION; H6- DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.50/+1.5/082 | OS |  | 06/14/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10242544&pc=QCB | 2023826-2020-01453 | 2020-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+5.5/065 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2020 BUT THE PROBLEM WAS NOT RESOLVED. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED IN EARLY JULY 2020 AND WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT CONDITION AT THE LAST VISIT WAS ALL PERFECT. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.00/+5.5/065 | OD |  | 06/16/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10245033&pc=QCB | 2023826-2020-01400 | 2019-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.50/2.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.50/2.0/081 | OS |  | 06/19/2020 | 2020 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10245035&pc=QCB | 2023826-2020-01399 | 2020-04-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -9.50/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/2.0/091 | OD |  | 06/19/2020 | 2020 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10245038&pc=QCB | 2023826-2020-01436 | 2020-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.50/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CORRECTED DATA: D10 - (PREVIOUSLY D10 WHICH IS NOW D9 SECTION OF THE NEW VERSION OF ESUBMITTER MEDWATCH REPORT): DEVICE WAS NOT AVAILABLE FOR EVALUATION AT THE TIME OF INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.5/092 | OD |  | 06/19/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10246661&pc=QCB | 2023826-2020-01464 | 2020-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.0/1.0/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS LATER EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.0/1.0/92 | OD |  | 06/16/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10247685&pc=QCB | 2023826-2020-01469 | 2020-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -18.0/1.5/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/DEFORMED AND RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -18.0/1.5/099 | OD |  | 06/21/2020 | 2020 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10254160&pc=QCB | 2023826-2020-01450 | 2020-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.0/+1.0/079 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.0/+1.0/079 | OD |  | 06/11/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10254176&pc=QCB | 2023826-2020-01451 | 2020-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5/+1.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.5/+1.0/089 | OS |  | 06/12/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10254178&pc=QCB | 2023826-2020-01446 | 2020-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+4.5/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+4.5/091 | OD |  | 06/11/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10254185&pc=QCB | 2023826-2020-01447 | 2020-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+4.5/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+4.5/092 | OS |  | 06/11/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10254216&pc=QCB | 2023826-2020-01448 | 2020-05-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -9.5/+4.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.5/+4.5/086 | OS |  | 06/11/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10254218&pc=QCB | 2023826-2020-01445 | 2020-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+1.5/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE AND PATIENT-RELATED FACTOR; "RECOMMENDED SIZE TOO SMALL FOR PATIENT'S EYE." Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -18.0/+1.5/072 | OS |  | 06/10/2020 | 2020 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10271734&pc=QCB | 2023826-2020-01472 | 2020-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11/1.0/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM RESOLVED. Manufacturer Narrative: CORRECTED DATA: H3 - "VISUAL INSPECTION FOUND VISIBLE DAMAGE TO THE LENS" SHOULD BE CORRECTED TO "VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS." CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11/1.0/93 | OS |  | 06/23/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10279891&pc=QCB | 2023826-2020-01458 | 2020-05-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.00/+4.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.00/+4.0/092 | OS |  | 06/23/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10279914&pc=QCB | 2023826-2020-01457 | 2020-05-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.00/+3.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND ONE HAPTIC TORN. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.00/+3.5/088 | OD |  | 06/23/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10285710&pc=QCB | 2023826-2020-01488 | 2020-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method; Adverse Event Without Identified Device or Use Problem | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/100 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME LENGTH, DIFFERENT SPHERE/CYLINDER/AXIS LENS DUE TO THE LENS FOR THE RIGHT EYE HAD BEEN IMPLANTED INTO THE LEFT EYE BY MISTAKE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: ADDITIONAL INFORMATION. H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+1.5/100 | OS |  | 06/23/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10285712&pc=QCB | 2023826-2020-01485 | 2020-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.5/+0.5/179 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. THERE WAS RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -7.5/+0.5/179 | OS |  | 06/22/2020 | 2020 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10285763&pc=QCB | 2023826-2020-01487 | 2020-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/097 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+2.0/097 | OD |  | 06/23/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10285823&pc=QCB | 2023826-2020-01482 | 2020-05-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+2.5/090 | OD |  | 06/19/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308793&pc=QCB | 2023826-2020-01514 | 2020-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/2.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/2.5/082 | OD |  | 06/29/2020 | 2020 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308798&pc=QCB | 2023826-2020-01515 | 2020-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.00/1.0/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/1.0/078 | OD |  | 06/29/2020 | 2020 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308828&pc=QCB | 2023826-2020-01521 | 2020-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.50/1.0/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND RESIDUE ON THE LENS AND THE LENS HAS A CURVED SHAPE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.0/108 | OD |  | 06/29/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308832&pc=QCB | 2023826-2020-01527 | 2020-05-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -6.50/1.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2019, BUT DID NOT RESOLVE THE PROBLEM. ON 25/MAY/2020 THE LENS WAS REMOVED AND EXCHANGED WITH SAME SIZE AND POWER LENS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.50/1.0/079 | OS |  | 07/01/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10308839&pc=QCB | 2023826-2020-01526 | 2020-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.00/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND EXCHANGED DURING A SECOND SURGERY ON (B)(6) 2020 WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/1.5/093 | OD |  | 06/30/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10310540&pc=QCB | 2023826-2020-01503 | 2019-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2020 BUT THE PROBLEM WAS NOT RESOLVED SO THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.00/+1.5/085 | OS |  | 06/29/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10310588&pc=QCB | 2023826-2020-01461 | 2020-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND THE HAPTICS TORN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.00/+1.5/086 | OD |  | 06/23/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10310646&pc=QCB | 2023826-2020-01460 | 2020-04-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.50/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND THE LENS OPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.50/+2.0/090 | OS |  | 06/23/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10310745&pc=QCB | 2023826-2020-01553 | 2016-03-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.00/+2.5/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD). THE REPORTER INDICATED THE ANTERIOR CHAMBER COLLAPSED DUE TO THE HAPTIC STUCK IN THE INCISION. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. THE REPORTER INDICATED THE "OVD-FREE" TECHNIQUE WAS DIFFERENT FROM THE STANDARD SURGERY. HAPTICS STUCK IN THE INCISION WOULD LEAD TO SUDDEN COLLAPSE OF THE ANTERIOR CHAMBER, WE DIDN'T HAVE SUCH EXPERIENCE BEFORE. THIS PATIENT WAS PART OF A STUDY COMPARING THE SAFETY OF IMPLANTABLE COLLAMER LENS (ICL) WITH AND WITHOUT OPHTHALMIC VISCOSURGICAL DEVICE (OVD). Manufacturer Narrative: B5: THE LENS WAS IMPLANTED ON (B)(6) 2016. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B1 - CHANGED TO PRODUCT PROBLEM. B2 - OTHER SERIOUS (IMPORTANT MEDICAL EVENTS) SHOULD BE REMOVED AS EVENT IS NOT SERIOUS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.00/+2.5/091 | OD |  | 06/26/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10332343&pc=QCB | 2023826-2020-01588 | 2020-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.0/1.5/124 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.0/1.5/124 | OD |  | 06/29/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10332731&pc=QCB | 2023826-2020-01590 | 2020-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -14.5/1.5/94 (SPHERE/CYINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.5/1.5/94 | OS |  | 06/29/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10332953&pc=QCB | 2023826-2020-01571 | 2019-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.00/2.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/2.5/097 | OD |  | 07/05/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10332962&pc=QCB | 2023826-2020-01575 | 2020-03-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.50/2.5/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 -12.6MM VTICMO12.6; -7.50/2.5/084 (SPHERE/CYLINDER/AXIS) SHOULD BE CORRECTED TO 12.1MM VTICMO12.1; -10/3.5/89 (SPHERE/CYLINDER/AXIS). CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.50/2.5/084 | OD |  | 07/07/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10332987&pc=QCB | 2023826-2020-01580 | 2020-06-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.00/3.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS DID NOT RESOVLE THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CORRECTED DATA B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.00/3.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRA-OPERATIVELY FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. H1- IN INITIAL MDR SERIOUS INJURY IS CORRECTED TO MALFUNCTION. H6- PATIENT CODE 3191 IN INITIAL MDR IS CORRECTED TO 2199. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/3.0/083 | OS |  | 07/10/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10332989&pc=QCB | 2023826-2020-01581 | 2020-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -9.00/1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "VAULT IS ADEQUATE NOW AND PATIENT IS DOING VERY WELL." Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: UNK, UNK, UNK. THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE: 3191 - SECONDARY SURGICAL INTERVENTION, LENS EXCHANGE. CLAIM# (B)(4). | -9.00/1.0/092 | OD |  | 07/10/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10332994&pc=QCB | 2023826-2020-01579 | 2020-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.00/3.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS DID NOT RESOLVE THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, SURGEON FOUND THE VALUED WAS TOO LOW, EXCHANGED WITH ANOTHER LONGER LENGTH LENS DURING THE (B)(6) 2020 SURGERY" TO RESOLVE THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT AND STRETCHED ALOT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: H10: IN INITIAL MDR SHOULD INCLUDE STATEMENT: H6: 4110: LENS WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/3.0/088 | OS |  | 07/10/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10333018&pc=QCB | 2023826-2020-01593 | 2019-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -14.0/2.0/087 (SPHERE/CYINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH DEBRIS ON THE LENS AND LENS WAS DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC STRETCHED OUT WITH DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.0/2.0/087 | OD |  | 06/30/2020 | 2020 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10333143&pc=QCB | 2023826-2020-01594 | 2020-04-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -15.0/5.0/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: UNK, UNK, UNK. THIS PRODUCT IS NOT MARKETED IN THE US. PATIENT CODE 3191: SECONDARY SURGICAL INTERVENTION, EXCHANGE. DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -15.0/5.0/098 | OS |  | 06/30/2020 | 2020 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10333722&pc=QCB | 2023826-2020-01596 | 2019-11-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -14.50/2.0/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.50/2.0/102 | OD |  | 06/30/2020 | 2020 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10334048&pc=QCB | 2023826-2020-01601 | 2020-03-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -11.5/3.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. IT WAS REPORTED THAT AFTER EXPLANTING THE LENS THE PROBLEM RESOLVED. CAUSE ON EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/3.0/090 | OD |  | 06/29/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10334130&pc=QCB | 2023826-2020-01600 | 2020-03-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -11.5/3.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. IT WAS REPORTED THAT AFTER EXPLANTING THE LENS THE PROBLEM RESOLVED. CAUSE ON EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC SPLIT ANND THE HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/3.0/084 | OS |  | 06/29/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335236&pc=QCB | 2023826-2020-01566 | 2020-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00/+2.5/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND SHALLOW ANTERIOR CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND THE HAPTIC STRETCHED OUT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -09.00/+2.5/087 | OS |  | 07/07/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335246&pc=QCB | 2023826-2020-01563 | 2020-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -14.00/+2.0/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.00/+2.0/084 | OD |  | 07/03/2020 | 2020 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335282&pc=QCB | 2023826-2020-01560 | 2020-06-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.50/+3.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A DIFFERENT TYPE OF LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.50/+3.0/091 | OD |  | 07/03/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335296&pc=QCB | 2023826-2020-01558 | 2020-06-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50/+0.5/112 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.50/+0.5/112 | OS |  | 07/01/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335315&pc=QCB | 2023826-2020-01555 | 2020-04-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.50/+1.5/088 | OD |  | 06/30/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335324&pc=QCB | 2023826-2020-01565 | 2020-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00/+2.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND SHALLOW ANTERIOR CHAMBER. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: AGE: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -09.00/+2.0/089 | OD |  | 07/07/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335326&pc=QCB | 2023826-2020-01543 | 2019-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.0/+1.0/076 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -6.0/+1.0/076 | OS |  | 06/29/2020 | 2020 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335328&pc=QCB | 2023826-2020-01535 | 2019-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+2.5/093 | OS |  | 06/29/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335333&pc=QCB | 2023826-2020-01549 | 2020-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.5/+4.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON THE SAME DATE, IN A SEPARATE SURGERY, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE PROBLEM WAS RESOLVED AFTER EXPLANT. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+4.0/089 | OS |  | 06/30/2020 | 2020 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335339&pc=QCB | 2023826-2020-01539 | 2020-01-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.0/080 | OD |  | 06/29/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335344&pc=QCB | 2023826-2020-01548 | 2020-03-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/018 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.5/+1.0/018 | OD |  | 06/30/2020 | 2020 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335351&pc=QCB | 2023826-2020-01537 | 2020-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.5/+1.5/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.5/+1.5/089 | OD |  | 06/29/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335517&pc=QCB | 2023826-2020-01605 | 2020-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.5/+3.5/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON THE SAME DATE, IN A SEPARATE SURGERY, THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE PROBLEM WAS RESOLVED AFTER EXPLANT. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+3.5/088 | OD |  | 07/01/2020 | 2020 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335672&pc=QCB | 2023826-2020-01597 | 2020-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.00/2.50/80 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A NON TORIC LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PROBLEM RESOLVED. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/2.50/80 | OS |  | 07/02/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10335730&pc=QCB | 2023826-2020-01598 | 2020-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.50/2.00/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A NON TORIC LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PROBLEM RESOLVED. Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.50/2.00/105 | OD |  | 07/02/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10353749&pc=QCB | 2023826-2020-01637 | 2020-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+0.5/074 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.0/+0.5/074 | OD |  | 07/06/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10353750&pc=QCB | 2023826-2020-01633 | 2020-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/101 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.5/101 | OD |  | 07/06/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10353751&pc=QCB | 2023826-2020-01634 | 2020-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+1.5/088 | OS |  | 07/06/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10353754&pc=QCB | 2023826-2020-01632 | 2020-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -10.0/+2.5/110 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE, IT FLIPPED UPSIDE DOWN. THIS OCCURRED ON (B)(6) 2020. REPORTEDLY, THE LENS WAS THEN REMOVED AND NO PATIENT INJURY OCCURRED. ATTEMPTS TO OBTAIN ADDITIONAL ELEVATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. THERE WAS RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.0/+2.5/110 | NA |  | 07/06/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10354269&pc=QCB | 2023826-2020-01621 | 2020-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.00/+2.0/088 | OS |  | 07/15/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10354271&pc=QCB | 2023826-2020-01620 | 2020-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.00/+1.0/088 | OD |  | 07/15/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10354275&pc=QCB | 2023826-2020-01610 | 2020-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.50/+1.0/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL, SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED (THE LENS WAS IMPLANTED VERTICALLY). THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINT TYPE EVENTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: D9 - RETURNED TO MANUFACTURER DATE CORRECTED TO (B)(6) 2020. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.50/+1.0/100 | OD |  | 07/09/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10354276&pc=QCB | 2023826-2020-01616 | 2020-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.00/+2.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC DEFORMED AND DEBRIS ON LENS. CLAM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.00/+2.0/094 | OD |  | 07/12/2020 | 2020 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10355551&pc=QCB | 2023826-2020-01652 | 2019-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -7.50/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H3 - PREVIOUS PRODUCT EVALUATION REPORT SHOULD BE CORRECTED TO: H3 LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. H10 - CLAIM#: (B)(4) SHOULD BE REMOVED AS IT WAS THE WRONG PRODUCT EVALUATION REPORT SUBMITTED, AND CLAIM NUMBER REFERENCED. H10 - PREVIOUS PRODUCT EVALUATION REPORT WAS INTENDED FOR CLAIM# (B)(4), SEE MFR# 2023826-2020-01844. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). THIS PRODUCT IS NOT MARKETED IN THE US | -7.50/2.0/095 | OD |  | 07/07/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10355672&pc=QCB | 2023826-2020-01656 | 2020-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -12.0/2.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.0/2.5/085 | OS |  | 07/10/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10359963&pc=QCB | 2023826-2020-01654 | 2020-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.0/2.5/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOITURE AND DEVRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.0/2.5/084 | OS |  | 07/08/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10360232&pc=QCB | 2023826-2020-01655 | 2020-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -5.5/0.5/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.5/0.5/92 | OD |  | 07/10/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10361985&pc=QCB | 2023826-2020-01624 | 2018-10-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -12.50/2.0/118 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.50/2.0/118 | OD |  | 07/15/2020 | 2020 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10361991&pc=QCB | 2023826-2020-01628 | 2020-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -8.50/1.0/162 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED ON (B)(6) 2020 DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A CASE/VIAL AND DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.50/1.0/162 | OS |  | 07/18/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10361995&pc=QCB | 2023826-2020-01629 | 2018-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.00/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2018 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION-LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/1.5/092 | OS | Physician | 07/18/2020 | 2020 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10364445&pc=QCB | 2023826-2020-01648 | 2020-04-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+3.0/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/+3.0/091 | OD |  | 07/10/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10364447&pc=QCB | 2023826-2020-01647 | 2020-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.5/+1.0/126 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH RESIDUE ON THE LENS. CORRECTED DATA: H10 - "H6 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT." SHOULD HAVE BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.5/+1.0/126 | OD |  | 07/09/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10364448&pc=QCB | 2023826-2020-01646 | 2020-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.0/+4.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#(B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -12.0/+4.0/094 | OS |  | 07/09/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10380125&pc=QCB | 2023826-2020-01659 | 2020-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.5/4.0/092 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE REPOSITIONING DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS EXCHANGED AND THE PROBLEM RESOLVED. Manufacturer Narrative: CORRECTED DATA: H6 - DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.5/4.0/092 | OS |  | 07/10/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10392912&pc=QCB | 2023826-2020-01669 | 2020-03-01 | Injury | STAAR SURGICAL COMPANY | 2020-08-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.0/003 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/+1.0/003 | OD | Physician | 07/22/2020 | 2020 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10392362&pc=QCB | 2023826-2020-01665 | 2020-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50/+1.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.50/+1.5/093 | OS |  | 07/21/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10392446&pc=QCB | 2023826-2020-01664 | 2020-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.50/+1.5/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.50/+1.5/086 | OD |  | 07/21/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10392715&pc=QCB | 2023826-2020-01661 | 2019-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -16.00/+3.0/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND FOREIGN RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -16.00/+3.0/082 | OS |  | 07/21/2020 | 2020 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10392750&pc=QCB | 2023826-2020-01671 | 2019-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.00/+2.0/094 | OD |  | 07/22/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10392854&pc=QCB | 2023826-2020-01675 | 2020-07-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.00/+2.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/+2.5/090 | OD |  | 07/24/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10398712&pc=QCB | 2023826-2020-01708 | 2020-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -06.50/+2.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING AND LENS ROTATION THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO SMALL DIAMETER ICL CAUSED LOW VAULTING AND TORIC ICL ROTATION. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | -06.50/+2.5/089 | OD |  | 07/15/2020 | 2020 | 49
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10401676&pc=QCB | 2023826-2020-01701 | 2020-02-06 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -09.00/1.0/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SLIGHT IRIS ATROPHY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -09.00/1.0/071 | OS | Physician | 07/24/2020 | 2020 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10401614&pc=QCB | 2023826-2020-01680 | 2019-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/2.0/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.00/2.0/101 | OD |  | 07/21/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10401617&pc=QCB | 2023826-2020-01685 | 2019-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -08.50/1.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -08.50/1.0/088 | OD |  | 07/21/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10401620&pc=QCB | 2023826-2020-01684 | 2019-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.00/3.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION:LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -18.00/3.5/088 | OD |  | 07/21/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10401621&pc=QCB | 2023826-2020-01686 | 2019-09-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.00/1.5/114 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/1.5/114 | OD |  | 07/21/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10401622&pc=QCB | 2023826-2020-01687 | 2019-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/3.0/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/3.0/102 | OD |  | 07/21/2020 | 2020 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10401628&pc=QCB | 2023826-2020-01688 | 2019-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.00/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2019 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.00/1.0/093 | OS |  | 07/21/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10401663&pc=QCB | 2023826-2020-01697 | 2020-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.50/3.0/076 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#: 716934. | -8.50/3.0/076 | OD |  | 07/22/2020 | 2020 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10403849&pc=QCB | 2023826-2020-01744 | 2019-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+1.5/102 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.5/+1.5/102 | OD |  | 07/21/2020 | 2020 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10403852&pc=QCB | 2023826-2020-01740 | 2020-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/081 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+1.5/081 | OS |  | 07/10/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10403857&pc=QCB | 2023826-2020-01737 | 2020-03-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.5/+2.0/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -14.5/+2.0/103 | OD |  | 07/21/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10403861&pc=QCB | 2023826-2020-01728 | 2020-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/104 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON 9B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.5/+1.0/104 | OS |  | 07/15/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10403869&pc=QCB | 2023826-2020-01729 | 2020-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -11.0/+2.0/093 | OD |  | 07/15/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10403877&pc=QCB | 2023826-2020-01731 | 2020-02-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5/+6.0/139 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -7.5/+6.0/139 | OD |  | 07/16/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10403901&pc=QCB | 2023826-2020-01745 | 2019-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.0/+1.0/050 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND RESIDUE ON THE LENS AND THE LENS HAS A CURVED SHAPE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -16.0/+1.0/050 | OS |  | 07/21/2020 | 2020 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10403915&pc=QCB | 2023826-2020-01727 | 2020-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/077 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.5/+1.0/077 | OD |  | 07/15/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10416008&pc=QCB | 2023826-2020-01722 | 2020-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.50/+1.5/179 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED LATER ON THE SAME DAY OF (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT, TORN AND DEFORMED WITH FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.50/+1.5/179 | OD |  | 07/26/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10416026&pc=QCB | 2023826-2020-01717 | 2019-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+1.0/076 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -09.00/+1.0/076 | OD |  | 07/21/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10416034&pc=QCB | 2023826-2020-01716 | 2020-03-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.00/+3.5/098 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.00/+3.5/098 | OD |  | 07/21/2020 | 2020 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10416042&pc=QCB | 2023826-2020-01711 | 2020-04-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.00/+4.0/102 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.00/+4.0/102 | OD |  | 07/18/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10416044&pc=QCB | 2023826-2020-01712 | 2020-04-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+3.5/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.00/+3.5/083 | OS |  | 07/18/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10418326&pc=QCB | 2023826-2020-01785 | 2020-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -12.5/4.5/80 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND RESIDUE ON THE LENS. CORRECTED DATA: G4: 18/JUL/2020 SHOULD BE CORRECTED TO 20/JUL/2020 IN INTIAL MEDWATCH REPORT. B4: SHOULD BE CORRECTED TO 20/JUL/2020. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.5/4.5/80 | OD |  | 07/20/2020 | 2020 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10426520&pc=QCB | 2023826-2020-01760 | 2020-06-23 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/1.5/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.50/1.5/079 | OS | Physician | 07/28/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10426528&pc=QCB | 2023826-2020-01765 | 2020-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.50/1.0/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.0/071 | OD |  | 07/29/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10426553&pc=QCB | 2023826-2020-01775 | 2020-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -7.00/2.0/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.00/2.0/073 | OS |  | 07/29/2020 | 2020 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10426572&pc=QCB | 2023826-2020-01756 | 2020-04-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/2.0/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION- LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/2.0/080 | OD |  | 07/27/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10426580&pc=QCB | 2023826-2020-01762 | 2020-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -13.00/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.00/1.0/091 | OS |  | 07/29/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10426582&pc=QCB | 2023826-2020-01767 | 2020-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -9.50/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDIE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.5/093 | OS |  | 07/29/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10426583&pc=QCB | 2023826-2020-01766 | 2020-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -2.00/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED ON (B)(6) 2020 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -2.00/1.5/091 | OD |  | 07/29/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10426598&pc=QCB | 2023826-2020-01780 | 2020-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -1.50/3.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: B5:"-1.50/3.5/092 (SPHERE/CYLINDER/AXIS)" SHOULD BE CORRECTED TO "-8.5/3.5/092 (SPHERE/CYLINDER/AXIS)" IN THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -1.50/3.5/092 | OD |  | 07/30/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10426601&pc=QCB | 2023826-2020-01774 | 2020-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.00/2.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2020 THE LENS WAS REMOVED AND EXCHANGED FOR A SPHERICAL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/2.5/090 | OD |  | 07/30/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10426603&pc=QCB | 2023826-2020-01768 | 2020-03-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/1.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: G3 - CORRECTED TO 20-NOV-2020 IN SUPPLEMENTAL MDR#1. H10 - CLAIM# (B)(4) IS N/A. CORRECTED TO CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/1.0/094 | OD |  | 07/29/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10426605&pc=QCB | 2023826-2020-01777 | 2020-04-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -7.00/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.00/1.5/090 | OS |  | 07/30/2020 | 2020 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10426612&pc=QCB | 2023826-2020-01781 | 2020-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -09.00/4.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -09.00/4.0/084 | OS |  | 07/30/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429081&pc=QCB | 2023826-2020-01843 | 2020-03-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/021 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE LENS, PLACED VERTICALLY, DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5 - "(B)(6) 2020" SHOULD BE CORRECTED TO "ON A UNKNOWN DATE". D7/D6B - EXPLANTED DATE: SHOULD BE CORRECTED TO UNK. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.5/+2.0/021 | OS |  | 07/27/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429082&pc=QCB | 2023826-2020-01827 | 2020-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -12.0/+2.0/072 | OS |  | 07/21/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429132&pc=QCB | 2023826-2020-01828 | 2019-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.0/+1.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -6.0/+1.0/089 | OS |  | 07/21/2020 | 2020 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429139&pc=QCB | 2023826-2020-01842 | 2020-01-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+1.0/090 | OD |  | 07/27/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429467&pc=QCB | 2023826-2020-01795 | 2019-04-19 | Injury | STAAR SURGICAL COMPANY | 2020-08-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -8.5/1.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/1.0/079 | OD | Physician | 07/21/2020 | 2020 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429440&pc=QCB | 2023826-2020-01797 | 2019-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.0/2.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/2.0/085 | OD |  | 07/21/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429462&pc=QCB | 2023826-2020-01798 | 2019-01-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2 -10.0/3.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.0/3.0/079 | OS |  | 07/21/2020 | 2020 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429468&pc=QCB | 2023826-2020-01787 | 2019-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.0/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.0/1.5/090 | OD |  | 03/26/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429490&pc=QCB | 2023826-2020-01792 | 2019-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.5/2.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/2.0/085 | OD |  | 07/21/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429562&pc=QCB | 2023826-2020-01794 | 2019-03-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.5/2.5/014 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PROBLEM RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN WITH RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.5/2.5/014 | OD |  | 07/21/2020 | 2020 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10429594&pc=QCB | 2023826-2020-01802 | 2020-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -9.5/2.0/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE DID NOT RESOLVE THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.5/2.0/111 | OS |  | 07/21/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10436171&pc=QCB | 2023826-2020-01776 | 2020-04-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_13.2, -8.5/2.0/85 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/2.0/85 | OD |  | 07/30/2020 | 2020 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10436173&pc=QCB | 2023826-2020-01771 | 2020-04-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -8.50/1.5/168 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.50/1.5/168 | OS |  | 07/29/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10440447&pc=QCB | 2023826-2020-01896 | 2019-09-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.5/1.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/1.5/080 | OD |  | 07/21/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10440593&pc=QCB | 2023826-2020-01807 | 2020-07-07 | Injury | STAAR SURGICAL COMPANY | 2020-08-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00/+4.0/056 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). | -12.00/+4.0/056 | OS | Physician | 07/29/2020 | 2020 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10440596&pc=QCB | 2023826-2020-01822 | 2020-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.00/+3.0/097 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -08.00/+3.0/097 | OD |  | 07/30/2020 | 2020 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10440599&pc=QCB | 2023826-2020-01821 | 2020-05-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.00/+4.0/109 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.00/+4.0/109 | OD |  | 07/30/2020 | 2020 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10446469&pc=QCB | 2023826-2020-01910 | 2020-05-05 | Injury | STAAR SURGICAL COMPANY | 2020-08-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO 12.6, -9.00/0.5/100(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/0.5/100 | OD | Physician | 07/30/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10445735&pc=QCB | 2023826-2020-01812 | 2020-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.50/+3.0/101 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.50/+3.0/101 | OD |  | 07/29/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10445767&pc=QCB | 2023826-2020-01818 | 2020-05-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/099 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.50/+1.0/099 | OS |  | 07/29/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10445776&pc=QCB | 2023826-2020-01823 | 2020-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.50/+2.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING . THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THERE WAS RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.50/+2.5/090 | OD |  | 07/30/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10445819&pc=QCB | 2023826-2020-01824 | 2020-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -13.00/+4.5/097 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING . THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.00/+4.5/097 | OS |  | 07/30/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10445824&pc=QCB | 2023826-2020-01813 | 2020-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+3.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/+3.0/089 | OD |  | 07/29/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10446465&pc=QCB | 2023826-2020-01902 | 2020-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO 12.6, -10.0/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.0/1.0/085 | OD |  | 07/29/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10446466&pc=QCB | 2023826-2020-01907 | 2020-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO 12.6, -12.5/1.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.5/1.5/088 | OD |  | 07/29/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10446467&pc=QCB | 2023826-2020-01908 | 2020-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO 12.6, -13.0/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.0/2.0/089 | OS |  | 07/29/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10446468&pc=QCB | 2023826-2020-01909 | 2020-04-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO 13.2, -10.0/1.5/.0/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.0/1.5/.0/89 | OD |  | 07/29/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10446470&pc=QCB | 2023826-2020-01906 | 2020-04-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO 12.6, -17.5/4.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -17.5/4.0/092 | OD |  | 07/29/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10446475&pc=QCB | 2023826-2020-01901 | 2020-06-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO 12.6, -9.00/1.5/88(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/1.5/88 | OS |  | 07/29/2020 | 2020 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10451070&pc=QCB | 2023826-2020-01848 | 2020-06-15 | Injury | STAAR SURGICAL COMPANY | 2020-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.50/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.50/2.0/091 | OD | Physician | 08/03/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10451126&pc=QCB | 2023826-2020-01859 |  | Injury | STAAR SURGICAL COMPANY | 2020-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICH13.2, -10.00/2.0/135 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS REMOVED ON (B)(6) 2020. A SHORTER LENGHT LENS WAS IMPLANTED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/2.0/135 | OS | Physician | 08/07/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10451061&pc=QCB | 2023826-2020-01849 | 2020-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.00/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/1.5/092 | OS |  | 08/03/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10451086&pc=QCB | 2023826-2020-01858 | 2020-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -11.00/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. REPORTEDLY, "LENS KEPT ROTATING 90 DEGREES." THE LENS WAS REMOVED ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. PER REPORTER ON (B)(6) 2020, DOCTOR "DECIDED NOT TO IMPLANT A NEW LENS. PATIENT WILL CONTINURE TO WEAR CONTACT LENSES AND GLASSES." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CORRECTED DATA: H6 - DEVICE CODE 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT) . CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. | -11.00/2.5/093 | OS |  | 08/06/2020 | 2020 | 62
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10451092&pc=QCB | 2023826-2020-01860 |  | Injury | STAAR SURGICAL COMPANY | 2020-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICH13.2, -10.00/2.0/050 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED. THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN AND RESIDUE COVERING THE WHOLE LENS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/2.0/050 | OD |  | 08/07/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10451116&pc=QCB | 2023826-2020-01847 | 2020-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/3.0/094 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/3.0/094 | OD |  | 08/02/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10451119&pc=QCB | 2023826-2020-01861 | 2020-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND WAS DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC DEFORMED AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/2.0/088 | OD |  | 08/07/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10451129&pc=QCB | 2023826-2020-01862 | 2020-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.50/2.0/087 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND STRETCHED OUT WITH RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/2.0/087 | OS |  | 08/07/2020 | 2020 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457345&pc=QCB | 2023826-2020-01877 | 2020-04-14 | Injury | STAAR SURGICAL COMPANY | 2020-08-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.0/+1.5/094 | OS | Physician | 07/29/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457336&pc=QCB | 2023826-2020-01889 | 2019-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+3.0/083 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN, AND RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+3.0/083 | OS |  | 08/04/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457338&pc=QCB | 2023826-2020-01885 | 2020-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+2.0/090 | OS |  | 08/03/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457340&pc=QCB | 2023826-2020-01886 | 2020-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+1.5/091 | OD |  | 08/03/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457341&pc=QCB | 2023826-2020-01890 | 2020-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+3.5/075 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THERE IS REPORTEDLY NO PLAN TO REPLACE THE LENS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -12.0/+3.5/075 | OS |  | 08/04/2020 | 2020 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457343&pc=QCB | 2023826-2020-01887 | 2020-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.0/+2.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. ON 30-JUL-2020 THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. (B)(4) Manufacturer Narrative: CORRECTED DATA: H3 - VISUAL INSPECTION SECTION OF THE DEVICE EVALUATION SHOULD BE CORRECTED TO: "VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH RESIDUE ON THE LENS" TO SUPPLEMENTAL MEDWATCH REPORT #1 (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+2.5/093 | OS |  | 08/03/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457346&pc=QCB | 2023826-2020-01874 | 2020-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.0/+1.0/071 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -16.0/+1.0/071 | OS |  | 07/29/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457348&pc=QCB | 2023826-2020-01881 | 2019-05-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/162 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND RESIDUE ON THE LENS AS WELL AS THE HAPTIC TORN AND BENT. CLAIM#: (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -13.5/+2.0/162 | OD |  | 07/31/2020 | 2020 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457350&pc=QCB | 2023826-2020-01878 | 2020-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.0/+4.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6)2020 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL, DRY WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US.. (B)(4). | -15.0/+4.5/091 | OS |  | 07/31/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457352&pc=QCB | 2023826-2020-01875 | 2020-04-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.5/+3.0/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.5/+3.0/080 | OS |  | 07/29/2020 | 2020 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457353&pc=QCB | 2023826-2020-01872 | 2020-03-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/102 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE PROBLEM WAS RESOLVED AFTER EXPLANT. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL, DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN. UNABLE TO PERFORM DIMENSIONAL INSPECTION SINCE BOTH HAPTIC BROKEN. CLAIM#(B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.5/+1.0/102 | OD |  | 07/29/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457354&pc=QCB | 2023826-2020-01871 | 2020-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+2.5/084 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -8.5/+2.5/084 | OS |  | 07/29/2020 | 2020 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457356&pc=QCB | 2023826-2020-01880 | 2020-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -9.0/+3.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -9.0/+3.5/086 | OD |  | 07/31/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457359&pc=QCB | 2023826-2020-01873 | 2020-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+2.5/103 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.5/+2.5/103 | OD |  | 07/29/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457360&pc=QCB | 2023826-2020-01888 | 2019-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+3.0/079 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+3.0/079 | OD |  | 08/04/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10457364&pc=QCB | 2023826-2020-01867 | 2020-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/172 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE UPSIDE DOWN. THIS OCCURRED ON (B)(6) 2020. THE LENS WAS EXPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL ELEVATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: ON (B)(6) 2020 THE LENS WAS SUCCESSFULLY EXCHANGED WITH AN ALTERNATE LENS AND THE PROBLEM IS RESOLVED. REPORTEDLY THE LENS WAS IN "TACO" CONFIGURATION WHEN IT WAS DEPLOYED IN THE ANTERIOR CHAMBER. DURING MANEUVERS TO UNFOLD, THE LENS FLIPPED. THE EXPLANTED LENS HAD SOME PIGMENT DEPOSITS AND WAS SLIGHTLY DISFIGURED, NECESSITATING AN EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED DRY IN A CARTRIDGE CASE. THERE WAS RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+1.5/172 | OD |  | 07/29/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10463778&pc=QCB | 2023826-2020-01913 | 2020-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.0/1.5/176, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.0/1.5/176 | OD |  | 07/29/2020 | 2020 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10463779&pc=QCB | 2023826-2020-01912 | 2020-04-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -13.0/1.0/011, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4) Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.0/1.0/011 | OD |  | 07/29/2020 | 2020 | 49
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10463782&pc=QCB | 2023826-2020-01918 | 2020-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.0/1.0/069 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.0/1.0/069 | OS |  | 07/31/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10463783&pc=QCB | 2023826-2020-01917 | 2020-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -4.0/0.5/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -4.0/0.5/077 | OD |  | 07/31/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476309&pc=QCB | 2023826-2020-01930 | 2020-06-30 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+2.0/083 | OD | Physician | 08/04/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476171&pc=QCB | 2023826-2020-01940 | 2020-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND DEBRIS ON LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.0/090 | OS |  | 08/07/2020 | 2020 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10476351&pc=QCB | 2023826-2020-01923 | 2020-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.00/+4.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -16.00/+4.0/091 | OD |  | 08/02/2020 | 2020 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10480707&pc=QCB | 2023826-2020-01999 | 2019-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -17.5/5.5/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -17.5/5.5/096 | OD |  | 08/04/2020 | 2020 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10480708&pc=QCB | 2023826-2020-02005 | 2019-11-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.0/2.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC TORN AND THE HAPTIC BROKEN. CORRECTED DATA: H6 -MEDICAL DEVICE PROBLEM CODE 1494: OFF-LABEL USE (ACD < 3.0MM). H6 - RESULT CODE: 114 SHOULD HAVE BEEN 213 IN INITIAL MEDWATCH REPORT, BUT WILL BE KEPT AS 114 IN CURRENT SUPPLEMENTAL DUE TO PRODUCT EVALUATION. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 - "THIS EXCHANGE RESOLVED THE PROBLEM" SHOULD BE CORRECTED TO "THIS EXPLANT RESOLVED THE PROBLEM". CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.0/2.5/088 | OS |  | 08/03/2020 | 2020 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10480711&pc=QCB | 2023826-2020-01997 | 2020-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.5/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND FOREIGN RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/2.5/092 | OD |  | 08/03/2020 | 2020 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10483341&pc=QCB | 2023826-2020-01944 | 2020-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -14.00/3.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.00/3.5/089 | OS |  | 08/10/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10483342&pc=QCB | 2023826-2020-01942 | 2020-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -9.00/2.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS ON THE LENS. H6 - MEDICAL DEVICE PROBLEM CODE: 2923 SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/2.0/082 | OD |  | 08/10/2020 | 2020 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10483343&pc=QCB | 2023826-2020-01950 |  | Injury | STAAR SURGICAL COMPANY | 2020-09-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -11.50/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN WITH RESIDUE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: HEALTH IMPACT CODE: 4625 - ENLARGEMENT OF INCISION SHOULD BE ADDED TO INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.50/1.5/089 | OS |  | 08/12/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10483345&pc=QCB | 2023826-2020-01943 | 2020-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -14.00/2.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2020 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.00/2.5/088 | OD |  | 08/10/2020 | 2020 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10483347&pc=QCB | 2023826-2020-01949 | 2020-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.00/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS AND EXCESSIVE VAULT WERE OBSERVED, THE LENS WAS REMOVED AND EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. (B)(4). Manufacturer Narrative: PMA/510K- THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/1.5/092 | OD |  | 08/12/2020 | 2020 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10491992&pc=QCB | 2023826-2020-01961 | 2020-08-04 | Injury | STAAR SURGICAL COMPANY | 2020-09-03 | QCB | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2017. THE SURGEON REPORTS A SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT (ECC). REPORTEDLY, THE ECC LOSS BEGAN PRIOR TO INITIAL IMPLANT: 2841 ON (B)(6) 2017 AND 2379 ON THE DATE OF SURGERY. Manufacturer Narrative: PMA/510K-THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/+1.5/082 | OS | Physician | 08/05/2020 | 2020 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10491979&pc=QCB | 2023826-2020-01964 | 2020-04-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/099 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE LENS. CLAIM#:(B)(4). Manufacturer Narrative: PMA/510K-THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+1.0/099 | OD |  | 08/07/2020 | 2020 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10491982&pc=QCB | 2023826-2020-01963 | 2020-08-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.5/+4.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THEL LENS. CLAIM# (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -10.5/+4.0/091 | OD |  | 08/06/2020 | 2020 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10491984&pc=QCB | 2023826-2020-01966 | 2020-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+2.0/088 | OS |  | 08/07/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10491989&pc=QCB | 2023826-2020-01956 |  | Injury | STAAR SURGICAL COMPANY | 2020-09-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -14.0/+4.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 LENS, -14.0/+4.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) EYE ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. (B)(4) Manufacturer Narrative: DATE OF EVENT-UNK. (B)(4). | -14.0/+4.0/093 | OD |  | 08/05/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10491991&pc=QCB | 2023826-2020-01959 | 2020-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS -6.5/+1.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE. THE SURGEON REPORTEDLY EXPLANTED THE LENS BECAUSE THE LENS IS "TOO BIG." Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE LENS WAS EXPLANTED FROM THE RIGHT (OD) EYE DUE TO EXCESSIVE VAULT. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "THE ICL WAS TOO LARGE FOR THE EYE DESPITE USING THE OPTIMIZED WTW MEASUREMENT." H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A VIAL. THERE WAS DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.5/+1.0/087 | OD |  | 08/10/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10492096&pc=QCB | 2023826-2020-02018 | 2020-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -10.5/2.0/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE IN A STERILITY POUCH. VISUAL INSPECTION FOUND ONE OF THE HAPTICS BENT AND RESIDUE ON LENS SURFACE. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | -10.5/2.0/108 | OS |  | 08/04/2020 | 2020 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10496556&pc=QCB | 2023826-2020-01973 | 2020-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -16.5/+1.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. ON 17-JUL-2020 THE LENS WAS EXCHANGED WITH A SAME SIZE, DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.5/+1.5/093 | OS |  | 08/10/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10496575&pc=QCB | 2023826-2020-01970 | 2020-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/105 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+1.5/105 | OD |  | 08/09/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10496602&pc=QCB | 2023826-2020-01971 | 2020-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/070 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+1.5/070 | OS |  | 08/09/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10496606&pc=QCB | 2023826-2020-01972 | 2020-01-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-09-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/098 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTIFUGE VIAL; DRY. VISUAL INSPECTION FOUND HAPTIC TORN/BENT/AND DEFORMED. (B)(4) Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. (B)(4). | -8.5/+1.0/098 | OS |  | 08/09/2020 | 2020 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8670840&pc=QCB | 2023826-2019-01008 | 2017-01-10 05:00 | Injury | STAAR SURGICAL COMPANY AG | 2019-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available | N/A | Event Description: THE REPORTER INDICATED THAT A 12.0MM, TICM120V4, -11.5/+1.5/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2014. LOW VAULT WITH ROTATION WAS OBSERVED. LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2017 (1ST) AND ON (B)(6) 2018 (2ND) BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. IT WAS REPORTED THAT THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: "TICM120V4" SHOULD BE CORRECTED TO "VTICM12.0" IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED DRY WITH RESIDUE IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTICS TORN, TORN PIECES ARE MISSING. (B)(4). Manufacturer Narrative: CORRECTED DATA: "VTICM12.0" SHOULD BE CORRECTED TO "TICM120V4" IN THE PREVIOUS MDR. (B)(4). Manufacturer Narrative: CORRECTED DATA: "VTICM12.0" SHOULD BE CORRECTED TO "TICM120V4" IN THE PREVIOUS MDR. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/+1.5/90 | OD |  | 05/09/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9257751&pc=QCB | 2023826-2019-01962 | 2019-09-15 04:00 | Injury | STAAR SURGICAL COMPANY AG | 2019-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.0MM TICM120V4 IMPLANTABLE COLLAMER LENS, -13.50/+4.00/110 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2009. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL/LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TEAR AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN (B)(4) BUT NOT MARKETED IN THE U.S. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.50/+4.00/110 | OD |  | 10/01/2019 | 2019 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9372989&pc=QCB | 2023826-2019-02246 | 2017-04-12 04:00 | Injury | STAAR SURGICAL COMPANY AG | 2019-11-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.5MM TICM125V4 IMPLANTABLE COLLAMER LENS, -17.50/+2.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2012. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS EXPLANTED THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: D2 - CORRECTED TO PHAKIC INTRAOCULAR LENS, PRODUCT CODE: QCB. D10 - 'NO' REPORTED IN SUPPLEMENTAL MDR#1 IS NOT APPLICABLE. CLAIM# (B)(4). Manufacturer Narrative: DATE OF EVENT: (B)(6) 2017. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(6) BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -17.50/+2.5/094 | OD |  | 11/06/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10214964&pc=QCB | 2023826-2020-01350 | 2020-05-05 | Injury | STAAR SURGICAL COMPANY AG | 2020-06-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Code Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+1.5/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2013. THE REPORTER INDICATED THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN (B)(6) BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.00/+1.5/100 | OD | Physician | 06/04/2020 | 2020 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10310437&pc=QCB | 2023826-2020-01509 | 2020-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+5.5/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B1- ADVERSE EVENT. B2- REQUIRES INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF -14.00/2.5/084 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE SURGEON NOTES EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF ICA. ON (B)(6) 2019 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. D6B- EXPLANT DATE (B)(6) 2019. H6-CLINICAL CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.00/+5.5/087 | OS |  | 06/29/2020 | 2020 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10393030&pc=QCB | 2023826-2020-01668 | 2020-06-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2020-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Contamination /Decontamination Problem | No Code Available; Appropriate Clinical Signs, Symptoms, Conditions Term / Code Not Available | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.00/+1.5/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO FILIFORM FOREIGN MATTER WAS NOTED ATTACHED TO THE LENS. THERE WAS NO PATIENT IMPACT. THE LENS WAS EXCHANGED FOR ANOTHER SAME LENGTH/MODEL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. HEALTH EFFECT CLINICAL CODE: 4581 - FILIFORM MATTER ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -08.00/+1.5/080 | OD |  | 07/22/2020 | 2020 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8323177&pc=QCB | 2023826-2019-00201 | 2018-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available; No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.5/+2.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2018 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.5/+2.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION AND THE PROBLEM DID NOT RESOLVE. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. A SHORT AND SMALL CILIARY PROCESS WITH SHALLOW CILIARY SULCUS AND POSTERIORLY POSITIONED CILIARY BODY WAS FOUND BY ULTRASOUND BIOMICROSCOPY. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. COMMON DEVICE NAME SHOULD BE CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS IN THE ORIGINAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+2.0/091 | OS |  | 01/10/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9869647&pc=QCB | 2023826-2020-00635 | 2020-02-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2020-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Code Available; No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12/+3/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2015. LOW VAULT WITH ROTAION WAS OBSERVED. REPOSITIONING OF THE LENS WAS PERFORMED SEVERAL TIME BUT IT DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED, SURGEON PLANS TO EXCHANGE LENS. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SAME SIZE DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. D6B: EXPLANT LENS (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12/+3/108 | OD |  | 02/24/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9924099&pc=QCB | 2023826-2020-00843 | 2020-02-20 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-04-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available; No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.0/0.5/035 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - DEVICE CODE 1494 IS NOT APPLICABLE IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: B1: UPDATED TO MALFUNCTION. B2: NO LONGER APPLICABLE. B5: UPDATED TO: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.0/0.5/035 (SPHERE / CYLINDER / AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. EXCESSIVE VAULT WAS REPORTED. THE LENS REMAINS IMPLANTED WITH PLANNED EXCHANGE. H1: UPDATED TO MALFUNCTION. H6: PATIENT CODE 3191 NO LONGER APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.0/0.5/035 | OS |  | 03/10/2020 | 2020 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9941037&pc=QCB | 2023826-2020-00869 |  | Injury | STAAR SURGICAL COMPANY | 2020-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Code Available; No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -8.50/6.0/4.0 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO PATIENTS LEFT EYE (OS) ON (B)(6) 2020. EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS OBSERVED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.50/6.0/4.0 | OS |  | 03/13/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=10451108&pc=QCB | 2023826-2020-01857 | 2020-08-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Code Available; No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -10.50/0.5/125 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS DURING INITIAL ON (B)(6) 2020 SURGERY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#:(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/0.5/125 | OD |  | 08/05/2020 | 2020 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8307843&pc=QCB | 2023826-2019-00225 | 2018-12-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-02-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+4.0/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY ON (B)(6) 2018. REPORTEDLY, THERE WAS NO PATIENT INJURY. ON (B)(6) 2018 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. Manufacturer Narrative: CORRECTED DATA: DESCRIBE EVENT OR PROBLEM - "-15.0/+4.0/085" SHOULD BE CORRECTED TO "-15.50/+4.0/085" IN ORIGINAL MDR. COMMON DEVICE NAME SHOULD BE CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS IN ORIGINAL MDR. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND BENT. THERE WAS RESIDUE ON LENS SURFACE. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.0/+4.0/085 | OS |  | 01/10/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8409312&pc=QCB | 2023826-2019-00383 | 2019-02-15 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Improper or Incorrect Procedure or Method; Ejection Problem | No Consequences Or Impact To Patient | P030016 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS ATTEMPTING TO IMPLANT A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -8.5/+3.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE. THERE WAS PATIENT CONTACT WITH THE LENS, WHICH WAS REMOVED INTRAOPERATIVELY ON (B)(6) 2019. LENS WAS NOT FULLY IMPLANTED. REPORTEDLY, THE SURGEON WAS NOT ADVANCING THE ICL TO THE TIP OF THE CARTRIDGE. WHEN HE PUSHED THE PLUNGER DOWN THE BARREL, IT OVERRODE THE ICL AND TRAPPED THE TRAILING HAPTICS IN THE TIP OF THE CARTRIDGE, TAKING A "BITE" OUT OF THE ICL. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED ON (B)(6) 2019. THERE WAS NO PATIENT INJURY. REFERENCE MFR REPORT# 2023826-2019-00379 FOR PRIMARY LENS. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MFR: THE LENS WAS RETURNED IN LIQUID, IN THE LENS VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.5/+3.0/092 | OD |  | 02/15/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8409335&pc=QCB | 2023826-2019-00379 | 2019-02-15 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Improper or Incorrect Procedure or Method; Ejection Problem | No Consequences Or Impact To Patient | P030016 | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON WAS ATTEMPTING TO IMPLANT A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -8.5/+3.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE. THERE WAS PATIENT CONTACT WITH THE LENS, WHICH WAS REMOVED INTRAOPERATIVELY ON (B)(6) 2019. LENS WAS NOT FULLY IMPLANTED. REPORTEDLY, THE SURGEON WAS NOT ADVANCING THE ICL TO THE TIP OF THE CARTRIDGE. WHEN HE PUSHED THE PLUNGER DOWN THE BARREL, IT OVERRODE THE ICL AND TRAPPED THE TRAILING HAPTICS IN THE TIP OF THE CARTRIDGE, TAKING A "BITE" OUT OF THE ICL. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED ON (B)(6) 2019. THERE WAS NO PATIENT INJURY. REFERENCE MFR REPORT# 2023826-2019-00383 FOR BACK-UP LENS. Manufacturer Narrative: ORIGINAL MDR SHOULD BE CORRECTED TO "VISUAL INSPECTION FOUND THE HAPTIC TORN." CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MFR: THE LENS WAS RETURNED IN LIQUID, IN THE LENS VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN. LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.5/+3.0/090 | OD |  | 02/15/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8414861&pc=QCB | 2023826-2019-00412 | 2019-02-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.5/+1.5/118 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TEARING DURING LOADING. THERE WAS NO PATIENT INJURY. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. COMMON DEVICE NAME SHOULD BE CORRECTED TO PHAKIC TORIC INTRAOCULAR LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.5/+1.5/118 | NA |  | 02/12/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8460886&pc=QCB | 2023826-2019-00460 | 2019-02-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-03-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2019 A VTICM5_12.1, -6.50/15.088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO THE PATIENT'S LEFT EYE (OS). THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY, THE LENS WAS REMOVED DURING THE SAME SURGERY. THE EXCHANGE LENS WAS IMPLANTED ON (B)(6) 2019. REPORTER STATES "SECOND LENS PROCEDURE WENT OFF WELL AND PATIENT IS HAPPY". PROBLEM RESOLVED. Manufacturer Narrative: EVENT DESCRIPTION: (-6.50/15.088) SHOULD BE CHANGED TO: -6.5/1.5/088 IN INITIAL MDR. RESULT CODE 114 SHOULD BE ADDED TO SUPPLEMENTAL MDR #1. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO CENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND FOREIGN RESIDUE ON THE LENS AND THE LENS WAS IN TWO PIECES. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.50/15.088 | OS |  | 03/12/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8487887&pc=QCB | 2023826-2019-00536 | 2019-01-23 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -13.50/2.0/047 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED WITTHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT WITH NO PATIENT INJURY. A REPLACEMENT LENS WAS IMPLANTED ON THE SAME DAY IN A SECOND SURGERY AND IT WAS REPORTED THAT THE PROBLEM RESOLVED. IT WAS LATER REPORTED THAT THE PATIENT RETURNED FOR A SECOND SURGERY ON (B)(6) 2019. ADDITIONAL INFORMATION/CLARIFICATION HAS BEEN REQUESTED FOR THIS CASE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A LENS CASE. VISUAL INSPECTION FOUND THAT A HAPTIC WAS TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.50/2.0/047 | OD |  | 03/06/2019 | 2019 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8497336&pc=QCB | 2023826-2019-00581 | 2019-02-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-04-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2019 A VTICM5_12.6, -13.00/1.0/173 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO THE EYE. THERE WAS NO PATIENT CONTACT. REPORTER INDICATED "THE LENS WAS NOTED TO BE DAMAGED AT TIME OF LENS EJECTION. SURGERY ABORTED, PATIENT WILL RETURN FOR ICL INSERTION." CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: DEVICE EVALUATION- THE LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND LENS HAPTIC TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.00/1.0/173 | NA |  | 02/25/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8571468&pc=QCB | 2023826-2019-00731 | 2019-02-28 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -15.50/3.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE WHEN IT WAS INSERTED. THERE WAS PATIENT CONTACT BUT NO INJURY. THE LENS WAS REMOVED AND REPLACED WITH A BACKUP LENS DURING INITIAL SURGERY. THIS REPLACEMENT RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS DUE TO USER ERROR/LOADING ISSUE. Manufacturer Narrative: LENS WAS RETURNED WRAPPED IN A GAUZE. VISUAL INSPECTION FOUND THAT A HAPTIC WAS TORN. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.50/3.0/085 | NA | Other Health Care Professional | 04/03/2019 | 2019 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8571534&pc=QCB | 2023826-2019-00704 | 2019-03-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.00/+1.50/090 (SPHERE/CYLINDER/AXIS) DURING THE SAME SURGERY DUE TO LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS DUE TO USER ERROR AND THE DEVICE. Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN A VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.00/+1.50/090 | NA |  | 03/11/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8609182&pc=QCB | 2023826-2019-00830 | 2018-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-05-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY ON (B)(6) 2018. WITHIN THE SAME SURGERY AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. REPORTEDLY, THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH CLEAR SURGICAL RESIDUE ON LENS IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND DEFORMED WITH FOREIGN RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/+1.5/090 | OS |  | 04/17/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8611922&pc=QCB | 2023826-2019-00823 | 2019-03-25 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2019 A HAPTIC OF A 13.2MM, TMICL13.2, -10.00/2.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMAR LENS TORE "WHILE REMOVING FROM VIAL OR WHILE LOADING, WITH NO PATIENT CONTACT AND NO PATIENT INJURY. BACK UP LENS WAS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). | -10.00/2.0/087 | NA | Administrator/Supervisor | 04/16/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8657933&pc=QCB | 2023826-2019-00976 | 2019-03-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-05-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -10/2.0/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING INJECTION/DELIVERY INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. A SPARE LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED STUCK IN A CARTRIDGE IN A SPECIMEN CUP. VISUAL INSPECTION FOUND: LENS STUCK IN CARTRIDGE, PLUNGER OVERRIDDEN AND FOREIGN MATERIAL ON CARTRIDGE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10/2.0/103 | OD |  | 05/03/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8671660&pc=QCB | 2023826-2019-01017 | 2019-05-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-06-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+5.5/078 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY ON (B)(6) 2019. REPORTEDLY, THERE WAS NO PATIENT INJURY. ON (B)(6) 2019 AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -14.5/+5.5/078 | OS |  | 05/09/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8675196&pc=QCB | 2023826-2019-01007 | 2019-05-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-06-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.00/2.0/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME INITIAL SURGERY AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS RETURNED DRY IN A MICRO CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THAT A HAPTIC WAS TORN. (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.00/2.0/90 | OD |  | 05/09/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8676878&pc=QCB | 2023826-2019-01037 | 2018-09-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-06-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event; Patient Device Interaction Problem | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/098 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE PATIENT WAS UNCOOPERATIVE DURING THE OPERATION AND ASKED TO STOP THE SURGERY. SO THE SURGEON REMOVED THE ICL LENS. THE SURGEON DOES NOT PLAN TO IMPLANT ANOTHER ICL LENS. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS PATIENT RELATED FACTOR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH RESIDUE ON THE LENS. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.0/+1.5/098 | OD |  | 05/16/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8692880&pc=QCB | 2023826-2019-01065 | 2019-03-09 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-06-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Unintended Movement | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT, WHILE EXPLANTING A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/072 (SPHERE/CYLINDER/AXIS) FROM THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. REPORTEDLY, THE PATIENT DID NOT COOPERATE WELL. DURING IMPLANT, THE PATIENT'S EYE MOVEMENTS CAUSED THE LENS TO TURN OVER AND THEN WHEN THE SURGEON REMOVED THE LENS, IT BECAME DAMAGED. THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY ON (B)(6) 2019. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AT A LATER DATE. REPORTEDLY, THERE WAS NO PATIENT INJURY. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.0/+2.0/072 | OS |  | 05/16/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8709756&pc=QCB | 2023826-2019-01092 | 2019-04-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-06-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICTED THAT A 13.2MM VTICM5_13.2, -10.5/2.5/093 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS TORE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. THE LENS WAS REMOVED AND A BACKUP LENS WAS IMPLANTED DURING INITIAL SURGERY. Manufacturer Narrative: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/2.5/093 | NA |  | 05/20/2019 | 2019 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8795445&pc=QCB | 2023826-2019-01239 | 2019-06-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6 MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.0/+3.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), IT TORE. REPORTEDLY, THE LENS WAS INSERTED HALF-WAY INTO THE EYE WHEN THE TEAR WAS NOTED. IT WAS REMOVED IMMEDIATELY AND AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED INTRAOPERATIVELY. THIS OCCURRED ON (B)(6) 2019. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE AND USER ERROR. NO PATIENT INJURY OCCURRED. Manufacturer Narrative: ADDITIONAL DETAILS: THE LENS GOT STUCK IN THE CARTRIDGE WHEN THE SURGEON WAS ATTEMPTING TO IMPLANT IT. AS THE LENS WAS GOING INTO THE EYE, THE SURGEON NOTICED THAT THE LENS HAD TORN. DEVICE EVALUATION: THE LENS WAS RETURNED STUCK IN THE CARTRIDGE. THERE WAS CLEAR SURGICAL RESIDUE ON THE CARTRIDGE. VISUAL INSPECTION FOUND THE PLUNGER WAS OVERRIDDEN. (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNKNOWN. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -9.0/+3.0/094 | OS |  | 06/20/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8795475&pc=QCB | 2023826-2019-01248 | 2019-05-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+3.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY ON (B)(6) 2019. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED ON (B)(6) 2019. REPORTEDLY, THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC HAD A TEAR AND WAS BENT. THERE WAS FOREIGN RESIDUE ON THE LENS AS WELL. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+3.0/087 | OS |  | 06/28/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8809456&pc=QCB | 2023826-2019-01254 | 2019-04-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.5/+2.0/105 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND ONE LENS HAPTIC TORN. (B)(4). Manufacturer Narrative: EVENT DESCRIPTION: THE LENS TORE/BROKE DURING DELIVERY INTO THE PATIENT'S RIGHT EYE (OD). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -17.5/+2.0/105 | OD |  | 06/24/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8817753&pc=QCB | 2023826-2019-01229 | 2018-11-06 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2 MM, VTICMO13.2, -16.0/5.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS BROKE/TORE DURING INJECTION/DELIVERY INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, THERE WAS NO PATIENT INJURY. ON (B)(6) 2018 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. H6 - DEVICE CODE 1494: OFF LABEL USE: ACD <3.0MM; PATIENT WAS GREATER THAN 45 YEARS OLD AT THE TIME OF IMPLANTATION. (B)(4). Manufacturer Narrative: DATE OF REPORT: CORRECTED AWARE DATE. DATE RECEIVED BY MANUFACTURER SHOULD BE REPORTED AS 24-JUN-2019 IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.0/5.5/092 | OS |  | 06/24/2019 | 2019 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8841430&pc=QCB | 2023826-2019-01364 | 2018-12-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-07-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -6.00/2.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) BUT NO INJURY, ANOTHER LENS WAS IMPLANTED. THE REPORTER STATED THAT THE PATIENT IS IN GOOD CONDITION. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISUTRE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: LENS TORE DURING DELIVERY INTO THE EYE SHOULD BE CORRECTED TO LENS WAS DAMAGED WHEN REMOVING FROM THE VIAL. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.00/2.5/090 | NA |  | 07/02/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8860003&pc=QCB | 2023826-2019-01249 | 2019-06-18 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2019 A VTICM5_13.2, -8.50/2.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO THE PATIENTS LEFT EYE (OS). THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY, THE LENS WAS REMOVED DURING INITIAL SURGERY. A NEW LENS WAS IMPLANTED ON (B)(6) 2019, RESOLVING THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS "HAPTIC BROKEN - MANUFACTURE FAULT." PER REPORTER, STATUS OF PATIENT'S EYE "ALL GOOD." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.50/2.5/080 | OS | Physician | 07/11/2019 | 2019 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8870924&pc=QCB | 2023826-2019-01408 | 2018-10-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.5/+5.0/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE. THIS OCCURRED ON (B)(6) 2018. THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED ON A LATER DATE AND THE PROBLEM WAS RESOLVED. NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.5/+5.0/103 | OD |  | 07/16/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8876518&pc=QCB | 2023826-2019-01391 | 2019-07-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, - 9.00/+2.0/094 (SPHERE/CYLINDER/AXIS), AND THE LENS HAPTIC TORE. THERE WAS NO PATIENT INJURY AND THE BACKUP LENS WAS USED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: THE REPORTER INDICATED THE LENS WAS INSERTED AND REMOVED DURING THE SAME SURGERY DUE TO THE LENS WAS NOTED TO BE TORN. THERE WAS NO PATIENT INJURY. THE BACKUP LENS WAS INSERTED AND THE PROBLEM WAS RESOLVED. THE PAITENT IS DOING WELL. THE CAUSE OF THE EVENT WAS UNKNOWN. (B)(4). Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | - 9.00/+2.0/094 | NA |  | 07/12/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8899505&pc=QCB | 2023826-2019-01459 | 2019-07-10 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2019 A 13.2MM, VTICM5_13.2, -12.5/0.5/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJETION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED, PATIENT IS HAPPY. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC AND THE OPTIC TORN AND THE LENS TORN IN TWO PIECES. CLAIM#: (B)(4). Manufacturer Narrative: DATE OF REPORT: TYPO IN INITIAL MDR. CORRECTED FROM 17-JUL-2109 TO 17-JUL-2019. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.5/0.5/071 | NA |  | 07/17/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8908404&pc=QCB | 2023826-2019-01483 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-08-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -3.0/+1.0/065 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS REMOVED DUE TO LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. A DIFFERENT MODEL BUT SAME LENGTH LENS WAS IMPLANTED AS A REPLACEMENT LENS ON (B)(6) 2019, AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: 'THE LENS WAS REMOVED' CORRECTED TO 'THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY' IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON LENS. VISUAL INSPECTION FOUND PIECES OF THE LENS OPTIC AND HAPTICS TORN OFF AND STUCK TOGETHER. UNABLE TO EVALUATE THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -3.0/+1.0/065 | OS |  | 07/26/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8912822&pc=QCB | 2023826-2019-01508 | 2019-05-10 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/077 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY ON (B)(6) 2019. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED ON (B)(6) 2019. REPORTEDLY, THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE SURGEON STATES THAT THE FIRST SURGERY WAS SUSPENDED BECAUSE ONE OF THE PROXIMAL HAPTICS WAS TORN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS CLEAR SURGICAL RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT BOTH THE HAPTIC AND OPTIC WAS TORN. THE TORN PIECE WAS NOT RETURNED. (B)(4). Manufacturer Narrative: ALTERNATE LENS IMPLANT DATE SHOULD BE CORRECTED TO (B)(6) 2019 IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.0/+2.0/077 | OS |  | 07/26/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8918801&pc=QCB | 2023826-2019-01489 | 2019-02-25 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/068 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. ANOTHER SAME MODEL LENS WAS IMPLANTED DURING THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+1.5/068 | NA |  | 07/22/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8918829&pc=QCB | 2023826-2019-01496 | 2019-07-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.0/089 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. ANOTHER SAME MODEL BUT DIFFERENT LENGTH LENS WAS IMPLANTED DURING THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN. UDI CORRECTED TO N/A IN INITIAL MDR. METHOD CODE 4111 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.5/+2.0/089 | NA |  | 07/23/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8918831&pc=QCB | 2023826-2019-01494 | 2019-07-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.00/+4.50/089 (SPHERE/CYLINDER/AXIS) WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. ANOTHER LENS WAS NOT IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A REPLACEMENT LENS (VTICM5_12.6) ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICRO-CENTRIFUGE VIAL, WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND PART OF THE LENS OPTIC AND HAPTIC TORN OFF. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.00/+4.50/089 | NA |  | 07/24/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8930936&pc=QCB | 2023826-2019-01465 | 2019-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2019 A VTICH12.6, -7.50/3.5/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO THE PATIENT'S LEFT EYE (OS). THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED AND EXCHANGED ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.50/3.5/102 | OS |  | 08/05/2019 | 2019 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8937589&pc=QCB | 2023826-2019-01554 | 2019-07-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-08-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/099 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THE SURGEON WHILE IMPLANTING A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -11.0/1.5/099 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT RIGHT EYE (OD) THE LENS TORE/BROKE DURING DELIVERY INTO THE EYE ON (B)(6) 2019. THE LENS WAS REMOVED WITHIN THE SAME SURGERY. THERE WAS NO PATIENT INJURY. ON (B)(6) 2019 ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED, AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. ADDITIONAL INFORMATION: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND OPTIC AND HAPTIC TORN. LENS RETURNED IN TWO PIECES. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+1.5/099 | OD |  | 07/31/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=8966517&pc=QCB | 2023826-2019-01642 | 2018-08-31 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/100 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE. THIS OCCURRED ON (B)(6) 2018. THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED ON A LATER DATE AND THE PROBLEM WAS RESOLVED. NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.0/+1.5/100 | OD |  | 08/08/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9013815&pc=QCB | 2023826-2019-01617 | 2019-07-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER STATED THAT ON (B)(6) 2019 THE SURGEON ATTEMPTED TO IMPLANT 13.2MM, VTICM5_13.2, -8.5/2.0/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS AS A REPLACEMENT LENS TO CORRECT LOW VAULT BUT THE LENS GOT DAMAGE WHILE INJECTING INTO THE RIGHT EYE (OD). THERE WAS PATIENT CONTACT WITH NO INJURY. ON (B)(6) 2019 A LENS OF THE SAME LENGTH WAS IMPLANTED SUCCESSFULLY AND THE ISSUE WAS RESOLVED. PATIENT IS HAPPY. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN WITH RESIDUE ON THE LENS. CORRECTED DATA: INITIAL MDR SHOULD ALSO HAVE STATED: 'THE LENS WAS REMOVED WITHIN THE SAME SURGERY.' INJECTOR SYSTEM MODEL: FTP, LOT# UNK SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. DEVICE CODE 1494 SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.5/2.0/108 | OD |  | 08/20/2019 | 2019 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9045309&pc=QCB | 2023826-2019-01668 | 2019-08-09 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -8.5/1.5/067 (SPHER/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WENT INTO THE PATIENTS EYE UPSIDE DOWN. THE LENS WAS REMOVED AND A BACKUP LENS WAS IMPLANTED. NO PATIENT INJURY WAS REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.5/1.5/067 | NA | Other Health Care Professional | 08/16/2019 | 2019 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9050323&pc=QCB | 2023826-2019-01687 | 2019-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-09-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, - 18.00/+1.0/044 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THE REPORTER INDICATED THE FOAM TIP PLUNGER DETACHED FROM THE INJECTOR, OVERRODE AND TORE THE LENS DURING DELIVERY. THERE WAS NO PATIENT INJURY OR ANY PRODUCT MALFUNCTION. THE SURGEON DID NOT IMPLANT ANOTHER LENS. Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | - 18.00/+1.0/044 | OS |  | 08/22/2019 | 2019 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9100711&pc=QCB | 2023826-2019-01682 | 2019-08-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, - 11.50/+2.0/072 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE PATIENTS EYE MOVED AND THE SURGEON REMOVED THE LENS. THERE WAS NO PATIENT INJURY AND THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: EXPLANTED DATE: (B)(6) 2019. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANTED: UNK. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. (B)(4). | - 11.50/+2.0/072 | OD |  | 08/23/2019 | 2019 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9111483&pc=QCB | 2023826-2019-01776 | 2019-07-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -16.0/2.5/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. PATIENT CONTACT WITH NO PATIENT INJURY WAS REPORTED. ON (B)(6) 2019 A REPLACEMENT LENS WAS IMPLANTED AND IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A SMALL PIECE OF THE LENS RETURNED. CLAIM# : (B)(4). Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE: AGE <21 YEARS OLD. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.0/2.5/098 | OD |  | 08/25/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9125514&pc=QCB | 2023826-2019-01770 | 2018-07-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-09-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2018 A VTICMO12.6, -6.50/1.0/130 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, ON (B)(6) 2018. A REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2018. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTIUON FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.50/1.0/130 | NA |  | 09/10/2019 | 2019 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9198677&pc=QCB | 2023826-2019-01941 | 2019-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2019-10-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -14.0/1.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. NO PATIENT INJURY WAS REPORTED. ON (B)(6) 2019 A REPLACEMENT LENS WAS IMPLANTED AND IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.0/1.5/086 | OS |  | 09/19/2019 | 2019 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9205136&pc=QCB | 2023826-2019-01936 | 2019-08-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2019 A VTICMO13.2, -8.00/1.5/168 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO THE PATIENT'S LEFT EYE (OS). THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED DURING INITIAL SURGERY ON (B)(6) 2019. A REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2019. THIS RESOLVED PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH CLEAR SURGICAL RESIDUE ON THE LENS. VISUAL INSPECTION FOUND LENS OPTIC AND HAPTIC TORN AND RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.00/1.5/168 | OS |  | 10/03/2019 | 2019 | 47
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9220593&pc=QCB | 2023826-2019-01989 | 2019-08-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-10-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.0/2/82, IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED WITHIN THE SAME SURGERY, NO PATIENT INJURY WAS REPORTED. ON (B)(6) 2019 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MFR: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN AS WELL AS RESIDUE ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.0/2/82 | OS |  | 09/26/2019 | 2019 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9252680&pc=QCB | 2023826-2019-02008 | 2019-10-03 | Malfunction | STAAR SURGICAL COMPANY | 2019-10-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, - 6.50/+2.0/088 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY, WITH NO PATIENT INJURY. THE BACKUP LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE SURGEON DID NOT GIVE A REASON FOR REMOVING THE LENS AND THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS WAS DISCARDED. Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | - 6.50/+2.0/088 | NA | Other Health Care Professional | 10/04/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9252669&pc=QCB | 2023826-2019-02005 | 2019-09-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | P030016 | Event Description: THE REPORTER INDICATED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, - 10.50/+4.0/087 (SPHERE/CYLINDER/AXIS), WAS DAMAGED/TORN DURING THE FOLDING OF THE LENS AND WHEN THE LENS WAS BEING PULLED THROUGH THE INJECTOR. THERE WAS PATIENT CONTACT BUT NO INJURY. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEMS. THE CAUSE OF THE EVENT WAS DUE TO A LOADING ERROR. Manufacturer Narrative: DATE AVAILABLE FOR EVALUATION: CORRECTED TO 30-OCT-2019. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A LENS VIAL. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | - 10.50/+4.0/087 | NA |  | 09/30/2019 | 2019 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9335812&pc=QCB | 2023826-2019-02148 | 2019-10-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -12.00/1.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LENS TORE DURING INJECTION INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. PER REPORTER, "THE LENS DID NOT FULLY EJECT FROM CARTRIDGE. NOT SURE IF PLUNGER AND/OR CARTRIDGE CAUSED LENS TO SHRED DURING INJECTION OF BACK HALF." Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A ZIPLOC BAG, IN A VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN, AND THE LENS WAS RETURNED IN THREE PIECES. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/1.0/088 | OD |  | 10/29/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9336524&pc=QCB | 2023826-2019-02178 | 2019-07-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO_12.6, -6.5/1.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE AND THE PROBLEM RESOLVED. THERE WAS NO PATIENT INJURY. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6- DEVICE CODE 1494: OFF-LABEL USE; AGE 21 YEARS OLD SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.5/1.5/087 | OD |  | 10/21/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9350267&pc=QCB | 2023826-2019-02177 | 2019-07-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.6/2.5/068 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. ON (B)(6) 2019 A REPLACEMENT LENS WAS IMPLANTED AND IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED. CAUSE OF EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3-DEVICE EVALUATION: RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS AND CLEAR SURGICAL RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND RESIDUE ON LENS. (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.6/2.5/068 | OD |  | 10/21/2019 | 2019 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9385346&pc=QCB | 2023826-2019-02270 | 2019-08-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-11-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 13.2MM VTICMO 13.2, -7.00/3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. SURGEON "NOTICED THAT A PART OF THE LENS" CAME OUT INTO PATIENT'S EYE AND THE REST OF THE LENS REMAINED IN THE CARTRIDGE. LENS WAS REMOVED, EYE WASHED WITH NO FURTHER COMPLICATIONS. NO PATIENT INJURY. A REPLACEMENT LENS WAS IMPLANTED ON 02/OCT/2019 AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS BLACK RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: AGE: UNK. WEIGHT: UNK. ETHNICITY:UNK. RACE UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.00/3.0/093 | OD |  | 08/29/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9412445&pc=QCB | 2023826-2019-02281 | 2019-08-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.5/+1.0/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY ON (B)(6) 2019. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED ON (B)(6) 2019. REPORTEDLY, THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.5/+1.0/072 | OS |  | 11/11/2019 | 2019 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9440834&pc=QCB | 2023826-2019-02331 | 2019-07-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.5/2.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. ON (B)(6) 2019 A REPLACEMENT LENS WAS IMPLANTED AND IT WAS REPORTED THAT THE PROBLEM WAS RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE; ACD< 3.00MM AND AGE < 21 YEARS OLD AT DATE OF IMPLANTATION. CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION : LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC AND ONE OF THE HAPTICS TORN. CLAIM # (B)(4). Manufacturer Narrative: THE PRODUCT IS MANUFACTURED IN THE US, BUT NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/2.5/094 | OD |  | 11/11/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9460114&pc=QCB | 2023826-2019-02364 | 2019-07-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -9.00/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS REMOVED AND REPLACED INTRAOPERATIVELY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE. Manufacturer Narrative: CORRECTED DATA: B5-THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 -9.00/1.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE DURING THE LOADING PROCESS WITH PATIENT CONTACT AND NO PATIENT INJURY. THE LENS WAS NOT IMPLANTED. A REPLACEMENT LENS WAS IMPLANTED AND REPORTER STATES "EVERYTHING IS GOOD NOW." CAUSE OF EVENT IS REPORTED AS DEVICE. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3- DEVICE EVALUATION- LENS RETURNED IN LIQUID IN LENS VIAL WITH CLEAR SURGICAL DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN WITH DEBRIS ON LENS. CLAIM# : (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.00/1.0/090 | OS |  | 11/21/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9476272&pc=QCB | 2023826-2019-02393 | 2019-11-21 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Material Split, Cut or Torn | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), IT TORE. IT WAS REMOVED INTRAOPERATIVELY AND THERE WAS NO PATIENT INJURY. THIS OCCURRED ON (B)(6) 2019. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR AND LOADING ERROR; HOWEVER THE SURGEON STATES THAT THE LENS WAS "DELIVERED WITH THE HOLE." Manufacturer Narrative: B5- PREVIOUSLY STATED THIS OCCURRED ON (B)(6) 2019. CORRECTED INFO: THIS OCCURRED ON (B)(6) 2019. ADDITIONAL INFO: THE LENS WAS IMPLANTED, REMOVED, AND REPLACED ALL WITHIN THE SAME SURGERY. THE PROBLEM IS RESOLVED. H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THAT THE OPTIC WAS TORN. H6-METHOD CODE 3331, DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: B5-UPDATED INFORMATION: A BACKUP LENS WAS IMPLANTED. REPORTEDLY, "EVERYTHING IS O.K." CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/+2.0/083 | OS |  | 11/26/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9477701&pc=QCB | 2023826-2019-02416 | 2019-07-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2019-12-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -11.5/0.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN. CORRECTED DATA: H6 - PATIENT CODE 1494: OFF-LABEL USE (AGE < 21 YEARS OLD) SHOULD HAVE BEEN INCLUDED IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/0.5/092 | OS |  | 11/18/2019 | 2019 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9492306&pc=QCB | 2023826-2019-02398 |  | Malfunction | STAAR SURGICAL COMPANY | 2019-12-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Difficult to Fold, Unfold or Collapse | No Consequences Or Impact To Patient | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, - 12.50/+2.0/106 (SPHERE/CYLINDER/AXIS), FROM THE PATIENTS RIGHT EYE (OD), WITHIN THE SAME SURGERY DUE TO THE LENS UNFOLDED INCORRECTLY IN THE EYE. THERE WAS NO PATIENT INJURY AND THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM, THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING WELL. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THE LENS WAS RETURNED IN LIQUID, IN LENS VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | - 12.50/+2.0/106 | OD | Nurse | 11/19/2019 | 2019
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9559340&pc=QCB | 2023826-2020-00040 | 2019-12-09 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Improper or Incorrect Procedure or Method; No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2019. THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED ON THE SAME DATE IN A SEPARATE SURGERY AND THE PROBLEM WAS RESOLVED. NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/+2.0/091 | OD |  | 12/15/2019 | 2019 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9565680&pc=QCB | 2023826-2020-00009 | 2019-07-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | N/A | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2019 A VTICMO13.2, -13.50/1.5/134 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING INJECTION INTO PATIENT'S RIGHT EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REMOVED INTRAOPERATIVELY. A REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2019, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - RESULT CODE: 114 SHOULD HAVE BEEN INCLUDED IN SUPPLEMENTAL MEDWATCH REPORT #1. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.50/1.5/134 | OD |  | 12/11/2019 | 2019 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=9583807&pc=QCB | 2023826-2020-00097 | 2019-07-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2020-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Consequences Or Impact To Patient | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+2.5/083 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE/BROKE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS WAS IMPLANTED ON (B)(6) 2019 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN/ BROKEN/ BENT/ DEFORMED. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.0/+2.5/083 | OS |  | 12/18/2019 | 2019 | 20