https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788214&pc=QCB | 2023826-2021-01343 | 2021-03-15 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.5/178 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LENS ROTATION. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -11.0/+2.5/178 | OS | Physician | 04/09/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788216&pc=QCB | 2023826-2021-01367 | 2021-04-12 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-06 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0/+0.5/142 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2021. THE LENS WAS REMOVED INTRA-OPERATIVELY. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AT A LATER DATE. NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -16.0/+0.5/142 | OD | Physician | 04/13/2021 | 2021 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788220&pc=QCB | 2023826-2021-01352 | 2021-02-10 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -12.0/+2.0/090 | OS | Physician | 04/12/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788221&pc=QCB | 2023826-2021-01359 | 2021-02-06 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-06 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS GOT STUCK IN THE CARTRIDGE OR INJECTION SYSTEM. THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY ON (B)(4) 2021. AN ALTERNATE LENS WAS IMPLANTED AT A LATER DATE AND THE PROBLEM WAS RESOLVED. REPORTEDLY, THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.5/+1.5/088 | OS | Physician | 04/13/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788222&pc=QCB | 2023826-2021-01351 | 2021-01-04 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -9.0/+1.0/092 | OD | Physician | 04/12/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788223&pc=QCB | 2023826-2021-01340 | 2020-12-13 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+1.5/068 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT MODEL LENS DUE TO BLURRED VISION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -15.0/+1.5/068 | OS | Physician | 04/08/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788225&pc=QCB | 2023826-2021-01354 | 2021-03-08 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -12.5/+1.5/086 | OD | Physician | 04/12/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788226&pc=QCB | 2023826-2021-01353 | 2021-02-19 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -17.5/+1.5/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -17.5/+1.5/083 | OD | Physician | 04/12/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788232&pc=QCB | 2023826-2021-01337 | 2021-04-06 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-06 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/092 (SPHERE/CYLINDER/AXIS) TORE DURING INJECTION INTO THE LEFT (OS) EYE. THERE WAS PATIENT CONTACT WITH THE LENS BUT NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY ON (B)(6) 2021. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.5/+1.5/092 | OS | Physician | 04/07/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788236&pc=QCB | 2023826-2021-01339 | 2020-11-09 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/+2.5/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFTT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -10.5/+2.5/073 | OS | Physician | 04/08/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788239&pc=QCB | 2023826-2021-01338 | 2021-01-30 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ON (B)(6) 2021 A SHORTER LENS WAS IMPLANTED AN THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY THE DOCTOR WANTED TO LET THE PATIENT OVERCORRECT A LITTLE BIT BECAUSE HE IS YOUNG AND THEN EXCHANGED IT WITH A STRONGER POWER LENS. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -8.5/+1.0/093 | OD | Physician | 04/08/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788243&pc=QCB | 2023826-2021-01342 | 2021-03-15 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.5/+2.5/001 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LENS ROTATION. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. G4-PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/+2.5/001 | OD | Physician | 04/09/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788245&pc=QCB | 2023826-2021-01334 | 2020-09-23 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.5/+1.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -16.5/+1.0/094 | OD | Physician | 04/07/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788218&pc=QCB | 2023826-2021-01365 | 2020-12-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.0/+2.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6- INVESTIGATION CODE 4110 IN PREVIOUS MDR IS NOT APPLICABLE. INVESTIGATION FINDINGS CODE- 213 CORRECTED TO 114. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+2.0/095 | OS | Physician | 04/13/2021 | 2021 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788219&pc=QCB | 2023826-2021-01366 | 2021-03-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -8.5/+2.5/057 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON REPORTS LENS ROTATION. REPORTEDLY, LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MMVTICM5 13.7 OF -8.5/2.5/057 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPOSITIONED 17° CCW, AND THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -8.5/+2.5/057 | OS |  | 04/13/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11788227&pc=QCB | 2023826-2021-01344 | 2020-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.0/+3.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT MODEL LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.0/+3.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT MODEL LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.0/+3.0/094 | OD |  | 04/09/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11795195&pc=QCB | 2023826-2021-01413 | 2021-02-24 | Injury | STAAR SURGICAL COMPANY | 2021-05-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: IT WAS REPORTED THAT A 13.7MM, VTICM5_13.7, -8.0/1.0/58 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT RIGHT EYE (OD) AS A REPLACEMENT LENS DUE TO LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THE PROBLEM RESOLVED. "PATIENT IS MYOPIC AGAIN" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). FOR ADDITIONAL INFORMATION THE REPORTED STATED " HE WILL UNDERWENT CORNEAL REFRACTIVE SURGERY WITH RELEX SMILE." SEE MFR # 2023826-2021-00134 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -8.0/1.0/58 | OD | Physician | 02/24/2021 | 2021 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11795198&pc=QCB | 2023826-2021-01412 | 2021-02-24 | Injury | STAAR SURGICAL COMPANY | 2021-05-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: IT WAS REPORTED THAT A 13.7MM, VTICM5_13.7, -8.50/1.0/27 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT LEFT EYE (OS) AS A REPLACEMENT LENS DUE TO LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THE PROBLEM RESOLVED. "PATIENT IS MYOPIC AGAIN" WAS REPORT FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). FOR ADDITIONAL INFORMATION THE REPORTED STATED " HE WILL UNDERWENT CORNEAL REFRACTIVE SURGERY WITH RELEX SMILE." SEE MFR # 2023826-2021-00135 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -8.50/1.0/27 | OS | Physician | 02/24/2021 | 2021 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11806407&pc=QCB | 2023826-2021-01417 | 2021-03-10 | Injury | STAAR SURGICAL COMPANY | 2021-05-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, DATE REC¿D BY MFR: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -12.50/+1.0/091 | OD | Physician | 04/12/2021 | 2021 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11806460&pc=QCB | 2023826-2021-01418 | 2021-04-05 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.50/+3.5/091 (SPHERE/CYLINDER/AXIS) WITHIN THE SAME SURGERY DUE TO THE LENS WAS STUCK IN THE PLUNGER AND TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED ON (B)(6) 2021 WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS QUIET AND THE PATIENT IS HAPPY. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.50/+3.5/091 | OD | Physician | 04/12/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11808980&pc=QCB | 2023826-2021-01423 | 2021-03-30 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2021 A 13.2MM, VTICM5_13.2, -14.00/4.00/97, IMPLANTABLE COLLAMER LENS TORE DURING LOADING IN CARTRIDGE. THERE WAS NO PATIENT CONTACT. A REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2021 AND THE PROBLEM RESOLVED. DOCTOR AND PATIENT ARE HAPPY. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -14.00/4.00/97 | NA | Physician | 04/12/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11809660&pc=QCB | 2023826-2021-01424 | 2021-02-04 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM ,VTICMO13.2, -6.00/1.0/106 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT. PROBLEM WAS RESOLVED. CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.00/1.0/106 | OD | Physician | 04/15/2021 | 2021 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11809741&pc=QCB | 2023826-2021-01426 | 2021-03-01 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM ,VTICMO12.6, -12.00/2.5/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT. PROBLEM WAS RESOLVED. CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/2.5/110 | OD | Physician | 04/15/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11809890&pc=QCB | 2023826-2021-01427 | 2021-02-23 | Injury | STAAR SURGICAL COMPANY | 2021-05-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM ,VTICMO12.6, -18.0/4.0/178 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATED "THE INITIAL LENS WAS DESIGNED IN THE VERTICAL POSITION (12 TO 6 O 'CLOCK), AND THE LARGER LENS WAS DESIGNED IN THE HORIZONTAL POSITION (3 TO 9 O 'CLOCK). IN ORDER TO BE PLACED IN THE DESIGNED POSITION, THE AXIS IS ADJUSTED". CAUSE OF EVENT WAS REPORTED TO BE UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -18.0/4.0/178 | OD | Physician | 04/13/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11809917&pc=QCB | 2023826-2021-01429 | 2021-02-13 | Injury | STAAR SURGICAL COMPANY | 2021-05-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM ,VTICMO12.6, -17.5/4.0/177 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATED "THE INITIAL LENS WAS DESIGNED IN THE VERTICAL POSITION (12 TO 6 O 'CLOCK), AND THE LARGER LENS WAS DESIGNED IN THE HORIZONTAL POSITION (3 TO 9 O 'CLOCK). IN ORDER TO BE PLACED IN THE DESIGNED POSITION, THE AXIS IS ADJUSTED". CAUSE OF EVENT WAS REPORTED TO BE UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -17.5/4.0/177 | OS | Physician | 04/13/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11810267&pc=QCB | 2023826-2021-01428 | 2021-03-02 | Injury | STAAR SURGICAL COMPANY | 2021-05-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM ,VTICMO13.7, -17.5/5.5/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED TO BE UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -17.5/5.5/099 | OD | Physician | 04/13/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11810279&pc=QCB | 2023826-2021-01430 | 2021-03-02 | Injury | STAAR SURGICAL COMPANY | 2021-05-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM ,VTICMO13.7, -17.0/5.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED TO BE UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -17.0/5.0/084 | OS | Physician | 04/13/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11806448&pc=QCB | 2023826-2021-01416 |  | Injury | STAAR SURGICAL COMPANY | 2021-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -07.00/+1.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REPORTED AS HAVING ROTATED AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL MEDWATCH WILL BE SUBMITTED. Manufacturer Narrative: D6B: UNK. B5 - REPORTEDLY THE LENS WAS REMOVED AND REPLACED WITH A LARGER ICL. AFTER SUBSEQUENT PRK THE PATIENT'S VISION IS 20/20 AND IS HAPPY. CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -07.00/+1.5/090 | OS | Physician | 04/20/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11813779&pc=QCB | 2023826-2021-01492 | 2021-01-22 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/099 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. ANOTHER LENS WAS NOT IMPLANTED. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.50/+2.0/099 | OD | Physician | 04/15/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11813801&pc=QCB | 2023826-2021-01493 | 2020-11-11 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/066 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED DURING THE SAME SURGERY WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, DATE REC¿D BY MFR: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -12.00/+2.0/066 | NA | Physician | 04/13/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11815895&pc=QCB | 2023826-2021-01500 | 2021-02-24 | Injury | STAAR SURGICAL COMPANY | 2021-05-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.50/+1.0/072 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.50/+1.0/072 | OS | Physician | 04/13/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11815908&pc=QCB | 2023826-2021-01501 | 2021-01-27 | Injury | STAAR SURGICAL COMPANY | 2021-05-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/079 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.00/+2.0/079 | OD | Physician | 04/13/2021 | 2021 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11815912&pc=QCB | 2023826-2021-01499 | 2020-12-03 | Injury | STAAR SURGICAL COMPANY | 2021-05-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -16.50/+1.0/060 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -16.50/+1.0/060 | OD | Physician | 04/12/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11816017&pc=QCB | 2023826-2021-01502 | 2021-01-13 | Injury | STAAR SURGICAL COMPANY | 2021-05-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.50/+1.5/091 | OD | Physician | 04/13/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11813723&pc=QCB | 2023826-2021-01488 | 2021-04-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON WAS LOADING A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.00/+2.0/111 (SPHERE/CYLINDER/AXIS), AND NOTED A DEFECT (GAP) OF THE ICL HAPTIC. THERE WAS NO PATIENT CONTACT OR INJURY. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN WITH DEBRIS AND RESIDUE ON HAPTIC AND RESIDUE ON OPTIC. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.00/+2.0/111 | NA | Physician | 04/13/2021 | 2021 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11815966&pc=QCB | 2023826-2021-01498 | 2020-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50/+3.5/070 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS OPTIC DEFORMED AND THE HAPTIC BENT, DEFORMED AND STRETCHED OUT. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, DATE REC¿D BY MFR: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -08.50/+3.5/070 | OS |  | 04/12/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11816167&pc=QCB | 2023826-2021-01494 | 2020-12-23 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.50/+1.0/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, DATE REC¿D BY MFR: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -13.50/+1.0/085 | OD | Physician | 04/15/2021 | 2021 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11816177&pc=QCB | 2023826-2021-01495 | 2020-12-28 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, DATE REC¿D BY MFR: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -10.50/+1.0/090 | OD | Physician | 04/15/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11816268&pc=QCB | 2023826-2021-01496 | 2021-01-15 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -05.50/+1.0/057 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, DATE REC¿D BY MFR: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -05.50/+1.0/057 | OD | Physician | 04/15/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11816284&pc=QCB | 2023826-2021-01497 | 2020-12-29 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+1.5/059 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, DATE REC¿D BY MFR: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -13.00/+1.5/059 | OS | Physician | 04/15/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11820833&pc=QCB | 2023826-2021-01403 | 2021-02-04 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -2.5/2.0/005 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K:THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -2.5/2.0/005 | OD | Physician | 04/21/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11820862&pc=QCB | 2023826-2021-01401 | 2021-02-05 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.00/1.5/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K : THIS PRODUCT IS NOT MARKETED IN THE US. OFF-LABEL USE, ACD<3.0MM. (B)(4). | -10.00/1.5/099 | OS | Physician | 04/19/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11820900&pc=QCB | 2023826-2021-01400 | 2021-03-23 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -16.00/3.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.00/3.0/095 | OS | Physician | 04/19/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11820904&pc=QCB | 2023826-2021-01404 | 2021-02-04 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -3.5/3.0/005 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -3.5/3.0/005 | OS | Physician | 04/21/2021 | 2021 | 52
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11822582&pc=QCB | 2023826-2021-01439 | 2021-02-26 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM ,VTICM5_13.7, -11.00/0.5/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. A SHORTER LENGTH LENS WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED TO BE OTHER: " WRONG LENGTH OF LENS. THE LENS WAS TOO LONG AND THE VAULT TOO HIGH". Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.00/0.5/103 | OD | Physician | 04/13/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11822954&pc=QCB | 2023826-2021-01435 | 2021-01-22 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM ,VTICM5_13.2, -13.50/2.50/80 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. A SHORTER LENGTH LENS WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED TO BE OTHER: " WRONG LENGTH OF LENS. YES BECAUSE THE INITAL LENS WAS TOO LONG FOR THE PATIENT AND WE RECEIVED HIGHT VAULT". Manufacturer Narrative: (B)(6). PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -13.50/2.50/80 | OD | Physician | 04/13/2021 | 2021 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823165&pc=QCB | 2023826-2021-01478 | 2021-01-29 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/156 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+1.5/156 | OS | Physician | 04/20/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823175&pc=QCB | 2023826-2021-01475 | 2021-01-01 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+3.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -15.0/+3.0/092 | OS | Physician | 04/18/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823178&pc=QCB | 2023826-2021-01463 | 2021-04-01 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -12.0/+1.0/083 | OS | Physician | 04/16/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823180&pc=QCB | 2023826-2021-01487 | 2021-01-27 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/078 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -9.5/+1.5/078 | OS | Physician | 04/22/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823205&pc=QCB | 2023826-2021-01447 | 2021-01-25 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -10.5/+1.5/092 | OD | Physician | 04/15/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823208&pc=QCB | 2023826-2021-01459 | 2021-01-03 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -14.0/+3.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LENS ROTATION. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY A REPLACEMENT LENS WON'T BE IMPLANTED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -14.0/+3.0/088 | OS | Physician | 04/16/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823211&pc=QCB | 2023826-2021-01458 | 2021-01-03 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -13.5/+2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LENS ROTATION. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.5/+2.5/091 | OD | Physician | 04/16/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823217&pc=QCB | 2023826-2021-01476 | 2021-02-23 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/069 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -10.0/+1.0/069 | OS | Physician | 04/20/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823247&pc=QCB | 2023826-2021-01461 | 2021-03-28 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/084 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -8.5/+1.5/084 | OS | Physician | 04/16/2021 | 2021 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823250&pc=QCB | 2023826-2021-01482 | 2019-10-24 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -10.0/+2.0/094 | OS | Physician | 04/20/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823264&pc=QCB | 2023826-2021-01485 | 2021-04-20 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -4.0/+4.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS EXCESSIVE VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -4.0/+4.5/088 | OD | Physician | 04/21/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823268&pc=QCB | 2023826-2021-01479 | 2021-01-19 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+0.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -12.0/+0.5/087 | OS | Physician | 04/19/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823270&pc=QCB | 2023826-2021-01462 | 2021-03-30 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2021. THE LENS WAS REMOVED INTRA-OPERATIVELY. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AT A LATER DATE. NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE CODE (B)(4): OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.5/+1.5/093 | OD | Physician | 04/16/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823279&pc=QCB | 2023826-2021-01474 | 2021-01-01 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.5/+2.5/101 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.5/+2.5/101 | OD | Physician | 04/18/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823282&pc=QCB | 2023826-2021-01451 | 2020-10-23 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -11.0/+1.0/092 | OS | Physician | 04/15/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823292&pc=QCB | 2023826-2021-01446 | 2020-09-11 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5/+1.0/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -7.5/+1.0/092 | OD | Physician | 04/15/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823324&pc=QCB | 2023826-2021-01434 | 2021-02-04 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM ,VTICM5_13.7, -4.0/6.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE ON (B)(6) 2021. THE REPORTER STATED " THE HAPTIC OF THE ICL BREAK OFF DURING THE IMPLANTATION. THE ICL REMOVAL WAS PERFORMED". PER UPDATED INFORMATION IT WAS REPORTED THAT THERE WAS NO CONTACT WITH THE PATIENT AND THAT A REPLACEMENT LENS WAS IMPLANTED 2 WEEKS LATER. PATIENT CONDITION 20/20; VAULT 460 UM. Manufacturer Narrative: PATIENT ID: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -4.0/6.0/095 | NA | Physician | 04/13/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823325&pc=QCB | 2023826-2021-01481 | 2021-03-17 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -11.0/+2.0/091 | OS | Physician | 04/21/2021 | 2021 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823335&pc=QCB | 2023826-2021-01465 | 2021-01-03 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+2.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B(4). | -14.5/+2.5/095 | OS | Physician | 04/16/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823344&pc=QCB | 2023826-2021-01448 | 2020-08-26 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -10.0/+1.0/086 | OS | Physician | 04/15/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823350&pc=QCB | 2023826-2021-01460 | 2021-03-28 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.0/+2.5/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+2.5/088 | OD | Physician | 04/16/2021 | 2021 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11820894&pc=QCB | 2023826-2021-01409 | 2021-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -14.00/1.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021, BUT DID NOT RESOLVE THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, A LENS REPLACEMENT IS PLANNED. Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -14.00/1.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021, BUT DID NOT RESOLVE THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, A LENS REPLACEMENT IS PLANNED. ON (B)(6) 2021 THE LENS WAS REPLACED WITH A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. REPORTEDLY, ALL FINE. PATIENT IS HAPPY! CORRECTION: B1 - CHANGED FROM PRODUCT PROBLEM TO ADVERSE EVENT B2 - UPDATED TO REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -14.00/1.0/094 | OD |  | 04/22/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823159&pc=QCB | 2023826-2021-01468 | 2020-11-26 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE CAUSE OF THE EVENT IS REPORTED AS OTHER. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -11.5/+1.5/092 | OD |  | 04/19/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823182&pc=QCB | 2023826-2021-01453 | 2020-12-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/086 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6_ 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+2.0/086 | OD |  | 04/15/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823256&pc=QCB | 2023826-2021-01477 | 2021-03-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY THE DOCTOR DOES NOT WANT TO IMPLANT A REPLACEMENT LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.0/+1.0/093 | OD |  | 04/20/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823314&pc=QCB | 2023826-2021-01445 | 2021-03-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Patient-Device Incompatibility; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER STATES A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS -7.0/+2.5/171 (SPHERE/CYLINDER/AXIS) WAS "VOLD TO FLAT, REMOVE." ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL AND BIOHAZARD BAG. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFO: ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE CAUSE IS REPORTED AS DEVICE AND A PATIENT-RELATED FACTOR. REPORTEDLY, THE PATIENT HAD HIGH ANXIETY DUE TO LOW VAULT AND WAS UNWILLING TO WAIT FOR MORE HEALING AND REQUESTED THE ICL BE EXPLANTED. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -7.0/+2.5/171 | NA |  | 04/20/2021 | 2021 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823333&pc=QCB | 2023826-2021-01470 | 2020-04-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -6.5/+2.0/076 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT AXIS LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE CODE (B)(4): OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -6.5/+2.0/076 | OD |  | 04/19/2021 | 2021 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11823354&pc=QCB | 2023826-2021-01467 | 2020-11-26 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE CAUSE OF THE EVENT IS REPORTED AS OTHER. Manufacturer Narrative: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -12.0/+2.0/090 | OS |  | 04/19/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11828237&pc=QCB | 2023826-2021-01425 | 2020-12-11 | Injury | STAAR SURGICAL COMPANY | 2021-05-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM ,VTICMO13.2, -13.0/2.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED TO BE UNKNOWN. Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -13.0/2.0/081 | OS | Physician | 04/15/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11826904&pc=QCB | 2023826-2021-01433 |  | Injury | STAAR SURGICAL COMPANY | 2021-05-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Unintended Movement | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM ,VTICM5_13.2, -4.5/3.5/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT CAUSED BY THE LENS BEING IMPLANTED UPSIDE DOWN. THERE WAS NO PATIENT INJURY. USER ERROR WAS REPORTED TO BE THE CAUSE OF THIS EVENT, THE REPORTER STATED " YES THAT WAS USER-SURGEON ERROR AND YES THE DEVICE FAIL TO PERFORM BECAUSE OF THAT". IT WAS REPORTED THAT ADDITION OF NEW PI (YAG) WAS PERFORMED BECAUSE THE SURGEON WAS WORRY OF CLOSURE OF THE ANGLE. A REPLACEMENT LENS OF THE SAME LENGTH WAS LATER IMPLANTED. IT WAS REPORTED THAT THE PROBLEM RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, INSIDE A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -4.5/3.5/92 | OD | Physician | 04/14/2021 | 2021 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11827874&pc=QCB | 2023826-2021-01438 | 2021-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P990013 | Event Description: IT WAS REPORTED THAT A 13.7MM, TMICL13.7, -16/1.0/172 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. LENS DISLOCATION/SUBLUXATION WAS OBSERVED, LENS REMAINS IMPLANTED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS THE DEVICE, THE REPORTER STATED " PER DR. (B)(6) LENS IS NOT STABLE. COULD CAUSE CATARACT IF IT CONTACTS WITH NATURAL LENS. PER MEDICAL MONITOR LENS HAS ROTATIONAL INSTABILITY." FOR STATUS OF THE EYE (SIGNS/SYMPTOMS): NO CHANGE IN SYMPTOMS. LENS CONTINUES TO NOT REMAIN IN POSITION. FOR ADDITIONAL INFORMATION THE REPORTER STATED " UNCERTAIN WHETHER LENS HAS LOW VAULT. LENS IS SCHEDULED TO BE EXPLANTED (B)(6) 2021". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION THE LENS WAS EXPLANTED ON (B)(6) 2021 AND THE PROBLEM RESOLVED. THE REPORTER STATED "AT 1-MONTH POST-EXPLANT, CORRECTED VISUAL ACUITY MEASURED 20/25. SUBSEQUENT EXAMINATIONS DOCUMENTED THAT THERE WERE NO UNTOWARD FINDINGS, AND THE VISUAL ACUITY AND MANIFEST REFRACTION REMAINED STABLE." H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTICS TORN. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -16/1.0/172 | OS | Other Health Care Professional | 04/14/2021 | 2021 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11830094&pc=QCB | 2023826-2021-01431 | 2021-02-04 | Injury | STAAR SURGICAL COMPANY | 2021-05-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM ,VTICMO12.6,-12.5/3.5/113 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATE TO LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR SPHERICAL LENS OF SAME LENGTH AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED TO BE UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.5/3.5/113 | OS | Physician | 04/15/2020 | 2020
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11830095&pc=QCB | 2023826-2021-01432 | 2020-12-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM ,VTICMO12.1, -9.0/1.5/85 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THIS EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED TO BE UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE OUT OF SPECIFICATION. H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -9.0/1.5/85 | OD | Physician | 04/15/2020 | 2020 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11839585&pc=QCB | 2023826-2021-01567 | 2021-04-13 | Injury | STAAR SURGICAL COMPANY | 2021-05-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -8.5/1.0/112 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTER TO BE UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/1.0/112 | OS | Physician | 04/22/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11839586&pc=QCB | 2023826-2021-01565 | 2021-04-13 | Injury | STAAR SURGICAL COMPANY | 2021-05-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -6.5/1.0/85 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTER TO BE UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.5/1.0/85 | OD | Physician | 04/22/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11839589&pc=QCB | 2023826-2021-01561 | 2021-01-27 | Injury | STAAR SURGICAL COMPANY | 2021-05-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.0/1.0/085, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTER TO BE UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -12.0/1.0/085 | OD | Physician | 04/19/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11839590&pc=QCB | 2023826-2021-01562 | 2021-04-04 | Injury | STAAR SURGICAL COMPANY | 2021-05-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.0/1.5/135, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTER TO BE UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -11.0/1.5/135 | OS | Physician | 04/19/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11839591&pc=QCB | 2023826-2021-01559 | 2020-12-28 | Injury | STAAR SURGICAL COMPANY | 2021-05-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM ,VTICMO12.6, -10.5/2.5/91, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTER TO BE UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/2.5/91 | OD | Physician | 04/19/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11845811&pc=QCB | 2023826-2021-01560 | 2021-01-28 | Injury | STAAR SURGICAL COMPANY | 2021-05-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM ,VTICMO12, -7.0/1.0/050 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020 AS REPLACEMENT LENS BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH SPHERICAL LENS DUE TO LOW VAULT WITH ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTER TO BE UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.0/1.0/050 | OS | Physician | 04/19/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11845962&pc=QCB | 2023826-2021-01551 | 2020-11-24 | Injury | STAAR SURGICAL COMPANY | 2021-05-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50/+1.5/046 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT IS HAPPY. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -08.50/+1.5/046 | OD | Physician | 04/23/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11845978&pc=QCB | 2023826-2021-01554 | 2021-01-30 | Injury | STAAR SURGICAL COMPANY | 2021-05-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+3.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHTT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.00/+3.5/092 | OD | Physician | 04/19/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11852366&pc=QCB | 2023826-2021-01541 | 2021-04-02 | Injury | STAAR SURGICAL COMPANY | 2021-05-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+2.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6)2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE CODE: (B)(4): THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -10.50/+2.5/090 | OS | Physician | 04/19/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11852371&pc=QCB | 2023826-2021-01542 | 2021-03-31 | Injury | STAAR SURGICAL COMPANY | 2021-05-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+3.5/066 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.00/+3.5/066 | OS | Physician | 04/20/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11852375&pc=QCB | 2023826-2021-01543 | 2020-04-22 | Injury | STAAR SURGICAL COMPANY | 2021-05-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00/+1.5/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING AND REFRACTIVE CHANGE OVERTIME. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) . CLAIM # (B)(4). | -15.00/+1.5/074 | OS | Physician | 04/20/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11852402&pc=QCB | 2023826-2021-01553 | 2021-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -04.50/+1.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS HIGH VAULT POST-OP. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2021. THE LENS EXCHANGE RESOLVED THE PROBLEM. THE PATIENT IS DOING WELL AND THERE WERE NO COMPLICATIONS. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. | -04.50/+1.5/092 | OS |  | 04/23/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855166&pc=QCB | 2023826-2021-01503 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-05-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -6.50/3.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS NOT IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -6.50/3.0/092 | NA | Physician | 04/26/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855167&pc=QCB | 2023826-2021-01504 | 2021-03-01 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.00/5.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS TORE DURING INJECTION/DELIVERY. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY ON (B)(6) 2021. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON 27/APR/2021, "NO OTHER LENS IMPLANTED". THE NEW EXCHANGE LENS IS PLANNED. MFR# 2023826-2020-02598. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.00/5.0/081 | OS | Physician | 04/26/2021 | 2021 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855178&pc=QCB | 2023826-2021-01507 | 2021-04-02 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.5/1.5/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. DATE REC¿D BY MFR- THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.5/1.5/079 | OD | Physician | 04/26/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855181&pc=QCB | 2023826-2021-01510 | 2020-12-20 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.50/2.0/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULT WITH ROTATION. A NEW LONGER LENGTH LENS WAS IMPLANTED ON (B)(6) 2021 AND THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS LENS SIZE. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. DATE REC¿D BY MFR- THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.50/2.0/105 | OD | Physician | 04/26/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855227&pc=QCB | 2023826-2021-01515 | 2020-11-19 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.0/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.0/1.5/089 | OD | Physician | 04/28/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855237&pc=QCB | 2023826-2021-01525 | 2021-02-26 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.00/3.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -15.00/3.5/087 | OD | Physician | 04/29/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855244&pc=QCB | 2023826-2021-01529 | 2021-04-28 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2021 A 13.2MM VTICMO 13.2, -17.50/1.5/095. (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS FOR THE PATIENT'S RIGHT EYE (OD). TORE/BROKE DURING LOADING IN CARTRIDGE. LENS STUCK IN CARTRIDGE/INJECTION SYSTEM DURING INJECTION. NO PATIENT CONTACT (LENS DID NOT TOUCH PATIENT'S EYE). REPORTEDLY, "BACK-UP LENS WAS USED. NO PROBLEM WITH THE LOADING. NO COMPLICATIONS WITH SURGERY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE, USER ERROR, "BODY IN OPTICAL PART IS QUITE THICK, PLUS OBJECTIVELY MORE VISCO MATERIAL IN CARTRIDGE AS NEEDED CAN OCCURED TEAR/BREAK MENTIONED IN THIS CASE. AND WITH COMBINATION WITH UNEXPERIENCED USER." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -17.50/1.5/095. | OD | Physician | 04/29/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855245&pc=QCB | 2023826-2021-01530 | 2021-04-24 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -14.00/2.0/076 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT IN A SECOND SURGERY ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. DATE REC¿D BY MFR- THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -14.00/2.0/076 | OS | Physician | 04/30/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855247&pc=QCB | 2023826-2021-01528 | 2021-04-12 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -13.50/3.5/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. DATE REC¿D BY MFR- THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.50/3.5/098 | OD | Physician | 04/29/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855259&pc=QCB | 2023826-2021-01536 | 2021-02-24 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -9.50/4.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. REPORTEDLY, LENS REPLACEMENT IS PLANNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.50/4.0/094 | OS | Physician | 04/30/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855260&pc=QCB | 2023826-2021-01537 | 2020-12-16 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.0/3.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.0/3.5/089 | OD | Physician | 03/20/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11860707&pc=QCB | 2023826-2021-01602 | 2019-10-24 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.5/074 | OD | Physician | 04/26/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11860709&pc=QCB | 2023826-2021-01596 | 2021-04-22 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-20 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT, WHILE ATTEMPTING TO IMPLANT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+2.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS NOT FULLY IMPLANTED BUT THERE WAS PATIENT CONTACT WITH THE LENS. THIS OCCURRED ON (B)(6) 2021. AN ALTERNATE LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. REPORTEDLY, THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT IS REPORTED AS "DEVICE?" AND "USER ERROR?" Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -15.0/+2.0/090 | OS | Physician | 04/23/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11860711&pc=QCB | 2023826-2021-01595 | 2018-12-17 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.5/098 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2018 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -11.0/+2.5/098 | OS | Physician | 04/28/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11860712&pc=QCB | 2023826-2021-01594 | 2020-09-03 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -17.0/+3.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.0/+3.5/089 | OD | Physician | 04/27/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11860713&pc=QCB | 2023826-2021-01592 | 2021-03-08 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+1.0/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+1.0/091 | OS | Physician | 04/27/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11860714&pc=QCB | 2023826-2021-01590 | 2019-09-03 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXPLANTED AND THEN REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4) | -11.0/+1.5/092 | OD | Physician | 04/27/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11860715&pc=QCB | 2023826-2021-01586 | 2020-12-23 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+1.0/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -7.0/+1.0/092 | OD | Physician | 04/28/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11860717&pc=QCB | 2023826-2021-01584 | 2020-12-31 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -3.5/+2.0/171 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -3.5/+2.0/171 | OD | Physician | 04/25/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11860721&pc=QCB | 2023826-2021-01580 | 2021-03-04 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+5.0/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THEN REPLACED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+5.0/085 | OS | Physician | 04/23/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11860722&pc=QCB | 2023826-2021-01578 | 2020-01-17 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 | QCB | IMPLANTABLE COLLAMER LENS | Patient-Device Incompatibility; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -13.0/+2.5/099 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT MODEL LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. REPORTEDLY, IT IS SPECULATED THAT THE GROOVE TO GROOVE DIAMETER OF THE PATIENT IS TOO LARGE, EXCEEDING THE MAXIMUM SIZE OF THE CRYSTAL AND THE AXIS OF TICL CANNOT BE FIXED. AFTER CHANGING TO ICL THE PATIENT CONSCIOUSLY ADAPTS WELL. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+2.5/099 | OD | Physician | 04/23/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11860725&pc=QCB | 2023826-2021-01599 | 2021-03-07 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-20 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/101 (SPHERE/CYLINDER/AXIS) GOT CAUGHT IN THE CARTRIDGE AND WAS DAMAGED DURING LOADING. THIS OCCURRED ON (B)(6) 2021. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. AN ALTERNATE LENS WAS IMPLANTED ON THE SAME DATE AND THE PROBLEM IS RESOLVED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.0/+1.5/101 | NA | Physician | 04/29/2021 | 2021 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11860726&pc=QCB | 2023826-2021-01576 | 2021-01-30 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -18.0/+1.0/144 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT SPHERE/CYLINDER/AXIS LENS DUE TO AN ERROR IN REFRACTION CALCULATION. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -18.0/+1.0/144 | OS | Physician | 04/23/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855168&pc=QCB | 2023826-2021-01508 | 2021-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -9.5/1.0/169 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE , HIGHER POWER (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN, DEFORMED, AND STRETCHED OUT. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN, DEFORMED, AND STRETCHED OUT. RE-HYDRATION DEMONSTRATED SIGNIFICANT DAMAGE OF LENS HAPTIC, THEREFORE LENGTH MEASUREMENT NOT POSSIBLE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.5/1.0/169 | OS | Physician | 04/26/2021 | 2021 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855177&pc=QCB | 2023826-2021-01408 |  | Injury | STAAR SURGICAL COMPANY | 2021-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -13.00/1.0/140 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). REPORTEDLY, LENS TO BE EXPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H1: SERIOUS INJURY. ADDITIONAL INFORMATION: B5 - PER INFOFLO THE LENS WAS EXPLANTED. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.00/1.0/140 | OD | Physician | 05/04/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855182&pc=QCB | 2023826-2021-01514 | 2021-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: HE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -6.5/1.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019.LOW VAULT WAS OBSERVED, THE LENS REMAINS IMPLANTED. REPORTEDLY, LENS REPLACEMENT IS PLANNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -6.50/1.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON 14/JUN/2019. THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON 20/MAY/2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT IS HAPPY." CLAIM#: (B)(4). Manufacturer Narrative: D6B EXPLANT: (B)(6) 2021. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -6.5/1.0/079 | OS |  | 04/27/2021 | 2021 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855183&pc=QCB | 2023826-2021-01512 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, TMICL12.6, -13.50/2.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS DID NOT RELEASE PROPERLY FROM THE PLUNGER. THE BACK UP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED STUCK IN CARTRIDGE, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND LENS RETURNED IN CARTRIDGE. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.50/2.0/086 | NA |  | 04/27/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855188&pc=QCB | 2023826-2021-01513 | 2021-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.7, -7.00/1.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. LOW VAULT WAS OBSERVED, THE LENS REMAINS IMPLANTED. REPORTEDLY, LENS REPLACEMENT IS PLANNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.00/1.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT IS HAPPY." CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -7.00/1.0/088 | OD |  | 04/27/2021 | 2021 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855230&pc=QCB | 2023826-2021-01517 | 2020-12-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.00/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.00/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.00/1.0/091 | OD |  | 04/27/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855232&pc=QCB | 2023826-2021-01518 | 2020-12-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.00/2.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.00/2.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. DATE REC¿D BY MFR- THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.00/2.0/084 | OS |  | 04/27/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855238&pc=QCB | 2023826-2021-01523 | 2020-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -3.00/1.5/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 -3.00/1.5/104 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 LENS WAS EXCHANGED WITH A SAME LENGTH/OPPOSITE TORIC AXIS LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -3.00/1.5/104 | OS | Physician | 04/28/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11855254&pc=QCB | 2023826-2021-01535 | 2021-02-24 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -08.00/3.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. REPORTEDLY, LENS REPLACEMENT IS PLANNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6- INVESTIGATION FINDINGS-114. INVESTIGATION CONCLUSION- 4315. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -08.00/3.5/094 | OD | Physician | 04/30/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11860731&pc=QCB | 2023826-2021-01579 | 2021-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE CODE (B)(4): OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -9.5/+1.0/088 | OD |  | 04/23/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11878386&pc=QCB | 2023826-2021-01608 | 2021-04-09 | Injury | STAAR SURGICAL COMPANY | 2021-05-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.00/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.00/1.5/093 | OS | Physician | 05/05/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11878397&pc=QCB | 2023826-2021-01607 | 2020-10-30 | Injury | STAAR SURGICAL COMPANY | 2021-05-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -17.50/1.0/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -17.50/1.0/099 | OD | Physician | 05/06/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11878405&pc=QCB | 2023826-2021-01609 | 2021-03-12 | Injury | STAAR SURGICAL COMPANY | 2021-05-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.0/2.5/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US CLAIM# (B)(4). | -11.0/2.5/078 | OS | Physician | 05/05/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11878458&pc=QCB | 2023826-2021-01603 | 2020-12-16 | Injury | STAAR SURGICAL COMPANY | 2021-05-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.00/3.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/3.0/090 | OS | Physician | 05/02/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11887117&pc=QCB | 2023826-2021-01627 | 2021-04-22 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON WAS LOADING A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.50/+0.5/101 (SPHERE/CYLINDER/AXIS), AND THE LENS WAS TORN. THERE WAS NO PATIENT CONTACT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -16.50/+0.5/101 | NA | Physician | 04/26/2021 | 2021 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11887151&pc=QCB | 2023826-2021-01619 | 2021-01-15 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -06.00/+2.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6)2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -06.00/+2.5/095 | OS | Physician | 04/26/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11890516&pc=QCB | 2023826-2021-01629 | 2021-04-28 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -14.00/+2.0/071 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.00/+2.0/071 | OD | Physician | 04/28/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11890522&pc=QCB | 2023826-2021-01618 | 2021-01-15 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -05.00/+1.0/077 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -05.00/+1.0/077 | OD | Physician | 04/26/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11891114&pc=QCB | 2023826-2021-01632 | 2021-01-04 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -08.50/+1.0/069 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -08.50/+1.0/069 | OD | Physician | 04/30/2021 | 2021 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11891130&pc=QCB | 2023826-2021-01620 | 2021-03-01 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.00/+1.0/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.00/+1.0/080 | OD | Physician | 04/26/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11891179&pc=QCB | 2023826-2021-01693 | 2020-11-10 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.5/2.5/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTER TO BE UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -8.5/2.5/100 | OD | Physician | 04/28/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11891183&pc=QCB | 2023826-2021-01692 |  | Injury | STAAR SURGICAL COMPANY | 2021-05-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -8.5/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTER TO BE UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/2.5/093 | OS | Physician | 04/28/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11891191&pc=QCB | 2023826-2021-01694 | 2020-12-02 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.0/1.5/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTER TO BE UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -11.0/1.5/075 | OS | Physician | 04/28/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11891208&pc=QCB | 2023826-2021-01691 | 2020-12-07 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM ,VTICMO13.2, -13.0/3.5/95 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTER TO BE UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.0/3.5/95 | OS | Physician | 04/28/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11891210&pc=QCB | 2023826-2021-01700 | 2020-08-28 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM ,VTICMO13.2, -9.0/1.5/98 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTER TO BE UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.0/1.5/98 | OD | Physician | 04/30/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11891218&pc=QCB | 2023826-2021-01702 | 2021-02-27 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM ,VTICMO12.6, -11.5/3.0/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT. IT WAS REPORTED THAT THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED TO BE UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE CODE (B)(4): OFF-LABEL USE (ACD < 3.0MM). CLAIM #: (B)(4). | -11.5/3.0/90 | OD | Physician | 04/30/2021 | 2021 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11891219&pc=QCB | 2023826-2021-01677 | 2021-03-31 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM ,VTICMO13.7, -15.5/3.0/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTER TO BE UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.5/3.0/098 | OD | Physician | 04/26/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11891220&pc=QCB | 2023826-2021-01703 | 2021-02-27 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM ,VTICMO12.6, -12.5/2.0/87 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH LENS WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT. IT WAS REPORTED THAT THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED TO BE UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.5/2.0/87 | OS | Physician | 04/30/2021 | 2021 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11891221&pc=QCB | 2023826-2021-01678 | 2021-03-31 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM ,VTICMO13.7, -14.5/2.5/104(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTER TO BE UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.5/2.5/104 | OS | Physician | 04/26/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11891257&pc=QCB | 2023826-2021-01633 | 2021-01-04 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/112 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -09.50/+1.0/112 | OS | Physician | 04/30/2021 | 2021 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11891280&pc=QCB | 2023826-2021-01625 | 2021-02-06 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+2.0/109 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.50/+2.0/109 | OD | Physician | 04/27/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11891285&pc=QCB | 2023826-2021-01638 | 2021-02-02 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50/+1.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -08.50/+1.5/090 | OD | Physician | 04/30/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11891288&pc=QCB | 2023826-2021-01640 | 2020-11-07 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.50/+1.5/136 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S.. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -12.50/+1.5/136 | OD | Physician | 04/30/2021 | 2021 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11890520&pc=QCB | 2023826-2021-01623 | 2021-04-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -01.00/+2.0/075 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINED IMPLANTED BUT THE PLAN IS TO EXPLANT THE LENS. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: B5 - THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. REPORTEDLY, "ALL GOOD. PATIENT IS HAPPY." CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -01.00/+2.0/075 | OD | Physician | 04/27/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11891168&pc=QCB | 2023826-2021-01676 | 2021-03-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM ,VTICM5_13.7, -10.50/6.0/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. IT WAS REPORTED THAT THE SURGEON SELECTED THE WRONG LENS POWER, USER ERROR WAS REPORTED TO BE THE CAUSE OF THIS EVENT. ON (B)(6) 2021 REMANIPULATION OF ICL HAPTIC WAS REPORTED. THE REPORTER STATED " THE SURGEON RE-MANIPULTED THE LENS TO DOUBLE CHECK THAT THERE IS NO LENS TILT AND ALL ICL HAPTICS ARE LYING PROPERLY IN THE SULCUS." IT WAS REPORTED THAT LENS EXCHANGE IS PLANNED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 13. 7MM, VTICM5_ 13. 7, -10.50/6.0/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IT WAS REPORTED THAT THE SURGEON SELECTED THE WRONG LENS POWER, USER ERROR WAS REPORTED TO BE THE CAUSE OF THIS EVENT. ON 11 MAR 2021 REMANIPULATION OF ICL HAPTIC WAS PERFORMED. THE REPORTER STATED, " THE SURGEON RE-MANIPULATED THE LENS TO DOUBLE CHECK THAT THERE IS NO LENS TILT AND ALL LCL HAPTICS ARE LYING PROPERLY IN THE SULCUS." ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS AND THE PROBLEM RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/6.0/92 | OD | Physician | 04/26/2021 | 2021 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11891170&pc=QCB | 2023826-2021-01697 | 2021-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -6.0/1.0/039 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS. THIS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATED "DUE TO THE SPECIAL ANATOMICAL STRUCTURE OF THE CILIARY SULCUS OF THIS PATIENT, THE SURGEON DECIDED TO CHANGE THE LENS POSITION FOR BETTER VAULT. THE ORIGINAL LENS IS:-6/+1/039, THE EXCHANGE LENS IS:-6/+1/135, THE AXIS OF LENS IS DIFFERENT." CAUSE WAS REPORTER TO BE UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -6.0/1.0/039 | OD | Physician | 04/29/2021 | 2021 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11891178&pc=QCB | 2023826-2021-01688 | 2020-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -4.0/0.5/164 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED DUE TO LOW VAULT WITH ROTATION, IT WAS REPORTED THAT THE PROBLEM RESOLVED. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH SPHERICAL LENS. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTER TO BE UNKNOWN. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -4.0/0.5/164 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED DUE TO LOW VAULT WITH ROTATION, BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS OF LONGER LENGTH AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED TO BE UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -4.0/0.5/164 | OS |  | 04/28/2021 | 2021 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11891283&pc=QCB | 2023826-2021-01630 | 2020-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -16.50/+2.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS ROTATION NOT ASSOCATED TO A LOW VAULT. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS REPOSITIONED ON (B)(6) 2020 BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS EXPLANTED. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -16.50/+2.5/090 | OS | Physician | 04/28/2021 | 2021 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11894529&pc=QCB | 2023826-2021-01726 | 2021-04-22 | Injury | STAAR SURGICAL COMPANY | 2021-05-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -9.0/+1.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EPLANTED AND THEN REPLACED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+1.5/091 | OD | Physician | 05/07/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11894533&pc=QCB | 2023826-2021-01715 | 2021-01-03 | Injury | STAAR SURGICAL COMPANY | 2021-05-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+0.5/094 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+0.5/094 | OS | Physician | 04/30/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11894536&pc=QCB | 2023826-2021-01725 | 2021-04-14 | Injury | STAAR SURGICAL COMPANY | 2021-05-27 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -11.0/+2.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT SPHERE/CYLINDER/AXIS LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/+2.0/090 | OD | Physician | 05/06/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11894539&pc=QCB | 2023826-2021-01722 | 2020-12-08 | Injury | STAAR SURGICAL COMPANY | 2021-05-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.5/090 | OD | Physician | 05/06/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11894542&pc=QCB | 2023826-2021-01724 | 2020-07-14 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-27 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/114 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2020. THE LENS WAS REMOVED AND SUCCESSFULLY REPLACED WITH AN ALTERNATE LENS INTRA-OPERATIVELY. NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.0/+1.0/114 | OD | Physician | 05/06/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11894549&pc=QCB | 2023826-2021-01714 | 2021-01-03 | Injury | STAAR SURGICAL COMPANY | 2021-05-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+0.5/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+0.5/089 | OD | Physician | 04/30/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11894550&pc=QCB | 2023826-2021-01718 | 2020-11-15 | Injury | STAAR SURGICAL COMPANY | 2021-05-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+2.0/091 | OS | Physician | 05/02/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11895231&pc=QCB | 2023826-2021-01635 | 2021-01-15 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+1.0/108 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -13.00/+1.0/108 | NA | Physician | 04/30/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11898399&pc=QCB | 2023826-2021-01683 | 2021-04-13 | Injury | STAAR SURGICAL COMPANY | 2021-05-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM ,VTICM5_13.2, -8.5/1.5/3 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTER TO BE UNKNOWN. Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/1.5/3 | OD | Physician | 04/27/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11894528&pc=QCB | 2023826-2021-01721 | 2020-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -14.0/+1.5/071 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO THE LENS BEING TOO LONG; SURGEON PREFERENCE CHANGE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE/DEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS AND RESIDUE ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.0/+1.5/071 | OS |  | 05/05/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11894535&pc=QCB | 2023826-2021-01717 | 2021-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/083 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN, BENT, DEFORMED, AND STRETCHED OUT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+2.0/083 | OS |  | 05/02/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11894538&pc=QCB | 2023826-2021-01723 | 2020-08-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+2.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THERE WAS PATIENT CONTACT WITH THE LENS BUT THE LENS WAS NOT FULLY IMPLANTED. THIS OCCURRED ON (B)(6) 2020. AN ALTERNATE LENS WAS IMPLANTED WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. REPORTEDLY, THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -10.0/+2.5/090 | OS |  | 05/06/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11894546&pc=QCB | 2023826-2021-01720 | 2020-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/094 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO THE LENS BEING TOO LONG; SURGEON PREFERENCE CHANGE. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE AND DEBRIS ON THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+1.5/094 | OD |  | 05/05/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11894555&pc=QCB | 2023826-2021-01716 | 2021-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WASBROKEN, DEFORMED, AND STRETCHED OUT. HERE WAS RESIDUE ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.5/+1.0/090 | OD |  | 05/02/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11895238&pc=QCB | 2023826-2021-01631 | 2021-01-11 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00/+0.5/079 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED DURING THE SAME SURGERY WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT, DEFORMED, AND STRETCHED OUT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -08.00/+0.5/079 | NA | Physician | 04/30/2021 | 2021 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11899091&pc=QCB | 2023826-2021-01681 | 2021-04-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P990013 | Event Description: IT WAS REPORTED THAT A 12.6MM, TMICL12.6, -11.5/1.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING DELIVERY. THE LENS WAS REMOVED AND REPLACED WITHIN THE SAME SURGERY WITH NO PATIENT INJURY AND THE PROBLEM RESOLVED. FOR CAUSE OF EVENT THE REPORTER STATED " FOOTPLATE RIPPED DURING DELIVERY, VERY THICK LENS, DIFFICULT TO PULL THROUGH CARTRIDGE. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. CLAIM#:(B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.5/1.5/082 | NA |  | 04/27/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11904719&pc=QCB | 2023826-2021-01696 | 2021-04-01 | Injury | STAAR SURGICAL COMPANY | 2021-05-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM513.2, -9.50/1.0/81 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT.THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTER TO BE UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.0/81 | OD | Physician | 04/28/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11917159&pc=QCB | 2023826-2021-01735 | 2020-12-30 | Injury | STAAR SURGICAL COMPANY | 2021-06-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -08.50/+1.5/134 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -08.50/+1.5/134 | OS | Physician | 05/06/2021 | 2021 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11917170&pc=QCB | 2023826-2021-01739 | 2021-03-23 | Injury | STAAR SURGICAL COMPANY | 2021-06-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00/+3.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.00/+3.0/089 | OD | Physician | 05/05/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11917175&pc=QCB | 2023826-2021-01740 | 2021-04-09 | Injury | STAAR SURGICAL COMPANY | 2021-06-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.00/+2.0/090 | OS | Physician | 05/05/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11925489&pc=QCB | 2023826-2021-01744 |  | Injury | STAAR SURGICAL COMPANY | 2021-06-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.50/1.5/023 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) AND HIGH VAULT WAS OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_ 13.2, -8.50/1.5/023 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021 AND HIGH VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT ID: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -8.50/1.5/023 | OS | Physician | 05/03/2021 | 2021 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11925684&pc=QCB | 2023826-2021-01745 |  | Injury | STAAR SURGICAL COMPANY | 2021-06-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM ,VTICM5_13.2, -9.00/2.0/170 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) AND HIGH VAULT WAS OBSERVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_ 13.2, -9.00/2.0/170 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021 AND HIGH VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT ID:(B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | -9.00/2.0/170 | OD | Physician | 05/03/2021 | 2021 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11931140&pc=QCB | 2023826-2021-01750 | 2021-05-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM ,VTICM5_12.6, -13.0/2.5/127 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHTS EYE (OD) ON (B)(6) 2021. IT WAS NOTED THAT THE LENSES WERE SWITCHED (LEFT EYE LENS WAS IMPLANTED INTO THE RIGHT EYE). CAUSE OF THIS EVENT WAS REPORTED TO BE USER ERROR. IT WAS REPORTED THAT LENS REPOSITIONING WAS PERFORMED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.6MM ,VTICM5_ 12.6, -13.0/2.5/127 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHTS EYE (OD) ON (B)(6) 2021. IT WAS NOTED THAT THE LENSES WERE SWITCHED (LEFT EYE LENS WAS IMPLANTED INTO THE RIGHT EYE ). CAUSE OF THIS EVENT WAS REPORTED TO BE USER ERROR. IT WAS REPORTED THAT LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2021 AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.0/2.5/127 | OD | Physician | 05/06/2021 | 2021 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11931520&pc=QCB | 2023826-2021-01788 | 2021-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.0/2.5/063 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. IT WAS NOTED THAT THE LENSES WERE SWITCHED (RIGHT EYE LENS WAS IMPLANTED INTO THE LEFT EYE). CAUSE OF THIS EVENT WAS REPORTED TO BE USER ERROR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.6MM ,VTICM5_ 12.6, -13.0/2.5/063 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. IT WAS NOTED THAT THE LENSES WERE SWITCHED (RIGHT EYE LENS WAS IMPLANTED INTO THE LEFT EYE ). CAUSE OF THIS EVENT WAS REPORTED TO BE USER ERROR. IT WAS REPORTED THAT LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2021 AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -13.0/2.5/063 | OS | Physician | 05/06/2021 | 2021 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11938957&pc=QCB | 2023826-2021-01650 | 2021-05-07 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -17.50/3.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS TORE DURING REMOVAL. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY ON (B)(6) 2021. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -17.50/3.5/092 | OS | Physician | 05/11/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11938961&pc=QCB | 2023826-2021-01655 | 2021-02-08 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -12.00/2.0/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHTFT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/2.0/111 | OD | Physician | 05/11/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11938965&pc=QCB | 2023826-2021-01653 | 2021-04-02 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.50/1.5/093 | OD | Physician | 05/11/2021 | 2021 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11938967&pc=QCB | 2023826-2021-01654 | 2021-02-20 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -14.00/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.00/2.0/090 | OS | Physician | 05/11/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11938969&pc=QCB | 2023826-2021-01659 | 2021-01-18 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 12.6MM VTICMO12.6, -5.00/1.5/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. A SHORTER LENGTH LENS WAS IMPLANTED DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: (B)(4): OFF-LABEL USE, ACD<3.0MM. CLAIM# (B)(4). | -5.00/1.5/074 | OD | Physician | 05/12/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11939012&pc=QCB | 2023826-2021-01663 | 2021-04-14 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6M, VTICM5_12.6, -13.50/3.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS REPLACED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.50/3.0/083 | OD | Physician | 05/14/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11939014&pc=QCB | 2023826-2021-01661 | 2021-03-28 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM512.6, -6.50/2.5/068 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.50/2.5/068 | OS | Physician | 05/14/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11939015&pc=QCB | 2023826-2021-01662 | 2021-04-07 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -11.00/4.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS REPLACED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.00/4.0/094 | OD | Physician | 05/14/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11939018&pc=QCB | 2023826-2021-01664 | 2021-04-14 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6M, VTICM5_12.6, -13.50/4.00/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS REPLACED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.50/4.00/086 | OS | Physician | 05/14/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11939019&pc=QCB | 2023826-2021-01667 | 2021-04-14 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -10.00/1.50/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS REPLACED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.00/1.50/093 | OD | Physician | 05/14/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11939020&pc=QCB | 2023826-2021-01665 | 2020-11-07 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.50/2.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2020, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. PER THE REPORTER, THERE ARE NO PLANS TO IMPLANT ANOTHER LENS Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4) | -10.50/2.0/083 | OS | Physician | 05/14/2020 | 2020 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11940994&pc=QCB | 2023826-2021-01761 | 2021-01-07 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+2.5/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE CODE: (B)(4) THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -11.50/+2.5/091 | OD | Physician | 05/11/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942671&pc=QCB | 2023826-2021-01787 | 2021-03-17 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM # (B)(4). | -9.5/+2.0/092 | OS | Physician | 05/10/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942672&pc=QCB | 2023826-2021-01776 | 2021-01-08 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -4.0/+6.0/077 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION; THE PROBLEM WAS NOT RESOLVED. ON (B)(6) 2021 THE LENS WAS EXPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE SURGEON STATES, " AFTER IMPLANTATION OF THE LENS INTO THE EYE, ROTATION AND ADJUSTMENT DID NOT SOLVE THE PROBLEM FOR MANY TIMES AND THE PATIENT GAVE UP THE OPERATION." Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -4.0/+6.0/077 | OD | Physician | 05/11/2021 | 2021 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942677&pc=QCB | 2023826-2021-01783 | 2021-04-26 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/015 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO THE ICL FELL INTO THE VITREOUS BODY DURING SURGERY. REPORTEDLY THE ICL WAS REMOVED DURING THE VITRECTOMY. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.5/+1.5/015 | OD | Physician | 05/13/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942679&pc=QCB | 2023826-2021-01775 | 2021-01-06 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.0/+1.0/171 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.0/+1.0/171 | OD | Physician | 05/11/2021 | 2021 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942680&pc=QCB | 2023826-2021-01785 |  | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+2.0/107 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6)2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN AND OTHER. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM# (B)(4). | -15.0/+2.0/107 | OS | Physician | 05/14/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942681&pc=QCB | 2023826-2021-01777 | 2021-01-30 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.5/+1.0/170 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.5/+1.0/170 | OD | Physician | 05/11/2021 | 2021 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942683&pc=QCB | 2023826-2021-01768 | 2021-03-15 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+1.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM # (B)(4). | -7.0/+1.0/088 | OD | Physician | 05/09/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942688&pc=QCB | 2023826-2021-01773 | 2021-04-29 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.0/+1.5/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+1.5/091 | OS | Physician | 05/11/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942690&pc=QCB | 2023826-2021-01784 | 2020-06-16 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.0/+2.0/104 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN AND OTHER. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -14.0/+2.0/104 | OD | Physician | 05/14/2021 | 2021 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942692&pc=QCB | 2023826-2021-01766 | 2021-04-10 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON THE SAME DATE, IN A SEPARATE SURGERY THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -9.0/+1.0/089 | OD | Physician | 05/09/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11938019&pc=QCB | 2023826-2021-01755 | 2021-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.50/+2.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED AND EXCHANGED ON (B)(6) 2021 FOR A SHORTER LENS. THE PROBLEM WAS RESOLVED AND THE PATIENT IS HAPPY. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.50/+2.0/086 | OD |  | 05/07/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11938966&pc=QCB | 2023826-2021-01656 | 2021-01-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -5.00/1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE, SIMILARE (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC DEFORMED AND STRETCHED OUT. CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTICATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.00/1.0/092 | OS |  | 05/11/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11941250&pc=QCB | 2023826-2021-01748 | 2021-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P990013 | Event Description: IT WAS REPORTED THAT A 13.2MM, TMICL13.2, -11.5/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGE FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED WITH RESIDUE/DEBRIS ON PRODUCT, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN AND RESIDUE ON LENS SURFACE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -11.5/2.5/093 | OS | Physician | 05/05/2021 | 2021 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942669&pc=QCB | 2023826-2021-01781 | 2021-05-10 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -2.0/+1.0/014 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS EXCESSIVE VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5: REPORTER STATED "1 YEAR FOLLOW-UP IS PLANNED" AND "PATIENT IS HAPPY." CLAIM#: (B)(4). Manufacturer Narrative: B5 - REPORTER STATED, "PATIENT IS HAPPY. HE DOESN'T WANT A LENS EXCHANGE. AN ANNUAL CHECK IS CARRIED OUT." CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -2.0/+1.0/014 | OD | Physician | 05/12/2021 | 2021 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11942693&pc=QCB | 2023826-2021-01767 | 2021-04-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THEN REPLACED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: RV THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS DIOPTER -9.50/1.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS LATER EXCHANGED ON (B)(6) 2021 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM # (B)(4). | -9.5/+1.0/088 | OS | Physician | 05/09/2021 | 2021 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11962106&pc=QCB | 2023826-2021-01798 | 2021-02-04 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2 -6.5/1.5/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITHIN THE SAME SURGERY. THIS REPLACEMENT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN, THE REPORTER STATED " THE LENS SIZE RECOMMENDED BY OCOS DID NOT FIT THE PATIENT." Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, DESCRIBE EVENT OR PROBLEM: UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -6.5/1.5/91 | OS | Physician | 05/10/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11962120&pc=QCB | 2023826-2021-01809 | 2021-02-01 | Injury | STAAR SURGICAL COMPANY | 2021-06-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.0/2.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021.ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/2.0/096 | OD | Physician | 05/12/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11962125&pc=QCB | 2023826-2021-01797 | 2021-02-04 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2 -6.5/1.5/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITHIN THE SAME SURGERY. THIS REPLACEMENT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN, THE REPORTER STATED " THE LENS SIZE RECOMMENDED BY OCOS DID NOT FIT THE PATIENT." Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, DESCRIBE EVENT OR PROBLEM: UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -6.5/1.5/89 | OD | Physician | 05/10/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11962147&pc=QCB | 2023826-2021-01800 | 2020-12-01 | Injury | STAAR SURGICAL COMPANY | 2021-06-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -10.0/1.0/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM # (B)(4). | -10.0/1.0/90 | OS | Physician | 05/10/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11962169&pc=QCB | 2023826-2021-01810 | 2021-02-01 | Injury | STAAR SURGICAL COMPANY | 2021-06-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.5/1.0/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021.ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/1.0/074 | OS | Physician | 05/12/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11962212&pc=QCB | 2023826-2021-01808 | 2021-01-04 | Injury | STAAR SURGICAL COMPANY | 2021-06-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.0/1.5/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021.ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/1.5/107 | OD | Physician | 05/12/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11962082&pc=QCB | 2023826-2021-01806 | 2021-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.5/1.5/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE DID NOT RESOLVE THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN, REPORTER STATED " THE LENS SIZE RECOMMENDED BY OCOS DID NOT FIT THE PATIENT." Manufacturer Narrative: DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, INSIDE A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS BROKEN. CLAIM#: (B)(4). Manufacturer Narrative: CODE: 213, 4310. CLAIM#(B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -9.5/1.5/91 | OD | Physician | 05/10/2021 | 2021 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11962100&pc=QCB | 2023826-2021-01807 | 2021-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -9.5/1.5/92(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021 AS A REPLACEMENT LENS TO CORRECT LOW VAULT BUT THE VAULT DID NOT IMPROVE. THE LENS WAS EXPLANTED WITHIN THE SAME SURGERY AND IT WAS REPORTED THAT THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN, REPORTER STATED " THE LENS SIZE RECOMMENDED BY OCOS DID NOT FIT THE PATIENT." SEE MFR # 2023826-2021-01806 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, INSIDE A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS, RESIDUE THROUGHOUT THE LENS WAS OBSERVED. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: VISUAL INSPECTION FOUND RESIDUE ON A HAPTIC TORNS LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. | -9.5/1.5/92 | OD | Physician | 05/10/2021 | 2021 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11962126&pc=QCB | 2023826-2021-01804 | 2021-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.0/1.0/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE DID NOT RESOLVE THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN, REPORTER STATED " THE LENS SIZE RECOMMENDED BY OCOS DID NOT FIT THE PATIENT." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID WITH RESIDUE/DEBRIS ON PRODUCT INSIDE A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.0/1.0/93 | OS | Physician | 05/10/2021 | 2021 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11962132&pc=QCB | 2023826-2021-01805 | 2021-02-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -10.0/1.5/86 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021 AS A REPLACEMENT LENS TO CORRECT LOW VAULT BUT THE VAULT DID NOT IMPROVE. THE LENS WAS EXPLANTED WITHIN THE SAME SURGERY AND IT WAS REPORTED THAT THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN, REPORTER STATED " THE LENS SIZE RECOMMENDED BY OCOS DID NOT FIT THE PATIENT." SEE MFR # 2023826-2021-01804 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, INSIDE A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS, RESIDUE THROUGHOUT THE LENS WAS OBSERVED. CLAIM#: (B)(4) Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.0/1.5/86 | OS | Physician | 05/10/2021 | 2021 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11962157&pc=QCB | 2023826-2021-01811 | 2021-03-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -10.0/2.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE EXCHANGE LENS WAS IMPLANTED VERTICALLY. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -10.0/2.5/105 | OD | Physician | 05/12/2021 | 2021 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11966816&pc=QCB | 2023826-2021-01815 | 2021-03-16 | Injury | STAAR SURGICAL COMPANY | 2021-06-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/1.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, DATE REC¿D BY MFR- THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.00/1.5/085 | OD | Physician | 05/18/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11966824&pc=QCB | 2023826-2021-01812 | 2021-05-08 | Injury | STAAR SURGICAL COMPANY | 2021-06-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -6.00/1.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS REPLACED FOR A LONGER LENGTH LENS, THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "PATIENT IS MAKING SATISFACTORY PROGRESS AFTER SURGERY." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.00/1.0/084 | OS | Physician | 05/17/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11966825&pc=QCB | 2023826-2021-01821 | 2020-07-24 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.50/3.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS TORE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK, ETHNICITY - UNK, RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -9.50/3.0/081 | OS | Physician | 05/19/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11966828&pc=QCB | 2023826-2021-01824 | 2021-01-04 | Injury | STAAR SURGICAL COMPANY | 2021-06-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.50/1.0/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -8.50/1.0/074 | OS | Physician | 05/20/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11967226&pc=QCB | 2023826-2021-01829 | 2021-04-02 | Injury | STAAR SURGICAL COMPANY | 2021-06-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.00/+0.50/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, DATE REC¿D BY MFR: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.00/+0.50/083 | OS | Physician | 05/11/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11967240&pc=QCB | 2023826-2021-01839 | 2021-04-08 | Injury | STAAR SURGICAL COMPANY | 2021-06-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50/+1.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -11.50/+1.0/088 | OS | Physician | 05/13/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11971160&pc=QCB | 2023826-2021-01845 | 2021-02-02 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/095 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS HAD EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -11.00/+1.5/095 | NA | Physician | 05/14/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11971173&pc=QCB | 2023826-2021-01835 | 2021-02-01 | Injury | STAAR SURGICAL COMPANY | 2021-06-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.00/+1.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM # (B)(4). | -08.00/+1.5/094 | OS | Physician | 05/12/2021 | 2021 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11971174&pc=QCB | 2023826-2021-01831 | 2020-07-24 | Injury | STAAR SURGICAL COMPANY | 2021-06-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/117 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM # (B)(4). | -09.50/+1.0/117 | OD | Physician | 05/11/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11971639&pc=QCB | 2023826-2021-01795 | 2021-02-04 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE DOCTOR FOUND A 12.6MM, VTICM5_12.6, -10.0/1.0/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS DAMAGE DURING LOADING THE LENS INTO THE CARTRIDGE. THE LENS GOT CAUGHT IN THE CARTRIDGE AND WAS DAMAGED. THERE WAS NO PATIENT CONTACT. A REPLACEMENT LENS WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, EVENT DESCRIPTION: UNK, THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.0/1.0/92 | NA | Physician | 05/10/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11971642&pc=QCB | 2023826-2021-01789 | 2019-12-19 | Injury | STAAR SURGICAL COMPANY | 2021-06-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -16.0/6.0/014 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE PROBLEM DID NOT RESOLVE. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR THE SAME LENGTH, DIFFERENT DIOPTER, NON-TORIC LENS. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -16.0/6.0/014 | OS | Physician | 05/10/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11966822&pc=QCB | 2023826-2021-01816 | 2021-04-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A IMPLANTABLE COLLAMER LENS WAS DAMAGED DURING SURGERY AND THE BACK- UP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION : B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -4.50/6.0/008 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION : B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -4.50/6.0/008 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: B5: REPORTEDLY NO PATIENT CONTACT WITH COMPLAINT LENS. D2: PRODUCT CODE: QCB. D4: T629610. H10: SUPPLEMENTAL 1 AND 2 ARE NOT APPLICABLE DUE TO SUBMISSION FOR ANOTHER CLAIM (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH - UNK, SEX - UNK, WEIGHT - UNK, ETHNICITY - UNK, RACE - UNK, DATE OF EVENT - UNK, PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). PRODUCT INFO. - UNK, IMPLANT DATE - UNK, THIS PRODUCT IS NOT MARKETED IN THE US, DEVICE MANUFACTURE DATE - UNK. CLAIM# (B)(4). | -4.50/6.0/008 | OS | Physician | 05/17/2021 | 2021 | 44 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11966840&pc=QCB | 2023826-2021-01827 |  | Injury | STAAR SURGICAL COMPANY | 2021-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6, -11.50/2.5/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS OPENED BUT NOT IMPLANTED DUE TO INTRAOPERATIVE INCIDENT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED, THAT A 12.6MM VTICM5 12.6, -11.50/2.5/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS OPENED, BUT NOT IMPLANTED, DUE TO INTRAOPERATIVE INCIDENT. TEAR OF THE ANTERIOR CAPSULE LENS. SURGERY TRANSFORMED IN PHACOEMULSIFICATION. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: LENS TORE/BROKE DURING INJECTION/DELIVERY INTO EYE. AS PER THE COMPLAINT QUESTIONNAIRE, THE REPORTER MARKED THE CAUSE OF THE EVENT AS OTHER. AND STATED "CRISTALLIN CAPSULAR TEAR DURING DELIVERY INTO THE EYE". CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.50/2.5/108 | NA | Physician | 05/21/2021 | 2021 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11967279&pc=QCB | 2023826-2021-01841 | 2021-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+1.5/161 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS WAS TOO LONG. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. IT WAS THE SURGEONS PREFERENCE TO CHANGE THE LENS. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS RESIDUE ON LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -16.50/+1.5/161 | OD |  | 05/13/2021 | 2021 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11971209&pc=QCB | 2023826-2021-01838 | 2021-01-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+2.0/177 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC TORN. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.00/+2.0/177 | OD |  | 05/11/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11971640&pc=QCB | 2023826-2021-01794 | 2020-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: IT WAS REPORTED THAT A 13.7MM, VTICM5_13.7, -4.50/2.5/160 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGE FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED TO BE THE DEVICE, THE REPORTER STATED " THE DEVICE OVERESTIMATES THE AUTOMAIC WHITE TO WHITE MEASUREMENT DOCTOR DID NOT CHECK IT MANUALLY." Manufacturer Narrative: IT WAS REPORTED THAT A 13.7MM, VTICM5_13.7, -4.50/2.5/169 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTED TO BE BOTH THE DEVICE AND USER ERROR, THE REPORTER STATED "THE DEVICE OVERESTIMATES THE AUTOMATIC WITHE TO WITHE MEASUREMENT DOCTOR DID NOT CHECK IT MANUALLY." CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US . CLAIM # (B)(4). | -4.50/2.5/160 | OD | Physician | 05/10/2021 | 2021 | 46 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11972287&pc=QCB | 2023826-2021-01793 | 2019-12-19 | Injury | STAAR SURGICAL COMPANY | 2021-06-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -16.0/6.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE PROBLEM DID NOT RESOLVE. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR THE SAME LENGTH LENS BUT THE PROBLEM DID NOT RESOLVE. THE REPORTER STATED "AS THE PATIENT HAS THE LENS VAULT OF 650 , WHICH IS ABOUT THE IDEAL LENS VAULT, THE DOCTOR DECIDED NOT TO CHANGE THE LENS LENGTH." CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -16.0/6.0/082 | OS | Physician | 05/11/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11972321&pc=QCB | 2023826-2021-01790 | 2021-04-01 | Injury | STAAR SURGICAL COMPANY | 2021-06-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -16.5/3.5/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH, DIFFERENT MODEL AND DIOPTER LENS DUE TO LOW VAULT WITHOUT ROTATION. THE REPORTER STATED " THE SURGEON DECIDED TO EXCHANGE THE TICL LENS WITH ICL LENS TO AVOID THE ROTATION OF THE LENS, SO THE DIOPTER FOR THE REPLACEMENT LENS CHANGED TO 14.5." IT WAS REPORTED THAT THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.5/3.5/102 | OS | Physician | 05/12/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11986234&pc=QCB | 2023826-2021-01843 | 2020-07-24 | Injury | STAAR SURGICAL COMPANY | 2021-06-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00/+1.0/108 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -09.00/+1.0/108 | OS | Physician | 05/14/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11986322&pc=QCB | 2023826-2021-01860 | 2021-01-12 | Injury | STAAR SURGICAL COMPANY | 2021-06-11 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/+1.5/084 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4): LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.0/+1.5/084 | OS | Physician | 05/18/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11986337&pc=QCB | 2023826-2021-01859 | 2020-09-18 | Injury | STAAR SURGICAL COMPANY | 2021-06-11 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.5/+3.0/098 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. ON 13-APR-2021 THE LENS WAS EXCHANGED WITH A LONGER LENS. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -15.5/+3.0/098 | OS | Physician | 05/18/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11986343&pc=QCB | 2023826-2021-01858 | 2021-03-02 | Injury | STAAR SURGICAL COMPANY | 2021-06-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM # (B)(4). | -12.5/+1.5/090 | OS | Physician | 05/18/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11986349&pc=QCB | 2023826-2021-01852 | 2021-01-10 | Injury | STAAR SURGICAL COMPANY | 2021-06-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+1.0/080 | OD | Physician | 05/16/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11986356&pc=QCB | 2023826-2021-01866 | 2021-04-29 | Injury | STAAR SURGICAL COMPANY | 2021-06-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.5/+4.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -16.5/+4.0/094 | OD | Physician | 05/19/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11986360&pc=QCB | 2023826-2021-01854 | 2021-04-09 | Injury | STAAR SURGICAL COMPANY | 2021-06-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+2.0/090 | OD | Physician | 05/17/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11986371&pc=QCB | 2023826-2021-01856 | 2021-03-27 | Injury | STAAR SURGICAL COMPANY | 2021-06-11 | QCB | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+4.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) UPSIDE DOWN ON (B)(6) 2021. ON (B)(6) 2021 AN ALTERNATE LENS WAS IMPLANTED AND THE PROBLEM IS RESOLVED. REPORTEDLY, NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -12.5/+4.0/094 | OS | Physician | 05/18/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11986382&pc=QCB | 2023826-2021-01869 | 2021-04-27 | Injury | STAAR SURGICAL COMPANY | 2021-06-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -13.5/+1.0/092 | OD | Physician | 05/19/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11986384&pc=QCB | 2023826-2021-01857 | 2021-03-02 | Injury | STAAR SURGICAL COMPANY | 2021-06-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/098 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S.. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -9.5/+2.5/098 | OD | Physician | 05/18/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11986387&pc=QCB | 2023826-2021-01850 | 2021-05-09 | Injury | STAAR SURGICAL COMPANY | 2021-06-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.5/170 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -10.0/+1.5/170 | OS | Physician | 05/17/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11986402&pc=QCB | 2023826-2021-01871 | 2021-04-23 | Injury | STAAR SURGICAL COMPANY | 2021-06-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/080 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -11.5/+2.0/080 | OS | Physician | 05/19/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11987105&pc=QCB | 2023826-2021-01791 | 2021-02-23 | Injury | STAAR SURGICAL COMPANY | 2021-06-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: IT WAS REPORTED THAT A 13.2MM, VTICM5_13.2, -5.50/2.0/011 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGE FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED TO BE THE DEVICE, THE REPORTER STATED "CALCULATION ERROR." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.50/2.0/011 | OD | Physician | 05/14/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11986409&pc=QCB | 2023826-2021-01846 | 2021-03-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER STATES A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS -13.0/+1.5/144 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED ON (B)(6) 2021. REPORTEDLY, THE LENS WAS DEFECTIVE, IMPLANTED THEN REMOVED. A BACKUP LENS WAS USED AND NO PATIENT INJURY OCCURRED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUEDEBRIS ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN AND THERE WAS DEBRIS ON THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.0/+1.5/144 | NA |  | 05/18/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11991430&pc=QCB | 2023826-2021-01899 | 2021-01-12 | Injury | STAAR SURGICAL COMPANY | 2021-06-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.5/1.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS EXCHANGED RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4). | -11.5/1.5/105 | OD | Physician | 05/17/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11996523&pc=QCB | 2023826-2021-01911 | 2020-10-04 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.0/2.0/81 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS WITHIN THE SAME SURGERY DUE TO LOW VAULT. PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -13.0/2.0/81 | OD | Physician | 05/19/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11996524&pc=QCB | 2023826-2021-01908 | 2021-04-02 | Injury | STAAR SURGICAL COMPANY | 2021-06-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.5/1.0/6 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). | -8.5/1.0/6 | OD | Physician | 05/19/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11996525&pc=QCB | 2023826-2021-01905 | 2021-05-10 | Injury | STAAR SURGICAL COMPANY | 2021-06-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.5/2.0/095, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021.ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/2.0/095 | OD | Physician | 05/18/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11996531&pc=QCB | 2023826-2021-01907 | 2021-04-26 | Injury | STAAR SURGICAL COMPANY | 2021-06-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -9.0/1.0/106 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.0/1.0/106 | OD | Physician | 05/19/2021 | 2021 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11996532&pc=QCB | 2023826-2021-01902 | 2021-02-12 | Injury | STAAR SURGICAL COMPANY | 2021-06-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -8.0/3.5/015 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). | -8.0/3.5/015 | OD | Physician | 05/18/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=11996534&pc=QCB | 2023826-2021-01900 | 2021-03-27 | Injury | STAAR SURGICAL COMPANY | 2021-06-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.0/3.5/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.0/3.5/90 | OD | Physician | 05/17/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12005896&pc=QCB | 2023826-2021-01930 | 2020-08-16 | Injury | STAAR SURGICAL COMPANY | 2021-06-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/107 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -10.50/+1.0/107 | OD | Physician | 05/16/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12005306&pc=QCB | 2023826-2021-01929 | 2021-05-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -09.00/+2.5/065 (SPHERE/CYLINDER/AXIS), WAS DAMAGED WHILE LOADING. THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE REPORTER INDICATED THE LENS WAS DAMAGED DUE TO SURGICAL TECH ERROR. THE CAUSE WAS DUE TO USER ERROR. THE PROBLEM WAS RESOLVED WITH THE IMPLANT OF THE BACKUP LENS. H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND A HAPTIC TORN/BROKEN. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.00/+2.5/065 | NA |  | 05/20/2021 | 2021 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12005448&pc=QCB | 2023826-2021-01949 | 2021-05-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00/+2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS LOW VAULTING AND REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED AND REPLACED FOR A LONGER LENS ON (B)(6) 2021. THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE CODE: (B)(4) THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | -14.00/+2.5/091 | OS |  | 05/21/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12005471&pc=QCB | 2023826-2021-01936 | 2021-04-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/104 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.50/+1.0/104 | OD |  | 05/18/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12006593&pc=QCB | 2023826-2021-01919 | 2020-03-27 | Injury | STAAR SURGICAL COMPANY | 2021-06-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.5/1.5/95 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -7.5/1.5/95 | OS | Physician | 05/19/2021 | 2021 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12006603&pc=QCB | 2023826-2021-01918 | 2020-03-27 | Injury | STAAR SURGICAL COMPANY | 2021-06-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.5/1.0/81 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -7.5/1.0/81 | OD | Physician | 05/19/2021 | 2021 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12006622&pc=QCB | 2023826-2021-01920 | 2021-03-07 | Injury | STAAR SURGICAL COMPANY | 2021-06-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -7.0/1.0/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -7.0/1.0/102 | OD | Physician | 05/19/2021 | 2021 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12009101&pc=QCB | 2023826-2021-01879 | 2021-04-08 | Injury | STAAR SURGICAL COMPANY | 2021-06-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.50/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.50/1.5/095 | OD | Physician | 05/23/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12009104&pc=QCB | 2023826-2021-01897 | 2021-05-10 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 12.6MM VTICMO12.6, -14.00/1.5/091 DIOPTER, IMPLANTABLE COLLAMER LENS INTO PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS STUCK IN CARTRIDGE/INJECTION SYSTEM DURING INJECTION. THERE WAS PATIENT CONTACT WITH NO PATIENT INJURY. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2021, THERE ARE NO PLANS TO IMPLANT AN OTHER LENS. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, DATE REC¿D BY MFR- THIS PRODUCT IS NOT MARKETED IN THE US, WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -14.00/1.5/091 | OD | Physician | 05/28/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12009106&pc=QCB | 2023826-2021-01888 | 2021-03-02 | Injury | STAAR SURGICAL COMPANY | 2021-06-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.00/4.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.00/4.0/095 | OS | Physician | 05/26/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12009133&pc=QCB | 2023826-2021-01889 | 2020-10-31 | Injury | STAAR SURGICAL COMPANY | 2021-06-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.00/1.5/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.00/1.5/098 | OS | Physician | 05/26/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12009141&pc=QCB | 2023826-2021-01880 | 2021-05-17 | Injury | STAAR SURGICAL COMPANY | 2021-06-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.00/4.0/008 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -18.00/4.0/008 | OD | Physician | 05/26/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12009178&pc=QCB | 2023826-2021-01892 |  | Injury | STAAR SURGICAL COMPANY | 2021-06-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/4.0/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.00/4.0/108 | OD | Physician | 05/27/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12013525&pc=QCB | 2023826-2021-01935 | 2021-03-21 | Injury | STAAR SURGICAL COMPANY | 2021-06-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/+0.5/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO THE LENS WAS IMPLANTED UPSIDE DOWN. THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.00/+0.5/091 | OS | Physician | 05/18/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12013597&pc=QCB | 2023826-2021-01947 | 2021-05-01 | Injury | STAAR SURGICAL COMPANY | 2021-06-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.00/+1.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -07.00/+1.5/088 | OD | Physician | 05/19/2021 | 2021 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12006571&pc=QCB | 2023826-2021-01903 | 2021-04-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -7.5/6.0/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE, INSIDE A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.5/6.0/073 | OD | Physician | 05/18/2021 | 2021 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12009113&pc=QCB | 2023826-2021-01881 | 2021-05-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -15.50/1.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED VERTICALLY WITH A SAME LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY. VISUAL INSPECTION FOUND VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.50/1.5/086 | OS | Physician | 05/26/2021 | 2021 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12009138&pc=QCB | 2023826-2021-01891 | 2021-05-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 13.2MM VTICM5_13.2, -10.00/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS STUCK IN CARTRIDGE/INJECTION SYSTEM DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. REPLACEMENT LENS IS PLANNED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, IRITIS, INFLAMMATION OF THE EYE IS OBSERVED. PER THE REPORTER ON (B)(6) 2021, INFLAMMATION IN THE EYE IS DUE TO MANIPULATION DURING SURGERY. IT IS BEING TREATED WITH ANTI-INFLAMMATORY. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED, THAT THE SURGEON INTRAOPERATIVELY IMPLANTED AND REMOVED A 13.2MM VTICM5, 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.00/1.5/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS GOT STUCK IN THE CARTRIDGE/INJECTION SYSTEM, DURING INJECTION. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. A REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2021. THE REPORTER STATED, THAT "THE LENS WAS NOT IMPLANTED, DUE TO BEING STUCK IN THE CARTRIDGE OR INJECTOR SYSTEM". CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, IRITIS, INFLAMMATION OF THE EYE IS OBSERVED. PER THE REPORTER, ON (B)(6) 2021, INFLAMMATION IN THE EYE IS, DUE TO MANIPULATION DURING SURGERY. IT IS BEING TREATED WITH ANTI-INFLAMMATORY. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.00/1.5/095 | OS |  | 05/27/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12015794&pc=QCB | 2023826-2021-01933 | 2021-05-05 | Injury | STAAR SURGICAL COMPANY | 2021-06-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+1.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -13.00/+1.5/093 | OD | Physician | 05/18/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12021853&pc=QCB | 2023826-2021-02004 | 2021-02-20 | Injury | STAAR SURGICAL COMPANY | 2021-06-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.0/+1.0/006 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021 . ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+1.0/006 | OS | Physician | 05/26/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12021856&pc=QCB | 2023826-2021-02009 | 2021-05-28 | Injury | STAAR SURGICAL COMPANY | 2021-06-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.5/+1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM # (B)(4). | -6.5/+1.0/090 | OD | Physician | 05/28/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12021857&pc=QCB | 2023826-2021-02008 | 2021-03-17 | Injury | STAAR SURGICAL COMPANY | 2021-06-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -13.0/+2.0/089 | OS | Physician | 05/28/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12021858&pc=QCB | 2023826-2021-02000 | 2021-04-08 | Injury | STAAR SURGICAL COMPANY | 2021-06-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.5/+3.5/076 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THEN REPLACED WITH A LARGER, DIFFERENT MODEL LENS DUE TO LENS ROTATION. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -12.5/+3.5/076 | OS | Physician | 05/27/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12021860&pc=QCB | 2023826-2021-02002 | 2021-05-10 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+3.0/025 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -14.5/+3.0/025 | OD | Physician | 05/27/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12021869&pc=QCB | 2023826-2021-01997 | 2021-04-15 | Injury | STAAR SURGICAL COMPANY | 2021-06-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.5/+3.5/001 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THEN REPLACED WITH A SAME SIZE DIFFERENT MODEL LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -9.5/+3.5/001 | OD | Physician | 05/27/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12021874&pc=QCB | 2023826-2021-01995 | 2020-02-27 | Injury | STAAR SURGICAL COMPANY | 2021-06-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+0.5/101 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "ICL WAS TOO LONG DESPITE STAAR SIZING RECOMMENDATION." Manufacturer Narrative: SEX, WEIGHT, ETHNICITY - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -17.5/+0.5/101 | OD | Physician | 05/26/2021 | 2021 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12021876&pc=QCB | 2023826-2021-02003 | 2021-03-03 | Injury | STAAR SURGICAL COMPANY | 2021-06-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.5/+0.5/029 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM# (B)(4). | -11.5/+0.5/029 | OD | Physician | 05/26/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12015790&pc=QCB | 2023826-2021-01934 | 2021-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+1.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN MICROCENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND ONE HAPTIC TORN AND BROKEN. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM #: (B)(4). Manufacturer Narrative: B5: PER EMAIL ON (B)(6) 2021, THE REPORTER STATED, THAT "THE SURGEON HAS NO WILLINGNESS TO IMPLANT ANOTHER LENS". CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.00/+1.0/092 | OS |  | 05/18/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12015801&pc=QCB | 2023826-2021-01937 | 2021-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/076 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -09.50/+1.0/076 | OS |  | 05/19/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12015807&pc=QCB | 2023826-2021-01950 | 2021-05-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.50/+4.0/106 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS LOW VAULTING AND REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS EXPLANTED AND REPLACED ON (B)(6) 2021 FOR A LONGER LENS. THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE CODE: (B)(4) THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | -15.50/+4.0/106 | OD |  | 05/21/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12021866&pc=QCB | 2023826-2021-01993 | 2021-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER STATES A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS -11.0/+2.0/071 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2021. REPORTEDLY, THE LENS WAS REMOVED ON THE SAME DATE AND REPLACED WITH AN ALTERNATE LENS BECAUSE THE PATIENT NEEDED A DIFFERENT POWER. Manufacturer Narrative: PLEASE DISREGARD THE CURRENT CLAIM AS IT IS A DUPLICATE OF MFR REP# 2023826-2021-01673. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -11.0/+2.0/071 | NA |  | 05/27/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12021870&pc=QCB | 2023826-2021-01990 | 2021-02-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -14.0/+1.5/078 (SPHERE/CYLINDER/AXIS) TORE DURING INSERTION. THERE WAS NO PATIENT INJURY AND A BACKUP LENS WAS IMPLANTED. THIS OCCURRED ON (B)(6) 2021. Manufacturer Narrative: H3. DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE ON THE LENS SURFACE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. CLAIM# (B)(4). Manufacturer Narrative: B5-CORRECTED INFO: REPORTEDLY THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY. THE ICL FOOTPLATE WAS STUCK UNDER THE PLUNGER OF THE INJECTOR SO IT GOT TORN OFF. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. EXPLANT DATE: UNK. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -14.0/+1.5/078 | NA |  | 05/24/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12029356&pc=QCB | 2023826-2021-01922 | 2021-05-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -15.00/1.5/141 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REMOVED BECAUSE IT GOT STUCK TO PLUNGER AND DOCTOR FELT IT MIGHT HAVE DAMAGED IT. NO PATIENT INJURY. USED BACKUP. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATE THAT THE A 13.7MM, TMICL13.7, -15.00/1.5/141 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED AND REPLACED WITHIN THE SAME SURGERY DUE TO LENS STUCK IN CARTRIDGE OR INJECTION SYSTEM. THERE WAS NO PATIENT INJURY AND PROBLEM RESOLVED. IN THE REPORTER'S OPINION THE DEVICE WAS THE CAUSE OF THE EVENT, FOR CAUSE DETAILS THE REPORTER STATED "LENS DID NOT RELEASE FROM PLUNGER". H6- HEALTH EFFECT- IMPACT CODE: 4627 TO 2199- PREVIOUS CODE NOT APPLICABLE, PATIENT DID NOT REQUIRE A CORRECTIVE PROCEDURE TO REMOVE LENS. H6- MEDICAL DEVICE PROBLEM CODE: 2993 TO 3189- PREVIOUS CODE NOT APPLICABLE, NO ADVERSE EVENT OCCURRED. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | -15.00/1.5/141 | NA | Physician | 05/19/2021 | 2021 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12030069&pc=QCB | 2023826-2021-02011 | 2020-10-29 | Injury | STAAR SURGICAL COMPANY | 2021-06-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -13.00/+3.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPORTED AS HAVING ROTATED ON (B)(6) 2020 AND WAS RE-POSITIONED ON (B)(6) 2020. THE RE-POSITIONING RESOLVED THE PROBLEM. THE PATIENT IS HAPPY AND DOING WELL SINCE TORIC ICL ROTATION. THE LENS REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2020-02281. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. EXPLANT DATE: NA. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.00/+3.5/094 | OS | Other Health Care Professional | 01/13/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12029789&pc=QCB | 2023826-2021-01924 | 2021-03-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -18.00/0.5/109(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. A REPLACEMENT LENS OF SHORTER LENGTH AND DIFFERENT MODEL (SPHERICAL) WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2 -18.00/0.5/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGE FOR A SPHERICAL LENS OF SHORTER LENGTH DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -18.00/0.5/109 | OD |  | 05/21/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12029790&pc=QCB | 2023826-2021-01923 | 2021-03-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: TH REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -17.50/0.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. A REPLACEMENT LENS OF SHORTER LENGTH AND DIFFERENT MODEL (SPHERICAL) WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -17.50/0.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS OF SHORTER LENGTH DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -17.50/0.5/087 | OS |  | 05/21/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12053771&pc=QCB | 2023826-2021-02048 | 2021-05-06 | Injury | STAAR SURGICAL COMPANY | 2021-06-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -05.50/+2.0/112 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -05.50/+2.0/112 | OS | Physician | 05/26/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12053790&pc=QCB | 2023826-2021-02049 | 2021-04-26 | Injury | STAAR SURGICAL COMPANY | 2021-06-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM # (B)(4). | -10.00/+1.0/089 | OD | Physician | 05/26/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12053804&pc=QCB | 2023826-2021-02056 | 2021-02-25 | Injury | STAAR SURGICAL COMPANY | 2021-06-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.50/+2.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM # (B)(4). | -13.50/+2.0/088 | OD | Physician | 05/27/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12053852&pc=QCB | 2023826-2021-02058 | 2021-05-20 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.00/+1.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REPORTED AS LOW VAULTING AND REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE REPORTER INDICATED THE PATIENT WILL BE MONITORED VERY CLOSELY WITHIN THE NEXT HALF YEAR BECAUSE PATIENT IS VERY HAPPY WITH THE VISION. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.00/+1.0/094 | OD | Physician | 05/27/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12055770&pc=QCB | 2023826-2021-02014 | 2021-01-29 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -6.0/0.5/011, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -6.0/0.5/011 | OS | Physician | 05/25/2021 | 2021 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12055818&pc=QCB | 2023826-2021-02013 | 2021-05-21 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A13.2MM, VTICM5_13.2, -8.0/2.5/094 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. IT WAS REPORTED THAT THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -8.0/2.5/094 | OD | Physician | 05/25/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060460&pc=QCB | 2023826-2021-01956 | 2021-05-05 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/1.0/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.00/1.0/110 | OD | Physician | 05/31/2021 | 2021 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060473&pc=QCB | 2023826-2021-01963 | 2021-05-24 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.00/2.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -9.00/2.5/088 | OD | Physician | 06/01/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060482&pc=QCB | 2023826-2021-01957 | 2021-05-05 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/2.0/063 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -10.50/2.0/063 | OS | Physician | 05/31/2021 | 2021 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060485&pc=QCB | 2023826-2021-01974 | 2021-03-10 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/4.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -11.00/4.5/089 | OD | Physician | 06/02/2021 | 2021 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060488&pc=QCB | 2023826-2021-01964 | 2021-05-24 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.50/2.0/093 | OD | Physician | 06/01/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060489&pc=QCB | 2023826-2021-01969 | 2021-05-19 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.00/4.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.00/4.5/091 | OS | Physician | 06/02/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060497&pc=QCB | 2023826-2021-01984 | 2021-01-24 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.50/2.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K- THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -17.50/2.5/081 | OS | Physician Assistant | 06/03/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060499&pc=QCB | 2023826-2021-01958 | 2020-12-10 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -6.50/2.5/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -6.50/2.5/078 | OD | Physician | 06/01/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060509&pc=QCB | 2023826-2021-01959 | 2020-12-10 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -5.50/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. PMA/510K- THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -5.50/2.0/090 | OD | Physician | 06/01/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060518&pc=QCB | 2023826-2021-01972 | 2021-01-14 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/2.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED IN A SECOND SURGERY WITH A SHORTER LENGTH LENS ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.50/2.0/086 | OD | Physician | 06/02/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060528&pc=QCB | 2023826-2021-01971 | 2021-02-22 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.00/1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2021 DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K- THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4. CLAIM# (B)(4). | -10.00/1.0/092 | OS | Physician | 06/02/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060546&pc=QCB | 2023826-2021-01970 | 2021-02-09 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED IN A SECOND SURGERY ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -10.50/2.0/088 | OD | Physician | 06/02/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060547&pc=QCB | 2023826-2021-01975 | 2021-03-10 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.50./4.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.50./4.5/093 | OS | Physician | 06/02/2021 | 2021 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12064150&pc=QCB | 2023826-2021-02047 | 2021-05-13 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.00/+1.5/031 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS SHOULD BE AT 14 DEGREES BUT AFTER A WEEK THE LENS ROTATED TO 70 DEGREES. THE LENS WAS RE-POSITIONED AND AFTER THE SURGERY, THE LENS ROTATED AGAIN. THE LENS WAS EXPLANTED ON (B)(6) 2021. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS ASYMPTOMATIC. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -07.00/+1.5/031 | OS | Physician | 05/26/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12064160&pc=QCB | 2023826-2021-02111 | 2021-04-13 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -06.50/+1.0/112 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO THE LENS TORE UPON INSERTION INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. THE LENS WAS DISCARDED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -06.50/+1.0/112 | OD | Other Health Care Professional | 05/26/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12064259&pc=QCB | 2023826-2021-02059 | 2021-05-20 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.50/+1.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REPORTED AS LOW VAULTING AND REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE REPORTER INDICATED THE PATIENT WILL BE MONITORED VERY CLOSELY WITHIN THE NEXT HALF YEAR BECAUSE PATIENT IS VERY HAPPY WITH THE VISION. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.50/+1.5/074 | OS | Physician | 05/27/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12064270&pc=QCB | 2023826-2021-02112 | 2021-05-13 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -05.00/+1.5/113 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION. THE LENS WAS REPOSITIONED ON (B)(6)2021 BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -05.00/+1.5/113 | OD | Physician | 06/23/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060486&pc=QCB | 2023826-2021-01965 | 2021-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO13.2, -7.50/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -7.50/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -7.50/1.5/091 | OD |  | 06/01/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060495&pc=QCB | 2023826-2021-01979 | 2021-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -6.50/3.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2021 DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -6.50/3.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: B5- CORRECTED "LONGER LENGTH LENS/LOW VAULT" TO "SHORTER LENGTH LENS/EXCESSIVE VAULT." CLAIM# (B)(4). | -6.50/3.5/081 | OS | Physician | 06/03/2021 | 2021 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060510&pc=QCB | 2023826-2021-01982 | 2021-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -18.00/3.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION NOT REQUIRED. CORRECTED DATA: D6B- CORRECTED TO (B)(6) 2021. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -18.00/3.5/092 | OD | Physician Assistant | 06/03/2021 | 2021 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12060537&pc=QCB | 2023826-2021-01962 | 2021-04-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -13.00/3.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -13.00/3.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -13.00/3.0/091 | OD |  | 06/01/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12064161&pc=QCB | 2023826-2021-02061 | 2021-05-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -07.50/+2.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINED IMPLANTED. AT THIS TIME NO SURGICAL INTERVENTION IS PLANNED. THE PATIENT WILL BE STRICTLY MONITORED AND IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6 - WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -07.50/+2.0/088 | OS |  | 05/28/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12064271&pc=QCB | 2023826-2021-02060 | 2021-05-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -07.50/+2.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE SURGEON PLANS NO SURGICAL INTERVENTION AT THIS TIME AND THE PAITENT WILL BE STRICTLY MONITORED. THE PATIENT IS HAPPY. Manufacturer Narrative: H6 - WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -07.50/+2.0/087 | OD |  | 05/28/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072237&pc=QCB | 2023826-2021-02093 | 2021-01-31 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/018 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -9.5/+1.0/018 | OS | Physician | 06/03/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072240&pc=QCB | 2023826-2021-02107 | 2020-11-02 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+3.5/075 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -12.0/+3.5/075 | OS | Physician | 06/04/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072241&pc=QCB | 2023826-2021-02108 | 2020-08-20 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+2.5/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED AND THEN REPLACED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM # (B)(4). | -10.0/+2.5/072 | OD | Physician | 06/04/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072242&pc=QCB | 2023826-2021-02077 | 2021-05-14 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -16.0/+2.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -16.0/+2.0/084 | OS | Physician | 05/31/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072244&pc=QCB | 2023826-2021-02101 | 2021-05-10 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+2.5/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM # (B)(4). | -8.5/+2.5/083 | OD | Physician | 06/04/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072246&pc=QCB | 2023826-2021-02106 | 2020-03-18 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.5/+3.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED AND THEN REPLACED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S.. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -14.5/+3.0/090 | OD | Physician | 06/04/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072251&pc=QCB | 2023826-2021-02100 | 2020-07-31 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.0/+1.0/105 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -6.0/+1.0/105 | OD | Physician | 06/04/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072253&pc=QCB | 2023826-2021-02098 | 2021-04-25 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.0/+4.0/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4): OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -13.0/+4.0/085 | OS | Physician | 06/03/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072257&pc=QCB | 2023826-2021-02082 | 2020-07-08 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+2.0/094 | OD | Physician | 06/01/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072259&pc=QCB | 2023826-2021-02091 | 2020-12-02 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM# (B)(4). | -10.5/+2.0/090 | OD | Physician | 06/03/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072261&pc=QCB | 2023826-2021-02074 | 2020-12-21 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/083 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -12.0/+2.0/083 | OS | Physician | 06/01/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072268&pc=QCB | 2023826-2021-02076 | 2021-05-14 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -16.0/+2.5/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -16.0/+2.5/084 | OD | Physician | 05/31/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072277&pc=QCB | 2023826-2021-02092 | 2021-01-31 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/173 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -9.0/+1.0/173 | OD | Physician | 06/03/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072238&pc=QCB | 2023826-2021-02084 | 2020-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+4.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT MODEL LENS DUE TO LOW VAULT. REFERENCE MFR FILE# (B)(4) FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: ADDITIONAL INFORMATION: THE REPORTER STATES, THE REASON FOR THE LENS REPLACEMENT WAS "LOW ARCH AND HIGH". THE PATIENT WAS CONCERNED WITH THE RISK OF SURGERY. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+4.0/087 | OD | Physician | 06/01/2021 | 2021 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072265&pc=QCB | 2023826-2021-02073 | 2021-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT AXIS LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHTIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -10.0/+1.0/090 | OS |  | 06/01/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072270&pc=QCB | 2023826-2021-02080 | 2020-05-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-06-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.0/+4.0/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#:(B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS DIOPTER -13.0/4.0/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS LATER EXCHANGED ON (B)(6) 2020 WITH A LONGER LENGTH LENS DUE TO LOW VAULT. PATIENT HAS A PRE-EXISTING CONDITION OF IRIS TREMORS. PROBLEM WAS NOT RESOLVED. REFERENCE MFR# (2023826-2021-02084) FOR REPLACEMENT LENS CLAIM. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4): OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -13.0/+4.0/085 | OD | Physician | 06/01/2021 | 2021 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072312&pc=QCB | 2023826-2021-02025 | 2021-04-27 | Injury | STAAR SURGICAL COMPANY | 2021-06-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -17.0/3.0/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -17.0/3.0/101 | OS | Physician | 05/26/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072335&pc=QCB | 2023826-2021-02020 | 2021-01-20 | Injury | STAAR SURGICAL COMPANY | 2021-06-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -11.5/2.5/97, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4): OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -11.5/2.5/97 | OS | Physician | 05/27/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072336&pc=QCB | 2023826-2021-02019 | 2021-01-27 | Injury | STAAR SURGICAL COMPANY | 2021-06-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -10.5/2.5/98, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -10.5/2.5/98 | OD | Physician | 05/27/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072471&pc=QCB | 2023826-2021-02021 | 2021-05-18 | Injury | STAAR SURGICAL COMPANY | 2021-06-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.0/1.5/91, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -13.0/1.5/91 | OD | Physician | 05/27/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12072505&pc=QCB | 2023826-2021-02016 | 2021-05-16 | Injury | STAAR SURGICAL COMPANY | 2021-06-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.0/2.0/87, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -10.0/2.0/87 | OD | Physician | 05/28/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12073018&pc=QCB | 2023826-2021-02024 | 2021-05-19 | Injury | STAAR SURGICAL COMPANY | 2021-06-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.5/1.5/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.5/1.5/90 | OD | Physician | 05/26/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12073137&pc=QCB | 2023826-2021-02034 | 2021-03-22 | Injury | STAAR SURGICAL COMPANY | 2021-06-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -3.50/1.50/5, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN, "TOO VAULT FOR THE LENGTH THAT OCOS INDICATED." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -3.50/1.50/5 | OD | Physician | 05/28/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12084581&pc=QCB | 2023826-2021-02167 | 2020-07-16 | Injury | STAAR SURGICAL COMPANY | 2021-06-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -08.00/+2.0/115 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -08.00/+2.0/115 | OD | Physician | 05/31/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12084844&pc=QCB | 2023826-2021-02178 | 2021-01-19 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/081 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.00/+2.0/081 | NA | Physician | 06/07/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12084855&pc=QCB | 2023826-2021-02170 | 2021-05-03 | Injury | STAAR SURGICAL COMPANY | 2021-06-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.00/+2.0/094 | OD | Physician | 06/01/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12084879&pc=QCB | 2023826-2021-02171 | 2021-05-10 | Injury | STAAR SURGICAL COMPANY | 2021-06-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.00/+2.0/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.00/+2.0/100 | OD | Physician | 06/01/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12084893&pc=QCB | 2023826-2021-02172 | 2021-04-17 | Injury | STAAR SURGICAL COMPANY | 2021-06-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.00/+1.5/093 | OD | Physician | 06/01/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12093381&pc=QCB | 2023826-2021-02173 | 2021-04-17 | Injury | STAAR SURGICAL COMPANY | 2021-06-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/100 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -11.50/+1.5/100 | OS | Physician | 06/01/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12096890&pc=QCB | 2023826-2021-02186 | 2020-12-23 | Injury | STAAR SURGICAL COMPANY | 2021-06-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -04.50/+1.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -04.50/+1.0/091 | OS | Physician | 06/03/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12096891&pc=QCB | 2023826-2021-02180 | 2020-11-25 | Injury | STAAR SURGICAL COMPANY | 2021-06-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -04.50/+1.0/108 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM # (B)(4). | -04.50/+1.0/108 | OS | Physician | 06/03/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12096892&pc=QCB | 2023826-2021-02181 | 2020-12-03 | Injury | STAAR SURGICAL COMPANY | 2021-06-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+3.0/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM # (B)(4). | -15.00/+3.0/082 | OD | Physician | 06/03/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12096922&pc=QCB | 2023826-2021-02182 | 2020-11-23 | Injury | STAAR SURGICAL COMPANY | 2021-06-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -11.50/+1.5/092 | OS | Physician | 06/03/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12097192&pc=QCB | 2023826-2021-02179 | 2020-11-04 | Injury | STAAR SURGICAL COMPANY | 2021-06-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -04.50/+4.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. . CLAIM # (B)(4). | -04.50/+4.0/088 | OS | Physician | 06/03/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12097196&pc=QCB | 2023826-2021-02177 | 2021-04-30 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -18.00/+3.0/070 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED DURING THE SAME SURGERY WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE, THE LENS GOT STUCK WITH PLUNGER AND DAMAGED DURING IMPLANTATION. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.00/+3.0/070 | OS | Physician | 06/02/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12097204&pc=QCB | 2023826-2021-02199 | 2021-02-26 | Malfunction | STAAR SURGICAL COMPANY | 2021-06-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+4.0/092 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED FOR A SHORTER LENS ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM # (B)(4). | -09.00/+4.0/092 | NA | Physician | 06/07/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12097206&pc=QCB | 2023826-2021-02193 | 2020-08-15 | Injury | STAAR SURGICAL COMPANY | 2021-06-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+2.5/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM # (B)(4). | -09.50/+2.5/096 | OS | Physician | 06/04/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12097208&pc=QCB | 2023826-2021-02192 | 2020-06-09 | Injury | STAAR SURGICAL COMPANY | 2021-06-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00/+1.0/067 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM # (B)(4). | -08.00/+1.0/067 | OS | Physician | 06/04/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102439&pc=QCB | 2023826-2021-02117 | 2020-11-14 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICH12.6, -2.00/2.5/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, EVERYTHING IS OK RIGHT NOW. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -2.00/2.5/096 | OS | Physician | 06/06/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102444&pc=QCB | 2023826-2021-02119 | 2021-01-28 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -5.50/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -5.50/1.5/091 | OS | Physician | 06/07/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102499&pc=QCB | 2023826-2021-02114 | 2021-04-05 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.00/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.00/1.0/093 | OD | Physician | 06/06/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102502&pc=QCB | 2023826-2021-02126 | 2021-01-15 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -5.50/2.0/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -5.50/2.0/074 | OD | Physician | 06/07/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102506&pc=QCB | 2023826-2021-02116 | 2021-01-04 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED UPSIDE DOWN A 13.2MM VTICMO13.2, -11.00/1.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, INTO THE PATIENT'S RIGHT EYE ON (B)(6) 2021. THE LENS TORE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REPLACED INTRAOPERATIVELY ON (B)(6) 2021 WITH A SHORTER LENGTH , SIMILAR (SPHERE/CYLINDER/AXIS) LENS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US.. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.00/1.5/088 | OD | Physician | 06/07/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102539&pc=QCB | 2023826-2021-02113 | 2021-04-05 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.50/1.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -8.50/1.0/094 | OS | Physician | 06/06/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102546&pc=QCB | 2023826-2021-02136 | 2021-06-07 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -6.50/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LENS OPACITY (ACS) WAS OBSERVED ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021. REPORTEDLY, MEDIUM HEAVY ASC WHICH AFFECTING VISUAL ACUITY. CLE SURGERY IS PLANNED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -6.50/1.5/089 | OS | Physician | 06/08/2021 | 2021 | 56
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102572&pc=QCB | 2023826-2021-02128 | 2021-01-27 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.00/2.0/088 | OD | Physician | 06/07/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102693&pc=QCB | 2023826-2021-02148 | 2021-02-20 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.00/2.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -8.00/2.0/094 | OS | Physician | 06/09/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102700&pc=QCB | 2023826-2021-02152 | 2021-04-20 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6, -14.50/2.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -14.50/2.0/084 | OD | Physician | 06/09/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102713&pc=QCB | 2023826-2021-02160 | 2021-03-21 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -12.00/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.00/2.5/093 | OD | Physician | 06/11/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102730&pc=QCB | 2023826-2021-02144 | 2020-08-25 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -17.00/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2020 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: (B)(4): OFF-LABEL USE, ACD<3.0MM. CLAIM# (B)(4). | -17.00/2.5/092 | OS | Physician | 06/09/2021 | 2021 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102741&pc=QCB | 2023826-2021-02158 | 2021-01-05 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -10.00/1.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED WITH A SHORTER LENGTH LENS INTRAOPERATIVELY ON (B)(6) 2021 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.00/1.5/087 | OS | Physician Assistant | 06/09/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102775&pc=QCB | 2023826-2021-02143 | 2020-06-13 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.00/1.5/003 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.00/1.5/003 | OD | Physician | 06/09/2021 | 2021 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102795&pc=QCB | 2023826-2021-02153 | 2021-01-12 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -8.50/2.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME SIZE, SIMILAR (SPHERE/CYLINDER/AXIS) LENS., THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -8.50/2.0/092 | OD | Physician | 06/09/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102812&pc=QCB | 2023826-2021-02151 | 2020-12-28 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -9.50/3.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.50/3.0/090 | OD | Physician | 06/09/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12103223&pc=QCB | 2023826-2021-02205 | 2020-12-24 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -7.0/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE CODE (B)(4): OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -7.0/1.5/095 | OS | Physician | 06/01/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12103246&pc=QCB | 2023826-2021-02222 | 2020-08-04 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -11.5/1.5/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS LATER EXPLANTED DUE TO EXCESSIVE VAULT. ON (B)(6) 2020 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED. UNKNOWN WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -11.5/1.5/108 | OD | Physician | 06/04/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12103290&pc=QCB | 2023826-2021-02203 | 2020-12-24 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -7.0/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -7.0/1.0/091 | OD | Physician | 06/01/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102515&pc=QCB | 2023826-2021-02118 | 2021-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.00/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: G4. EXP DATE: 30-JUN-2023. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -7.00/1.5/092 | OD |  | 06/07/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102523&pc=QCB | 2023826-2021-02123 | 2021-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND DISLOCATION/SUBLUXATION WAS OBSERVED. THE LENS WAS REPOSITIONED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS REPLACE WITH A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND DISLOCATION/SUBLUXATION. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: B4- CORRECTED TO 07-JUN-2021. G3- CORRECTED TO 07-JUN-2021 IN INITIAL MDR. G4- THIS PRODUCT IS NOT MARKETED IN THE US. ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS WITH DIOPTER -10.5/+2.0/087 WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2021 WITH A LENS OF DIFFERENT TORIC MODEL, SIMILAR POWER, AND LONGER LENGTH DUE TO LENS ROTATION NOT ASSOCIATED WITH A LOW VAULT AND DISLOCATION/SUBLUXATION. IT IS UNCLEAR WHETHER THIS WAS DONE WITHIN THE SAME PROCEDURE. PER EMAIL ON 10-DEC-2021 IT WAS REPORTED THAT "THE TASS WAS OBSERVED IN THE REPLACEMENT LENS WHICH SUBSIDED WITH THE MEDICAL THERAPY." POST OPERATIVELY, IT WAS REPORTED THAT THE "PATIENT IS DOING FINE WITH ADEQUATE VAULT AND NO ROTATION POST SURGERY". Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -10.5/+2.0/087 | OS | Physician | 06/07/2021 | 2021 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102528&pc=QCB | 2023826-2021-02127 | 2021-01-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.50/2.0/178 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.50/2.0/178 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. DUE TO LOW VAULT THE LENS WAS EXCHANGED ON (B)(6) 2021 FOR A LENS OF SAME MODEL, POWER, AND LENGTH BUT WITH OPPOSITE TORIC AXIS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNKNOWN. MANUFACTURER INFORMATION: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -11.50/2.0/178 | OD | Physician | 06/07/2021 | 2021 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102757&pc=QCB | 2023826-2021-02141 | 2020-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.00/4.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -16.00/4.0/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -16.00/4.0/087 | OD | Physician | 06/09/2021 | 2021 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102763&pc=QCB | 2023826-2021-02142 | 2020-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.00/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE: 2923 - LENS ROTATION. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.00/2.5/091 | OS | Physician | 06/09/2021 | 2021 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102802&pc=QCB | 2023826-2021-02155 | 2021-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -1.00/5.5/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SAME SIZE, DIFFERENT POWER (SPHERE/CYLINDER/AXIS) LENS. DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS WRONG VAULT CALCULATION. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTICT AND HAPTIC TORN, WITH CELLOPHANE ON LENS. CLAIM# (B)(4). Manufacturer Narrative: B5 - CORRECTION: THE LENS WAS REPLACED WITH A SHORTER LENGTH, DIFFERENT POWER (SPHERE/AXIS/CYLINDER) LENS DUE TO EXCESSIVE VAULT ON 21/JAN/2021. CLAIM# (B)(4). Manufacturer Narrative: H3: DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -1.00/5.5/098 | OD | Physician | 06/09/2021 | 2021 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102808&pc=QCB | 2023826-2021-02159 | 2020-12-03 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -13.0/1.5/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED WITH A SHORTER LENGTH LENS INTRAOPERATIVELY ON (B)(6) 2020 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H1- CHANGED "SERIOUS INJURY" TO "MALFUNCTION" H6- HEALTH EFFECT-IMPACT CODE: CHANGED TO 2199. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.0/1.5/101 | OS | Physician Assistant | 06/09/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12102818&pc=QCB | 2023826-2021-02157 | 2020-12-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -7.50/1.0/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED WITH A SHORTER LENGTH LENS INTRAOPERATIVELY ON (B)(6) 2020 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H1- CHANGED "SERIOUS INJURY" TO "MALFUNCTION". H6- HEALTH EFFECT-IMPACT CODE: CHANGED TO 2199. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: (B)(4) OFF-LABEL USE, ACD<3.0MM. CLAIM# (B)(4). | -7.50/1.0/110 | OD | Physician | 06/09/2021 | 2021 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12103165&pc=QCB | 2023826-2021-02213 | 2020-09-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2 MM, VTICMO13.2, -14.5/2.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. PATIENT RELATED FACTOR WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: B5: PER UPDATED INFORMATION CAUSE OF THIS EVENT IS UNKNOWN. CLAIM #(B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.5/2.0/082 | OS |  | 06/03/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12103226&pc=QCB | 2023826-2021-02211 | 2020-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2 MM, VTICMO13.2, -10.5/2.0/86 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. PATIENT RELATED FACTOR WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: B5: PER UPDATED INFORMATION CAUSE OF THIS EVENT IS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/2.0/86 | OS |  | 06/03/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12103253&pc=QCB | 2023826-2021-02215 | 2020-11-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.5/1.5/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. PATIENT RELATED FACTOR WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: PER UPDATED INFORMATION CAUSE OF THIS EVENT IS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -10.5/1.5/107 | OS |  | 06/03/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12103254&pc=QCB | 2023826-2021-02217 | 2020-11-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -8.5/2.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. PATIENT RELATED FACTOR WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE CODE (B)(4): OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -8.5/2.0/087 | OS |  | 06/03/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12103255&pc=QCB | 2023826-2021-02219 | 2020-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.0/1.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER SHORTER LENS DUE TO EXCESSIVE VAULT . THIS EXCHANGE RESOLVED THE PROBLEM. PATIENT RELATED FACTOR WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: B5: PER UPDATED INFORMATION CAUSE OF THIS EVENT IS UNKNOWN. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.0/1.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF THIS EVENT IS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -8.0/1.5/105 | OD | Physician | 06/03/2021 | 2021 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12104732&pc=QCB | 2023826-2021-02221 | 2021-05-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.0/2.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THE REPORTER INDICATED THAT THE LENS GOT STUCK BETWEEN THE CARTRIDGE AND THE FOAM TIP. THERE WAS NO PATIENT INJURY. Manufacturer Narrative: B5- PER UPDATED INFORMATION A REPLACEMENT LENS WAS IMPLANTED ON (B)(6) 2021 AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. FOR CAUSE DETAILS THE REPORTER STATED " LENS LOADING WAS PERFORMED AS USUAL, BUT LENS GOT STUCK IN THE CARTRIDGE WHILE INJECTING IN THE EYE". CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.0/2.5/083 | OS | Physician | 06/02/2021 | 2021 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12108801&pc=QCB | 2023826-2021-02197 | 2020-05-29 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/172 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING IN THE PATIENTS LEFT EYE (OS). THE LENS WAS REPLACED WITHIN THE SAME SURGERY WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.00/+1.5/172 | OS | Physician | 06/04/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12108833&pc=QCB | 2023826-2021-02184 | 2020-12-17 | Injury | STAAR SURGICAL COMPANY | 2021-07-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50/+1.0/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -08.50/+1.0/083 | OS | Physician | 06/03/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12108871&pc=QCB | 2023826-2021-02183 | 2020-12-16 | Injury | STAAR SURGICAL COMPANY | 2021-07-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/107 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -10.00/+1.0/107 | OD | Physician | 06/03/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12112712&pc=QCB | 2023826-2021-02195 | 2020-05-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.00/+2.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: B3: DATE OF EVENT: 29-MAY-2020. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -13.00/+2.5/088 | OD |  | 06/04/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119227&pc=QCB | 2023826-2021-02327 | 2021-03-04 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-05 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/108 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY ON (B)(6) 2021. AN ALTERNATE LENS WAS IMPLANTED AT A LATER DATE AND THE PROBLEM WAS RESOLVED. REPORTEDLY, THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE CODE (B)(4): OFF-LABEL USE (ACD < 3.0MM). H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -8.5/+1.0/108 | OS | Physician | 06/10/2021 | 2021 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119233&pc=QCB | 2023826-2021-02329 | 2021-01-10 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-05 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+0.5/132 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2021. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED WITH AN ALTERNATE LENS INTRA-OPERATIVELY. THE PROBLEM WAS RESOLVED. NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.0/+0.5/132 | OD | Physician | 06/10/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119237&pc=QCB | 2023826-2021-02318 | 2021-01-27 | Injury | STAAR SURGICAL COMPANY | 2021-07-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+2.5/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+2.5/092 | OS | Physician | 06/09/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119238&pc=QCB | 2023826-2021-02330 | 2021-03-13 | Injury | STAAR SURGICAL COMPANY | 2021-07-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -6.0/+1.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM# (B)(4). | -6.0/+1.0/094 | OD | Physician | 06/11/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119241&pc=QCB | 2023826-2021-02287 | 2020-11-17 | Injury | STAAR SURGICAL COMPANY | 2021-07-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM # (B)(4). | -11.0/+2.0/085 | OD | Physician | 06/07/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119249&pc=QCB | 2023826-2021-02315 | 2021-06-01 | Injury | STAAR SURGICAL COMPANY | 2021-07-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/098 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM # (B)(4). | -10.5/+1.0/098 | OS | Physician | 06/09/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119250&pc=QCB | 2023826-2021-02309 | 2019-05-14 | Injury | STAAR SURGICAL COMPANY | 2021-07-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+1.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6)2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM# (B)(4). | -12.5/+1.0/087 | OD | Physician | 06/09/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119254&pc=QCB | 2023826-2021-02316 | 2021-06-05 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/108 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENS INTRAOPERATIVELY DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -12.0/+1.5/108 | OS | Physician | 06/09/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119255&pc=QCB | 2023826-2021-02312 | 2021-05-31 | Injury | STAAR SURGICAL COMPANY | 2021-07-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/054 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4) . | -11.0/+1.0/054 | OS | Physician | 06/09/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119256&pc=QCB | 2023826-2021-02288 | 2020-10-02 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.5/+1.5/109 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY DUE TO EXCESSIVE VAULT. ON (B)(6) 2020 A SHORTER LENS WAS IMPLANTED AND THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE CODE (B)(4): OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -7.5/+1.5/109 | OD | Physician | 06/07/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119257&pc=QCB | 2023826-2021-02304 | 2020-12-15 | Injury | STAAR SURGICAL COMPANY | 2021-07-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -9.0/+1.5/094 | OD | Physician | 06/09/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119258&pc=QCB | 2023826-2021-02295 | 2020-10-02 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/113 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM# (B)(4). | -10.0/+2.0/113 | OD | Physician | 06/07/2021 | 2021 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119259&pc=QCB | 2023826-2021-02298 | 2021-01-05 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/082 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REMOVED INTRAOPERATIVELY DUE TO EXCESSIVE VAULT. ON (B)(6) 2021 A SHORTER LENS WAS IMPLANTED AND THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -8.0/+1.5/082 | OS | Physician | 06/07/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119260&pc=QCB | 2023826-2021-02297 | 2020-09-29 | Injury | STAAR SURGICAL COMPANY | 2021-07-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM # (B)(4). | -11.5/+2.0/091 | OD | Physician | 06/07/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119261&pc=QCB | 2023826-2021-02306 | 2020-08-17 | Injury | STAAR SURGICAL COMPANY | 2021-07-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.0/+2.0/179 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -9.0/+2.0/179 | OS | Physician | 06/09/2021 | 2021 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119264&pc=QCB | 2023826-2021-02289 | 2020-10-02 | Injury | STAAR SURGICAL COMPANY | 2021-07-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+3.5/083 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM # (B)(4). | -9.5/+3.5/083 | OS | Physician | 06/07/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119266&pc=QCB | 2023826-2021-02292 | 2020-11-13 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-05 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/090 (SPHERE/CYLINDER/AXIS) TORE BEFORE/DURING LOADING. IT WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) THEN REMOVED AND REPLACED INTRA-OPERATIVELY ON (B)(6) 2020. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -11.0/+1.5/090 | OS | Physician | 06/07/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119270&pc=QCB | 2023826-2021-02293 | 2020-10-27 | Injury | STAAR SURGICAL COMPANY | 2021-07-05 | QCB | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.0/+2.5/090 (SPHERE/CYLINDER/AXIS) TORE/BROKE DURING LOADING. IT WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH AN ALTERNATE LENS AND THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -6.0/+2.5/090 | OS | Physician | 06/07/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119273&pc=QCB | 2023826-2021-02286 | 2020-11-17 | Injury | STAAR SURGICAL COMPANY | 2021-07-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/071 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM # (B)(4). | -10.5/+2.0/071 | OS | Physician | 06/07/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119223&pc=QCB | 2023826-2021-02291 |  | Injury | STAAR SURGICAL COMPANY | 2021-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.0/+2.0/085 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LOW VAULT. ON (B)(6) 2020 A LONGER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISIBLE INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: B5: ON (B)(6) 2020, A LONGER LENS WAS IMPLANTED BUT THE PROBLEM WAS NOT RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM # (B)(4). | -9.0/+2.0/085 | OS | Physician | 06/07/2021 | 2021 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119224&pc=QCB | 2023826-2021-02302 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+2.0/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. INTRAOPERATIVELY THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. A LONGER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+2.0/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. INTRAOPERATIVELY THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. A SHORTER LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE CODE (B)(4): OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -9.0/+2.0/091 | OS | Physician | 06/07/2021 | 2021 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119234&pc=QCB | 2023826-2021-02326 |  | Injury | STAAR SURGICAL COMPANY | 2021-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.5/126 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LENS DISLOCATION/SUBLUXATION AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: RECEIVED TRANSLATION TO INCLUDE IN INTAL MDR B5- SECTION. "FIRST ICL VAULT 356 BUT ASYMMETRICAL AND AQUAPORT DECENTERED TEMPORAL IN PUPILLARY EDGE. SECOND ICL CENTERED AND STABLE, POSSIBLE ZONULAR DEFECT. " CLAIM#: (B)(4). Manufacturer Narrative: DATE OF EVENT - UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+2.5/126 | OS |  | 06/10/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119247&pc=QCB | 2023826-2021-02317 | 2021-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.5/+1.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THEN REPLACED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-PREVIOUSLY STATED THE LENS WAS EXPLANTED AND THEN REPLACED WITH A SHORTER LENS. UPDATED INFORMATION: THE LENS WAS EXCHANGED WITH A SHORTER LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM # (B)(4). | -6.5/+1.0/087 | OD | Physician | 06/09/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12119248&pc=QCB | 2023826-2021-02305 | 2020-12-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -9.0/+1.0/069 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-PREVIOUSLY STATED THE LENS WAS EXCHANGED ON (B)(6) 2020. CORRECTION: ON (B)(6) 2020 THE LENS WAS EXCHANGED. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM # (B)(4). | -9.0/+1.0/069 | OS |  | 06/09/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12125530&pc=QCB | 2023826-2021-02340 | 2020-11-26 | Injury | STAAR SURGICAL COMPANY | 2021-07-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50/+1.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -08.50/+1.0/088 | OD | Physician | 06/07/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12125536&pc=QCB | 2023826-2021-02337 | 2020-12-15 | Injury | STAAR SURGICAL COMPANY | 2021-07-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | -10.50/+1.0/093 | OD | Physician | 06/07/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12125541&pc=QCB | 2023826-2021-02348 | 2021-05-18 | Injury | STAAR SURGICAL COMPANY | 2021-07-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00/+1.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM # (B)(4). | -15.00/+1.0/095 | OD | Physician | 06/06/2021 | 2021 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12125549&pc=QCB | 2023826-2021-02345 | 2021-05-02 | Injury | STAAR SURGICAL COMPANY | 2021-07-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -05.00/+1.0/156 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | -05.00/+1.0/156 | OD | Physician | 06/06/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12125599&pc=QCB | 2023826-2021-02353 | 2021-05-07 | Injury | STAAR SURGICAL COMPANY | 2021-07-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. BIOPTICS WAS PERFORMED TO CORRECT RESIDUAL ASTIGMATISM. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - BIOPTICS PERFORMED TO CORRECT RESIDUAL ASTIGMATISM. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.00/+2.0/095 | OS | Physician | 06/08/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12125540&pc=QCB | 2023826-2021-02335 | 2021-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50/+1.5/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC BENT, DEFORMED AND STRETCHED OUT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -08.50/+1.5/082 | OD |  | 06/07/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12125584&pc=QCB | 2023826-2021-02352 | 2021-04-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -16.00/+3.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THIS WAS THE REPLACEMENT LENS FOR MFR REPORT # 2023826-2021-02350. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -16.00/+3.5/092 | OS |  | 06/06/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12125591&pc=QCB | 2023826-2021-02350 | 2021-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.50/+3.5/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS BUT THE PROBLEM WAS NOT RESOLVED. SEE MFR REPORT # 2023826-2021-02352 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -16.50/+3.5/091 | OS |  | 06/06/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12125593&pc=QCB | 2023826-2021-02351 | 2021-04-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -17.50/+4.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THIS WAS THE REPLACEMENT LENS FOR MFR REPORT # 2023826-2021-02349. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL. WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED, AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -17.50/+4.5/092 | OD |  | 06/06/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12125608&pc=QCB | 2023826-2021-02349 | 2021-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+5.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS BUT THE PROBLEM WAS NOT RESOLVED. SEE MFR REPORT # 2023826-2021-02351 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM #(B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | -18.00/+5.0/093 | OD |  | 06/06/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12132567&pc=QCB | 2023826-2021-02341 | 2020-11-24 | Injury | STAAR SURGICAL COMPANY | 2021-07-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.00/+1.5/178 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020, DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -15.00/+1.5/178 | OS | Physician | 06/07/2021 | 2021 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12132733&pc=QCB | 2023826-2021-02378 | 2021-01-19 | Injury | STAAR SURGICAL COMPANY | 2021-07-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -8.5/1.0/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGED RESOLVED THE PROBLEM. UNKNOWN WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -8.5/1.0/100 | OD | Physician | 06/07/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12132735&pc=QCB | 2023826-2021-02379 | 2020-10-01 | Injury | STAAR SURGICAL COMPANY | 2021-07-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -7.0/0.5/121 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGED RESOLVED THE PROBLEM. UNKNOWN WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.0/0.5/121 | OD | Physician | 06/07/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12132736&pc=QCB | 2023826-2021-02356 | 1998-08-30 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.5/2.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT WITHIN THE SAME SURGERY. THIS EXPLANT RESOLVED THE PROBLEM. UNKNOWN WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -10.5/2.5/088 | OS | Physician | 06/07/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12132738&pc=QCB | 2023826-2021-02371 | 2021-02-05 | Injury | STAAR SURGICAL COMPANY | 2021-07-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -11.0/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGED RESOLVED THE PROBLEM. UNKNOWN WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -11.0/1.5/090 | OD | Physician | 06/07/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12132741&pc=QCB | 2023826-2021-02382 | 2020-09-18 | Injury | STAAR SURGICAL COMPANY | 2021-07-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -9.0/2.0/79 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGED RESOLVED THE PROBLEM. UNKNOWN WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -9.0/2.0/79 | OS | Physician | 06/07/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12132742&pc=QCB | 2023826-2021-02370 | 2020-05-08 | Injury | STAAR SURGICAL COMPANY | 2021-07-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.0/2.5/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REMOVED AND LATER REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGED RESOLVED THE PROBLEM. UNKNOWN WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -15.0/2.5/100 | OD | Physician | 06/07/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12132744&pc=QCB | 2023826-2021-02358 | 2020-12-22 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -8.0/2.0/88 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED WITHIN THE SAME SURGERY AND LATER REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. IT WAS REPORTED THAT THE PROBLEM RESOLVED. UNKNOWN WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -8.0/2.0/88 | OD | Physician | 06/07/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12132748&pc=QCB | 2023826-2021-02366 | 2020-08-30 | Injury | STAAR SURGICAL COMPANY | 2021-07-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -6.0/2.5/176 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGED RESOLVED THE PROBLEM. UNKNOWN WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -6.0/2.5/176 | OS | Physician | 06/07/2021 | 2021 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12132749&pc=QCB | 2023826-2021-02357 | 2020-12-20 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.5/2.0/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT WITHIN THE SAME SURGERY. THIS EXPLANT RESOLVED THE PROBLEM. UNKNOWN WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -9.5/2.0/94 | OS | Physician | 06/07/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12132750&pc=QCB | 2023826-2021-02380 | 2020-10-01 | Injury | STAAR SURGICAL COMPANY | 2021-07-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -7.5/1.0/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGED RESOLVED THE PROBLEM. UNKNOWN WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM # (B)(4). | -7.5/1.0/89 | OS | Physician | 10/16/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12132568&pc=QCB | 2023826-2021-02342 | 2021-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED LATER THAT SAME DAY DUE TO LOW VAULTING. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN, SPLIT AND DEFORMED AND THE HAPTIC TORN, BROKEN, BENT, AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.00/+1.5/087 | OD |  | 06/07/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12132732&pc=QCB | 2023826-2021-02367 | 2020-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.5/3.0/001 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGED RESOLVED THE PROBLEM. UNKNOWN WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: SERIAL# (B)(6). CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -15.5/3.0/001 | OD |  | 06/07/2021 | 2021 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12132734&pc=QCB | 2023826-2021-02359 | 2020-12-24 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -8.0/3.0/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED WITHIN THE SAME SURGERY AND LATER REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. IT WAS REPORTED THAT THE PROBLEM RESOLVED. UNKNOWN WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: B3: CORRECTED TO (B)(6) 2020 IN INITIAL MDR PER UPDATED COMPLAINT QUESTIONNAIRE RECEIVED ON (B)(6) 2021. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -8.0/3.0/93 | OS | Physician | 06/07/2021 | 2021 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12132745&pc=QCB | 2023826-2021-02368 | 2020-05-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.0/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REMOVED AND LATER REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGED RESOLVED THE PROBLEM. UNKNOWN WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: H3: DEVICE EVALUATION IS NOT REQUIRED. ADDITIONAL INFORMATION: B3: DATE CHANGED TO 08/MAY/2020. B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS DIOPTER -15.0/2.5/100 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON 08/05/2020. THE LENS WAS EXPLANTED ON 12/05/2020 AND A NEW LENS OF SHORTER LENGTH WAS IMPLANTED, ON THE SAME DAY BUT DIFFERENT SURGERY, DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. D6A: DATE CHANGED TO 08/MAY/2020. D6B: DATE CHANGED TO 12/MAY/2020. CLAIM#: (B)(4). Manufacturer Narrative: MFR REPORT # 2023826-2021-02368 SUPPLEMENTAL # 1 IS NOT APPLICABLE TO THIS CLAIM. CLAIM #(B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -13.0/1.5/090 | OD | Physician | 06/07/2021 | 2021 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12133198&pc=QCB | 2023826-2021-02377 | 2020-10-27 | Injury | STAAR SURGICAL COMPANY | 2021-07-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.0/1.0/97 , IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. A REPLACEMENT LENS WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. THERE WAS NO PATIENT INJURY. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.0/1.0/97 | OS | Physician | 06/07/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12138348&pc=QCB | 2023826-2021-02415 | 2020-03-23 | Injury | STAAR SURGICAL COMPANY | 2021-07-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.5/2.0/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGED RESOLVED THE PROBLEM. UNKNOWN WAS REPORTED TO BE THE CAUSE OF THIS EVENT. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -8.5/2.0/108 | OD | Physician | 06/08/2021 | 2021 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12139672&pc=QCB | 2023826-2021-02418 | 2020-12-22 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.0/1.5/090, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT. IT WAS REPORTED THAT THE PROBLEM RESOLVED. THE CAUSE OF THIS EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -11.0/1.5/090 | OD | Physician | 12/22/2020 | 2020 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12140235&pc=QCB | 2023826-2021-02414 | 2020-06-02 | Injury | STAAR SURGICAL COMPANY | 2021-07-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.5/2.0/91, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS LATER REMOVED DUE TO EXCESSIVE VAULT. A SHORTER LENGTH LENS WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. THE CAUSE OF THIS EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -12.5/2.0/91 | OD | Physician | 06/08/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12140305&pc=QCB | 2023826-2021-02390 | 2021-05-03 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.00/+2.5/087 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS WAS INSERTED UPSIDE DOWN IN THE PATIENT'S RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.00/+2.5/087 | OD | Physician | 06/08/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12140308&pc=QCB | 2023826-2021-02391 | 2021-05-24 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -09.00/+2.5/093 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENT'S RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR BUT THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -09.00/+2.5/093 | OD | Physician | 06/08/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12140319&pc=QCB | 2023826-2021-02392 | 2021-05-18 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50/+1.5/109 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM# (B)(4). | -08.50/+1.5/109 | NA | Physician | 06/08/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12140320&pc=QCB | 2023826-2021-02393 | 2021-05-11 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/091 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED FOR A SHORTER LENS WITHIN THE SAME SURGERY AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM # (B)(4). | -11.00/+1.5/091 | NA | Physician | 06/08/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12140363&pc=QCB | 2023826-2021-02402 | 2020-12-26 | Injury | STAAR SURGICAL COMPANY | 2021-07-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50/+4.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -13.50/+4.5/088 | OS | Physician | 06/09/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12140364&pc=QCB | 2023826-2021-02403 | 2020-12-10 | Injury | STAAR SURGICAL COMPANY | 2021-07-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | -10.50/+1.0/091 | OS | Physician | 06/09/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12133153&pc=QCB | 2023826-2021-02361 | 2021-05-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -9.5/2.0/78 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. LOW VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. UNKNOWN WAS REPORTED TO BE THE CAUSE OF THIS EVENT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER UPDATED INFORMATION, THE PROBLEM HASN'T RESOLVED. AND THE DOCTOR CONSIDER TO CONTINUE TO OBSERVE. IT WAS REPORTED, THAT NO TREATMENT OR INTERVENTION HAS BEEN PERFORMED. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -9.5/2.0/78 | OD |  | 06/07/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12133164&pc=QCB | 2023826-2021-02363 | 2021-04-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.0/1.0/89 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. EXCESSIVE VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. UNKNOWN WAS REPORTED TO BE THE CAUSE OF THIS EVENT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6: WORK ORDER SEARCH - ONE SIMILAR COMPLAINT TYPE EVENT(S) WAS REPORTED FOR UNITS WITHIN THE SAME LOT. Manufacturer Narrative: B5 - THE LENS WAS REPLACED ON (B)(6) 2021 WITH A SHORTER LENGTH LENS AND PROBLEM RESOLVED. REPORTEDLY, "BECAUSE THE ARCH HEIGHT OF THE LEFT EYE AFTER LENS IMPLANTATION WAS TOO HIGH, IT WAS SUGGESTED TO OBSERVE FOR A PERIOD OF TIME AFTER THE SURGERY DOCTOR'S FACE TO FACE CONSULTATION. THE RESULTS OF POST-OPERATIVE REVIEW SHOWED THAT THE ARCH HEIGHT WAS STILL HIGH. AFTER COMMUNICATING WITH THE PATIENT, THE SURGERY DOCTOR DECIDED TO REPLACE THE SMALLER LENS". CAUSE OF EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.0/1.0/89 | OS | Physician | 06/07/2021 | 2021 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12140189&pc=QCB | 2023826-2021-02412 | 2020-12-22 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.5/2.0/093, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT. IT WAS REPORTED THAT THE PROBLEM RESOLVED. THE CAUSE OF THIS EVENT WAS PATIENT RELATED FACTOR. Manufacturer Narrative: CORRECTED DATA: A2- (B)(6) 2001. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -11.5/2.0/093 | OS | Physician | 12/22/2020 | 2020 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12140314&pc=QCB | 2023826-2021-02394 | 2020-12-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+2.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: B3: DATE OF EVENT: (B)(6) 2020 CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+2.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. D6A: (B)(6) 2020. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND ONE HAPTIC BENT. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.50/+2.5/084 | OD |  | 06/08/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12140346&pc=QCB | 2023826-2021-02395 | 2020-12-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, 10.50/+2.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: B3: DATE OF EVENT: (B)(6) 2020. CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+2.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. D6A: (B)(6) 2020. CLAIM #(B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT /DEFORMED/STRETCHED OUT. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.50/+2.0/093 | OS |  | 06/08/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12140359&pc=QCB | 2023826-2021-02398 | 2020-12-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.00/+2.0/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -14.00/+2.0/083 | OS |  | 06/09/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12140361&pc=QCB | 2023826-2021-02399 | 2020-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.00/+4.0/076 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. REOPTOMETRY WAS PERFORMED AFTER THE REFRACTIVE NUMBER WAS STABILIZED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B3: DATE OF EVENT: ON (B)(6) 2020. CORRECTED DATA: E1: DR. (B)(6). CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -08.00/+4.0/076 | OS |  | 06/09/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12143777&pc=QCB | 2023826-2021-02486 | 2019-10-04 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -6.5/0.5/065, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. THE LENS WA REMOVED AND REPLACED WITH A SHORTER LENGTH LENS WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULT. IT WAS REPORTED THAT THE PROBLEM RESOLVED. CAUSE OF EVET IS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.5/0.5/065 | OS | Physician | 06/09/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147090&pc=QCB | 2023826-2021-02407 | 2020-11-11 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50/+1.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -06.50/+1.5/090 | OS | Physician | 06/11/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147133&pc=QCB | 2023826-2021-02503 | 2020-12-23 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+1.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.00/+1.0/090 | OD | Physician | 06/09/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147143&pc=QCB | 2023826-2021-02406 | 2020-11-11 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -05.50/+1.0/059 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -05.50/+1.0/059 | OD | Physician | 06/11/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147182&pc=QCB | 2023826-2021-02409 | 2021-03-31 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -09.00/+2.0/021 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO THE PHYSICIAN PUT WRONG CYLINDER AXIS IN THE OCOS CALCULATION. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE DEVICE PERFORMED AS INTENDED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.00/+2.0/021 | OD | Physician | 06/10/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147462&pc=QCB | 2023826-2021-02234 | 2021-01-27 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -12.00/2.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.00/2.5/088 | OS | Physician | 06/14/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147464&pc=QCB | 2023826-2021-02229 | 2021-02-23 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.50/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REMOVED IN A SECOND SURGERY ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -7.50/1.5/093 | OD | Physician | 06/14/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147465&pc=QCB | 2023826-2021-02238 | 2021-03-22 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.00/2.0/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.00/2.0/100 | OD | Physician | 06/14/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147468&pc=QCB | 2023826-2021-02237 | 2019-10-01 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.50/2.5/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.50/2.5/108 | OS | Physician | 06/14/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147469&pc=QCB | 2023826-2021-02231 | 2020-06-09 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.00/2.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -7.00/2.0/092 | OD | Physician | 06/14/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147471&pc=QCB | 2023826-2021-02243 | 2021-03-19 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -6.00/2.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.00/2.0/092 | OD | Physician | 06/14/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147474&pc=QCB | 2023826-2021-02242 | 2021-03-18 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.00/1.5/157 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.00/1.5/157 | OD | Physician | 06/14/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147475&pc=QCB | 2023826-2021-02236 | 2021-03-26 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/2.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.00/2.0/094 | OD | Physician | 06/14/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147476&pc=QCB | 2023826-2021-02250 | 2021-03-27 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/1.0/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/1.0/100 | OD | Physician | 06/14/2021 | 2021 | 17
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147480&pc=QCB | 2023826-2021-02256 | 2021-02-02 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.00/1.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/1.5/088 | OD | Physician | 06/14/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147481&pc=QCB | 2023826-2021-02239 | 2020-05-06 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/3.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.50/3.0/081 | OD | Physician | 06/14/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147484&pc=QCB | 2023826-2021-02248 | 2021-02-04 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/1.0/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.50/1.0/109 | OD | Physician | 06/14/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147486&pc=QCB | 2023826-2021-02255 | 2021-02-02 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.00/2.0/089 | OS | Physician | 06/14/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147487&pc=QCB | 2023826-2021-02254 | 2021-02-02 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.00/2.0/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.00/2.0/112 | OD |  | 06/14/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147492&pc=QCB | 2023826-2021-02245 | 2021-01-27 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/2.0/131 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(6). | -12.50/2.0/131 | OS | Physician | 06/14/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147493&pc=QCB | 2023826-2021-02251 | 2021-05-25 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.00/0.5/156 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -8.00/0.5/156 | OD | Physician | 06/13/2021 | 2021 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147532&pc=QCB | 2023826-2021-02260 | 2021-02-16 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.00/1.5/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.00/1.5/070 | OD | Physician | 06/14/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147541&pc=QCB | 2023826-2021-02272 | 2021-06-14 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.50/1.5/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) . EXCESSIVE VAULT WAS OBSERVED, THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -7.50/1.5/073 | OD | Physician | 06/15/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147546&pc=QCB | 2023826-2021-02273 | 2021-06-14 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.00/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) . EXCESSIVE VAULT WAS OBSERVED, THE LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -8.00/1.5/089 | OS | Physician | 06/15/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147547&pc=QCB | 2023826-2021-02271 | 2019-11-06 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Improper or Incorrect Procedure or Method; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO_12.6, -16.00/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019 .THE LENS TORE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED WITH A SHORTER LENGTH LENS INTRAOPERATIVELY ON (B)(6) 2019. PROBLEM RESOLVED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. OFF-LABEL USE, ACD<3.0MM. CLAIM# (B)(4). | -16.00/2.0/093 | OS | Physician | 06/14/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147551&pc=QCB | 2023826-2021-02279 | 2021-03-25 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2021 A VTICM5_13.7, -10.0/1.0/118 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE BEFORE/DURING LOADING. A BACKUP LENS WAS IMPLANTED INTRAOPERATIVELY AND THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE POSTOPERATIVE CONDITION IS GOOD. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.0/1.0/118 | NA | Physician | 06/18/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147554&pc=QCB | 2023826-2021-02269 | 2020-04-16 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -7.00/1.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2020 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT.THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -7.00/1.0/082 | OS | Physician | 06/14/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147812&pc=QCB | 2023826-2021-02494 | 2020-12-21 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.5/1.0/092, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -12.5/1.0/092 | OD | Physician | 06/09/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147852&pc=QCB | 2023826-2021-02487 | 2021-06-02 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.0/1.5/87, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS LATER EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. IT WAS REPORTED THAT THE PROBLEM RESOLVED. CAUSE OF EVET IS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.0/1.5/87 | OS | Physician | 06/11/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147858&pc=QCB | 2023826-2021-02489 | 2021-01-21 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -8.5/1.5/109, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -8.5/1.5/109 | OD | Physician | 06/09/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147859&pc=QCB | 2023826-2021-02499 | 2018-12-01 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -11.0/1.5/084, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT - RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -11.0/1.5/084 | OS | Physician | 06/09/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147861&pc=QCB | 2023826-2021-02498 | 2020-12-08 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -4.5/1.0/84, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. OFF-LABEL USE (ACD < 3.0MM). CLAIM # (B)(4). | -4.5/1.0/84 | OD | Physician | 06/09/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147863&pc=QCB | 2023826-2021-02500 | 2020-12-08 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.5/2.0/090, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM #: (B)(4). | -10.5/2.0/090 | OD | Physician | 06/09/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147873&pc=QCB | 2023826-2021-02488 | 2021-01-22 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -11.5/1.0/070, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/1.0/070 | OS | Physician | 06/11/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147878&pc=QCB | 2023826-2021-02502 | 2021-05-08 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.00/2.50/92, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -9.00/2.50/92 | OS | Physician | 06/11/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147879&pc=QCB | 2023826-2021-02491 | 2021-01-02 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -9.5/2.0/093, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -9.5/2.0/093 | OD | Physician | 06/09/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147103&pc=QCB | 2023826-2021-02401 | 2020-06-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.00/+2.5/165 (SPHERE/CYLINDER/AXIS), WAS TORN BEFORE/DURING LOADING. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED ON (B)(6) 2020. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THERE WAS NO PATIENT CONTACT AND THE PROBLEM WAS RESOLVED WITH THE IMPLANT OF THE REPLACEMENT LENS. CLAIM # (B)(4). Manufacturer Narrative: SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. IMPLANT: UNK. EXPLANT: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.00/+2.5/165 | NA |  | 06/09/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147467&pc=QCB | 2023826-2021-02235 | 2021-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.00/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2021 DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.00/1.5/092 | OD |  | 06/14/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147488&pc=QCB | 2023826-2021-02247 | 2019-09-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.62MM VTICMO12.6, -9.00/1.0/064 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.00/1.0/064 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.00/1.0/064 | OS |  | 06/14/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147494&pc=QCB | 2023826-2021-02257 | 2021-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.5/1.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.5/2.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/1.5/086 | OS |  | 06/14/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147545&pc=QCB | 2023826-2021-02261 | 2021-03-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4) - OFF-LABEL USE, UNDER 21 YRS OF AGE AT DATE OF IMPLANT. CLAIM# (B)(4). | -12.50/2.0/089 | OS |  | 06/14/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147548&pc=QCB | 2023826-2021-02264 | 2021-05-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO 12.6, -10.50/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS REPLACED FOR A SAME LENGTH SPHERICAL LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.50/1.0/093 | OD | Physician | 06/14/2021 | 2021 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147549&pc=QCB | 2023826-2021-02276 | 2021-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM513.7, -4.00/1.0/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.REPORTEDLY, A LENS REPLACEMENT TO A SHORTER LENS IS PLANNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: INCISION ENLARGEMENT WAS PERFORMED. LENS WAS REPLACED ON (B)(6) 2021 WITH A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. REPORTEDLY, "EXCELLENT QUALITY OF VISION, VAULT ISSUE RESOLVED." CAUSE WAS REPORTED AS DEVICE: "VERY HIGH VAULT IMMEDIATELY POST-OP, NEEDED TO REPLACE WITH SMALLER LENS." CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -4.00/1.0/072 | OS | Physician | 06/15/2021 | 2021 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147557&pc=QCB | 2023826-2021-02275 | 2021-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM513.7, -4.00/1.0/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.REPORTEDLY, A LENS REPLACEMENT TO A SHORTER LENS IS PLANNED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: INCISION ENLARGEMENT WAS PERFORMED. LENS WAS REPLACED ON 24 OCT 2021 WITH A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. REPORTEDLY, "ECELLENT QUALITY OF VISION, VAULT ISSUE RESOLVED." CAUSE WAS REPORTED AS DEVICE: "VERY HIGH VAULT IMMEDIATELY POST-OP, NEEDED TO REPLACE WITH SMALLER LENS." CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -4.00/1.0/103 | OD | Physician | 06/15/2021 | 2021 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147813&pc=QCB | 2023826-2021-02479 | 2020-03-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.5/3.5/097, IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REMOVED WITHIN THE SAME SURGERY DUE TO A RUPTURE OF THE LENS SUSPENSORY LIGAMENT THAT WAS FOUND DURING THE OPERATION. IT WAS REPORTED THAT MEDICAL OR SURGICAL INTERVENTION WAS NOT NECESSARY, " NO LENS IMPLANTED." CAUSE WAS REPORTED TO BE PATIENT RELATED FACTOR. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED, THAT A 12.6MM, VTICMO12.6, -12.5/3.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE, DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS IMPLANTED AND REMOVED WITHIN THE SAME SURGERY. THERE WAS NO PATIENT INJURY. AND THE PROBLEM RESOLVED. IT WAS REPORTED, THAT NO SECONDARY INTERVENTION WAS NECESSARY "NO IMPLANTATION OF LENS" WAS REPORTED. CAUSE OF EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: B5 - CORRECTED TO: "THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.5/3.5/097, IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REMOVED WITHIN THE SAME SURGERY DUE TO LENS TEAR/ BREAK DURING INJECTION/DELIVERY INTO PATIENT'S RIGHT EYE (OD). IT WAS REPORTED THAT 'BECAUSE OF THE PATIENT'S TIME, THE LENS WAS NO IMPLANTED AGAIN'. IT WAS REPORTED THAT MEDICAL OR SURGICAL INTERVENTION WAS NOT NECESSARY, 'NO LENS IMPLANTED'. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. PROBLEM WAS NOT RESOLVED." H6- MEDICAL DEVICE PROBLEM CODE- 1494 OFF-LABEL USE (ACD<3.0MM) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.5/3.5/097 | OD | Physician | 06/09/2021 | 2021 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147851&pc=QCB | 2023826-2021-02483 | 2020-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.50/2.5/082, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS REMOVED DUE TO LOW VAULT. A REPLACEMENT LENS OF LONGER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE WAS UNKNOWN. Manufacturer Narrative: B5 - "LENS ROTATION WAS ALSO REPORTED" SHOULD BE INCLUDED IN THE INITAL MDR. H6 - MEDICAL DEVICE PROBLEM CODE 2923 SHOULD BE INCLUDED IN INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -10.50/2.5/082 | OD | Physician | 06/11/2021 | 2021 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12147856&pc=QCB | 2023826-2021-02501 | 2020-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -7.0/2.0/089, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: B5: PER UPDATED INFORMATION CAUSE OF THIS EVENT IS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -7.0/2.0/089 | OD |  | 06/09/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148269&pc=QCB | 2023826-2021-02434 | 2021-02-17 | Injury | STAAR SURGICAL COMPANY | 2021-07-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.0/+3.5/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -17.0/+3.5/103 | OS | Physician | 06/14/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148273&pc=QCB | 2023826-2021-02472 | 2021-06-12 | Injury | STAAR SURGICAL COMPANY | 2021-07-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THEN REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -11.0/+1.5/092 | OS | Physician | 06/15/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148276&pc=QCB | 2023826-2021-02467 | 2019-10-25 | Injury | STAAR SURGICAL COMPANY | 2021-07-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.5/081 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -11.5/+1.5/081 | OD | Physician | 06/14/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148278&pc=QCB | 2023826-2021-02468 | 2019-08-16 | Injury | STAAR SURGICAL COMPANY | 2021-07-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/078 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK.. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -11.0/+1.0/078 | OD | Physician | 06/14/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148281&pc=QCB | 2023826-2021-02426 | 2020-12-07 | Injury | STAAR SURGICAL COMPANY | 2021-07-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -12.5/+1.5/087 | OD | Physician | 06/14/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148285&pc=QCB | 2023826-2021-02427 | 2021-03-15 | Injury | STAAR SURGICAL COMPANY | 2021-07-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+2.0/080 | OD | Physician | 06/14/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148286&pc=QCB | 2023826-2021-02430 | 2020-06-20 | Injury | STAAR SURGICAL COMPANY | 2021-07-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+5.5/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -17.0/+5.5/095 | OD | Physician | 06/14/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148288&pc=QCB | 2023826-2021-02433 | 2021-02-17 | Injury | STAAR SURGICAL COMPANY | 2021-07-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.0/+4.0/081 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -16.0/+4.0/081 | OD | Physician | 06/14/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148290&pc=QCB | 2023826-2021-02460 | 2021-05-26 | Injury | STAAR SURGICAL COMPANY | 2021-07-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/099 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -11.5/+2.0/099 | OD | Physician | 06/14/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148291&pc=QCB | 2023826-2021-02477 | 2021-04-29 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-11 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -6.5/+1.5/000 (SPHERE/CYLINDER/AXIS) GOT CAUGHT IN THE CARTRIDGE AND WAS DAMAGED DURING LOADING. THIS OCCURRED ON (B)(6) 2021. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. AN ALTERNATE LENS WAS IMPLANTED ON THE SAME DATE AND THE PROBLEM IS RESOLVED. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -6.5/+1.5/000 | NA | Physician | 06/18/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148272&pc=QCB | 2023826-2021-02473 | 2020-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.0/+1.5/115 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -14.0/+1.5/115 | OS |  | 06/15/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148277&pc=QCB | 2023826-2021-02466 | 2020-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH AN ALTERNATE LENS, THE SURGEON STATING THAT THE CENTRAL HOLE OF THE LENS WAS NOT PENETRATED ALL THE WAY THROUGH. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS OTHER. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THAT THE PERFORATING HOLE IN THE CENTER OF THE LENS DIDN'T PENETRATE THROUGH. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THE LENS TO BE OUT OF SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6- METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -8.5/+1.0/092 | OD | Physician | 06/16/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148283&pc=QCB | 2023826-2021-02428 | 2021-02-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+3.5/102 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+3.5/102 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REMOVED AND A LONGER LENGTH LENS WAS IMPLANTED VERTICALLY WITHIN THE SAME SURGERY DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/+3.5/102 | OD |  | 06/14/2021 | 2021 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148284&pc=QCB | 2023826-2021-02436 | 2020-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/084 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED AND THEN REPLACED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.0/084 | OS |  | 06/14/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148287&pc=QCB | 2023826-2021-02435 | 2020-05-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/060 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -10.0/+1.0/060 | OS |  | 06/14/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12148289&pc=QCB | 2023826-2021-02422 | 2021-06-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -8.0/+2.0/094 (SPHERE/CYLINDER/AXIS) TORE DURING INSERTION. THERE WAS NO PATIENT INJURY. THIS OCCURRED ON (B)(6) 2021. Manufacturer Narrative: B1: CORRECTED; B2: PREVIOUSLY REPORTED IN ERROR; B5: ADDITIONAL INFORMATION: THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY. THE HAPTIC TORE DURING INSERTION; H6: DEVICE CODE 1494. (AGE 21 YRS AT DATE OF IMPLANT). CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. EXPLANT DATE - UNK. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -8.0/+2.0/094 | NA |  | 06/14/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12150395&pc=QCB | 2023826-2021-02543 | 2020-11-13 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/105 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PREMARKET IDENTIFICATION: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -12.00/+1.5/105 | OD | Physician | 06/14/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12150420&pc=QCB | 2023826-2021-02538 | 2020-07-17 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON WAS INSERTING A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50/+1.5/177 (SPHERE/CYLINDER/AXIS), AND THE LENS TORE DURING DELIVERY. THERE WAS NO PATIENT CONTACT OR INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS ON (B)(6) 2020 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.50/+1.5/177 | NA | Physician | 06/14/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12155580&pc=QCB | 2023826-2021-02462 | 2021-05-28 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -10.5/+1.5/083 | OD | Physician | 06/14/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12155585&pc=QCB | 2023826-2021-02464 | 2021-04-28 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+2.5/096 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -7.0/+2.5/096 | OD | Physician | 06/14/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12155586&pc=QCB | 2023826-2021-02455 | 2021-04-23 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -10.0/+1.0/074 | OS | Physician | 06/14/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12155588&pc=QCB | 2023826-2021-02453 | 2021-04-22 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+3.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/+3.0/093 | OD | Physician | 06/14/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12155589&pc=QCB | 2023826-2021-02447 | 2021-03-15 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM # (B)(4). | -8.5/+1.5/093 | OS | Physician | 06/14/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12155590&pc=QCB | 2023826-2021-02451 | 2021-04-22 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.0/112 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -9.5/+2.0/112 | OD | Physician | 06/14/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12155592&pc=QCB | 2023826-2021-02448 | 2021-03-12 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -3.0/+0.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -3.0/+0.5/074 | OS | Physician | 06/14/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12155598&pc=QCB | 2023826-2021-02442 | 2020-10-08 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -11.5/+2.5/088 | OD | Physician | 06/14/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12155606&pc=QCB | 2023826-2021-02440 | 2020-08-26 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.0/+3.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -12.0/+3.0/091 | OD | Physician | 06/14/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12155820&pc=QCB | 2023826-2021-02546 | 2020-10-16 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/113 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM # (B(4). | -12.00/+1.5/113 | OD | Physician | 06/14/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12155823&pc=QCB | 2023826-2021-02544 | 2020-11-13 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/078 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021, DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -11.50/+1.5/078 | OS | Physician | 06/14/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12155829&pc=QCB | 2023826-2021-02553 | 2021-03-06 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.00/+2.0/009 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS REPORTED AS LOW VAULTING AND REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2021-02552. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH A PREVIOUS OR PRE-EXISTING OCULAR DISEASE THAT WOULD PRECLUDE POST-OPERATIVE VISUAL ACUITY OF 20/60 OR BETTER (RETINITIS PIGMENTOSA). CLAIM # (B)(4). | -07.00/+2.0/009 | OD | Physician | 06/14/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12155835&pc=QCB | 2023826-2021-02536 | 2020-07-07 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50/+2.0/103 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -14.50/+2.0/103 | OD | Physician | 06/14/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12155840&pc=QCB | 2023826-2021-02537 | 2020-03-12 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM # (B)(4). | -10.00/+1.0/084 | OD | Physician | 06/14/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12155842&pc=QCB | 2023826-2021-02540 | 2020-12-17 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED LATER THAT SAME DAY DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.00/+1.5/094 | OD | Physician | 06/14/2021 | 2021 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12155851&pc=QCB | 2023826-2021-02548 | 2021-03-15 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00/+1.5/160 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -14.00/+1.5/160 | OS | Physician | 06/14/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12155857&pc=QCB | 2023826-2021-02549 | 2021-01-26 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00/+1.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -08.00/+1.0/093 | OS | Physician | 06/14/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12155863&pc=QCB | 2023826-2021-02551 | 2020-05-20 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.00/+1.0/109 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | -16.00/+1.0/109 | OS | Physician | 06/14/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12150408&pc=QCB | 2023826-2021-02552 | 2021-01-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -08.00/+2.0/179 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS BUT THE PROBLEM WAS NOT RESOLVED. SEE MFR. REPORT # 2023826-2021-02553 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH A PREVIOUS OR PRE-EXISTING OCULAR DISEASE THAT WOULD PRECLUDE POST-OPERATIVE VISUAL ACUITY OF 20/60 OR BETTER (RETINITIS PIGMENTOSA). CLAIM # (B)(4). | -08.00/+2.0/179 | OD |  | 06/14/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12155581&pc=QCB | 2023826-2021-02456 | 2021-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.0/+2.0/176 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT MODEL LENS DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO DAMAGE TO THE LENS WITH RESIDUE ON OPTIC AND HAPTIC. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -10.0/+2.0/176 | OS | Physician | 06/14/2021 | 2021 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12155583&pc=QCB | 2023826-2021-02457 | 2021-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/168 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT AXIS LENS DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN, BROKEN, BENT, AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -9.5/+1.5/168 | OS |  | 06/14/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12155597&pc=QCB | 2023826-2021-02439 | 2020-03-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+2.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). Manufacturer Narrative: SEX, WEIGHT, ETHNICITY, RACE - UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -10.0/+2.5/094 | OD |  | 06/14/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12167088&pc=QCB | 2023826-2021-02507 | 2019-08-20 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Improper or Incorrect Procedure or Method; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO 13.2, -9.50/1.0/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019.THE LENS TORE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED WITH A SHORTER LENGTH LENS INTRAOPERATIVELY ON (B)(6) 2019. PROBLEM RESOLVED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.0/105 | OD | Physician | 06/21/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12167094&pc=QCB | 2023826-2021-02509 | 2019-10-09 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -11.50/2.5/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED WITH A LONGER LENGTH LENS INTRAOPERATIVELY ON (B)(6) 2019 FROM PATIENT'S RIGHT EYE (OD) DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.50/2.5/099 | OD | Physician | 06/21/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12167095&pc=QCB | 2023826-2021-02506 | 2019-10-02 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -9.50/1.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED WITH A LONGER LENGTH LENS INTRAOPERATIVELY ON (B)(6) 2020 FROM PATIENT'S RIGHT EYE (OD) DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.50/1.5/082 | OD | Physician | 06/21/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12167096&pc=QCB | 2023826-2021-02508 | 2019-10-02 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -9.50/2.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED WITH A SHORTER LENGTH LENS INTRAOPERATIVELY ON (B)(6) 2019 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/2.5/080 | OD | Physician | 06/21/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12167117&pc=QCB | 2023826-2021-02525 | 2021-06-14 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -14.00/1.5/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021.THE LENS TORE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED WITH A SHORTER LENGTH LENS INTRAOPERATIVELY ON (B)(6) 2021. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, LENS TEAR DURING INJECTION. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -14.00/1.5/084 | OS | Physician | 06/24/2021 | 2021 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12167120&pc=QCB | 2023826-2021-02526 | 2021-06-12 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -5.50/5.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REMOVED FROM THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS TORE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. ON (B)(6) 2021 A SAME LENGTH LENS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, SECOND LENS IMPLANTATION WAS SMOOTH AND PATIENT IS DOING VERY WELL. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K (2021 REPORTS) - THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -5.50/5.0/089 | OD | Physician | 06/24/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12167137&pc=QCB | 2023826-2021-02527 | 2021-06-12 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.00/2.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.00/2.5/094 | OD | Physician | 06/24/2021 | 2021 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170140&pc=QCB | 2023826-2021-02580 | 2021-02-18 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.0/1.0/86 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -9.0/1.0/86 | OD | Physician | 06/14/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170331&pc=QCB | 2023826-2021-02616 | 2020-09-08 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.5/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -13.5/1.5/090 | OD | Physician | 06/14/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170422&pc=QCB | 2023826-2021-02570 | 2021-02-05 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -08.00/+1.0/098 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -08.00/+1.0/098 | OD | Physician | 06/14/2021 | 2021 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170433&pc=QCB | 2023826-2021-02601 | 2021-04-22 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.0/5.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -10.0/5.0/086 | OD | Physician | 06/14/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170453&pc=QCB | 2023826-2021-02564 | 2021-01-28 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.50/+2.5/077 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITHIN THE SAME SURGERY WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.50/+2.5/077 | OD | Physician | 06/14/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170498&pc=QCB | 2023826-2021-02588 | 2021-03-10 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.0/2.0/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -12.0/2.0/93 | OD | Physician | 06/14/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170500&pc=QCB | 2023826-2021-02559 | 2019-07-11 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -05.50/+1.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED SO THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR BUT IT WAS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -05.50/+1.5/094 | OD | Physician | 06/14/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170526&pc=QCB | 2023826-2021-02571 | 2020-10-21 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.00/+3.5/102 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -17.00/+3.5/102 | OD | Physician | 06/14/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170528&pc=QCB | 2023826-2021-02629 | 2021-02-17 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50/+2.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -07.50/+2.0/092 | OS | Physician | 06/14/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170538&pc=QCB | 2023826-2021-02572 | 2021-04-13 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/098 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -11.50/+1.5/098 | OD | Physician | 06/14/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170541&pc=QCB | 2023826-2021-02630 | 2021-04-06 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+1.5/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4). | -13.00/+1.5/091 | OS | Physician | 06/14/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170549&pc=QCB | 2023826-2021-02596 | 2021-04-15 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.5/3.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -10.5/3.5/082 | OS | Physician | 06/14/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170550&pc=QCB | 2023826-2021-02631 | 2021-02-06 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00/+1.5/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -09.00/+1.5/083 | OS | Physician | 06/14/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170560&pc=QCB | 2023826-2021-02632 | 2021-02-09 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+2.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -15.00/+2.0/088 | OD | Physician | 06/14/2021 | 2021 | 17
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170584&pc=QCB | 2023826-2021-02634 | 2021-02-09 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -05.00/+1.5/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -05.00/+1.5/080 | OD | Physician | 06/14/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170593&pc=QCB | 2023826-2021-02592 | 2021-02-26 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -11.5/1.5/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -11.5/1.5/94 | OS | Physician | 06/14/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170653&pc=QCB | 2023826-2021-02586 | 2021-03-03 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -8.5/1.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -8.5/1.0/095 | OS | Physician | 06/14/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170691&pc=QCB | 2023826-2021-02583 | 2020-12-01 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.00/4.5/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -13.00/4.5/074 | OS | Physician | 06/14/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170731&pc=QCB | 2023826-2021-02589 | 2021-02-20 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.00/1.5/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS WITHIN THE SAME SURGERY DUE TO LOW VAULT. PROBLEM RESOLVED. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -13.00/1.5/070 | OS | Physician | 06/14/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170749&pc=QCB | 2023826-2021-02593 | 2021-03-20 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.00/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -10.00/2.0/090 | OS | Physician | 06/14/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170759&pc=QCB | 2023826-2021-02591 | 2021-03-14 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.0/4.5/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -14.0/4.5/91 | OS | Physician | 06/14/2021 | 2021 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170776&pc=QCB | 2023826-2021-02594 | 2021-03-04 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -14.0/2.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS WITHIN THE SAME SURGERY DUE TO LOW VAULT. PROBLEM RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.0/2.5/080 | OS | Physician | 06/14/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170791&pc=QCB | 2023826-2021-02599 | 2021-02-01 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -7.5/3.0/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.5/3.0/077 | OD | Physician | 06/14/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170798&pc=QCB | 2023826-2021-02613 | 2020-11-16 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.0/1.0/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). WEIGHT - RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -11.0/1.0/109 | OS | Physician | 06/14/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170999&pc=QCB | 2023826-2021-02579 | 2021-02-24 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -11.0/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY, PROBLEM RESOLVED. CAUSE WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.0/2.0/091 | OS | Physician | 06/14/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12171056&pc=QCB | 2023826-2021-02587 | 2021-03-08 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -11.5/2.0/88 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. CAUSE OF THIS EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT ID: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.5/2.0/88 | OS | Physician | 06/14/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12167127&pc=QCB | 2023826-2021-02528 | 2021-05-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.00/1.5/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED, INCISION ENLARGEMENT AND ADDITIONAL SUTURES WERE PERFORMED DUE TO EXCESSIVE VAULT ON (B)(6) 2021. CAUSE OF THE EVENT IS REPORTED AS DEVICE. REPORTEDLY, THERE ARE PLANS TO IMPLANT A SMALLER LENS. Manufacturer Narrative: H6: WORK ORDER SEARCH FOUND ONE SIMILAR EVENT - CLAIM# (B)(4); EXCESSIVE VAULT; INTRA-OPERATIVE REMOVAL AND REPLACEMENT; PROBLEM RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: H3: THE LENS RETURNED IN A LIQUID WITH RESIDUE/DEBRIS MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THREAD/FIBRE ON A NO VISIBLE DAMAGE LENS. DIMENSION INSPECTION FOUND LENS TO BE WITHIN SPECIFICAITONS. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -8.00/1.5/073 | OD | Physician | 06/24/2021 | 2021 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12167136&pc=QCB | 2023826-2021-02529 | 2021-05-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.50/1.00/114 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED, INCISION ENLARGEMENT AND ADDITIONAL SUTURES WERE PERFORMED DUE TO EXCESSIVE VAULT ON (B)(6) 2021. CAUSE OF THE EVENT IS REPORTED AS DEVICE. REPORTEDLY, THERE ARE PLANS TO IMPLANT A SMALLER LENS. Manufacturer Narrative: H3: LENS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PREMARKET IDENTIFICATION: THIS PRODUCT IS NOT MARKETED IN THE US. ADVERSE EVENT PROBLEM: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -7.50/1.00/114 | OS | Physician | 06/24/2021 | 2021 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170458&pc=QCB | 2023826-2021-02563 | 2021-02-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/+2.0/094 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITHIN THE SAME SURGERY WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6- INVESTIGATION TYPE: 4110- LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | -11.00/+2.0/094 | OS | Physician | 06/14/2021 | 2021 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170490&pc=QCB | 2023826-2021-02560 | 2021-04-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.00/+1.5/091 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC BROKEN. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.00/+1.5/091 | NA |  | 06/14/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170566&pc=QCB | 2023826-2021-02633 | 2021-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.50/+1.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN, BROKEN, BENT AND STRETCHED OUT. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | -17.50/+1.0/088 | OD |  | 06/14/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12170595&pc=QCB | 2023826-2021-02635 | 2021-05-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -09.50/+1.5/175 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC BENT. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | -09.50/+1.5/175 | OS |  | 06/14/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12171405&pc=QCB | 2023826-2021-02619 | 2021-03-15 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.6, -11.50/2.00/74, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -11.50/2.00/74 | OS | Physician | 06/16/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12171426&pc=QCB | 2023826-2021-02621 | 2021-04-07 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -10.0/2.00/94, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -10.0/2.00/94 | OS | Physician | 06/16/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12177055&pc=QCB | 2023826-2021-02623 | 2021-05-11 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.1, -9.5/1.0/091, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO " LENS TOO SHORT. SURGEON PREFERENCE CHANGE." IT WAS REPORTED THAT THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -9.5/1.0/091 | OD | Physician | 06/15/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12177487&pc=QCB | 2023826-2021-02726 | 2021-06-14 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6; -15.0/3.5/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -15.0/3.5/095 | OD | Physician | 07/14/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178415&pc=QCB | 2023826-2021-02626 |  | Injury | STAAR SURGICAL COMPANY | 2021-07-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A TICL IMPLANTABLE COLLAMER LENS, -13.50/+2.0/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS). THE LENS WAS REPORTED AS HAVING ROTATED AND WAS REPOSITIONED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -13.50/+2.0/091 | OS | Other Health Care Professional | 06/18/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178425&pc=QCB | 2023826-2021-02628 | 2021-06-11 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -10.50/+0.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THERE IS A GOOD MEDIUM VAULT AND PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -10.50/+0.5/091 | OD | Physician | 06/16/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178432&pc=QCB | 2023826-2021-02707 | 2020-12-04 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. STATUS OF THE EYE: "CONSIDERING THAT THE ARCH HEIGHT OF THE PATIENT'S REPLACEMENT LENS MAY STILL BE LOW, THE LENS IS TAKEN OUT INSTEAD OF REPLACMENT." Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -8.5/+1.5/073 | OS | Physician | 06/21/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178436&pc=QCB | 2023826-2021-02715 | 2021-05-20 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+1.5/072 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE SURGEON REPORTS EXCESSIVE VAULT. ON (B)(6) 2021 IRIDECTOMIES WERE PERFORMED AND THE VAULT IMPROVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -15.0/+1.5/072 | OS | Physician | 06/23/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178437&pc=QCB | 2023826-2021-02716 | 2021-06-03 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -9.5/+2.5/092 | OS | Physician | 06/24/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178439&pc=QCB | 2023826-2021-02710 | 2021-03-09 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. INTRAOPERATIVELY THE LENS WAS IMPLANTED, REMOVED, AND REPLACED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -8.5/+1.5/083 | OS | Physician | 06/21/2021 | 2021 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178445&pc=QCB | 2023826-2021-02699 | 2021-06-22 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-15 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -13.0/+2.5/088 (SPHERE/CYLINDER/AXIS) TORE DURING INSERTION. THERE WAS PATIENT CONTACT WITH THE LENS HOWEVER THE LENS WAS NOT FULLY IMPLANTED. REPORTEDLY THE LENS GOT STUCK BETWEEN THE PLUNGER AND CARTRIDGE DURING IMPLANT. THIS OCCURRED ON (B)(6) 2021. Manufacturer Narrative: PT INFO: UNK. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.0/+2.5/088 | NA | Physician | 06/22/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178446&pc=QCB | 2023826-2021-02704 | 2020-12-02 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -12.5/+2.0/089 | OD | Physician | 06/21/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178450&pc=QCB | 2023826-2021-02720 | 2021-06-15 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -17.00/+1.5/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED AND EXCHANGED LATER THAT SAME DAY DUE TO LOW VAULTING. THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -17.00/+1.5/083 | OS | Physician | 06/17/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178453&pc=QCB | 2023826-2021-02705 | 2020-12-02 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+1.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -15.0/+1.5/074 | OD | Physician | 06/21/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178464&pc=QCB | 2023826-2021-02719 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-07-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.00/+1.0/095 (SPHERE/CYLINDER/AXIS), TORE BEFORE LOADING. THERE WAS NO PATIENT CONTACT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -08.00/+1.0/095 | NA | Physician | 06/17/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178503&pc=QCB | 2023826-2021-02727 | 2021-06-16 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6; -15.5/4.5/84 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -15.5/4.5/84 | OS | Physician | 07/14/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12171199&pc=QCB | 2023826-2021-02612 | 2021-05-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -2.0/2.0/89, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. EXCESSIVE VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -2.0/2.0/89, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021.ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -2.0/2.0/89 | OS |  | 06/14/2021 | 2021 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12171239&pc=QCB | 2023826-2021-02611 | 2021-05-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -2.0/2.0/86 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT RIGHT EYE (OD) ON (B)(6) 2021. EXCESSIVE VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -2.0/2.0/86, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -2.0/2.0/86 | OD |  | 06/14/2021 | 2021 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178430&pc=QCB | 2023826-2021-02706 | 2020-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B3- PER EMAIL ON 15-DEC-2021 THE DATE OF EVENT IS (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -8.5/+1.0/089 | OD | Physician | 06/21/2021 | 2021 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178434&pc=QCB | 2023826-2021-02717 | 2021-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+2.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT AXIS LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -14.5/+2.0/090 | OD |  | 06/24/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178440&pc=QCB | 2023826-2021-02703 | 2020-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.0/+4.0/104 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: D6B: PER EMAIL RECEIVED ON 07-DEC-2021, THE LENS WAS EXPLANTED ON (B)(6) 2020. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.0/+4.0/104 | OD | Physician | 06/21/2021 | 2021 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12178454&pc=QCB | 2023826-2021-02718 | 2021-07-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+3.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: LENS WAS REPOSITIONED ON (B)(6) 2021 AND PROBLEM RESOLVED. REPORTEDLY, "PATIENT IS HAPPY WITH BINOCULAR VISION." CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.50/+3.0/094 | OS | Physician | 06/17/2021 | 2021 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12184123&pc=QCB | 2023826-2021-02731 | 2021-05-06 | Injury | STAAR SURGICAL COMPANY | 2021-07-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.0/1.5/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -12.0/1.5/95 | OS | Physician | 07/14/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12193940&pc=QCB | 2023826-2021-02793 | 2021-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.5/2.5/88 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: G3 - DATE IN INITIAL MEDWATCH REPORT SHOULD HAVE BEEN 18-JUL-2021. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -11.5/2.5/88 | OD |  | 07/18/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12196829&pc=QCB | 2023826-2021-02794 | 2021-06-11 | Injury | STAAR SURGICAL COMPANY | 2021-07-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.0/3.5/81 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -12.0/3.5/81 | OS | Physician | 07/18/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12200113&pc=QCB | 2023826-2021-02739 | 2020-04-28 | Injury | STAAR SURGICAL COMPANY | 2021-07-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -11.5/1.5/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON(B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT- RACE: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/1.5/075 | OD | Physician | 06/21/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12200116&pc=QCB | 2023826-2021-02755 | 2021-06-09 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -7.5/1.0/89, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS WITHIN THE SAME SURGERY DUE TO LOW VAULT. PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -7.5/1.0/89 | OS | Physician | 06/22/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12200117&pc=QCB | 2023826-2021-02736 | 2021-05-20 | Injury | STAAR SURGICAL COMPANY | 2021-07-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.0/3.5/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.0/3.5/94 | OS | Physician | 06/20/2021 | 2021 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12200119&pc=QCB | 2023826-2021-02757 | 2021-06-18 | Injury | STAAR SURGICAL COMPANY | 2021-07-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -10.0/1.5/100, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS LATER EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTER AS OTHER. Manufacturer Narrative: WEIGHT-RACE: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.0/1.5/100 | OD | Physician | 06/23/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12200120&pc=QCB | 2023826-2021-02752 | 2021-06-13 | Injury | STAAR SURGICAL COMPANY | 2021-07-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.0/4.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.0/4.0/090 | OD | Physician | 06/21/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12200122&pc=QCB | 2023826-2021-02740 | 2020-08-03 | Injury | STAAR SURGICAL COMPANY | 2021-07-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -11.5/1.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -11.5/1.5/083 | OD | Physician | 06/21/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12200124&pc=QCB | 2023826-2021-02737 | 2021-02-01 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -5.50/1.00/92, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS WITHIN THE SAME SURGERY DUE TO LOW VAULT. PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN, THE PATIENT STATED " THE LENS SIZE RECOMMENDED BY OCOS DID NOT FIT THE PATIENT." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -5.50/1.00/92 | OS | Physician | 06/20/2021 | 2021 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12200519&pc=QCB | 2023826-2021-02773 | 2020-05-13 | Injury | STAAR SURGICAL COMPANY | 2021-07-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+2.0/109 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -12.50/+2.0/109 | OD | Physician | 06/21/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12199230&pc=QCB | 2023826-2021-02795 | 2021-04-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -13.0/1.5/93 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS)ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2021 WITH A SAME LENGTH/MODEL LENS; BUT IMPLANTED VERTICALLY DUE TO PATIENT EXPERIENCED EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/1.5/93 | OS | Physician | 07/18/2021 | 2021 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12200488&pc=QCB | 2023826-2021-02534 | 2021-06-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -09.00/+3.0/087 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS WAS INSERTED UPSIDE DOWN IN THE PATIENTS LEFT EYE (OS).THERE WAS NO PATIENT INJURY. THE REPORTER INDICATED THE LENS APPEARED NORMAL UPON LOADING BUT THE HAPTICS AND LENS WERE NOT AS FLEXIBLE AND DIFFICULT TO MOVE. THE SURGEON COULD NOT VISUALIZE THE MARKINGS ON THE HAPTICS TO ENSURE APPROPRIATE PLACEMENT AND THEREFORE COULD NOT BE SURE THE LENS WAS NOT FLIPPED. THE LENS WAS REPLACED WITHIN THE SAME SURGERY WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. POST-OP, VISION WAS ONLY HM AND THERE WAS SIGNIFICANT CORNEAL CLOUDING. SEE MFR. REPORT # 2023826-2021-02790 FOR REPLACEMENT LENS. THE LENS WAS DISCARDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS. AND THERE IS NO INDICATION, THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS -09.00/+3.0/087 (SPHERE/CYLINDER/AXIS) UPSIDE DOWN INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THERE WAS NO PATIENT INJURY. THE LENS WAS INTRAOPERATIVELY IMPLANTED; REMOVED AND REPLACED WITH ANOTHER 12.6MM LENS. THIS RESOLVED THE PROBLEM. POST-OP VISION WAS ONLY HM AND THERE WAS SIGNIFICANT CORNEAL CLOUDING. (B)(4). Manufacturer Narrative: WEIGHT: UNK, RACE: UNK, ETHNICITY: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.00/+3.0/087 | OS | Physician | 06/23/2021 | 2021 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12200499&pc=QCB | 2023826-2021-02771 | 2021-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -03.00/+6.0/013 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REMOVED ON 12-OCT-2021. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE CODE: (B)(4) THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -03.00/+6.0/013 | OS | Physician | 06/24/2021 | 2021 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12200503&pc=QCB | 2023826-2021-02791 | 2021-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS IN THE PATIENTS LEFT EYE (OS). THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO VAULT WAS >1.5MM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS -9.0/3.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS). REPORTEDLY ORIGINAL 13.2 LENS - VAULT>1.5MM. THE LENS WAS EXCHANGED ON (B)(6) 2021 FOR A SHORTER LENGTH LENS. THE BACKUP LENS WAS IMPLANTED UPSIDE DOWN. SEE MFR REPORT #: 2023826-2021-02534 FOR BACUP LENS. CLAIM # (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | -9.0/3.0/093 | OS | Physician | 06/23/2021 | 2021 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12200513&pc=QCB | 2023826-2021-02786 | 2021-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+1.5/121 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFO: B5: LENS WAS REMOVED AND REPLACED FOR A SAME LENGTH, DIFFERENT POWER LENS ON (B)(6) 2021. REPORTEDLY, "THE INITIAL PROBLEMS WITH THE VAULT IS SOLVED. THE REFRACTION IS NOT STABLE. IF STABILIZED, POST-OP CORRECTION BY CORNEAL LASER PROCEDURE IS INTENDED." CLAIM# (B)(4). Manufacturer Narrative: B5: LENS WAS REMOVED AND REPLACED FOR A SHORTER LENGTH, DIFFERENT POWER LENS ON (B)(6) 2021. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -16.50/+1.5/121 | OD | Physician | 06/24/2021 | 2021 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12204465&pc=QCB | 2023826-2021-02762 | 2021-05-16 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.7 MM, VTICMO13.7, -10.0/2.5/067 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -10.0/2.5/067 | OS | Physician | 06/24/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12204578&pc=QCB | 2023826-2021-02743 | 2020-11-06 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.0/1.5/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6)2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -10.0/1.5/104 | OD | Physician | 06/21/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12204599&pc=QCB | 2023826-2021-02744 | 2020-06-19 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.0/2.0/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -12.0/2.0/92 | OD | Physician | 06/21/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12204638&pc=QCB | 2023826-2021-02742 | 2020-10-03 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -11.5/2.5/091(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PATIENT INFORMATIONS :UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -11.5/2.5/091 | OS | Physician | 06/21/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12204766&pc=QCB | 2023826-2021-02756 | 2021-06-16 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -16.0/1.5/109(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -16.0/1.5/109 | OS | Physician | 06/22/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12207624&pc=QCB | 2023826-2021-02776 | 2020-12-13 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00/+2.0/091 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITHIN THE SAME SURGERY WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK , ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -08.00/+2.0/091 | NA | Physician | 06/21/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12207631&pc=QCB | 2023826-2021-02774 | 2019-11-20 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+2.0/111 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -09.50/+2.0/111 | OD | Physician | 06/21/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12207637&pc=QCB | 2023826-2021-02775 | 2020-12-08 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.50/+1.5/103 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -13.50/+1.5/103 | OD | Physician | 06/21/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12207644&pc=QCB | 2023826-2021-02778 | 2020-11-19 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+2.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -10.50/+2.5/090 | OS | Physician | 06/21/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12207656&pc=QCB | 2023826-2021-02789 | 2021-01-04 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00/+1.0/031 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM # (B)(4). | -09.00/+1.0/031 | OD | Physician | 06/25/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12207984&pc=QCB | 2023826-2021-02766 | 2021-05-18 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -8.5/1.0/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT ETHNICITY RACE: UNK. PREMARKET IDENTIFICATION: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -8.5/1.0/93 | OD | Physician | 06/24/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12207987&pc=QCB | 2023826-2021-02761 | 2021-05-16 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -9.5/2.0/076 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -9.5/2.0/076 | OD | Physician | 06/24/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12207990&pc=QCB | 2023826-2021-02763 | 2021-05-28 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6 MM, VTICMO12.6, -8.5/1.0/87, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -8.5/1.0/87 | OD | Physician | 06/24/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12207991&pc=QCB | 2023826-2021-02765 | 2021-06-05 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.0/3.0/93 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021.ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -13.0/3.0/93 | OD | Physician | 06/24/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12207653&pc=QCB | 2023826-2021-02780 | 2020-04-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/098 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A HAPTIC TORN AND BROKEN. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.50/+2.0/098 | OD |  | 06/21/2021 | 2021 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214164&pc=QCB | 2023826-2021-02640 | 2021-04-27 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -10.50/3.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.50/3.0/095 | OD | Physician | 06/29/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214174&pc=QCB | 2023826-2021-02646 | 2021-05-24 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.50/2.5/162 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE SURGEON DECIDED TO IMPLANT THE LENS VERTICALLY. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.50/2.5/162 | OD | Physician | 06/29/2021 | 2021 | 69
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214187&pc=QCB | 2023826-2021-02862 |  | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.00/1.00/072, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGELS. A SHORTER LENGTH LENS WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. THE REPORTER INDICATED THAT THE CAUSE OF THIS EVENT WAS "USER ERROR", INCORRECT WTW ENTERED. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -8.00/1.00/072 | OS | Physician | 06/22/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214188&pc=QCB | 2023826-2021-02661 | 2021-03-03 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/1.5/160 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.00/1.5/160 | OD | Physician | 06/29/2021 | 2021 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214196&pc=QCB | 2023826-2021-02650 | 2021-03-12 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -5.00/2.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -5.00/2.0/079 | OS | Physician | 06/29/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214203&pc=QCB | 2023826-2021-02671 | 2021-03-01 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.50/3.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -10.50/3.5/088 | OD | Physician | 06/30/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214223&pc=QCB | 2023826-2021-02672 | 2021-06-04 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.00/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.00/2.0/088 | OD | Physician | 06/30/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214235&pc=QCB | 2023826-2021-02662 | 2021-02-06 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.00/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/1.0/085 | OS | Physician | 06/29/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214238&pc=QCB | 2023826-2021-02664 | 2021-04-30 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.50/3.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.50/3.0/090 | OD | Physician | 06/29/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214313&pc=QCB | 2023826-2021-02690 | 2021-02-21 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.50/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. DATE REC¿D BY MFR (2020 REPORTS) - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.50/1.5/092 | OS | Physician | 07/01/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214314&pc=QCB | 2023826-2021-02677 | 2021-03-10 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.00/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.00/1.0/091 | OS | Physician | 07/01/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214320&pc=QCB | 2023826-2021-02684 | 2021-03-31 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.00/1.5/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/1.5/075 | OD | Physician | 07/01/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214323&pc=QCB | 2023826-2021-02682 | 2020-05-05 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -12.50/1.5/003 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.50/1.5/003 | OS | Physician | 07/01/2021 | 2021 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214328&pc=QCB | 2023826-2021-02687 | 2021-03-23 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.50/1.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -8.50/1.5/087 | OD | Physician | 07/01/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214330&pc=QCB | 2023826-2021-02681 | 2021-02-22 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.50/1.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -15.50/1.5/105 | OD | Physician | 07/01/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214334&pc=QCB | 2023826-2021-02689 | 2021-06-04 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/1.0/0 93(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.00/1.0/0 93 | OS | Physician | 07/01/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214336&pc=QCB | 2023826-2021-02676 | 2021-06-11 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.50/3.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS OTHER, PAITIENT'S MEASUREMENTS WERE JUST WITHIN THE CRITICAL RANGE. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -15.50/3.5/092 | OS | Physician | 07/01/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214344&pc=QCB | 2023826-2021-02683 | 2020-05-05 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -6.50/1.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.50/1.0/084 | OS | Physician | 07/01/2021 | 2021 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214357&pc=QCB | 2023826-2021-02688 | 2021-06-10 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.50/1.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -7.50/1.5/087 | OD | Physician | 07/01/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12215312&pc=QCB | 2023826-2021-02884 | 2021-07-05 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.5/2.5/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -10.5/2.5/085 | OD | Physician | 07/21/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214067&pc=QCB | 2023826-2021-02863 | 2021-06-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM,TMICL13.2,-9.00/3.0/094, IMPLANTABLE COLLAMER LENS TORE DURING LOADING. IT WAS REPORTED THAT THERE WAS NO PATIENT CONTACT. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED, THAT A 13.2 MM,TMICL13.2,-9.00/3.0/094 , IMPLANTABLE COLLAMER LENS TORE DURING LOADING. IT WAS REPORTED, THAT THERE WAS NO PATIENT CONTACT. PER UPDATED INFORMATION, A BACKUP LENS WAS IMPLANTED SUCCESSFULLY. CAUSE OF EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -9.00/3.0/094 | NA |  | 06/22/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214178&pc=QCB | 2023826-2021-02651 | 2021-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO13.2, -16.00/1.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "LENS ARCH WAS TOO HIGH AND TOO LOW. PATIENT WORRIED ABOUT THE ARCH HEIGHT IS UNWILLING TO REPLACE WITH A NEW LENS." LENS REMOVED AT PATIENTS REQUEST. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 -16.00/1.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. H3: DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.00/1.5/087 | OD | Physician | 06/29/2021 | 2021 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214192&pc=QCB | 2023826-2021-02644 | 2021-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.50/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO LOW VAULT ON (B)(6) 2021. THIS DID NOT RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON 30/JUN/2021, "THERE ARE NO PLANS TO IMPLANT ANOTHER LENS, DUE TO THE LOW VAULT. " SEE MFR # 2021-02643 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC TORN, BROKEN, BENT, DEFORMED AND STRECHED OUT. CLAIM# (B)(4). Manufacturer Narrative: B4 CORRECTED AWARE DATE: 28-JUN-2021. CLAIM # (B)(4). Manufacturer Narrative: H3: DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.5/093 | OD | Physician | 06/28/2021 | 2021 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214217&pc=QCB | 2023826-2021-02643 | 2021-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.50/1.5/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS DID NOT RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. SEE MFR # 2021-02644 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: B4: AWARE DATE CORRECTION: 28-JUN-2021. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.5/099 | OD | Physician | 06/28/2021 | 2021 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214218&pc=QCB | 2023826-2021-02645 | 2021-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.50/2.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON 30/JUN/2021, "THERE ARE NO PLANS TO IMPLANT ANOTHER LENS, DUE TO THE LOW VAULT. " Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND OPTIC SPLIT AND HAPTIC BROKEN. CLAIM# (B)(4). Manufacturer Narrative: B4 CORRECTED AWARE DATE: 28-JUN-2021. CLAIM # (B)(4). Manufacturer Narrative: H3: DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/2.0/082 | OS | Physician | 06/28/2021 | 2021 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214339&pc=QCB | 2023826-2021-02691 | 2021-05-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -10.00/1.0/177 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SAME SIZE, SIMILAR (SPHERE/CYLINDER/AXIS) LENS ON (B)(6) 2021 FROM PATIENT'S LEFT RIGHT EYE (OS) DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE REPLACEMENT LENS WAS PLACED VERTICALLY INSTEAD OF HORIZONTALLY. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.00/1.0/177 | OS | Physician | 07/02/2021 | 2021 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214351&pc=QCB | 2023826-2021-02685 | 2021-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/1.5/113 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B3: UNK. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. DATE REC¿D BY MFR (2020 REPORTS) - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/1.5/113 | OS | Physician | 07/01/2021 | 2021 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12214866&pc=QCB | 2023826-2021-02861 | 2021-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -16.0/2.0/87 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -16.0/2.0/87 | OD |  | 07/20/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12218765&pc=QCB | 2023826-2021-02246 | 2019-09-19 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.50/1.0/117 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2019. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.0/117 | OD | Physician | 06/14/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12219713&pc=QCB | 2023826-2021-02885 | 2021-06-08 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.0/3.0/179 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -13.0/3.0/179 | OS | Physician | 07/21/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12220967&pc=QCB | 2023826-2021-02886 | 2021-07-15 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8.00/+1.50/83 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND IT WAS INDICATED A NEW REPLACEMENT LENS WOULD BE IMPLANTED AT A LATER DATE. THE PROBLEM WAS NOT RESOLVED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -8.00/+1.50/83 | OD | Physician | 07/15/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12221015&pc=QCB | 2023826-2021-02887 | 2021-07-15 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8.00/+1.50/80 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING AND IT WAS INDICATED A NEW REPLACEMENT LENS WOULD BE IMPLANTED AT A LATER DATE. THE PROBLEM WAS NOT RESOLVED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -8.00/+1.50/80 | OS | Physician | 07/15/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222269&pc=QCB | 2023826-2021-02859 | 2021-05-29 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM# (B)(4). | -10.5/+1.0/092 | OD | Physician | 07/02/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222273&pc=QCB | 2023826-2021-02855 | 2021-04-24 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.5/+1.5/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED WITH A SHORTER LENS INTRAOPERATIVELY DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -17.5/+1.5/080 | OS | Physician | 07/01/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222274&pc=QCB | 2023826-2021-02847 | 2020-10-12 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.0/+2.5/081 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -15.0/+2.5/081 | OD | Physician | 07/01/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222275&pc=QCB | 2023826-2021-02849 | 2021-04-02 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.0/+3.5/096 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.0/+3.5/096 | OD | Physician | 07/01/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222281&pc=QCB | 2023826-2021-02846 | 2021-02-08 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.5/+1.0/115 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -7.5/+1.0/115 | OD | Physician | 07/01/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222283&pc=QCB | 2023826-2021-02848 | 2021-03-22 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE CODE (B)(4): OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -10.5/+1.5/080 | OD | Physician | 07/01/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222285&pc=QCB | 2023826-2021-02841 | 2021-06-12 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.5/076 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY ON (B)(6) 2021. AN ALTERNATE LENS WAS IMPLANTED AT A LATER DATE AND THE PROBLEM WAS RESOLVED. REPORTEDLY, THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE (B)(4): LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.5/+2.5/076 | OS | Physician | 07/01/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222288&pc=QCB | 2023826-2021-02843 | 2021-04-10 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.5/+3.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT SPHERE/CYLINDER/AXIS LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+3.0/089 | OS | Physician | 06/30/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222293&pc=QCB | 2023826-2021-02835 | 2021-04-24 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+1.5/092 | OS | Physician | 06/30/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222294&pc=QCB | 2023826-2021-02824 | 2021-04-01 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -7.0/+1.0/041 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -7.0/+1.0/041 | OS | Physician | 06/29/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222299&pc=QCB | 2023826-2021-02825 | 2020-12-31 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.0/+2.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THEN REPLACED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+2.5/087 | OD | Physician | 06/29/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222305&pc=QCB | 2023826-2021-02821 | 2021-03-01 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -12.0/+2.5/089 | OD | Physician | 06/29/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222306&pc=QCB | 2023826-2021-02822 | 2021-04-27 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/+2.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THEN REPLACED WITH A LONGER LENS DUE TO LENS ROTATION. REPORTEDLY THE "TICL REPLACEMENT FOR ICL COMBINED WITH CORNEAL LASER SURGERY." THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/+2.5/088 | OS | Physician | 06/29/2021 | 2021 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222308&pc=QCB | 2023826-2021-02816 | 2020-12-03 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.5/+1.0/110 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -6.5/+1.0/110 | OS | Physician | 06/29/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222311&pc=QCB | 2023826-2021-02820 | 2020-12-14 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.5/+1.0/106 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+1.0/106 | OD | Physician | 06/29/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222312&pc=QCB | 2023826-2021-02823 | 2021-04-20 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.5/+1.0/106 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.5/+1.0/106 | OD | Physician | 06/29/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222314&pc=QCB | 2023826-2021-02807 | 2021-03-27 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.5/+3.0/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. INTRAOPERATIVELY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -17.5/+3.0/095 | OD | Physician | 06/29/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222318&pc=QCB | 2023826-2021-02813 | 2021-03-04 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.0/+2.0/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -8.0/+2.0/090 | OS | Physician | 06/29/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222321&pc=QCB | 2023826-2021-02804 | 2021-05-02 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -6.0/+4.0/105 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR: "THE OS TICL IS 13.2MM AND HAS GOOD STABILITY AND VAULT, BUT THAT THE SULCUS IN THE OD IS MORE NARROW THAN THE OS." Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -6.0/+4.0/105 | OD | Physician | 06/28/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222322&pc=QCB | 2023826-2021-02798 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/075 (SPHERE/CYLINDER/AXIS) TORE WHILE PULLING THROUGH THE CARTRIDGE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: SEX, WEIGHT, ETHNICITY, RACE-UNK. DATE OF EVENT-UNK. CLAIM # (B)(4). | -8.5/+1.0/075 | NA | Physician | 06/29/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222323&pc=QCB | 2023826-2021-02806 | 2020-12-27 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -7.0/+1.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH, DIFFERENT AXIS LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -7.0/+1.0/089 | OS | Physician | 06/29/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222328&pc=QCB | 2023826-2021-02800 | 2021-05-20 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.0/+3.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -9.0/+3.5/087 | OS | Physician | 06/27/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222272&pc=QCB | 2023826-2021-02858 | 2021-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/178 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -9.5/+1.5/178 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR AN ICL LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. STATE OF THE LENS RETURNED IS DAMAGED BEYOND INSPECTION. VISUAL INSPECTION FOUND THE OPTIC TORN AND HAPTIC BROKEN TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -9.5/+1.5/178 | OS | Physician | 07/02/2021 | 2021 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222280&pc=QCB | 2023826-2021-02842 | 2020-12-26 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/146 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) UPSIDE DOWN. THIS OCCURRED ON (B)(6) 2020. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. IT WAS REPLACED WITH AN ALTERNATE LENS AT A LATER DATE AND THE PROBLEM IS RESOLVED. REPORTEDLY, NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS OTHER. Manufacturer Narrative: B5: ADDITIONAL INFORMATION: REPORTEDLY THE LENS WAS DAMAGED DURING LOADING. (B)(4). Manufacturer Narrative: WEIGHT ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -8.5/+1.0/146 | OD |  | 07/01/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222298&pc=QCB | 2023826-2021-02844 | 2021-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -18.0/+2.5/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT CYLINDER/AXIS LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/+2.5/083 | OD | Physician | 07/01/2021 | 2021 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222309&pc=QCB | 2023826-2021-02839 | 2021-06-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+2.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2021. THE LENS WAS REMOVED INTRA-OPERATIVELY. AN ALTERNATE LENS WAS IMPLANTED AT A LATER DATE AND THE PROBLEM WAS RESOLVED. NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. THERE WAS RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. THERE WAS RESIDUE ON THE LENS SURFACE. CLAIM# (B)(4). Manufacturer Narrative: H3: LEASE DISREGARD PREVIOUS H3 STATEMENT REGARDING DEVICE EVALUATION. REPORTED IN ERROR. H6: INVESTIGATION CODE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -8.0/+2.0/095 | OD | Physician | 07/01/2021 | 2021 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222327&pc=QCB | 2023826-2021-02803 | 2019-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.5/+3.5/110 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THEN REPLACED WITH A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE AND A PATIENT-RELATED FACTOR: "SEEMS LIKE THE DEVICE WASN'T ACCURATE, AFTER REMEASURING THE WTW WE GOT A 12.5MM BUT THE OCOS LENS BASED ON THE ACD (3.07MM) + WTW WAS 13.2MM, SO I HAD TO CHANGE THE SIZE TO 13.7MM. SO IT'S BOTH A DEVICE ERROR AND A PATIENT RELATED FACTOR (BUT WAS DISCOVERED LATER)." Manufacturer Narrative: B5- CORRECTION: PREVIOUSLY REPORTED A CYLINDER VALUE OF +3.5. THE ACTUAL VALUE IS +3.0. PREVIOUSLY REPORTED THAT THE LENS WAS EXPLANTED AND THEN REPLACED. THE LENS WAS ACTUALLY EXCHANGED. CLAIM# (B)(4). Manufacturer Narrative: AGE, WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -10.5/+3.5/110 | OD |  | 06/28/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12222772&pc=QCB | 2023826-2021-02890 | 2021-05-18 | Injury | STAAR SURGICAL COMPANY | 2021-07-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.5/2.0/87, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -10.5/2.0/87 | OS | Physician | 06/24/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12229811&pc=QCB | 2023826-2021-02917 | 2021-07-14 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7; -8.5/+1.0/79 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE DURING INJECTION/DELIVERY INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THERE WAS NO PATIENT CONTACT/INJURY. IT WAS NOTED A REPLACEMENT LENS WAS ORDERED WITH A PLANNED FUTURE SURGERY AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM#: (B)(4). | -8.5/+1.0/79 | OD | Other Health Care Professional | 07/14/2021 | 2021 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12233827&pc=QCB | 2023826-2021-02870 | 2021-02-26 | Injury | STAAR SURGICAL COMPANY | 2021-07-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM513.2, -9.00/1.5/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.00/1.5/098 | OD | Physician | 07/06/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12233844&pc=QCB | 2023826-2021-02871 | 2021-02-26 | Injury | STAAR SURGICAL COMPANY | 2021-07-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM513.2, -11.50/2.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.50/2.5/090 | OS | Physician | 07/06/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12233850&pc=QCB | 2023826-2021-02879 | 2021-06-30 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.50/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REMOVED FROM THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021.THE LENS TORE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. A SAME SIZE, SIMILAR (SPHERE/CYLINDER/AXIS) LENS WAS IMPLANTED ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.50/1.0/089 | OS | Physician | 07/09/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12235547&pc=QCB | 2023826-2021-02933 | 2021-01-14 | Injury | STAAR SURGICAL COMPANY | 2021-07-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.0/2.5/110 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -13.0/2.5/110 | OD | Physician | 07/25/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12233889&pc=QCB | 2023826-2021-02875 | 2021-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). REPORTEDLY, A LENS REPOSITIONING IS PLANNED. LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT A 13.2MM, VTICM5 13.2, -11.00/6.0/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. LENS WAS REPOSTIONED ON (B)(6) 2021. THIS DID NOT RESOLVE THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, REFRACTIVE ERROR IN FORM OS ASTIGMATISM IS STILL EXISTING AFTER REPOSITIONING. LASER SURGICAL TREATMENT IS PLANNED. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: AGE - UNK. SEX - UNK. WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. DATE OF EVENT - UNK. PRODUCT CODE: UNK (ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK OR 'UNK' ENTRY). IMPLANT DATE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE MFG DATE - UNK. CLAIM# (B)(4). | -11.00/6.0/074 | OS |  | 07/05/2021 | 2021 | 47
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12233906&pc=QCB | 2023826-2021-02872 | 2020-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.00/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. REPORTEDLY, THERE IS NO PLAN TO IMPLANT ANOTHER LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: H3 - DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATON. CLAIM #: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.00/1.0/091 | OD | Physician | 07/06/2021 | 2021 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12236055&pc=QCB | 2023826-2021-02935 | 2021-07-13 | Injury | STAAR SURGICAL COMPANY | 2021-07-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.0/2.0/68 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/2.0/68 | OS | Physician | 07/26/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12240188&pc=QCB | 2023826-2021-02940 | 2021-07-07 | Injury | STAAR SURGICAL COMPANY | 2021-07-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED 12.6MM VTICMO12.6; -11.0/2.5/001 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/2.5/001 | OD | Physician | 07/26/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12240628&pc=QCB | 2023826-2021-02941 | 2021-07-17 | Injury | STAAR SURGICAL COMPANY | 2021-07-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.5/2.5/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/2.5/90 | OS | Physician | 07/26/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12241558&pc=QCB | 2023826-2021-02906 | 2021-06-15 | Injury | STAAR SURGICAL COMPANY | 2021-07-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -07.50/+3.0/132 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED WITH A LONGER LENS. THE CAUSE OF THE EVENT WAS DUE TO PATIENT-RELATED FACTOR BUT THE DEVICE ALSO FAILED TO PERFORM AS INTENDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | -07.50/+3.0/132 | OS | Physician | 06/28/2021 | 2021 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12241789&pc=QCB | 2023826-2021-02901 | 2021-06-02 | Injury | STAAR SURGICAL COMPANY | 2021-07-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+0.5/106 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -09.50/+0.5/106 | OD | Physician | 06/28/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12241792&pc=QCB | 2023826-2021-02913 | 2021-05-04 | Injury | STAAR SURGICAL COMPANY | 2021-07-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -07.00/+1.5/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE REPLACEMENT LENS WAS IMPLANTED VERTICALLY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | -07.00/+1.5/086 | OD | Physician | 06/28/2021 | 2021 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12241830&pc=QCB | 2023826-2021-02894 | 2021-06-28 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A BOTTOM PIECE OF A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/135 (SPHERE/CYLINDER/AXIS), WAS MISSING DUE TO LOADING AND THE SURGEON IMPLANTED THE LENS. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: AGE: UNK, SEX: UNK, WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, EXPLANT DATE: NA. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.00/+1.5/135 | NA | Other Health Care Professional | 06/28/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12242101&pc=QCB | 2023826-2021-02926 | 2021-02-17 | Injury | STAAR SURGICAL COMPANY | 2021-07-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.00/+2.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -14.00/+2.5/088 | OS | Physician | 06/29/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12243196&pc=QCB | 2023826-2021-02927 | 2021-02-26 | Injury | STAAR SURGICAL COMPANY | 2021-07-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.50/+1.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -12.50/+1.5/092 | OD | Physician | 06/29/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12243198&pc=QCB | 2023826-2021-02936 | 2021-02-19 | Injury | STAAR SURGICAL COMPANY | 2021-07-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.50/+3.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -12.50/+3.5/093 | OD | Physician | 06/29/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12241523&pc=QCB | 2023826-2021-02904 | 2020-12-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.50/+2.0/081 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS EXCHANGED WITH A SHORTER LENS BUT THE PROBLEM WAS NOT RESOLVED. SEE MFR. REPORT # 2023826-2021-02905 FOR THE REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS DUE TO PATIENT-RELATED FACTOR BUT THE DEVICE ALSO FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN LENS VIAL, WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS RESIDUE ON LENS. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -08.50/+2.0/081 | OS | Physician | 06/28/2021 | 2021 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12241931&pc=QCB | 2023826-2021-02905 | 2021-06-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -07.00/+2.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. THIS LENS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2021-02904. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: APPROXIMATELY 1 MONTH LATER POST-OP THE PROBLEM WAS RESOLVED AND THE PATIENT HAD NO COMPLAINT. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK, RACE: UNK, ETHNICITY: UNK, EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -07.00/+2.0/090 | OS |  | 06/28/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12246257&pc=QCB | 2023826-2021-02948 | 2021-02-18 | Injury | STAAR SURGICAL COMPANY | 2021-07-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -03.50/+1.0/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -03.50/+1.0/085 | OD | Physician | 06/29/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12246291&pc=QCB | 2023826-2021-02947 | 2021-04-01 | Injury | STAAR SURGICAL COMPANY | 2021-07-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+2.5/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO THE LENS TORE DURING DELIVERY. THERE WAS NO PATIENT INJURY. THE LENS WAS EXCHANGED FOR ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.00/+2.5/087 | OD | Physician | 06/29/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12246297&pc=QCB | 2023826-2021-02922 | 2021-05-14 | Injury | STAAR SURGICAL COMPANY | 2021-07-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -09.00/+1.0/142 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED FOR A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PREMARKET IDENTIFICATION: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -09.00/+1.0/142 | OD | Physician | 06/29/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12247491&pc=QCB | 2023826-2021-02988 | 2021-07-10 | Injury | STAAR SURGICAL COMPANY | 2021-07-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -10.0/1.0/73 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -10.0/1.0/73 | OS | Physician | 07/26/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12247808&pc=QCB | 2023826-2021-02989 | 2021-07-08 | Injury | STAAR SURGICAL COMPANY | 2021-07-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -12.0/2.0/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/2.0/89 | OS | Physician | 07/26/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12249566&pc=QCB | 2023826-2021-02993 | 2021-06-29 | Injury | STAAR SURGICAL COMPANY | 2021-07-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -7.5/1.5/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.5/1.5/91 | OS | Physician | 07/26/2021 | 2021 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12249642&pc=QCB | 2023826-2021-02991 | 2021-06-28 | Injury | STAAR SURGICAL COMPANY | 2021-07-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -12.0/2.0/084 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/2.0/084 | OS | Physician | 07/26/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12250440&pc=QCB | 2023826-2021-02949 | 2021-03-17 | Injury | STAAR SURGICAL COMPANY | 2021-07-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50/+3.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021 . THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.50/+3.5/090 | OD | Physician | 06/29/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12250447&pc=QCB | 2023826-2021-02950 | 2021-06-11 | Injury | STAAR SURGICAL COMPANY | 2021-07-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.5/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -10.50/+1.5/074 | OS | Physician | 06/29/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12250516&pc=QCB | 2023826-2021-02978 | 2021-04-28 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/101 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE AFTER OPENING THE VIAL. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.00/+2.0/101 | NA | Physician | 07/01/2021 | 2021 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12250499&pc=QCB | 2023826-2021-02952 | 2020-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/+5.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS EXPLANTED. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL/DIOPTER BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. AFTER STABILIZATION, THE UNCORRECTED ASTIGMATISM WAS TREATED WITH FS-LASIK. Manufacturer Narrative: H6: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.00/+5.0/088 | OD | Physician | 06/29/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12250500&pc=QCB | 2023826-2021-02986 | 2021-03-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.50/+2.5/093 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITHIN THE SAME SURGERY FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5: SHORTLY AFTER THE REPLACEMENT LENS WAS IMPLANTED, IT WAS NOTED THE LENS HAD A LOW VAULT AND WAS EXPLANTED. SEE MFR. REPORT # 2023826-2021-02985 FOR REPLACEMENT LENS. H3: THE DEVICE EVALUATION IS NOT REQUIRED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.50/+2.5/093 | OS |  | 06/29/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12250512&pc=QCB | 2023826-2021-02985 | 2021-03-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+2.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION NOT REQUIRED. THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+2.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-02986. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. D6A: (B)(6) 2021. D6B: (B)(6) 2021. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.50/+2.5/089 | OS |  | 06/30/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12255830&pc=QCB | 2023826-2021-02764 | 2021-05-31 | Injury | STAAR SURGICAL COMPANY | 2021-07-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.5/1.5/95 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -12.5/1.5/95 | OD | Physician | 06/24/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12255931&pc=QCB | 2023826-2021-02973 | 2021-04-08 | Injury | STAAR SURGICAL COMPANY | 2021-07-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+2.5/099 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.00/+2.5/099 | OD | Physician | 06/30/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12256673&pc=QCB | 2023826-2021-03010 | 2021-06-03 | Injury | STAAR SURGICAL COMPANY | 2021-07-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.5/1.5/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -9.5/1.5/90 | OS | Physician | 07/27/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12257023&pc=QCB | 2023826-2021-02980 | 2021-02-28 | Injury | STAAR SURGICAL COMPANY | 2021-07-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+1.0/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -15.00/+1.0/086 | OD | Physician | 07/01/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12257025&pc=QCB | 2023826-2021-02982 | 2021-02-10 | Injury | STAAR SURGICAL COMPANY | 2021-07-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/073 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -11.00/+1.5/073 | OS | Physician | 07/01/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12255902&pc=QCB | 2023826-2021-02997 | 2021-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -05.00/+0.50/019 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5: THE LENS WAS REPLACED ON (B)(6) 2021 WITH A LONGER LENS. AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. H3: DEVICE EVALUATION NOT REQUIRED. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -05.00/+0.50/019 | OD |  | 07/02/2021 | 2021 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12257020&pc=QCB | 2023826-2021-02984 | 2021-01-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00/+1.5/089 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS HAD LOW VAULTING. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION:THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -08.00/+1.5/089 | NA | Physician | 07/01/2021 | 2021 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12257021&pc=QCB | 2023826-2021-03003 | 2021-06-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -07.00/+1.0/005 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS HAD EXCESSIVE VAULTING. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN, BROKEN AND BENT. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -07.00/+1.0/005 | NA |  | 07/01/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12257024&pc=QCB | 2023826-2021-02981 | 2021-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-07-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+3.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR ANOTHER ICL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION OF WIDTH FOUND THE LENS TO BE WITHIN SPECIFICATIONS. DIMENSIONAL INSPECTION OF LENGTH FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. IT WAS COMMENTED THE LENS RETURNED IS NOT THE LENGTH STATED IN THE CLAIM. CLAIM # (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -10.0/+2.0/089 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2021 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. H3 - DEVICE EVALUATION: " DIMENSIONAL INSPECTION OF LENGTH FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING" SHOULD BE CHANGED TO "DIMENSIONAL INSPECTION OF LENGTH FOUND THE LENS TO BE WITHIN SPECIFICATIONS." CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.50/+3.0/091 | OD | Physician | 07/01/2021 | 2021 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12257203&pc=QCB | 2023826-2021-02996 | 2021-06-04 | Injury | STAAR SURGICAL COMPANY | 2021-07-31 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.00/+1.0/003 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT WAS SATISFIED. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -14.00/+1.0/003 | OS | Physician | 07/01/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12257205&pc=QCB | 2023826-2021-03002 | 2021-03-16 | Injury | STAAR SURGICAL COMPANY | 2021-07-31 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50/+2.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS EXCHANGED FOR A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -09.50/+2.0/093 | OS | Physician | 07/01/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12257220&pc=QCB | 2023826-2021-02962 | 2021-02-08 | Injury | STAAR SURGICAL COMPANY | 2021-07-31 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -9.5/+1.0/072 | OD | Physician | 07/05/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12257223&pc=QCB | 2023826-2021-02958 | 2021-06-18 | Injury | STAAR SURGICAL COMPANY | 2021-07-31 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS -11.0/+1.5/118 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -11.0/+1.5/118 | OD | Physician | 07/05/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12257225&pc=QCB | 2023826-2021-02961 | 2021-04-10 | Injury | STAAR SURGICAL COMPANY | 2021-07-31 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5/+2.0/075 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.5/+2.0/075 | OD | Physician | 07/05/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12257230&pc=QCB | 2023826-2021-02959 | 2021-06-20 | Injury | STAAR SURGICAL COMPANY | 2021-07-31 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/132 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.0/132 | OD | Physician | 07/05/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12257237&pc=QCB | 2023826-2021-03005 | 2021-05-28 | Malfunction | STAAR SURGICAL COMPANY | 2021-07-31 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00/+1.5/078 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS HAD EXCESSIVE VAULTING. THE LENS WAS REPLACED WITHIN THE SAME SURGERY WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -08.00/+1.5/078 | NA | Physician | 07/02/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12264413&pc=QCB | 2023826-2021-03042 | 2021-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -11.5/2.0/106 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE PATIENT EXPERIENCED EXCESSIVE VAULTING. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5 - THE REPORTER INDICATED, THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2; -11.5/2.0/106 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. AND ON (B)(6) 2021, THE LENS WAS EXPLANTED. AND LATER ON THE SAME DATE, A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/2.0/106 | OD |  | 07/16/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12264448&pc=QCB | 2023826-2021-03043 | 2021-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -11.5/2.0/81 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED EXCESSIVE VAULTING. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -11.5/2.0/81 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/2.0/81 | OS |  | 07/16/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12271662&pc=QCB | 2023826-2021-03016 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-08-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+1.0/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS). THE SURGEON FOUND THE TORIC LENS HAD NO ALIGNMENT MARKS, SO HE ALIGNED THE LENS ACCORDING TO THE PERI-OPTICAL HOLES WITH HIS EXPERIENCE. THE LENS REMAINED IMPLANTED. AT THE LAST VISIT, THE PATIENTS VA WAS 20/20. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: SEX: UNK, WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. DATE OF EVENT: UNK. IMPLANT DATE : UNK. EXPLANT DATE : NA. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). | -09.00/+1.0/088 | OS | Physician | 07/04/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12271696&pc=QCB | 2023826-2021-03020 | 2021-05-14 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.00/+3.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINED IMPLANTED. CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.00/+3.5/095 | OD | Physician | 07/05/2021 | 2021 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12278258&pc=QCB | 2023826-2021-03021 | 2021-06-01 | Injury | STAAR SURGICAL COMPANY | 2021-08-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.00/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -16.00/+2.0/090 | OD | Physician | 07/05/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12278265&pc=QCB | 2023826-2021-03026 | 2021-04-10 | Injury | STAAR SURGICAL COMPANY | 2021-08-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+1.0/008 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -11.50/+1.0/008 | OD | Physician | 07/05/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12278276&pc=QCB | 2023826-2021-03017 | 2021-04-27 | Injury | STAAR SURGICAL COMPANY | 2021-08-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+2.5/097 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -11.50/+2.5/097 | OD | Physician | 07/05/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12278289&pc=QCB | 2023826-2021-03051 | 2021-04-15 | Injury | STAAR SURGICAL COMPANY | 2021-08-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+1.5/092 | OD | Physician | 07/14/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12278290&pc=QCB | 2023826-2021-03045 | 2021-06-01 | Injury | STAAR SURGICAL COMPANY | 2021-08-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -11.5/+2.0/089 | OS | Physician | 07/11/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12278294&pc=QCB | 2023826-2021-03049 | 2021-06-16 | Injury | STAAR SURGICAL COMPANY | 2021-08-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+4.5/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -11.5/+4.5/085 | OS | Physician | 07/14/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12278296&pc=QCB | 2023826-2021-03047 | 2021-01-12 | Injury | STAAR SURGICAL COMPANY | 2021-08-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH, DIFFERENT SPHERE/AXIS LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.5/+1.0/088 | OS | Physician | 07/13/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12278316&pc=QCB | 2023826-2021-03027 | 2021-04-10 | Injury | STAAR SURGICAL COMPANY | 2021-08-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+2.0/177 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -11.00/+2.0/177 | OS | Physician | 07/05/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12278320&pc=QCB | 2023826-2021-03022 | 2021-06-01 | Injury | STAAR SURGICAL COMPANY | 2021-08-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.50/+1.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | -16.50/+1.5/092 | OS | Physician | 07/05/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12278342&pc=QCB | 2023826-2021-03038 | 2021-04-28 | Injury | STAAR SURGICAL COMPANY | 2021-08-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/130 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -09.50/+1.0/130 | OS | Physician | 07/14/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12278287&pc=QCB | 2023826-2021-03048 | 2021-02-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Contamination; Off-Label Use | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/094 (SPHERE/CYLINDER/AXIS) WAS DISCOVERED TO HAVE A FOREIGN BODY ADHESION. THIS OCCURRED ON (B)(6) 2021. THERE WAS NO PATIENT CONTACT WITH THE LENS. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND RESIDUE/DEBRIS ON THE LENS SURFACE. A FIBER-LIKE PARTICULATE WAS NOTED ON THE LENS OPTIC. (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER, WEIGHT, ETHNICITY, RACE-UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+1.5/094 | NA | Physician | 07/13/2021 | 2021 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12278293&pc=QCB | 2023826-2021-03052 | 2020-07-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -3.5/+2.0/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION AND LOW VAULT. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -3.5/+2.0/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED BUT DID NOT RESOLVE THE PROBLEM. THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS ON (B)(6) 2021 DUE TO LOW VAULT WITH ROTATION. THIS RESOLVED THE PROBLEM. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -3.5/+2.0/102 | OD |  | 07/14/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12278338&pc=QCB | 2023826-2021-03033 | 2021-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -07.00/+5.0/112 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL, WITH RESIDUE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND ONE HAPTIC BROKEN. DIMENSIONAL VERIFICATION WAS PERFORMED. AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -07.00/+5.0/112 | OD |  | 07/13/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12278383&pc=QCB | 2023826-2021-03034 | 2021-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -09.50/+5.5/087 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND BENT. DIMENSIONAL VERIFICATION WAS PERFORMED. AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.50/+5.5/087 | OS |  | 07/13/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12281706&pc=QCB | 2023826-2021-03068 | 2021-06-25 | Injury | STAAR SURGICAL COMPANY | 2021-08-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.50/4.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS TORE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS REPLACED ON (B)(6) 2021 WITH SAME LENGTH LENS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.50/4.0/095 | OD | Physician | 07/14/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12281707&pc=QCB | 2023826-2021-03063 | 2021-03-03 | Injury | STAAR SURGICAL COMPANY | 2021-08-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -1.00/1.0/113 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -1.00/1.0/113 | OS | Physician | 07/13/2021 | 2021 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12281716&pc=QCB | 2023826-2021-03070 | 2021-01-21 | Injury | STAAR SURGICAL COMPANY | 2021-08-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.00/1.5/117 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT WITH ROTATION WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2020, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS REPLACED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.00/1.5/117 | OD | Physician | 07/15/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12281717&pc=QCB | 2023826-2021-03071 | 2021-01-21 | Injury | STAAR SURGICAL COMPANY | 2021-08-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.50/2.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LOW VAULT WITH ROTATION WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2020, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS REPLACED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.50/2.5/087 | OS | Physician | 07/15/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12281695&pc=QCB | 2023826-2021-03062 | 2021-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -10.5/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SAME SIZE, SIMILAR (SPHERE/CYLINDER/AXIS) LENS LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY, RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN, BENT, AND DEFORMED. CLAIM# (B)(4). Manufacturer Narrative: H6- INVESTIGATION TYPE: 4110- LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.5/2.0/093 | OS | Physician | 07/11/2021 | 2021 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12281703&pc=QCB | 2023826-2021-03059 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/2.0/060 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. INFLAMMATION, DRY EYE, AND ANTIBIOTIC DROPS PRESCRIBED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS; -13.00/+2.0/60 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. EXCESSIVE VAULTING; SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES WAS OBSERVED; AND ELEVATED IOP WAS RECORDED. LENS REMAINS IMPLANTED. NO INJURY HAS BEEN REPORTED. CLAIM #: (B)(4). Manufacturer Narrative: D6A: IMPLANT DATE: (B)(6) 2021. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.00/2.0/060 | OS | Physician | 07/12/2021 | 2021 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12281708&pc=QCB | 2023826-2021-03064 | 2021-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.00/4.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, THE LENS WAS REPOSITIONED ON (B)(6) 2021, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS REPLACED FOR A SPHERICAL LENS AND THE PROBLEM WAS RESOLVED. REPORTEDLY, RESIDUAL ASTIGMATISM WILL BE TREATED BY LASER. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: H3: DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.00/4.0/093 | OD | Physician | 07/13/2021 | 2021 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12305261&pc=QCB | 2023826-2021-03116 | 2021-06-22 | Injury | STAAR SURGICAL COMPANY | 2021-08-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.0/2.5/91, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -12.0/2.5/91 | OS | Physician | 07/11/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12308401&pc=QCB | 2023826-2021-03078 | 2021-06-01 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -6.00/2.0/117 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING INJECTION/DELIVERY INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. ON (B)(6) 2021, A SAME SIZE, SIMILAR (SPHERE/CYLINDER/AXIS) LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -6.00/2.0/117 | OD | Physician | 07/19/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12308407&pc=QCB | 2023826-2021-03081 | 2021-06-01 | Injury | STAAR SURGICAL COMPANY | 2021-08-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.50/1.0/171 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -8.50/1.0/171 | OD | Physician | 07/21/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12308415&pc=QCB | 2023826-2021-03087 | 2021-07-16 | Injury | STAAR SURGICAL COMPANY | 2021-08-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.50/1.0/091 | OD | Physician | 07/22/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12308420&pc=QCB | 2023826-2021-03088 | 2021-07-16 | Injury | STAAR SURGICAL COMPANY | 2021-08-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -7.00/1.5/061 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY- UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -7.00/1.5/061 | OS | Physician | 07/22/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12308422&pc=QCB | 2023826-2021-03091 | 2021-06-03 | Injury | STAAR SURGICAL COMPANY | 2021-08-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. IN A SECOND SURGERY THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY- UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -10.50/2.5/092 | OS | Physician | 07/23/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12308225&pc=QCB | 2023826-2021-00920 | 2021-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.50/2.0/067 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR, PATIENT WAS WORRIED ABOUT LOW ARCH AND STRONGLY REQUESTED THE LENS REMOVAL. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.50/2.0/067 | OD | Physician | 07/19/2021 | 2021 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12308236&pc=QCB | 2023826-2021-03072 | 2021-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -6.50/0.5/167 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, LENS REPLACEMENT IS PLANNED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL AND CORRECTION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -6.50/0.5/167 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. H6 IMPACT CODE REPORTED AS 2199, NEW INFORMATION REPORT AS 4629. CORRECTION: D6B - REPORTED AS ON (B)(6) 2021, CORRECTION TO ON (B)(6) 2021. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -6.50/0.5/167 | OD |  | 07/17/2021 | 2021 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12308243&pc=QCB | 2023826-2021-00921 | 2021-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/2.5/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR, PATIENT WAS WORRIED ABOUT LOW ARCH AND STRONGLY REQUESTED THE LENS REMOVAL. Manufacturer Narrative: DEVICE EVALUATION: DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/2.5/101 | OS | Physician | 07/19/2021 | 2021 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12308269&pc=QCB | 2023826-2021-03073 | 2021-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -4.50/0.5/117 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. REPORTEDLY, LENS REPLACEMENT IS PLANNED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -4.50/0.5/117 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. H6 IMPACT CODE REPORTED AS 2199, NEW INFORMATION REPORT AS 4629. REPORTED AS (B)(6) 2021, CORRECTION TO (B)(6) 2021. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -4.50/0.5/117 | OS |  | 07/17/2021 | 2021 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12308346&pc=QCB | 2023826-2021-03113 | 2021-07-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7, -9.50/3.0/101, IMPLANTABLE COLLAMER LENS WAS DAMAGED DURING LOADING. THERE WAS NO PATIENT CONTACT. CAUSE OF EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 13.7MM, -9.50/3.0/101, IMPLANTABLE COLLAMER LENS WAS DAMAGED DURING LOADING. THERE WAS NO PATIENT CONTACT. A REPLACEMENT LENS WAS IMPLANTED SUCCESSFULLY, PROBLEM RESOLVED. IT WAS REPORTED THAT THE CAUSE OF THIS EVENT WAS USER ERROR AND THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. FOR CAUSE DETAILS THE REPORTER STATED, "LENS GOT DAMAGED WHILE LOADING IT WITH THE LOADING FORCEPS INTO THE CARTRIDGE." CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -9.50/3.0/101 | NA | Physician | 07/12/2021 | 2021 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12308402&pc=QCB | 2023826-2021-03082 | 2021-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.00/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED IN A SECOND SURGERY ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. PER THE REPORTER ON (B)(6) 2021, A SMALLER LENGTH LENS WAS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, DRY. VISUAL INSPECTION FOUND OPTIC SPLIT AND HAPTIC BROKEN. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: PER EMAIL ON (B)(6) 2021 A REPLACEMENT LENS WAS NOT IMPLANTED IN THE PATIENT'S EYE. H3 - DEVICE EVALUATION: A DIMENSIONAL INSPECTION IS REQUIRED BUT COULD NOT BE PERFORMED DUE TO THE LENS BEING SIGNIFICANTLY DAMAGED; PER PRODUCT EVAULATION ON (B)(6) 2021. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -7.00/2.0/089 | OD | Physician | 07/21/2021 | 2021 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12308411&pc=QCB | 2023826-2021-03086 | 2021-07-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -11.50/1.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING INJECTION/DELIVERY INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THERE WAS PATIENT CONTACT (LENS TOUCHED THE EYE) WITH NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE, TEAR IN LENS, STOPPED INJECTION, PUSHED LENS OUT ABOVE TABLE, LENS TORE COMPLETELY. Manufacturer Narrative: B5: REMOVE: THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY ON (B)(6) 2021. ADD: "LENS WAS NOT IMPLANTED. IMPLANTATION OF A NEW LENS HAS BEEN PLANNED". CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.50/1.0/094 | OS | Physician | 07/23/2021 | 2021 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12315509&pc=QCB | 2023826-2021-03151 | 2021-04-28 | Injury | STAAR SURGICAL COMPANY | 2021-08-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -7.50/+1.5/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -7.50/+1.5/087 | OS | Physician | 07/28/2021 | 2021 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12315789&pc=QCB | 2023826-2021-03152 | 2021-06-11 | Injury | STAAR SURGICAL COMPANY | 2021-08-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -13.0/1.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -13.0/1.0/089 | OS | Physician | 07/27/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12316226&pc=QCB | 2023826-2021-03154 | 2021-07-02 | Injury | STAAR SURGICAL COMPANY | 2021-08-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.0/3.0/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4): OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | -13.0/3.0/90 | OD | Physician | 07/27/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12316392&pc=QCB | 2023826-2021-03155 | 2021-07-02 | Injury | STAAR SURGICAL COMPANY | 2021-08-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.5/3.0/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. MEDICAL DEVICE PROBLEM CODE: (B)(4): OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | -11.5/3.0/89 | OS | Physician | 07/27/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12316430&pc=QCB | 2023826-2021-03139 | 2021-06-06 | Injury | STAAR SURGICAL COMPANY | 2021-08-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.0/2.0/95, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -12.0/2.0/95 | OD | Physician | 07/16/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12316431&pc=QCB | 2023826-2021-03140 | 2021-06-07 | Injury | STAAR SURGICAL COMPANY | 2021-08-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -8.0/2.0/175, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -8.0/2.0/175 | OS | Physician | 07/16/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12316435&pc=QCB | 2023826-2021-03150 | 2021-05-07 | Injury | STAAR SURGICAL COMPANY | 2021-08-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -12.0/1.0/087, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -12.0/1.0/087 | OD | Physician | 07/14/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12316439&pc=QCB | 2023826-2021-03137 | 2021-06-02 | Injury | STAAR SURGICAL COMPANY | 2021-08-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -12.0/2.0/92, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.0/2.0/92 | OD | Physician | 07/15/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12319346&pc=QCB | 2023826-2021-03175 | 2021-05-14 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -4.0/0.5/99 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. DURING THE SAME INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -4.0/0.5/99 | OS | Physician | 07/27/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12319610&pc=QCB | 2023826-2021-03176 | 2021-07-20 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -8.5/1.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. DURING THE SAME INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -8.5/1.5/089 | OS | Physician | 07/27/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12321645&pc=QCB | 2023826-2021-03141 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-08-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER STATED THAT A 12.6MM,TMICL12.6, -14.00/3.0/119, IMPLANTABLE COLLAMER LENS TORE DURING LOADING. THERE WAS NO PATIENT CONTACT, BACK UP LENS WAS USED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -14.00/3.0/119 | NA | Physician | 07/15/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12321650&pc=QCB | 2023826-2021-03138 | 2021-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -3.50/1.0/104, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- IT WAS LATER REPORTED THAT THE SURGEON DECIDED TO PERFORM LRI (LIMBAL RELAXING INCISIONS) AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -3.50/1.0/104 | OD | Physician | 07/15/2021 | 2021 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12321652&pc=QCB | 2023826-2021-03149 | 2021-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM5_13.7, -7.50/1.5/105, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 13.7MM, VTICM513.7, -7.50/1.5/105, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 13.7MM, VTICM513.7, -7.50/1.5/105, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -7.50/1.5/105 | OD | Physician | 07/14/2021 | 2021 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12322357&pc=QCB | 2023826-2021-03104 | 2021-07-02 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+2.0/086 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITHIN THE SAME SURGERY WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM # (B)(4). | -13.00/+2.0/086 | OS | Physician | 07/23/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12322364&pc=QCB | 2023826-2021-03100 | 2021-07-09 | Injury | STAAR SURGICAL COMPANY | 2021-08-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -06.50/+3.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS EXPLANTED. THE LENS WAS REPLACED WITH A DIFFERENT MODEL/DIOPTER BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS WAS DISCARDED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -06.50/+3.0/092 | OD | Physician | 07/21/2021 | 2021 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12322365&pc=QCB | 2023826-2021-03101 | 2021-07-16 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON WAS LOADING A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50/+2.5/094 (SPHERE/CYLINDER/AXIS) AND THE ICL HAPTICS WERE BROKEN WHILE PULLING THE LENS IN THE CARTRIDGE. THERE WAS NO PATIENT CONTACT. THE LENS WAS REPLACED WITH A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT AND SURGEON ARE HAPPY THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -08.50/+2.5/094 | NA | Physician | 07/23/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12322372&pc=QCB | 2023826-2021-03133 | 2021-02-02 | Injury | STAAR SURGICAL COMPANY | 2021-08-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+1.5/089 | OD | Physician | 07/23/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12322377&pc=QCB | 2023826-2021-03134 | 2021-07-08 | Injury | STAAR SURGICAL COMPANY | 2021-08-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -4.5/+1.0/081 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -4.5/+1.0/081 | OS | Physician | 07/23/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12322378&pc=QCB | 2023826-2021-03128 | 2021-05-31 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-15 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.0/+1.0/129 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) UPSIDE DOWN. THIS OCCURRED ON (B)(6) 2021. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH AN ALTERNATE LENS. THE PROBLEM IS RESOLVED. REPORTEDLY, NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE (B)(4): LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -7.0/+1.0/129 | OS | Physician | 07/21/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12322383&pc=QCB | 2023826-2021-03121 | 2021-06-28 | Injury | STAAR SURGICAL COMPANY | 2021-08-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+4.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -13.0/+4.0/089 | OD | Physician | 07/18/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12322387&pc=QCB | 2023826-2021-03122 | 2021-06-08 | Injury | STAAR SURGICAL COMPANY | 2021-08-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -5.5/+2.5/002 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE SURGEON STATES THE ORIGINAL LENS WAS IMPLANTED VERTICALLY AND DECIDED TO IMPLANT THE REPLACEMENT LENS HORIZONTALLY BECAUSE OF HIGH VAULT. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.5/+2.5/002 | OD | Physician | 07/18/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12322439&pc=QCB | 2023826-2021-03108 | 2021-02-21 | Injury | STAAR SURGICAL COMPANY | 2021-08-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/+2.5/113 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | -11.00/+2.5/113 | OD | Physician | 07/20/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12322444&pc=QCB | 2023826-2021-03102 | 2021-07-16 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.50/+1.5/094 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS GOT STUCK IN THE INJECTOR AND TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. THE PATIENT AND SURGEON ARE HAPPY. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -08.50/+1.5/094 | OS | Physician | 07/23/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12322361&pc=QCB | 2023826-2021-03097 | 2021-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS RIGHT EYE (OD). THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -06.00/+1.5/080 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. H3: DEVICE EVALUATION NOT REQUIRED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -06.00/+1.5/080 | OD |  | 07/20/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12322363&pc=QCB | 2023826-2021-03095 | 2020-10-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -13.50/+3.0/052 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.50/+3.0/052 | OD |  | 07/19/2021 | 2021 | 50
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12322371&pc=QCB | 2023826-2021-03126 | 2021-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT AXIS LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: B5: THE LENS WAS IMPLANTED VERTICALLY AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H3- DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+1.0/093 | OS | Physician | 07/21/2021 | 2021 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12322375&pc=QCB | 2023826-2021-03125 | 2021-05-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/135 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THEN REPLACED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: PREVIOUSLY REPORTED THE LENS WAS EXPLANTED AND THEN REPLACED WITH A SHORTER LENS. CORRECTED INFO: THE LENS WAS EXCHANGED WITH A SHORTER LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -9.0/+1.5/135 | OS |  | 07/21/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12322386&pc=QCB | 2023826-2021-03124 | 2021-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.0/+4.0/096 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZED DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: CORRECTED INFORMATION: THE LENS WAS REPOSITIONED ON (B)(6) 2021, NOT ON (B)(6) 2021 AS PREVIOUSLY REPORTED. ADDITIONAL INFORMATION: THE LENS ROTATED AGAIN AFTER REPOSITIONING AND SO IT WAS REPLACED WITH A NON-TORIC LENS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -8.0/+4.0/096 | OS |  | 07/20/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12322434&pc=QCB | 2023826-2021-03096 | 2020-10-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -14.00/+3.0/106 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -14.00/+3.0/106 | OS |  | 07/19/2021 | 2021 | 50
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12322436&pc=QCB | 2023826-2021-03098 | 2021-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS LEFT EYE (OS). THE SURGEON PLANS TO EXPLANT AND EXCHANGE THE LENS. THE LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE REPORTER INDICATED, THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS, -05.00/+1.0/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021, DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS. AND THE PROBLEM WAS RESOLVED. H3: DEVICE EVALUATION NOT REQUIRED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -05.00/+1.0/085 | OS |  | 07/20/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12328515&pc=QCB | 2023826-2021-03184 | 2021-06-26 | Injury | STAAR SURGICAL COMPANY | 2021-08-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.5/5.0/082 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND LATER THE SAME DAY IN A SEPARATE SURGERY A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/5.0/082 | OS | Physician | 07/29/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12329160&pc=QCB | 2023826-2021-03214 | 2021-07-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -11.00/+2.5/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REMAINED IMPLANTED. A VECTOR ANALYSIS WAS REQUESTED AND NO ICL ROTATION WAS RECOMMENDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE REPORTER INDICATED THE PATIENTS CURRENT CONDITION IS SHE IS BEING TREATED FOR DRY EYES AND IS SEEING GREAT IMPROVEMENT. (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.00/+2.5/090 | OD | Other Health Care Professional | 07/23/2021 | 2021 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12331372&pc=QCB | 2023826-2021-03188 | 2021-04-24 | Injury | STAAR SURGICAL COMPANY | 2021-08-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.5/1.5/112, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/1.5/112 | OS | Physician | 07/19/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12331376&pc=QCB | 2023826-2021-03185 | 2020-06-10 | Injury | STAAR SURGICAL COMPANY | 2021-08-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -7.5/1.5/87, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -7.5/1.5/87 | OS | Physician | 07/18/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12331379&pc=QCB | 2023826-2021-03186 | 2020-11-19 | Injury | STAAR SURGICAL COMPANY | 2021-08-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.5/4.5/095, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/4.5/095 | OD | Physician | 07/18/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12331381&pc=QCB | 2023826-2021-03190 | 2021-05-19 | Injury | STAAR SURGICAL COMPANY | 2021-08-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.0/1.5/89, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT - RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -9.0/1.5/89 | OS | Physician | 07/19/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12331382&pc=QCB | 2023826-2021-03195 | 2001-05-01 | Injury | STAAR SURGICAL COMPANY | 2021-08-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -9.5/1.5/090, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -9.5/1.5/090 | OD | Physician | 07/19/2021 | 2021 | 5
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12331383&pc=QCB | 2023826-2021-03199 | 2021-06-24 | Injury | STAAR SURGICAL COMPANY | 2021-08-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -6.5/1.0/087, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -6.5/1.0/087 | OS | Physician | 07/19/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12331384&pc=QCB | 2023826-2021-03193 | 2021-03-01 | Injury | STAAR SURGICAL COMPANY | 2021-08-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.0/1.5/097, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE CODE (B)(4). CLAIM # (B)(4). | -12.0/1.5/097 | OD | Physician | 07/19/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12331385&pc=QCB | 2023826-2021-03197 | 2021-06-21 | Injury | STAAR SURGICAL COMPANY | 2021-08-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.5/1.0/99, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -14.5/1.0/99 | OD | Physician | 07/19/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12331387&pc=QCB | 2023826-2021-03201 | 2021-07-10 | Injury | STAAR SURGICAL COMPANY | 2021-08-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -6.0/1.0/094, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). WEIGHT - RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM # (B)(4). | -6.0/1.0/094 | OS | Physician | 07/19/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12331394&pc=QCB | 2023826-2021-03203 | 2021-04-28 | Injury | STAAR SURGICAL COMPANY | 2021-08-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -11.0/1.5/091, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -11.0/1.5/091 | OD | Physician | 07/19/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12331395&pc=QCB | 2023826-2021-03205 | 2021-01-29 | Injury | STAAR SURGICAL COMPANY | 2021-08-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.5/3.5/083, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.5/3.5/083 | OD | Physician | 07/19/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12331397&pc=QCB | 2023826-2021-03210 | 2020-12-29 | Injury | STAAR SURGICAL COMPANY | 2021-08-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -8.0/2.0/86, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT ETHNICITY AND RACE: UNK. PREMARKET IDENTIFICATION: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.0/2.0/86 | OS | Physician | 07/21/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12331399&pc=QCB | 2023826-2021-03211 | 2021-04-01 | Injury | STAAR SURGICAL COMPANY | 2021-08-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -12.0/3.0/94, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -12.0/3.0/94 | OS | Physician | 07/21/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12332804&pc=QCB | 2023826-2021-03187 | 2021-03-31 | Injury | STAAR SURGICAL COMPANY | 2021-08-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1 MM, VTICMO12.1 -18.0/5.0/112, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -18.0/5.0/112 | OD | Physician | 07/18/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12333685&pc=QCB | 2023826-2021-03261 | 2021-07-22 | Injury | STAAR SURGICAL COMPANY | 2021-08-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -10.5/2.0/84 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -10.5/2.0/84 | OS | Physician | 07/30/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12334187&pc=QCB | 2023826-2021-03262 | 2020-12-21 | Injury | STAAR SURGICAL COMPANY | 2021-08-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -12.50/+3.50/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -12.50/+3.50/90 | OD | Physician | 07/30/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12331393&pc=QCB | 2023826-2021-03207 | 2021-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6M, VTICMO12.6, -6.5/1.0/142, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THIS EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE, INSIDE A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -6.5/1.0/142 | OD | Physician | 07/20/2021 | 2021 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12338365&pc=QCB | 2023826-2021-03277 | 2021-05-04 | Injury | STAAR SURGICAL COMPANY | 2021-08-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -12.5/1.5/89 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: G4-PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -12.5/1.5/89 | OD | Physician | 07/31/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12338859&pc=QCB | 2023826-2021-03278 | 2021-04-11 | Injury | STAAR SURGICAL COMPANY | 2021-08-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.0/1.5/109 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -9.0/1.5/109 | OD | Physician | 07/31/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12339314&pc=QCB | 2023826-2021-03279 | 2021-03-26 | Injury | STAAR SURGICAL COMPANY | 2021-08-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.5/2.0/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM#: (B)(4). | -12.5/2.0/094 | OS | Physician | 07/31/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12339690&pc=QCB | 2023826-2021-03280 | 2021-05-28 | Injury | STAAR SURGICAL COMPANY | 2021-08-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -10.5/2.0/93 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -10.5/2.0/93 | OD | Physician | 07/31/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12340116&pc=QCB | 2023826-2021-03215 | 2021-07-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND REMAINED IMPLANTED. A VECTOR ANALYSIS WAS REQUESTED AND NO ICL ROTATION WAS RECOMMENDED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE REPORTER INDICATED THE PATIENTS CURRENT CONDITION IS SHE IS BEING TREATED FOR DRY EYES AND IS SEEING GREAT IMPROVEMENT. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.00/+1.5/091 | OS | Other Health Care Professional | 07/23/2021 | 2021 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12341283&pc=QCB | 2023826-2021-03156 | 2021-05-22 | Injury | STAAR SURGICAL COMPANY | 2021-08-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.50/2.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -9.50/2.0/079 | OS | Physician | 07/26/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12341285&pc=QCB | 2023826-2021-03157 | 2021-06-18 | Injury | STAAR SURGICAL COMPANY | 2021-08-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -16.00/2.5/009 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -16.00/2.5/009 | OD | Physician | 07/26/2021 | 2021 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12341287&pc=QCB | 2023826-2021-03160 | 2021-05-19 | Injury | STAAR SURGICAL COMPANY | 2021-08-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.00/1.0/060 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.00/1.0/060 | OD | Physician | 07/25/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12341288&pc=QCB | 2023826-2021-03159 | 2021-05-03 | Injury | STAAR SURGICAL COMPANY | 2021-08-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.50/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.0/089 | OS | Physician | 07/26/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12341289&pc=QCB | 2023826-2021-03164 | 2021-06-09 | Injury | STAAR SURGICAL COMPANY | 2021-08-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -8.00/0.5/179 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK; ETHNICITY: UNK; RACE: UNK; THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.00/0.5/179 | OD | Physician | 07/27/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12341290&pc=QCB | 2023826-2021-03162 | 2021-02-08 | Injury | STAAR SURGICAL COMPANY | 2021-08-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.5/1.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -7.5/1.0/096 | OD | Physician | 07/26/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12341291&pc=QCB | 2023826-2021-03161 | 2021-05-08 | Injury | STAAR SURGICAL COMPANY | 2021-08-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.50/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -8.50/1.0/093 | OD | Physician | 07/25/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12341292&pc=QCB | 2023826-2021-03163 | 2021-02-08 | Injury | STAAR SURGICAL COMPANY | 2021-08-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -3.5/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -3.5/1.0/089 | OS | Physician | 07/26/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12341293&pc=QCB | 2023826-2021-03166 | 2021-07-02 | Injury | STAAR SURGICAL COMPANY | 2021-08-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -9.0/2.5/093 (SPHERE/ CYLINDER/ AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.0/2.5/093 | OS | Physician | 07/27/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12341299&pc=QCB | 2023826-2021-03172 | 2021-07-19 | Injury | STAAR SURGICAL COMPANY | 2021-08-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.00/1.5/086 (SPHERE/ CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED IN A SECOND SURGERY ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. IN A THIRD SURGERY ON (B)(6) 2021 A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -8.00/1.5/086 | OS | Physician | 07/29/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12341300&pc=QCB | 2023826-2021-03171 | 2021-07-01 | Injury | STAAR SURGICAL COMPANY | 2021-08-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/1.5/113 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -11.00/1.5/113 | OS | Physician | 07/29/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12341301&pc=QCB | 2023826-2021-03169 | 2021-06-28 | Injury | STAAR SURGICAL COMPANY | 2021-08-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -12.50/4.50/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.50/4.50/100 | OS | Physician | 07/28/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12341305&pc=QCB | 2023826-2021-03174 | 2021-07-27 | Injury | STAAR SURGICAL COMPANY | 2021-08-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/1.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.50/1.0/095 | OS | Physician | 07/30/2021 | 2021 | 17
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12344344&pc=QCB | 2023826-2021-03299 | 2021-05-29 | Injury | STAAR SURGICAL COMPANY | 2021-08-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -9.5/2.0/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -9.5/2.0/89 | OS | Physician | 07/31/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12344791&pc=QCB | 2023826-2021-03301 | 2021-07-07 | Injury | STAAR SURGICAL COMPANY | 2021-08-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -7.0/1.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.0/1.0/089 | OD | Physician | 07/31/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12355616&pc=QCB | 2023826-2021-03313 | 2021-07-13 | Injury | STAAR SURGICAL COMPANY | 2021-08-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.5/2.0/74 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -10.5/2.0/74 | OS | Physician | 08/03/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12355792&pc=QCB | 2023826-2021-03314 | 2021-07-20 | Injury | STAAR SURGICAL COMPANY | 2021-08-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.5/1.5/93 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -9.5/1.5/93 | OS | Physician | 08/03/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12355881&pc=QCB | 2023826-2021-03315 | 2021-07-28 | Injury | STAAR SURGICAL COMPANY | 2021-08-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.0/1.0/098 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -11.0/1.0/098 | OD | Physician | 08/03/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12355968&pc=QCB | 2023826-2021-03316 | 2021-06-23 | Injury | STAAR SURGICAL COMPANY | 2021-08-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.0/1.5/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -9.0/1.5/90 | OD | Physician | 08/04/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12356249&pc=QCB | 2023826-2021-03317 | 2021-04-29 | Injury | STAAR SURGICAL COMPANY | 2021-08-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.0/2.5/93 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -11.0/2.5/93 | OD | Physician | 08/04/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12356719&pc=QCB | 2023826-2021-03319 | 2021-07-28 | Injury | STAAR SURGICAL COMPANY | 2021-08-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.5/3.0/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -12.5/3.0/088 | OD | Physician | 08/04/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12359489&pc=QCB | 2023826-2021-03264 | 2021-07-27 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -11.50/+1.0/105 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING AND THE LENS REMAINED IMPLANTED. THE PATIENT WILL BE MONITORED AND IS DOING WELL. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, EXPLANT DATE: NA. (B)(4). CLAIM # (B)(4). | -11.50/+1.0/105 | OD | Physician | 07/28/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12359553&pc=QCB | 2023826-2021-03270 |  | Injury | STAAR SURGICAL COMPANY | 2021-08-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50/+5.0/087 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS EYE IN 2015. THE LENS WAS EXPLANTED DUE TO PROBLEMS THE PATIENT HAD WITH THE ICL LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.50/+5.0/087 | NA | Physician | 07/26/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12359561&pc=QCB | 2023826-2021-03271 | 2021-05-28 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Unintended Movement | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/090 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS WAS INSERTED UPSIDE DOWN INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM # (B)(4). | -09.50/+1.0/090 | OD | Physician | 07/25/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12359610&pc=QCB | 2023826-2021-03272 | 2021-07-01 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.00/+1.5/095 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.00/+1.5/095 | OD | Physician | 07/27/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12359619&pc=QCB | 2023826-2021-03273 | 2020-06-30 | Injury | STAAR SURGICAL COMPANY | 2021-08-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50/+1.5/113 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE INITIAL LENS WAS IMPLANTED VERTICALLY AND THE REPLACEMENT LENS WAS IMPLANTED HORIZONTALLY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -08.50/+1.5/113 | OS | Physician | 07/26/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12359634&pc=QCB | 2023826-2021-03305 | 2021-07-06 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/075 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -09.50/+1.0/075 | OD | Physician | 07/27/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12359660&pc=QCB | 2023826-2021-03306 | 2021-06-30 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+1.5/086 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -09.00/+1.5/086 | OS | Physician | 07/27/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12363345&pc=QCB | 2023826-2021-03292 | 2021-05-18 | Injury | STAAR SURGICAL COMPANY | 2021-08-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.0/+3.0/064 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LENS ROTATION. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM# (B)(4). | -10.0/+3.0/064 | OD | Physician | 07/29/2021 | 2021 | 71
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12359518&pc=QCB | 2023826-2021-03268 | 2021-08-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/085 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TURNED. THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THERE WAS NO PATIENT INJURY. THE REPLACEMENT OF THE BACKUP LENS RESOLVED THE PROBLEM AND THE EYE IS STABLE. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.50/+2.0/085 | NA | Other Health Care Professional | 08/05/2021 | 2021 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12363340&pc=QCB | 2023826-2021-03294 | 2021-07-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.5/078 (SPHERE/CYLINDER/AXIS) TORE DURING INJECTION INTO THE LEFT (OS) EYE. THERE WAS PATIENT CONTACT WITH THE LENS BUT NO PATIENT INJURY. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY ON (B)(6) 2021. Manufacturer Narrative: ADDITIONAL INFORMATION: PER THE FINAL COMPLAINT QUESTIONNAIRE RECEIVED ON 25-FEB-2022, REMOVAL OF THE INITIAL LENS DID RESOLVE THE PROBLEM AND "THE PATIENT DOESN'T WISH A FURTHER SURGICAL INTERVENTION". CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+2.5/078 | OS | Physician | 07/29/2021 | 2021 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12363341&pc=QCB | 2023826-2021-03298 | 2021-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.0/+2.0/083 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZED DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "THE FEELING OF PUSHING IS HEAVY AND THE SLIPPERINESS OF THE CARTRIDGE IS REDUCED." Manufacturer Narrative: B5-ADDITIONAL INFORMATION: REPORTEDLY AFTER THE INITIAL LENS ROTATED AGAIN IT WAS REPLACED WITH A NON-TORIC LENS. A MONTH LATER LASIK WAS USED TO CORRECT THE ASTIGMATISM. H6-INVESTIGATION CODE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H6-INVESTIGATION CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW-BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. CLAIM# (B)(4). | -8.0/+2.0/083 | OD | Physician | 07/30/2021 | 2021 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12363343&pc=QCB | 2023826-2021-03291 | 2021-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -5.0/+1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT AXIS LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM# (B)(4). | -5.0/+1.0/092 | OS |  | 07/28/2021 | 2021 | 47
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12363344&pc=QCB | 2023826-2021-03297 | 2021-07-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5/+1.5/025 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2021. THE LENS WAS REMOVED INTRA-OPERATIVELY. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AT A LATER DATE. NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE (B)(4): LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -6.5/+1.5/025 | OD | Physician | 07/30/2021 | 2021 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12363346&pc=QCB | 2023826-2021-03290 | 2021-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.5/+2.0/122 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE (B)(4): LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -15.5/+2.0/122 | OD | Physician | 07/27/2021 | 2021 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12363348&pc=QCB | 2023826-2021-03287 | 2021-07-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.0/+1.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2021. THE LENS WAS REMOVED INTRA-OPERATIVELY. NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.0/+1.5/088 | OD | Physician | 07/27/2021 | 2021 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12363350&pc=QCB | 2023826-2021-03288 | 2021-06-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/111 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS EXCESSIVE VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5-UPDATED INFORMATION: ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS. THE PROBLEM IS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+1.5/111 | OD |  | 07/28/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12364611&pc=QCB | 2023826-2021-03222 | 2021-07-19 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -8.50/1.0/015 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -8.50/1.0/015 | OS | Physician | 08/01/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12364612&pc=QCB | 2023826-2021-03216 | 2021-07-05 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.50/3.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.50/3.0/087 | OD | Physician | 08/02/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12364613&pc=QCB | 2023826-2021-03219 | 2021-06-21 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.00/2.0/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -13.00/2.0/105 | OD | Physician | 08/01/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12364614&pc=QCB | 2023826-2021-03220 | 2021-07-17 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/1.5/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.50/1.5/109 | OD | Physician | 08/01/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12364615&pc=QCB | 2023826-2021-03221 | 2021-07-23 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/1.5/127 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -10.50/1.5/127 | OS | Physician | 08/01/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12364618&pc=QCB | 2023826-2021-03228 | 2021-07-19 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.50/2.0/162 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/2.0/162 | OD | Physician | 08/01/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12364619&pc=QCB | 2023826-2021-03225 | 2021-06-12 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.50/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.50/1.5/090 | OS | Physician | 08/01/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12364622&pc=QCB | 2023826-2021-03229 | 2021-01-08 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -15.00/2.5/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -15.00/2.5/077 | OS | Physician | 08/04/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12364623&pc=QCB | 2023826-2021-03235 | 2021-03-01 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.50/1.0/000 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.50/1.0/000 | OD | Physician | 08/04/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12364625&pc=QCB | 2023826-2021-03234 | 2020-11-16 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.50/1.0/092 | OD | Physician | 08/04/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12364627&pc=QCB | 2023826-2021-03233 | 2020-11-17 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.50/1.0/089 | OS | Physician | 08/04/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12364628&pc=QCB | 2023826-2021-03236 | 2021-03-02 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.50/1.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-UNK. ETHNICITY-UNK. RACE-UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.50/1.0/095 | OS | Physician | 08/04/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12364631&pc=QCB | 2023826-2021-03242 | 2021-05-15 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/1.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.00/1.0/095 | OD | Physician | 08/04/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12364633&pc=QCB | 2023826-2021-03238 | 2021-03-04 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.00/1.5/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.00/1.5/101 | OD | Physician | 08/04/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12364634&pc=QCB | 2023826-2021-03217 | 2021-03-30 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -17.00/1.5.0/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -17.00/1.5.0/107 | OD | Physician | 08/02/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12364635&pc=QCB | 2023826-2021-03243 | 2019-06-03 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/4.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2019. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/4.5/093 | OD | Physician | 08/04/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12364636&pc=QCB | 2023826-2021-03249 | 2020-12-29 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.00/2.5/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.00/2.5/074 | OS | Physician | 08/04/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12364637&pc=QCB | 2023826-2021-03248 | 2021-01-18 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/1.5/090 (SPHERE/ CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.50/1.5/090 | OS | Physician | 08/04/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12364641&pc=QCB | 2023826-2021-03247 | 2021-02-22 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.50/1.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.5/085 | OS | Physician | 08/04/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12364642&pc=QCB | 2023826-2021-03246 | 2020-12-22 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/0.5/045 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.50/0.5/045 | OD | Physician | 08/04/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12364644&pc=QCB | 2023826-2021-03253 | 2021-04-29 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.50/2.0/093 | OD | Physician | 08/05/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12364645&pc=QCB | 2023826-2021-03230 | 2021-06-01 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -4.00/1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -4.00/1.0/092 | OS | Physician | 08/04/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12364646&pc=QCB | 2023826-2021-03254 | 2021-07-27 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.00/2.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/2.5/094 | OS | Physician | 08/05/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12364649&pc=QCB | 2023826-2021-03252 | 2021-04-21 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.50/1.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.50/1.5/094 | OD | Physician | 08/05/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12364660&pc=QCB | 2023826-2021-03244 | 2021-01-25 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -5.50/2.0/117 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.50/2.0/117 | OS | Physician | 08/04/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12364668&pc=QCB | 2023826-2021-03259 | 2021-05-22 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -6.00/1.0/066 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.00/1.0/066 | OS | Physician | 08/06/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12366462&pc=QCB | 2023826-2021-03265 | 2021-07-27 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS HAVING EXCESSIVE VAULTING AND THE LENS REMAINED IMPLANTED. THE PATIENT WILL BE MONITORED AND IS DOING WELL. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: (B)(4). | -11.00/+1.5/089 | OS | Physician | 07/28/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12366472&pc=QCB | 2023826-2021-03310 | 2021-07-22 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50/+2.0/093, INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS OBSERVED TO HAVE BEEN IMPLANTED UPSIDE DOWN AND THERE WAS NO PATIENT INJURY. THE LENS WAS EXPLANTED ON (B)(6) 2021 AND WAS RE-IMPLANTED INTO THE EYE. THERE WAS NO PATIENT INJURY REPORTED WITH THIS SECOND INTERVENTION AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: WEIGHT: UNK, RACE: UNK, ETHNICITY : UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -09.50/+2.0/093 | OS | Physician | 07/29/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12367470&pc=QCB | 2023826-2021-03353 | 2021-03-30 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.5/2.0/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -11.5/2.0/92 | OD | Physician | 08/04/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12367616&pc=QCB | 2023826-2021-03354 | 2021-05-28 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -13.5/3.5/128 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/3.5/128 | OD | Physician | 08/04/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12365637&pc=QCB | 2023826-2021-03307 | 2021-06-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -11.50/+3.5/096 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE SURGEON AND PATIENT HAVE DECIDED TO LEAVE THE LENS IMPLANTED IN THE EYE. THE IOP IS NORMAL AND THE PATIENT IS HAPPY WITH THE OUTCOME. THE PATIENT WILL BE MONITORED. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT CLAIM # (B)(4). | -11.50/+3.5/096 | OD |  | 07/28/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12366480&pc=QCB | 2023826-2021-03309 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50/+2.0/174 (SPHERE/CYLINDER/AXIS), WAS RETURNED TO THE MANUFACTURER. THERE WAS PATIENT CONTACT. THE CAUSE FOR THE RETURN WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -14.50/+2.0/174 | NA | Physician | 07/27/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12366488&pc=QCB | 2023826-2021-03311 |  | Injury | STAAR SURGICAL COMPANY | 2021-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00/+3.5/133 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODE/LENGTH LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. REFRACTIVE VERIFICATION WAS PERFORMED AND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.00/+3.5/133 | OD | Physician | 07/29/2021 | 2021 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12366495&pc=QCB | 2023826-2021-03312 | 2021-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.00/+2.0/094 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2021. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2021. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND ONE HAPTIC BENT. WIDTH AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. LENGTH (DIAMETER) VERIFICATION WAS PERFORMED AND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H6: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, RACE: UNK ETHNICITY: UNK, EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -14.00/+2.0/094 | OD | Physician | 07/30/2021 | 2021 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12375184&pc=QCB | 2023826-2021-03381 | 2021-05-04 | Injury | STAAR SURGICAL COMPANY | 2021-08-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -9.0/1.0/117 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -9.0/1.0/117 | OS | Physician | 08/04/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12375278&pc=QCB | 2023826-2021-03382 | 2021-06-07 | Injury | STAAR SURGICAL COMPANY | 2021-08-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -3.0/1.5/101 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM#: (B)(4). | -3.0/1.5/101 | OS | Physician | 08/04/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12375391&pc=QCB | 2023826-2021-03322 | 2021-05-12 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50/+1.0/072 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS WAS INSERTED UPSIDE DOWN INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITHIN THE SAME SURGERY WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -06.50/+1.0/072 | OS | Physician | 07/31/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12375395&pc=QCB | 2023826-2021-03324 | 2021-06-04 | Injury | STAAR SURGICAL COMPANY | 2021-08-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50/+2.5/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION NOT REQUIRED. (B)(4). CLAIM # (B)(4). | -11.50/+2.5/093 | OD | Physician | 07/30/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12375399&pc=QCB | 2023826-2021-03325 | 2020-07-30 | Injury | STAAR SURGICAL COMPANY | 2021-08-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -14.00/+2.5/077 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED SO THE LENS WAS EXPLANTED. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION NOT REQUIRED. (B)(4). CLAIM # (B)(4). | -14.00/+2.5/077 | OS | Physician | 07/30/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12375420&pc=QCB | 2023826-2021-03358 | 2021-02-28 | Injury | STAAR SURGICAL COMPANY | 2021-08-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.00/+2.0/095 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO THE LENS TORE DURING DELIVERY AND LENS DISLOCATION/SUBLUXATION. THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS QUIET. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -16.00/+2.0/095 | OS | Physician | 07/30/2021 | 2021 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12375387&pc=QCB | 2023826-2021-03323 | 2021-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.00/+3.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED AND THE PROBLEM WAS RESOLVED BUT THE SURGEON FELT THE LENS MAY STILL ROTATE SO THE LENS WAS EXPLANTED. THE LENS WAS REPLACED WITH A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -11.00/+3.0/094 | OD | Physician | 07/29/2021 | 2021 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12375390&pc=QCB | 2023826-2021-03321 | 2021-04-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -06.00/+3.5/092 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITHIN THE SAME SURGERY WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6- INVESTIGATION TYPE: 4110- LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -06.00/+3.5/092 | OS | Physician | 08/02/2021 | 2021 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12375403&pc=QCB | 2023826-2021-03308 | 2021-06-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -09.00/+3.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE SURGEON AND PATIENT HAVE DECIDED TO LEAVE THE LENS IMPLANTED IN THE EYE. THE IOP IS NORMAL AND THE PATIENT IS HAPPY WITH THE OUTCOME. THE PATIENT WILL BE MONITORED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -09.00/+3.5/089 | OS |  | 07/28/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12378572&pc=QCB | 2023826-2021-03371 |  | Injury | STAAR SURGICAL COMPANY | 2021-08-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -10.0/+3.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LENS DISLOCATION/SUBLUXATION. IN A SEPARATE SURGERY A LONGER LENS WAS IMPLANTED. REPORTEDLY THE LENS WAS "TURNED DOWN IN THE EYE AFTER SURGERY BECAUSE OF THE SIZE OF THE RECOMMENDED LENS WAS SMALL." Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT -UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CARTRIDGE WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.0/+3.0/093 | OD | Physician | 08/02/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12378628&pc=QCB | 2023826-2021-03378 | 2021-07-26 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-27 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/084 (SPHERE/ CYLINDER/ AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) UPSIDE DOWN. THIS OCCURRED ON (B)(6) 2021. INTRAOPERATIVELY THE LENS WAS REMOVED. REPORTEDLY, NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.5/084 | OS | Physician | 08/06/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12380886&pc=QCB | 2023826-2021-03386 | 2021-01-26 | Injury | STAAR SURGICAL COMPANY | 2021-08-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.5/1.0/107 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM#: (B)(4). | -12.5/1.0/107 | OS | Physician | 08/04/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12380930&pc=QCB | 2023826-2021-03387 | 2021-01-25 | Injury | STAAR SURGICAL COMPANY | 2021-08-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.5/1.5/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM#: (B)(4). | -9.5/1.5/086 | OS | Physician | 08/04/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12378544&pc=QCB | 2023826-2021-03373 | 2021-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND HAPTIC BROKEN. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO HAPTIC FOOTPLATES DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.5/+2.5/089 | OD | Physician | 08/04/2021 | 2021 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12378549&pc=QCB | 2023826-2021-03359 | 2021-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS -10.0/+2.0/078 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE. REPORTEDLY, THE LENS WAS EXPLANTED DUE TO VAULT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH, SEX, WEIGHT, ETHNICITY, RACE : UNK. IMPLANT DATE,EXPLANT DATE: UNK. CLAIM# (B)(4). | -10.0/+2.0/078 | OD | Physician | 08/03/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12378592&pc=QCB | 2023826-2021-03365 | 2019-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/079 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REPOSITIONED DUE TO LOW VAULT AND LENS ROTATION. ON (B)(6) 2019 THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-CORRECTED INFORMATION: LENS REPOSITIONING AND LENS ROTATION REPORTED PREVIOUSLY IN ERROR. SEE MFR REP# 2023826-2021-03694 FOR REPLACEMENT LENS. H6-DEVICE CODE 2923 PREVIOUSLY REPORTED IN ERROR. INVESTIGATION TYPE 4110 AND WORK ORDER SEARCH RESULTS PREVIOUSLY REPORTED IN ERROR. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE CODE (B)(4): OFF-LABEL USE (ACD < 3.0MM).INVESTIGATION TYPE (B)(4): LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -9.5/+2.5/079 | OS |  | 08/01/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12378594&pc=QCB | 2023826-2021-03366 | 2021-07-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS -12.5/+1.5/090 (SPHERE/ CYLINDER/ AXIS) INTO THE PATIENT'S RIGHT EYE (OD), IT WAS TORN. THIS OCCURRED ON (B)(6) 2021. REPORTEDLY THE LENS REMAINS IMPLANTED AND NO PATIENT INJURY OCCURRED. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: REPORTEDLY THE PROBLEM IS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+1.5/090 | OD |  | 08/02/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12378623&pc=QCB | 2023826-2021-03372 | 2021-08-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -5.5/+2.0/117 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) IN 2014. THE LENS WAS REPOSITIONED. REPORTEDLY THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: CORRECTED INFORMATION: PREVIOUSLY REPORTED THE LENS WAS REPOSITIONED. THE LENS HAS NOT BEEN REPOSITIONED BUT IT IS PLANNED. ADDITIONAL INFORMATION: THE LENS WAS IMPLANTED (B)(6) 2014. THE LENS ROTATED. THE CAUSE IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -5.5/+2.0/117 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) IN (B)(6) 2014. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS REPORTED. THE LENS REMAINS IMPLANTED. THE PATIENT DOES NOT WANT ANY FURTHER SURGERY AND IS HAPPY. H10: MALFUNCTION. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT -UNK. IMPLANT DATE -UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -5.5/+2.0/117 | OD | Physician | 08/03/2021 | 2021 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12378630&pc=QCB | 2023826-2021-03374 | 2020-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/169 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT AXIS LENS DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY THE PATIENT HAD CYSTS FROM HORIZONTAL SULCUS TO SULCUS. THE ICL WAS IMPLANTED VERTICALLY THE FIRST TIME. "BECAUSE OF THE HEIGHT OF THE SUPPLY, HORIZONTAL IMPLANTATION WAS SELECTED FOR THE SECOND TIME." Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND FIBERS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+1.0/169 | OS | Physician | 08/04/2021 | 2021 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12381178&pc=QCB | 2023826-2021-03263 |  | Injury | STAAR SURGICAL COMPANY | 2021-08-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -07.50/+3.5/072 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS HAD A GOOD VAULT BUT HAD ROTATED SO WAS EXPLANTED AND AN MFIOL WAS IMPLANTED. THIS LENS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2021-01756. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. DATE OF EVENT: UNK. EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM # (B)(4). | -07.50/+3.5/072 | OS | Physician | 07/30/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12389112&pc=QCB | 2023826-2021-03412 | 2021-07-16 | Injury | STAAR SURGICAL COMPANY | 2021-08-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.5/0.5/108 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER". Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -11.5/0.5/108 | OD | Physician | 08/09/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12389113&pc=QCB | 2023826-2021-03414 | 2021-03-10 | Injury | STAAR SURGICAL COMPANY | 2021-08-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED 12.6MM VTICMO12.6; -17.0/2.5/78 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -17.0/2.5/78 | OD | Physician | 08/08/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12389114&pc=QCB | 2023826-2021-03405 | 2021-07-16 | Injury | STAAR SURGICAL COMPANY | 2021-08-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -7.5/2.0/71 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -7.5/2.0/71 | OS | Physician | 08/05/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12389115&pc=QCB | 2023826-2021-03407 | 2021-06-28 | Injury | STAAR SURGICAL COMPANY | 2021-08-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -13.0/3.0/93 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON( B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -13.0/3.0/93 | OS | Physician | 08/05/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12389119&pc=QCB | 2023826-2021-03404 | 2021-06-22 | Injury | STAAR SURGICAL COMPANY | 2021-08-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -10.5/2.0/075 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -10.5/2.0/075 | OS | Physician | 08/05/2021 | 2021 | 17
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12389121&pc=QCB | 2023826-2021-03406 | 2021-07-22 | Injury | STAAR SURGICAL COMPANY | 2021-08-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -7.5/2.0/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. MEDICAL DEVICE PROBLEM CODE: (B)(4): OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | -7.5/2.0/89 | OS | Physician | 08/05/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12389127&pc=QCB | 2023826-2021-03411 | 2021-08-06 | Injury | STAAR SURGICAL COMPANY | 2021-08-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -7.0/1.5/87 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. LATER ON THE SAME DATE IN A SEPARATE SURGERY; THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -7.0/1.5/87 | OD | Physician | 08/09/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12389116&pc=QCB | 2023826-2021-03413 | 2021-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.0/2.0/86 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B4 - DATE SHOULD BE CORRECTED TO (B)(6) 2021. G3 - DATE ENTERED IN INITIAL MEDWATCH REPORT SHOULD BE CORRECTED TO 08-AUG-2021. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM#: (B)(4). | -11.0/2.0/86 | OD |  | 08/08/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12389118&pc=QCB | 2023826-2021-03408 | 2021-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -9.0/1.5/94 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH/MODEL LENS BUT DIFFERENT AXIS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -9.0/1.5/94 | OD | Physician | 08/09/2021 | 2021 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12399254&pc=QCB | 2023826-2021-03394 | 2021-07-12 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-31 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD), BUT THE LENS WAS STUCK IN THE CARTRIDGE. THERE WAS PATIENT CONTACT BUT NO INJURY. THE LENS WAS REPLACED WITH A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING WELL. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # : (B)(4). | -12.00/+1.5/093 | OD | Physician | 08/05/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12399405&pc=QCB | 2023826-2021-03399 | 2021-04-13 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-31 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+2.0/091 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS WAS INSERTED UPSIDE DOWN INTO THE PATIENTS RIGHT EYE (OD). THE LENS WAS REPLACED WITHIN THE SAME SURGERY WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.50/+2.0/091 | OD | Physician | 08/05/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12399683&pc=QCB | 2023826-2021-03391 | 2021-07-28 | Malfunction | STAAR SURGICAL COMPANY | 2021-08-31 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.00/+3.0/062 (SPHERE/CYLINDER/AXIS), TORE DURING LOADING AND THERE WAS NO PATIENT CONTACT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS FINE. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -15.00/+3.0/062 | NA | Physician | 08/03/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12406738&pc=QCB | 2023826-2021-03335 | 2021-03-14 | Injury | STAAR SURGICAL COMPANY | 2021-09-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -13.00/1.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.00/1.0/083 | OD | Physician | 08/10/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12406741&pc=QCB | 2023826-2021-03340 | 2021-08-02 | Injury | STAAR SURGICAL COMPANY | 2021-09-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.50/1.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -9.50/1.5/094 | OD | Physician | 08/10/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12406780&pc=QCB | 2023826-2021-03329 | 2021-07-10 | Injury | STAAR SURGICAL COMPANY | 2021-09-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.50/1.0/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -8.50/1.0/072 | OS | Physician | 08/09/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12406785&pc=QCB | 2023826-2021-03337 | 2021-03-13 | Injury | STAAR SURGICAL COMPANY | 2021-09-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -09.50/2.50/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY- UNK. RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4) . | -09.50/2.50/093 | OS | Physician | 08/10/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12406789&pc=QCB | 2023826-2021-03344 | 2021-08-05 | Injury | STAAR SURGICAL COMPANY | 2021-09-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -13.50/1.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, VAULT POS OP DAY TWO 350 MICRON VISION 6/6 IOP 13. PATIENT HAPPY. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.50/1.5/097 | OD | Physician | 08/12/2021 | 2021 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12406791&pc=QCB | 2023826-2021-03345 | 2021-07-07 | Injury | STAAR SURGICAL COMPANY | 2021-09-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6, -13.5/3.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.5/3.0/092 | OS | Physician | 08/12/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12406856&pc=QCB | 2023826-2021-03332 | 2021-08-07 | Injury | STAAR SURGICAL COMPANY | 2021-09-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -6.00/0.5/120 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -6.00/0.5/120 | OS | Physician | 08/09/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12406873&pc=QCB | 2023826-2021-03348 | 2021-07-19 | Injury | STAAR SURGICAL COMPANY | 2021-09-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -09.00/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -09.00/2.5/093 | OS | Physician | 08/13/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12406881&pc=QCB | 2023826-2021-03346 | 2021-01-10 | Injury | STAAR SURGICAL COMPANY | 2021-09-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -12.00/3.0/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.00/3.0/111 | OD | Physician | 08/13/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12406931&pc=QCB | 2023826-2021-03333 | 2021-08-06 | Injury | STAAR SURGICAL COMPANY | 2021-09-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.00/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -7.00/1.0/090 | OD | Physician | 08/09/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12406936&pc=QCB | 2023826-2021-03343 | 2021-07-07 | Injury | STAAR SURGICAL COMPANY | 2021-09-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 13.2MM VTICMO13.2, -09.00/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. A NEW SHORTER LENGTH LENS WAS IMPLANTED ON (B)(6)2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -09.00/1.5/093 | OS | Physician | 08/10/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12406977&pc=QCB | 2023826-2021-03341 | 2021-07-12 | Injury | STAAR SURGICAL COMPANY | 2021-09-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.00/3.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -15.00/3.0/089 | OD | Physician | 08/10/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12407192&pc=QCB | 2023826-2021-03336 | 2021-05-24 | Injury | STAAR SURGICAL COMPANY | 2021-09-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.50/2.00/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -11.50/2.00/093 | OD | Physician | 08/10/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12411357&pc=QCB | 2023826-2021-03458 | 2021-07-12 | Injury | STAAR SURGICAL COMPANY | 2021-09-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -10.5/1.0/77 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -10.5/1.0/77 | OD | Physician | 08/12/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12411358&pc=QCB | 2023826-2021-03437 | 2021-07-30 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.0/1.0/24 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. INTRAOPERATIVELY THE LENS WAS REMOVED AND A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -9.0/1.0/24 | OS | Physician | 08/10/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12411037&pc=QCB | 2023826-2021-03401 | 2021-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -08.50/+1.0/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED LATER THAT SAME DAY DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (THE LENS WAS IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -08.50/+1.0/085 | OD | Physician | 08/02/2021 | 2021 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12411049&pc=QCB | 2023826-2021-03402 | 2021-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.5/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (THE LENS WAS IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.50/+1.5/074 | OS | Physician | 08/02/2021 | 2021 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12418647&pc=QCB | 2023826-2021-03462 | 2020-12-13 | Injury | STAAR SURGICAL COMPANY | 2021-09-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -9.0/1.0/113 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -9.0/1.0/113 | OD | Physician | 08/13/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12418650&pc=QCB | 2023826-2021-03463 |  | Injury | STAAR SURGICAL COMPANY | 2021-09-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -5.5/2.0/175 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -5.5/2.0/175 | OS | Physician | 08/13/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12418644&pc=QCB | 2023826-2021-03464 | 2021-04-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -8.5/2.5/012 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. INTRAOPERATIVELY THE LENS WAS REMOVED AND AT A LATER DATE A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE WIDTH WITHIN SPECIFICATIONS BUT THE OVERALL LENGTH (DIAMETER) FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -8.5/2.5/012 | OS | Physician | 08/13/2021 | 2021 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12425751&pc=QCB | 2023826-2021-03466 | 2021-04-16 | Injury | STAAR SURGICAL COMPANY | 2021-09-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -8.5/1.5/079 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -8.5/1.5/079 | OS | Physician | 08/14/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12425752&pc=QCB | 2023826-2021-03467 | 2021-07-30 | Injury | STAAR SURGICAL COMPANY | 2021-09-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -9.5/1.00/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -9.5/1.00/95 | OD | Physician | 08/18/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12425753&pc=QCB | 2023826-2021-03465 | 2021-07-24 | Injury | STAAR SURGICAL COMPANY | 2021-09-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -8.0/1.5/113 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -8.0/1.5/113 | OS | Physician | 08/14/2021 | 2021 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12425756&pc=QCB | 2023826-2021-03469 | 2021-08-06 | Injury | STAAR SURGICAL COMPANY | 2021-09-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -10.00/1.0/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -10.00/1.0/90 | OS | Physician | 08/18/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12425757&pc=QCB | 2023826-2021-03468 | 2021-07-30 | Injury | STAAR SURGICAL COMPANY | 2021-09-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -10.00/+1.00/98 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -10.00/+1.00/98 | OS | Physician | 08/18/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12425758&pc=QCB | 2023826-2021-03470 | 2021-07-11 | Injury | STAAR SURGICAL COMPANY | 2021-09-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -9.5/2.5/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -9.5/2.5/90 | OD | Physician | 08/18/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12426105&pc=QCB | 2023826-2021-03450 | 2021-02-18 | Injury | STAAR SURGICAL COMPANY | 2021-09-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.5/+2.0/081 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION WHICH DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM IS RESOLVED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE (B)(4): LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.5/+2.0/081 | OS | Physician | 08/10/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12426107&pc=QCB | 2023826-2021-03451 | 2021-05-26 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-03 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/098 (SPHERE/CYLINDER/AXIS) TORE BEFORE/DURING LOADING. REPORTEDLY, THE LENS WAS TORN WHEN CHECKED UNDER A MICROSCOPE BEFORE THE SURGERY. THIS OCCURRED ON (B)(6) 2021. THERE WAS NO PATIENT CONTACT WITH THE LENS AND THERE ARE NO PLANS TO IMPLANT A REPLACEMENT LENS. Manufacturer Narrative: AGE, WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -9.0/+1.5/098 | NA | Physician | 08/10/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12426110&pc=QCB | 2023826-2021-03448 | 2021-07-26 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-03 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.5/102 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2021. THE LENS WAS REMOVED INTRA-OPERATIVELY. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AT A LATER DATE AND THE PROBLEM IS RESOLVED. NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -11.0/+2.5/102 | OD | Physician | 08/10/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12426114&pc=QCB | 2023826-2021-03457 | 2021-01-18 | Injury | STAAR SURGICAL COMPANY | 2021-09-03 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -6.0/+1.0/136 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -6.0/+1.0/136 | OD | Physician | 08/13/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12426118&pc=QCB | 2023826-2021-03449 | 2021-02-18 | Injury | STAAR SURGICAL COMPANY | 2021-09-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.0/+3.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION WHICH DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM IS RESOLVED. Manufacturer Narrative: AGE OR DOB, WEIGHT, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. INVESTIGATION TYPE (B)(4): LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.0/+3.0/095 | OD | Physician | 08/10/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12426104&pc=QCB | 2023826-2021-03446 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-09-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER STATES A 12.1MM TMICL12.1 IMPLANTABLE COLLAMER LENS -8.0/+2.5/099 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2021. THE SURGEON REPORTS THE LENS IS TOO SMALL AND KEEPS ROTATING. REPORTEDLY THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5: PER EMAIL ON (B)(6) 2021. IT WAS REPORTED, THAT THE "LENS (B)(6) WILL NOT BE REMOVED AND RETURNED. THE LENS IS NOT ROTATING". CLAIM# (B)(4). Manufacturer Narrative: SEX, WEIGHT, ETHNICITY, RACE - UNK. DATE OF EVENT - UNK. (B)(4). CLAIM# (B)(4). | -8.0/+2.5/099 | OD | Physician | 08/13/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12426106&pc=QCB | 2023826-2021-03445 |  | Injury | STAAR SURGICAL COMPANY | 2021-09-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER STATES A 12.1MM TMICL12.1 IMPLANTABLE COLLAMER LENS -10.5/+3.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2021. THE SURGEON REPORTS THE LENS IS TOO SMALL AND KEEPS ROTATING. REPORTEDLY THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5- THE REPORTER STATES A 12.1MM TMICL12.1 IMPLANTABLE COLLAMER LENS -10.5/+3.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2021. THE LENS WAS EXPLANTED ON AN UNKNOWN DATE. D6B- UNK. H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND THE OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: SEX, WEIGHT, ETHNICITY - UNK. DATE OF EVENT - UNK. (B)(4). CLAIM# (B)(4). | -10.5/+3.0/090 | OS | Physician | 08/13/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12433311&pc=QCB | 2023826-2021-03425 | 2021-05-19 | Injury | STAAR SURGICAL COMPANY | 2021-09-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -11.00/1.0/093 | OD | Physician | 08/19/2021 | 2021 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12433312&pc=QCB | 2023826-2021-03422 | 2021-04-21 | Injury | STAAR SURGICAL COMPANY | 2021-09-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.00/3.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -12.00/3.0/090 | OS | Physician | 08/19/2021 | 2021 | 51
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12433313&pc=QCB | 2023826-2021-03421 | 2021-08-18 | Injury | STAAR SURGICAL COMPANY | 2021-09-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/2.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -11.00/2.0/092 | OD | Physician | 08/19/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12433315&pc=QCB | 2023826-2021-03423 | 2021-08-18 | Injury | STAAR SURGICAL COMPANY | 2021-09-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -08.00/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -08.00/1.5/092 | OS | Physician | 08/19/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12433316&pc=QCB | 2023826-2021-03424 | 2021-04-14 | Injury | STAAR SURGICAL COMPANY | 2021-09-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/1.5/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.50/1.5/108 | OS | Physician | 08/19/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12433317&pc=QCB | 2023826-2021-03415 | 2021-01-06 | Injury | STAAR SURGICAL COMPANY | 2021-09-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.50/1.5/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -8.50/1.5/112 | OS | Physician | 08/15/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12433323&pc=QCB | 2023826-2021-03428 | 2021-08-05 | Injury | STAAR SURGICAL COMPANY | 2021-09-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.50/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -7.50/1.5/091 | OS | Physician | 08/19/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12433325&pc=QCB | 2023826-2021-03427 | 2021-08-05 | Injury | STAAR SURGICAL COMPANY | 2021-09-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/3.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.00/3.5/089 | OS | Physician | 08/19/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12433327&pc=QCB | 2023826-2021-03417 | 2021-07-04 | Injury | STAAR SURGICAL COMPANY | 2021-09-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/1.5/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK, ETHNICITY- UNK, RACE- UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -10.00/1.5/070 | OD | Physician | 08/18/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12433328&pc=QCB | 2023826-2021-03431 | 2021-07-12 | Injury | STAAR SURGICAL COMPANY | 2021-09-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK, RACE- UNK, ETHNICITY- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -10.00/1.0/090 | OD | Physician | 08/20/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12433329&pc=QCB | 2023826-2021-03432 | 2021-05-06 | Injury | STAAR SURGICAL COMPANY | 2021-09-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -8.50/1.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED WITH A SHORTER LENGTH LENS INTRAOPERATIVELY ON (B)(6) 2021 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE PROGRESS IS GOOD. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -8.50/1.5/080 | OD | Physician | 08/20/2021 | 2021 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12433332&pc=QCB | 2023826-2021-03426 | 2021-05-21 | Injury | STAAR SURGICAL COMPANY | 2021-09-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -03.50/0.5/063 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -03.50/0.5/063 | OS | Physician | 08/19/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12433336&pc=QCB | 2023826-2021-03433 | 2021-05-06 | Injury | STAAR SURGICAL COMPANY | 2021-09-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -09/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED WITH A SHORTER LENGTH LENS INTRAOPERATIVELY ON (B)(6) 2021 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE PROGRESS IS GOOD. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -09/1.5/091 | OS | Physician | 08/20/2021 | 2021 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12437058&pc=QCB | 2023826-2021-03479 | 2021-07-22 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/091 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PT INFO: UNK. REPORT SOURCE: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | -10.00/+2.0/091 | OS | Physician | 08/08/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12437075&pc=QCB | 2023826-2021-03472 | 2021-08-05 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON WAS ATTEMPTING TO LOAD A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/087 (SPHERE/CYLINDER/AXIS), AND THE LENS WOULD NOT LOAD INTO THE CARTRIDGE AND TORE. THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED WITH NO PROBLEM. THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS WAS DISCARDED. Manufacturer Narrative: PT INFO: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.50/+1.0/087 | NA | Physician | 08/11/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12433310&pc=QCB | 2023826-2021-03420 | 2021-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/1.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL/CORRECTION INFORMATION: D10: LOT NUMBERS ADDED. H3: LENS RETURNED - DATE ADDED, NO EVALUATION REQUIRED. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.50/1.0/079 | OS | Physician | 08/18/2021 | 2021 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12433319&pc=QCB | 2023826-2021-03419 | 2021-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/1.0/068 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL/CORRECTION INFORMATION: D10: LOT NUMBERS ADDED. H3: LENS RETURNED - DATE ADDED, NO EVALUATION REQUIRED. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -11.00/1.0/068 | OD | Physician | 08/18/2021 | 2021 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12437008&pc=QCB | 2023826-2021-03473 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/132 (SPHERE/CYLINDER/AXIS), IS BEING RETURNED TO THE MANUFACTURER AND THE REASON IS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: D6A: (B)(6) 2021. CORRECTED DATA: B5: A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/132 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENTS RIGHT EYE (OD) AND REMAINED IMPLANTED. THERE WAS NO PROBLEM WITH THE LENS AND THE LENS WILL NOT BE RETURNED. THIS COMPLAINT WAS REPORTED IN ERROR. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. CLAIM # (B)(4). | -10.50/+1.0/132 | OD |  | 08/10/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12437066&pc=QCB | 2023826-2021-03478 | 2021-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+1.0/009 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISIBLE INSPECTION FOUND HAPTIC BROKEN. UNABLE TO PERFORM DIMENSIONAL/REFRACTIVE VERIFICATION BECAUSE LENS RETURNED SIGNIFICANTLY DAMAGED. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.00/+1.0/009 | OS | Physician | 08/09/2021 | 2021 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12437071&pc=QCB | 2023826-2021-03474 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, 09.50/+1.0/066 (SPHERE/CYLINDER/AXIS), FOR THE LEFT EYE (OS), IS BEING RETURNED TO THE MANUFACTURER AND THE REASON IS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: D6A: (B)(6) 2021. CORRECTED DATA: B5: A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/066 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE LEFT EYE (OS) AND REMAINED IMPLANTED. THERE WAS NO PROBLEM WITH THE LENS AND THE LENS WILL NOT BE RETURNED. THIS COMPLAINT WAS REPORTED IN ERROR. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. CLAIM # (B)(4). | -09.50/+1.0/066 | OS |  | 08/10/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12443200&pc=QCB | 2023826-2021-03528 | 2021-05-28 | Injury | STAAR SURGICAL COMPANY | 2021-09-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -10.0/1.5/162 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/1.5/162 | OD | Physician | 08/19/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12443201&pc=QCB | 2023826-2021-03532 | 2021-07-23 | Injury | STAAR SURGICAL COMPANY | 2021-09-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.0/1.5/94 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -12.0/1.5/94 | OD | Physician | 08/19/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12443203&pc=QCB | 2023826-2021-03527 | 2021-07-12 | Injury | STAAR SURGICAL COMPANY | 2021-09-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -10.0/3.0/085 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -10.0/3.0/085 | OD | Physician | 08/18/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12443371&pc=QCB | 2023826-2021-03539 | 2021-07-08 | Injury | STAAR SURGICAL COMPANY | 2021-09-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -7.5/1.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS DUE EXCESSIVE VAULT. IT WAS REPORTED THAT AFTER EXCHANGE PATIENT HAD A 100% CT VAULT. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS): "DOING WELL, HAPPY WITH VISION, VAULT CORRECT" WAS REPORTED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. CLAIM # (B)(4). | -7.5/1.0/088 | OD | Physician | 08/13/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12443372&pc=QCB | 2023826-2021-03540 | 2021-07-08 | Injury | STAAR SURGICAL COMPANY | 2021-09-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -10.00/1.5/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. IT WAS REPORTED THAT AFTER EXCHANGE PATIENT HAD A 100% CT VAULT. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS): "DOING WELL, HAPPY WITH VISION, VAULT CORRECT" WAS REPORTED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. CLAIM # (B)(4). | -10.00/1.5/075 | OS | Physician | 08/13/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12443626&pc=QCB | 2023826-2021-03483 | 2021-07-16 | Injury | STAAR SURGICAL COMPANY | 2021-09-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.00/+1.5/082 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION NOT REQUIRED. CLAIM # (B)(4). | -07.00/+1.5/082 | OS | Physician | 08/10/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12443507&pc=QCB | 2023826-2021-03543 | 2021-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6. -11.0/2.0/142 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTICS BENT/STRETCHED OUT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.0/2.0/142 | OD | Physician | 08/14/2021 | 2021 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12443635&pc=QCB | 2023826-2021-03485 | 2021-08-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.50/+5.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2021. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B4: CORRECTED TO (B)(6) 2021.; G3: CORRECTED TO 12-AUG-2021 IN INITIAL MDR. CLAIM (B)(4). Manufacturer Narrative: PT INFO: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -08.50/+5.0/093 | OD | Physician | 08/12/2021 | 2021 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12451419&pc=QCB | 2023826-2021-03545 | 2021-08-02 | Injury | STAAR SURGICAL COMPANY | 2021-09-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -8.5/1.0/070 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM#: (B)(4). | -8.5/1.0/070 | OD | Physician | 08/18/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12451420&pc=QCB | 2023826-2021-03546 | 2021-05-24 | Injury | STAAR SURGICAL COMPANY | 2021-09-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -8.5/1.0/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -8.5/1.0/086 | OS | Physician | 08/18/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12451421&pc=QCB | 2023826-2021-03544 | 2021-05-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -8.5/1.0/073 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS DID NOT RESOLVE THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. SEE MFR# 2023826-2021-03545 FOR REPLACEMENT LENS CLAIM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#:(B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -8.5/1.0/073 | OD | Physician | 08/18/2021 | 2021 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12454427&pc=QCB | 2023826-2021-03489 | 2021-08-10 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00/+2.5/113 (SPHERE/CYLINDER/AXIS), BUT THE LENS BROKE IN HALF DURING INJECTION INTO THE PATIENTS RIGHT EYE (OD). THERE WAS PATIENT CONTACT BUT NO INJURY. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -18.00/+2.5/113 | OD | Physician | 08/12/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12456990&pc=QCB | 2023826-2021-03560 | 2021-07-16 | Injury | STAAR SURGICAL COMPANY | 2021-09-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO_12.1; -8.0/1.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -8.0/1.0/093 | OD | Physician | 08/20/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12456993&pc=QCB | 2023826-2021-03561 | 2021-07-16 | Injury | STAAR SURGICAL COMPANY | 2021-09-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO_12.1; -5.5/1.0/68 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -5.5/1.0/68 | OS | Physician | 08/20/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12457496&pc=QCB | 2023826-2021-03550 | 2021-03-11 | Injury | STAAR SURGICAL COMPANY | 2021-09-10 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -7.0/+4.0/122 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED AND THE PATIENT IS REPORTEDLY HAPPY WITH THE OUTCOME. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. INVESTIGATION TYPE (B)(4): LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -7.0/+4.0/122 | OD | Physician | 07/23/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12454429&pc=QCB | 2023826-2021-03490 | 2021-08-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+3.0/026 (SPHERE/CYLINDER/AXIS), BUT THE LENS BECAME STUCK IN THE CARTRIDGE AND WAS DAMAGED DURING DELIVERY. THERE WAS NO PATIENT CONTACT. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. CAUSE OF THE EVENT WAS DUE TO USER ERROR AND THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON FELT THE CAUSE OF THE EVENT WAS DUE TO THE CARTRIDGE. H6 - CARTRIDGE LOT NUMBER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H6: CARTRIDGE LOT NUMBER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK, RACE: UNK, ETHNICITY: UNK, THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.00/+3.0/026 | NA | Physician | 08/12/2021 | 2021 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12454439&pc=QCB | 2023826-2021-03488 | 2021-08-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/079 (SPHERE/CYLINDER/AXIS), WAS BEING LOADED AND THE LENS WOULD NOT TRAVEL THROUGH THE CARTRIDGE AND BECAME STUCK. ONE OF THE EDGES OF THE LENS WAS BROKEN. THERE WAS NO PATIENT CONTACT. ANOTHER ICL LENS WAS NOT IMPLANTED. CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: CAUSE OF EVENT WAS REPORTED AS OTHER AND DEVICE FAILED TO PERFORM AS INTENDED. THE LENS DID NOT TRAVEL THROUGH THE CARTRIDGE AND IT STUCK TO IT. IT DID NOT MOVE WITH ANY MANEUVERING. ONE OF THE EDGES WAS BROKEN. H6 - WORK ORDER SEARCH: NO ADDITIONAL SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, RACE: UNK, ETHNICITY: UNK, THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -10.00/+1.0/079 | NA | Physician | 08/12/2021 | 2021 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12457495&pc=QCB | 2023826-2021-03556 | 2021-06-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS -13.5/+5.5/148 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD),IT WAS TORN. LENS STUCK IN CARTRIDGE OR INJECTION SYSTEM IS ALSO REPORTED. THERE WAS NO PATIENT CONTACT WITH THE LENS. THE CAUSE OF THE EVENT IS REPORTED AS BOTH USER ERROR AND UNKNOWN, THE REPORTER SAYING THE LENS WAS TORN IN THE DELIVERY SYSTEM. AN ALTERNATE LENS WAS LATER IMPLANTED AND THE PROBLEM IS RESOLVED. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE REPORTER STATES THE LENS WAS IMPLANTED (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REMOVED. AT A LATER DATE AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: B5-CORRECTED INFORMATION: PER REPORTER THE LENS WAS NOT IMPLANTED. IT TORE PRIOR TO IMPLANT. D6A, 6B-PREVIOUSLY REPORTED IN ERROR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.5/+5.5/148 | OD | Physician | 08/19/2021 | 2021 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12465244&pc=QCB | 2023826-2021-03580 | 2021-07-27 | Injury | STAAR SURGICAL COMPANY | 2021-09-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -12.5/1.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. IT WAS ALSO NOTED THE REPLACEMENT LENS WAS IMPLANTED VERTICALLY. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. MEDICAL DEVICE PROBLEM CODE: (B)(4): OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -12.5/1.0/089 | OD | Physician | 08/23/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12467192&pc=QCB | 2023826-2021-03500 | 2021-04-24 | Injury | STAAR SURGICAL COMPANY | 2021-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -12.50/1.0/172 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON 24/APR/2021 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.50/1.0/172 | OS | Physician | 08/23/2021 | 2021 | 63
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12467194&pc=QCB | 2023826-2021-03501 | 2021-05-12 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -8.50/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -8.50/1.0/089 | OD | Physician | 08/23/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12467195&pc=QCB | 2023826-2021-03499 | 2021-04-24 | Injury | STAAR SURGICAL COMPANY | 2021-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -14.00/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -14.00/1.5/090 | OD | Physician | 08/23/2021 | 2021 | 63
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12467196&pc=QCB | 2023826-2021-03495 | 2021-07-20 | Injury | STAAR SURGICAL COMPANY | 2021-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -8.50/4.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -8.50/4.0/089 | OD | Physician | 08/23/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12467197&pc=QCB | 2023826-2021-03508 | 2021-08-13 | Injury | STAAR SURGICAL COMPANY | 2021-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.50/0.5/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.50/0.5/102 | OD | Physician | 08/24/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12467198&pc=QCB | 2023826-2021-03507 | 2020-07-31 | Injury | STAAR SURGICAL COMPANY | 2021-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -13.00/2.00/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.00/2.00/089 | OD | Physician | 08/24/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12467200&pc=QCB | 2023826-2021-03506 | 2021-07-15 | Injury | STAAR SURGICAL COMPANY | 2021-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.00/3.00/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -9.00/3.00/090 | OD | Physician | 08/24/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12467204&pc=QCB | 2023826-2021-03512 | 2021-07-07 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.5/1.5/093 | OS | Physician | 08/26/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12467205&pc=QCB | 2023826-2021-03504 | 2021-08-16 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -4.50/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -4.50/1.5/092 | OD | Physician | 08/24/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12467207&pc=QCB | 2023826-2021-03502 | 2021-05-12 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -9.00/1.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.00/1.5/082 | OS | Physician | 08/23/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12467208&pc=QCB | 2023826-2021-03519 | 2021-07-10 | Injury | STAAR SURGICAL COMPANY | 2021-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.50/1.5/092 (SPHERE/ CYLINDER/ AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT ON (B)(6) 2021. A SHORTER LENGTH LENS WAS IMPLANTED ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.5/092 | OS | Physician | 08/27/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12467213&pc=QCB | 2023826-2021-03509 | 2021-08-03 | Injury | STAAR SURGICAL COMPANY | 2021-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -6.00/1.0/167 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PREMARKET IDENTIFICATION: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -6.00/1.0/167 | OD | Physician | 08/24/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12467214&pc=QCB | 2023826-2021-03518 | 2020-12-23 | Injury | STAAR SURGICAL COMPANY | 2021-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.00/1.0/090 | OD | Physician | 08/27/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12467217&pc=QCB | 2023826-2021-03514 | 2021-07-23 | Injury | STAAR SURGICAL COMPANY | 2021-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.50/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -8.50/1.0/090 | OS | Physician | 08/26/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12467218&pc=QCB | 2023826-2021-03524 | 2021-08-16 | Injury | STAAR SURGICAL COMPANY | 2021-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO12.1, -14.00/2.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -14.00/2.5/089 | OS | Physician | 08/27/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12467219&pc=QCB | 2023826-2021-03523 | 2021-08-27 | Injury | STAAR SURGICAL COMPANY | 2021-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.00/2.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.00/2.5/095 | OD | Physician | 08/27/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12467220&pc=QCB | 2023826-2021-03525 | 2021-07-28 | Injury | STAAR SURGICAL COMPANY | 2021-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.00/1.0/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT ON (B)(6) 2021. A SHORTER LENGTH LENS WAS IMPLANTED ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.00/1.0/080 | OS | Physician | 08/27/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12467222&pc=QCB | 2023826-2021-03526 | 2021-04-09 | Injury | STAAR SURGICAL COMPANY | 2021-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.00/1.5/125 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -14.00/1.5/125 | OS | Physician | 08/27/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12467227&pc=QCB | 2023826-2021-03515 | 2021-07-07 | Injury | STAAR SURGICAL COMPANY | 2021-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.50/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.5/091 | OS | Physician | 08/26/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12467228&pc=QCB | 2023826-2021-03520 | 2021-07-19 | Injury | STAAR SURGICAL COMPANY | 2021-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/1.0/113 (SPHERE/ CYLINDER/ AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/1.0/113 | OS | Physician | 08/27/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12471485&pc=QCB | 2023826-2021-03547 | 2021-07-09 | Injury | STAAR SURGICAL COMPANY | 2021-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -10.0/+3.0/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -10.0/+3.0/094 | OD | Physician | 08/20/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12471486&pc=QCB | 2023826-2021-03641 | 2021-07-29 | Injury | STAAR SURGICAL COMPANY | 2021-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -12.0/1.5/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -12.0/1.5/90 | OS | Physician | 08/23/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12471487&pc=QCB | 2023826-2021-03637 | 2021-07-23 | Injury | STAAR SURGICAL COMPANY | 2021-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO_12.6; -10.0/2.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. MEDICAL DEVICE PROBLEM CODE: (B)(4): OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -10.0/2.0/091 | OD | Physician | 08/23/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12471488&pc=QCB | 2023826-2021-03640 | 2021-07-20 | Injury | STAAR SURGICAL COMPANY | 2021-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.0/2.0/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. MEDICAL DEVICE PROBLEM CODE: (B)(4): OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -9.0/2.0/094 | OD | Physician | 08/23/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12471494&pc=QCB | 2023826-2021-03638 | 2021-07-14 | Injury | STAAR SURGICAL COMPANY | 2021-09-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -8.5/1.5/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. MEDICAL DEVICE PROBLEM CODE: (B)(4): OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | -8.5/1.5/094 | OS | Physician | 08/23/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12477834&pc=QCB | 2023826-2021-03585 | 2021-08-17 | Injury | STAAR SURGICAL COMPANY | 2021-09-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.5/+1.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -8.5/+1.0/095 | OS | Physician | 08/17/2021 | 2021 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12478191&pc=QCB | 2023826-2021-03643 | 2021-07-29 | Injury | STAAR SURGICAL COMPANY | 2021-09-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -6.0/1.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -6.0/1.5/089 | OS | Physician | 08/23/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12478194&pc=QCB | 2023826-2021-03548 | 2020-10-15 | Injury | STAAR SURGICAL COMPANY | 2021-09-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -13.0/+2.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. IT WAS ALSO NOTED THE PATIENT IS EXTREMELY HAPPY AND VAULT REDUCED TO NORMAL. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -13.0/+2.0/092 | OD | Physician | 08/20/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12478196&pc=QCB | 2023826-2021-03656 | 2021-06-07 | Injury | STAAR SURGICAL COMPANY | 2021-09-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO_13.2; -13.0/2.5/85 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/2.5/85 | OD | Physician | 08/24/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12478201&pc=QCB | 2023826-2021-03642 | 2021-07-12 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -12.5/1.5/085 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. DURING THE SAME SURGERY INTRAOPERATIVELY THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -12.5/1.5/085 | OS | Physician | 08/23/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12478203&pc=QCB | 2023826-2021-03644 | 2021-07-08 | Injury | STAAR SURGICAL COMPANY | 2021-09-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO_12.1; -10.5/2.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -10.5/2.5/090 | OD | Physician | 08/23/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12478490&pc=QCB | 2023826-2021-03572 | 2021-05-21 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/089 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO THE LENS WAS STUCK IN THE CARTRIDGE OR INJECTION SYSTEM AND WAS DAMAGED. THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED ON (B)(6) 2021 WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION NOT REQUIRED. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.50/+2.0/089 | NA | Physician | 08/24/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12475986&pc=QCB | 2023826-2021-03581 | 2021-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.5/0.5/81 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE, LOW VAULT WITH ROTATION AND POOR VISION. THIS EXCHANGE RESOLVED THE PROBLEM. THE REPORTER INDICATED THAT AFTER LENS WAS INJECTED INTO THE EYE AND THE DOCTOR FOUND THE LENS WAS BROKEN. SINCE THE TEAR WAS NOT IN THE OPTICAL ZONE THE DOCTOR DECIDED NOT TO REMOVE IT. BUT, LATER FOUND THAT THE LENS HAD ROTATED, AND THE VAULT HAD DECREASED WITH POOR VISION. THE REPORTER STATED THAT THE DOCTOR CONSIDERED THAT THE ROTATION WAS DUE TO A BROKEN FOOTPLATE, SO IT WAS NOT REPLACED WITH A LONGER LENGTH LENS. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC TORN TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). WEIGHT, ETHNICITY, RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.5/0.5/81 | OS | Physician | 08/16/2021 | 2021 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12478497&pc=QCB | 2023826-2021-03567 | 2021-06-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -08.00/+1.0/180 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. REFRACTIVE VERIFICATION WAS PERFORMED AND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM # (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -08.00/+1.0/180 | OS | Physician | 08/18/2021 | 2021 | 42 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12478509&pc=QCB | 2023826-2021-03569 | 2023-02-20 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON WAS INSERTING A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/089 (SPHERE/CYLINDER/AXIS), AND THE LENS TORE DURING DELIVERY. THERE WAS NO PATIENT CONTACT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE LENS WAS DISCARDED. Manufacturer Narrative: H6: WORK ORDER SEARCH - NO SIMILAR COMPLAINTS, CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: "THE REPORTER INDICATED THE SURGEON ATTEMPTED TO IMPLANT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+5.0/089 (SPHERE/CYLINDER/AXIS) WHEN THE LENS TORE/BROKE DURING THE INJECTION/DELIVERY INTO THE EYE. THE LENS WAS NOT IMPLANTED. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. ON (B)(6) 2023 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN." CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -10.00/+1.0/089 | NA | Physician | 08/18/2021 | 2021 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12478522&pc=QCB | 2023826-2021-03570 | 2021-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/073 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO THE LENS TORE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -10.50/+1.0/073 | OS | Physician | 08/19/2021 | 2021 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12478526&pc=QCB | 2023826-2021-03571 | 2021-08-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED AN ICL IMPLANTABLE COLLAMER LENS, DUE TO THE LENS WAS DAMAGED. THE BACKUP LENS WAS IMPLANTED SUCCESSFULLY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: B5: THE REPORTER INDICATED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+1.5/087 (SPHERE/CYLINDER/AXIS) WAS TORN DURING LOADING AND THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING WELL AND IS PLEASED. THE CAUSE OF THE EVENT WAS THE DEVICE AND USER ERROR. THE LENS SEEMED TO JAM IN THE CARTRIDGE DURING LOADING. D2: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB. H6: HEALTH EFFECT - IMPACT CODE: 2645. CLAIM # (B)(4). Manufacturer Narrative: AGE: UNK. WEIGHT: UNK. SEX: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. CLAIM # (B)(4). | -18.00/+1.5/087 | NA |  | 08/23/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12479189&pc=QCB | 2023826-2021-03597 | 2021-07-27 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.0/5.0/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS ENDED UP IN FOLDED CONDITION UPON INJECTION AND DID NOT OPEN WITH MANIPULATION. LENS WAS IMPLANTED AND REMOVED WITHIN SAME SURGERY. THE LENS WAS DAMAGED WHILE REMOVING FROM EYE. NO PATIENT INJURY WAS REPORTED. ON (B)(6) 2021 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -8.0/5.0/112 | NA | Physician | 08/20/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12479210&pc=QCB | 2023826-2021-03596 | 2021-07-23 | Injury | STAAR SURGICAL COMPANY | 2021-09-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.00/3.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT. CAUSE OF EVENT WAS UNKNOWN. THE REPORTER INDICATED THAT THE PROBLEM WAS NOT RESOLVED. FOR ADDITIONAL INFORMATION THE REPORTER STATED " VAULT WAS TOO LOW AND THE PATIENT WAS UNHAPPY WITH THE OUTCOME. HAD TO EXPLANT THE LENS. EXCHANGE WILL BE DONE, ONCE PATIENT DECIDED ABOUT IT." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -10.00/3.0/089 | OS | Physician | 08/20/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12479260&pc=QCB | 2023826-2021-03589 | 2021-06-28 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -8.5/1.0/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE ON (B)(6) 2021. THE LENS WAS IMPLANTED, REMOVED AND REPLACED WITHIN THE SAME SURGERY. THERE WAS NO PATIENT CONTACT. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -8.5/1.0/110 | NA | Physician | 08/19/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12485424&pc=QCB | 2023826-2021-03660 | 2020-09-23 | Injury | STAAR SURGICAL COMPANY | 2021-09-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -11.00/0.50/60 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND LATER ON THE SAME DATE REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS NOTED AS "OTHER". Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -11.00/0.50/60 | OS | Physician | 08/25/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12485427&pc=QCB | 2023826-2021-03664 | 2021-03-21 | Injury | STAAR SURGICAL COMPANY | 2021-09-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -14.0/2.0/82 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -14.0/2.0/82 | OD | Physician | 08/26/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12485428&pc=QCB | 2023826-2021-03665 | 2021-03-21 | Injury | STAAR SURGICAL COMPANY | 2021-09-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -13.0/1.5/124 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -13.0/1.5/124 | OS | Physician | 08/26/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12485431&pc=QCB | 2023826-2021-03659 | 2021-08-06 | Injury | STAAR SURGICAL COMPANY | 2021-09-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THE THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.0/1.5/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. MEDICAL DEVICE PROBLEM CODE: (B)(4): OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | -10.0/1.5/088 | OD | Physician | 08/25/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12485432&pc=QCB | 2023826-2021-03662 | 2021-06-19 | Injury | STAAR SURGICAL COMPANY | 2021-09-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -15.5/1.5/075 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -15.5/1.5/075 | OD | Physician | 08/26/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12485433&pc=QCB | 2023826-2021-03668 | 2021-07-16 | Injury | STAAR SURGICAL COMPANY | 2021-09-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -8.0/1.0/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/1.0/92 | OD | Physician | 08/26/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12485434&pc=QCB | 2023826-2021-03666 | 2021-04-09 | Injury | STAAR SURGICAL COMPANY | 2021-09-16 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.2; -12.0/1.5/71 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -12.0/1.5/71 | OD | Physician | 08/26/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12479030&pc=QCB | 2023826-2021-03583 | 2021-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.0/2.5/91, (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND LENS DISLOCATION/SUBLUXATION. THIS EXCHANGE RESOLVED THE PROBLEM. FOR CAUSE DETAILS THE REPORTER STATED " WTW MEASUREMENT DOES NOT CORRELATE WITH SULCUS TO SULCUS MEASUREMENT." Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.0/2.5/91 | OD | Physician | 08/18/2021 | 2021 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12479055&pc=QCB | 2023826-2021-03588 | 2021-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -15.0/4.0/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 AND (B)(6) 2021 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS REMOVED, REPLACED AND FIXED VERTICALLY. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -15.0/4.0/91 | OD | Physician | 08/18/2021 | 2021 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12492851&pc=QCB | 2023826-2021-03646 | 2021-08-17 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-17 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. INTRAOPERATIVELY THE LENS WAS REMOVED DUE TO THE LENS BEING IMPLANTED UPSIDE DOWN. AN ALTERNATE LENS WAS IMPLANTED AT A LATER DATE AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE AND OTHER: "I THINK THE REASON WHY THE ICL WAS INSERTED IN REVERSE IS DUE TO THE INJECTOR." NO PATIENT INJURY IS REPORTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -11.5/+2.0/093 | OS | Physician | 08/24/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12489325&pc=QCB | 2023826-2021-03573 | 2021-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+2.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H3: DEVICE EVALUATON: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISIUAL INSPECTION FOUND A HAPTIC TORN AND BROKEN LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.00/+2.5/092 | OS | Physician | 08/24/2021 | 2021 | 47 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12492855&pc=QCB | 2023826-2021-03647 | 2021-08-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT, DURING LOADING A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.0/+1.0/089 (SPHERE/CYLINDER/AXIS), IT TORE. THIS WAS DISCOVERED AFTER OPENING THE VIAL AND THERE WAS NO PATIENT CONTACT WITH THE LENS. THIS OCCURRED ON (B)(6) 2021. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AT A LATER DATE. THE SURGEON STATES, "THE ICL CRACKED WHEN I GRABBED IT WITH PAC-MAN TWEEZERS AND PULLED IT INTO THE CARTRIDGE. I THINK THERE IS A PROBLEM WITH THE STRENGTH OF THE ICL." Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -12.0/+1.0/089 | NA | Physician | 08/24/2021 | 2021 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12493515&pc=QCB | 2023826-2021-03575 | 2021-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.00/+1.5/088 | OS | Physician | 08/25/2021 | 2021 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12493516&pc=QCB | 2023826-2021-03576 | 2021-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS RIGHT EYE (OD). THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -2.50/+2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4) Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. . B5 - "ON 08/11/2022 THE LENS WAS EXCHANGED WITH A SAME MODEL/ SHORTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE 08-NOV-2022 SHOULD BE ADDED TO INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR . H6 - HEALTH EFFECT - IMPACT CODE: 4629 SHOULD BE ADDED TO THE INITAL MDR. CLAIM#(B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND AN OPTIC DEFORMED AND HAPTIC BROKEN/DEFORMED TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H4 - DEVICE MANUFACTURE DATE: 28-OCT-2020 SHOULD BE ADDED TO THE INITAL MDR. CLAIM # (B)(4). Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. CLAIM # (B)(4). | -2.50/+2.0/091 | OD | Physician | 08/25/2021 | 2021 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12502285&pc=QCB | 2023826-2021-03685 | 2021-08-12 | Injury | STAAR SURGICAL COMPANY | 2021-09-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -8.0/1.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM#: (B)(4). | -8.0/1.5/091 | OS | Physician | 08/30/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12502291&pc=QCB | 2023826-2021-03687 | 2021-05-24 | Injury | STAAR SURGICAL COMPANY | 2021-09-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -14.0/2.0/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) (UNDER 21YRS OF AGE AT DATE OF IMPLANT) CLAIM#: (B)(4). | -14.0/2.0/91 | OD | Physician | 08/29/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12502311&pc=QCB | 2023826-2021-03681 | 2021-07-01 | Injury | STAAR SURGICAL COMPANY | 2021-09-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -16.5/2.0/096 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -16.5/2.0/096 | OS | Physician | 08/30/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12502716&pc=QCB | 2023826-2021-03577 | 2021-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS LEFT EYE (OS). THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -3.00/+2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REFRACTIVE SURPRISE WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. . B5 - "ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME MODEL/ SHORTER LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. . H6 - HEALTH EFFECT - IMPACT CODE: 4629 SHOULD BE ADDED TO THE INITAL MDR. CLAIM# (B)(4). Manufacturer Narrative: D4 - EXPIRATION DATE: 30-SEP-2023 SHOULD BE ADDED TO THE INITIAL MDR. H4 - DEVICE MANUFACTURE DATE: 28-OCT-2020 SHOULD BE ADDED TO THE INTIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -3.00/+2.0/091 | OS | Physician | 08/25/2021 | 2021 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12510738&pc=QCB | 2023826-2021-03731 | 2021-07-02 | Injury | STAAR SURGICAL COMPANY | 2021-09-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -13.0/2.0/93 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -13.0/2.0/93 | OD | Physician | 08/29/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12512079&pc=QCB | 2023826-2021-03599 | 2021-08-28 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/1.0/129 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.00/1.0/129 | OD | Physician | 08/30/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12512080&pc=QCB | 2023826-2021-03602 | 2021-08-21 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -7.50/1.5/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.50/1.5/108 | OD | Physician | 08/30/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12512081&pc=QCB | 2023826-2021-03600 | 2021-08-27 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.50/1.0/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -8.50/1.0/109 | OD | Physician | 08/30/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12512083&pc=QCB | 2023826-2021-03603 | 2021-08-28 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/1.5/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.50/1.5/111 | OD | Physician | 08/30/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12512084&pc=QCB | 2023826-2021-03606 | 2021-06-29 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -8.00/2.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.00/2.0/081 | OD | Physician | 08/30/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12512085&pc=QCB | 2023826-2021-03605 | 2021-07-27 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.00/1.5/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -13.00/1.5/108 | OS | Physician | 08/30/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12512087&pc=QCB | 2023826-2021-03604 | 2021-08-23 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -8.50/2.0/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A LONGER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S RIGHT EYE (OD) DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT,-ETHNICITY, RACE- UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -8.50/2.0/073 | OD | Physician | 08/30/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12512090&pc=QCB | 2023826-2021-03613 | 2021-03-19 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.50/2.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -14.50/2.0/085 | OS | Physician | 08/29/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12512092&pc=QCB | 2023826-2021-03612 | 2021-07-28 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 13.2MM VTICMO 13.2, -14.00/1.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS FROM THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021 DUE TO LOW VAULT . A LONGER LENGTH LENS WAS IMPLANTED ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -14.00/1.5/087 | OS | Physician | 08/29/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12512093&pc=QCB | 2023826-2021-03608 | 2021-03-03 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.00/1.5/113 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/1.5/113 | OD | Physician | 08/29/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12512096&pc=QCB | 2023826-2021-03620 | 2020-09-04 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -8.00/2.0/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REPLACED WITH A LONGER LENGTH SPHERICAL LENS DUE TO LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -8.00/2.0/099 | OS | Physician | 08/30/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12512097&pc=QCB | 2023826-2021-03621 | 2021-07-26 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -10.0/1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.0/1.0/092 | OS | Physician | 09/02/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12512102&pc=QCB | 2023826-2021-03618 | 2021-07-27 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.00/3.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT IN A SECOND SURGERY ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/3.5/089 | OD | Physician | 08/31/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12512103&pc=QCB | 2023826-2021-03627 | 2021-07-10 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -14.00/5.00/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -14.00/5.00/091 | OS | Physician | 09/02/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12512105&pc=QCB | 2023826-2021-03622 | 2021-07-15 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -10.5/1.5/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/1.5/070 | OS | Physician | 09/02/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12512107&pc=QCB | 2023826-2021-03625 | 2021-08-13 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -9.00/1.50/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.00/1.50/105 | OD | Physician | 09/02/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12512109&pc=QCB | 2023826-2021-03626 | 2021-07-10 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -14.00/5.00/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -14.00/5.00/089 | OD | Physician | 09/02/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12512112&pc=QCB | 2023826-2021-03630 | 2021-06-29 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.50/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.50/1.5/095 | OD | Physician | 09/03/2021 | 2021 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12512113&pc=QCB | 2023826-2021-03635 | 2021-06-04 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.50/3.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). OFF-LABEL USE, UNDER 21 YRS OF AGE AT DATE OF IMPLANT. CLAIM# (B)(4). | -17.50/3.5/091 | OS | Physician | 09/03/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12512115&pc=QCB | 2023826-2021-03634 | 2021-05-23 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.80/1.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.80/1.5/082 | OS | Physician | 09/03/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12512118&pc=QCB | 2023826-2021-03629 | 2021-07-17 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.50/2.50/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -12.50/2.50/093 | OS | Physician | 09/03/2021 | 2021 | 41
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12512125&pc=QCB | 2023826-2021-03636 | 2021-07-02 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.50/0.5/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.50/0.5/104 | OS | Physician | 09/03/2021 | 2021 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12516287&pc=QCB | 2023826-2021-03706 | 2021-07-01 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO -12.50/+3.0/X100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE PROBLEM WAS NOT RESOLVED. " PATIENT RELATED FACTOR" WAS REPORTED TO BE THE CAUSE OF THIS EVENT. IT WAS ALSO REPORTED THAT THE DEVICE FAILED TO PERFORM AS INTENDED, THE REPORTER STATED, " SULCUS TO SULCUS MEASURED BY UBM AND FOUND TO BE LARGER THAN EXPECTED." LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS FOUND. CLAIM # (B)(4). | -12.50/+3.0/x100 | OD | Physician | 08/23/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12516310&pc=QCB | 2023826-2021-03707 | 2021-07-01 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO -13.50/+3.0/X079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE PROBLEM WAS NOT RESOLVED. " PATIENT RELATED FACTOR" WAS REPORTED TO BE THE CAUSE OF THIS EVENT. IT WAS ALSO REPORTED THAT THE DEVICE FAILED TO PERFORM AS INTENDED, THE REPORTER STATED, " SULCUS TO SULCUS MEASURED BY UBM AND FOUND TO BE LARGER THAN EXPECTED." LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS FOUND. CLAIM # (B)(4). | -13.50/+3.0/x079 | OS | Physician | 08/23/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12516505&pc=QCB | 2023826-2021-03752 | 2021-07-17 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -9.5/1.0/110 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM#: (B)(4). | -9.5/1.0/110 | OD | Physician | 08/29/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12516554&pc=QCB | 2023826-2021-03749 | 2021-02-14 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -9.0/1.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -9.0/1.5/091 | OD | Physician | 08/30/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12516562&pc=QCB | 2023826-2021-03748 | 2021-07-31 | Injury | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -9.0/1.0/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. MEDICAL DEVICE PROBLEM CODE: (B)(4): OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | -9.0/1.0/89 | OS | Physician | 08/29/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12517004&pc=QCB | 2023826-2021-03727 | 2021-08-07 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.0/1.0/175 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE ON (B)(6) 2021. THE LENS WAS IMPLANTED,REMOVED, AND REPLACED WITHIN THE SAME SURGERY AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT THERE WAS NO PATIENT INJURY. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS FOUND. CLAIM # (B)(4). | -8.0/1.0/175 | NA | Physician | 08/24/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12517027&pc=QCB | 2023826-2021-03726 | 2021-08-05 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -14.0/1.0/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE ON (B)(6) 2021. THE LENS WAS IMPLANTED AND REMOVED WITHIN THE SAME SURGERY. ON (B)(6) 2021 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT THERE WAS NO PATIENT INJURY. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US.A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS FOUND. CLAIM # (B)(4). | -14.0/1.0/105 | NA | Physician | 08/24/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12523330&pc=QCB | 2023826-2021-03756 | 2021-05-17 | Injury | STAAR SURGICAL COMPANY | 2021-09-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -4.5/1.0/172 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM#: (B)(4). | -4.5/1.0/172 | OD | Physician | 08/31/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12523374&pc=QCB | 2023826-2021-03762 | 2021-07-23 | Injury | STAAR SURGICAL COMPANY | 2021-09-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.0/1.0/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/1.0/91 | OS | Physician | 09/02/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12523393&pc=QCB | 2023826-2021-03760 | 2021-08-20 | Injury | STAAR SURGICAL COMPANY | 2021-09-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.5/3.0/179 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -9.5/3.0/179 | OS | Physician | 09/02/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12523410&pc=QCB | 2023826-2021-03757 | 2021-08-24 | Injury | STAAR SURGICAL COMPANY | 2021-09-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -16.0/2.0/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -16.0/2.0/90 | OD | Physician | 08/31/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12523440&pc=QCB | 2023826-2021-03761 | 2021-07-21 | Injury | STAAR SURGICAL COMPANY | 2021-09-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.5/1.0/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -12.5/1.0/92 | OD | Physician | 09/02/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12523446&pc=QCB | 2023826-2021-03758 | 2021-08-24 | Injury | STAAR SURGICAL COMPANY | 2021-09-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -14.0/1.0/87 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -14.0/1.0/87 | OS | Physician | 08/31/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12523450&pc=QCB | 2023826-2021-03763 | 2021-07-23 | Injury | STAAR SURGICAL COMPANY | 2021-09-23 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -9.5/1.5/84 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4): OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | -9.5/1.5/84 | OD | Physician | 09/02/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12523535&pc=QCB | 2023826-2021-03711 | 2021-04-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -8.5/2.0/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS GOT STUCK IN THE CARTRIDGE OR INJECTION SYSTEM. THE LENS WAS IMPLANTED AND REMOVED WITHIN THE SAME SURGERY. ON (B)(6) 2021 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT THERE WAS NO PATIENT INJURY. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: B5- PER UPDATED INFORMATION A 12.6MM, VTICMO12.6, -8.5/2.0/080 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION DELIVERY INTO THE EYE ON (B)(6) 2021. THE LENS WAS IMPLANTED AND REMOVE WITHIN THE SAME SURGERY. THERE WAS NO PATIENT INJURY. A REPLACEMENT LENS OF IMPLANTED ON (B)(6) 2021 AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/501K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM #: (B)(4). | -8.5/2.0/080 | NA | Physician | 08/24/2021 | 2021 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12523539&pc=QCB | 2023826-2021-03691 | 2021-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -6.0/+1.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT AXIS LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISIBLE INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND LENS TO BE IN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.0/+1.5/093 | OS | Physician | 08/30/2021 | 2021 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12523542&pc=QCB | 2023826-2021-03694 |  | Injury | STAAR SURGICAL COMPANY | 2021-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.5/+2.5/079 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS REPOSITIONED, THE SURGEON REPORTING A VAULT OF 64UM. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: REPORTEDLY THE LENS ROTATED PRIOR TO REPOSITIONING. THE PROBLEM WAS RESOLVED WITH REPOSITIONING. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM# (B)(4). | -9.5/+2.5/079 | OS |  | 08/29/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12523840&pc=QCB | 2023826-2021-03724 | 2021-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -16.5/6.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE PROBLEM DID NOT RESOLVE. THE LENS WAS EXPLANTED AND A LONGER LENGTH LENS WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -16.5/6.0/083 | OS | Physician | 08/24/2021 | 2021 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12523886&pc=QCB | 2023826-2021-03723 | 2021-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -9.0/1.5/67 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT IMPLANTED HORIZONTALLY. PROBLEM RESOLVED. CAUSE WAS UNKNOWN. IT WAS REPORTED THAT THE INITIAL LENS WAS IMPLANTED VERTICALLY. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS FOUND. CLAIM # (B)(4). | -9.0/1.5/67 | OS | Physician | 08/24/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12523912&pc=QCB | 2023826-2021-03710 | 2021-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.00/2.5/178 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS AND PROBLEM RESOLVED. CAUSE WAS UNKNOWN. Manufacturer Narrative: B5- PER UPDATED INFORMATION THE REPLACEMENT LENS WAS IMPLANTED VERTICALLY. H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, INSIDE A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/2.5/178 | OD | Physician | 08/24/2021 | 2021 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12527941&pc=QCB | 2023826-2021-03764 | 2021-07-23 | Injury | STAAR SURGICAL COMPANY | 2021-09-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.0/1.0/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/1.0/90 | OS | Physician | 09/02/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12530671&pc=QCB | 2023826-2021-03716 |  | Injury | STAAR SURGICAL COMPANY | 2021-09-24 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: IT WAS REPORTED THAT A 12.6MM,TMICL12.6, -6.50/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS REMOVED FROM THE PATIENTS LEFT EYE (OS) AND REPLACED WITH A SHORTER LENGTH LENS DUE TO SIZING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT INFORMATION: UNK, EVENT DATE: UNK, IMPLANT DATE: UNK. A LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINTS FOUND. CLAIM # (B)(4). | -6.50/2.5/091 | OS | Physician | 08/26/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12524230&pc=QCB | 2023826-2021-03712 | 2020-08-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.5/1.5/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE ON (B)(6) 2020. THE LENS WAS IMPLANTED, REMOVED AND REPLACED WITHIN THE SAME SURGERY AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT THERE WAS NO PATIENT INJURY. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: A PIECE OF THE LENS WAS RETURNED IN LIQUID, IN MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN/SPLIT AND THE HAPTIC BROKEN. CLAIM # (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). WEIGHT, ETHNICITY, RACE: UNK. PMA/501K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS FOUND. CLAIM # (B)(4). | -15.5/1.5/104 | NA | Physician | 08/26/2021 | 2021 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12530669&pc=QCB | 2023826-2021-03715 |  | Injury | STAAR SURGICAL COMPANY | 2021-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: IT WAS REPORTED THAT A 12.6MM,TMICL12.6, -6.50/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS REMOVED FROM THE PATIENTS RIGHT EYE (OD) AND REPLACED WITH A SHORTER LENGTH LENS DUE TO SIZING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTION: D4 SERIAL NUMBER: CORRECTED FROM (B)(6) TO (B)(6). CLAIM # (B)(4). Manufacturer Narrative: AGE, SEX, WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. A LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINTS FOUND. CLAIM # (B)(4). | -6.50/2.5/092 | OD |  | 08/26/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12533422&pc=QCB | 2023826-2021-03743 | 2021-08-14 | Injury | STAAR SURGICAL COMPANY | 2021-09-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.50/3.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS DEVICE, ORBSCAN HAD ERROR. ORBSCAN TO BE RECALIBRATED. Manufacturer Narrative: DATE OF BIRTH- (B)(6), WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -17.50/3.5/097 | OD | Physician | 09/09/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12533424&pc=QCB | 2023826-2021-03738 | 2021-05-28 | Injury | STAAR SURGICAL COMPANY | 2021-09-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.00/1.0/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -8.00/1.0/072 | OD | Physician | 09/08/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12533425&pc=QCB | 2023826-2021-03740 | 2021-08-25 | Injury | STAAR SURGICAL COMPANY | 2021-09-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -09.50/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -09.50/1.0/089 | OD | Physician | 09/08/2021 | 2021 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12533427&pc=QCB | 2023826-2021-03737 | 2021-07-26 | Injury | STAAR SURGICAL COMPANY | 2021-09-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.00/2.0/085 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -14.00/2.0/085 | OS | Physician | 09/07/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12533431&pc=QCB | 2023826-2021-03744 | 2021-08-17 | Injury | STAAR SURGICAL COMPANY | 2021-09-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/1.0/123 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.00/1.0/123 | OD | Physician | 09/08/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12538174&pc=QCB | 2023826-2021-03786 | 2021-04-22 | Injury | STAAR SURGICAL COMPANY | 2021-09-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -4.5/1.0/82 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -4.5/1.0/82 | OS | Physician | 09/03/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12538175&pc=QCB | 2023826-2021-03783 | 2021-05-31 | Injury | STAAR SURGICAL COMPANY | 2021-09-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -8.0/1.0/084 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM#: (B)(4). | -8.0/1.0/084 | OS | Physician | 09/03/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12538176&pc=QCB | 2023826-2021-03780 | 2021-07-05 | Injury | STAAR SURGICAL COMPANY | 2021-09-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -12.0/1.0/93 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. IT WAS ALSO NOTED "DUE TO THE LACK OF THE TICL; THE SURGEON DECIDED TO IMPLANT ICL AND CORRECT ASTIGMATISM WITH CORNEA INCISION". THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM#: (B)(4). | -12.0/1.0/93 | OD | Physician | 09/01/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12538178&pc=QCB | 2023826-2021-03782 | 2021-05-27 | Injury | STAAR SURGICAL COMPANY | 2021-09-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -8.5/1.0/085 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM#: (B)(4). | -8.5/1.0/085 | OD | Physician | 09/03/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12538179&pc=QCB | 2023826-2021-03787 | 2021-07-05 | Injury | STAAR SURGICAL COMPANY | 2021-09-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -3.0/0.5/072 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM#: (B)(4). | -3.0/0.5/072 | OS | Physician | 09/03/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12538180&pc=QCB | 2023826-2021-03686 | 2021-07-10 | Injury | STAAR SURGICAL COMPANY | 2021-09-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -9.0/1.0/69 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM#: (B)(4). | -9.0/1.0/69 | OD | Physician | 08/29/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12538182&pc=QCB | 2023826-2021-03784 | 2021-04-09 | Injury | STAAR SURGICAL COMPANY | 2021-09-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.5/1.5/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -9.5/1.5/92 | OD | Physician | 09/03/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12537702&pc=QCB | 2023826-2021-03698 | 2021-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/106 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.00/+2.0/106 | OS | Physician | 08/30/2021 | 2021 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12544573&pc=QCB | 2023826-2021-03699 | 2021-06-30 | Injury | STAAR SURGICAL COMPANY | 2021-09-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+2.0/137 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO THE SURGEON HAD SELECTED THE WRONG LENS POWER. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -09.00/+2.0/137 | OD | Physician | 08/30/2021 | 2021 | 44
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12544643&pc=QCB | 2023826-2021-03814 | 2021-05-07 | Injury | STAAR SURGICAL COMPANY | 2021-09-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -9.5/1.0/2 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. MEDICAL DEVICE PROBLEM: (B)(4): OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -9.5/1.0/2 | OD | Physician | 09/05/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12544644&pc=QCB | 2023826-2021-03798 | 2021-07-27 | Injury | STAAR SURGICAL COMPANY | 2021-09-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.50/+1.0/106 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. MEDICAL DEVICE PROBLEM CODE: (B)(4): OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | -10.50/+1.0/106 | OS | Physician | 09/03/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12544647&pc=QCB | 2023826-2021-03801 | 2021-05-11 | Injury | STAAR SURGICAL COMPANY | 2021-09-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.5/1.5/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -11.5/1.5/093 | OS | Physician | 09/05/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12543503&pc=QCB | 2023826-2021-03789 | 2021-07-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -13.5/2.0/91 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN ON (B)(6) 2021. THE LENS WAS IMPLANTED, REMOVED AND REPLACED WITHIN THE SAME SURGERY AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED TO BE USER ERROR. Manufacturer Narrative: B5 - PER EMAIL ON 08/OCT/2021, THE REPORTER STATED THAT "THE PATIENT DOESN'T HAVE INJURY" AND THE CAUSE OF THE EVENT WAS DUE TO THE "DOCTOR'S INSTALLATION ERROR, IT HAD NOTHING TO DO WITH THE DEVICE." CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS FOUND. CLAIM#: (B)(4). | -13.5/2.0/91 | NA |  | 08/29/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12544299&pc=QCB | 2023826-2021-03790 | 2021-05-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.5/1.0/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.5/1.0/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -10.5/1.0/108 | OD | Physician | 08/29/2021 | 2021 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12544565&pc=QCB | 2023826-2021-03700 | 2021-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -05.50/+4.0/078 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE REPLACEMENT LENS WAS IMPLANTED IN A VERTICAL POSITION. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE CODE: (B)(4): THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4): | -05.50/+4.0/078 | OD | Physician | 08/30/2021 | 2021 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12544586&pc=QCB | 2023826-2021-03701 | 2021-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -02.00/+6.0/104 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2021. THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS EXPLANTED. THE PROBLEM WAS RESOLVED. THE PATIENT IS USING HER SPECTACLES AND HER EYES ARE IN PERFECT CONDITION. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -2.0/6.0/104 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED WITH A LOW VAULT AND REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT THIS DID NOT RESOLVE THE PROBLEM. ON THE SAME DAY THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "QUIET EYE". REPORTEDLY "NO, A NEW IMPLANT IS NOT PLANNED. NOW SHE IS USING HER SPECTACLES; HER EYES ARE IN PERFECT CONDITIONS; EPITHELIUM, CORNEA AND CRISTALLINE LENS ARE OK)."THE CAUSE OF THE EVENT WAS "OTHER". CAUSE DETAILS: "LENS MISALIGNMENT". CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. LENS RETURNED DAMAGED. VISUAL INSPECTION FOUND OPTIC TORN/DEFORMED AND HAPTIC BROKEN/BENT/DEFORMED. DUE TO DAMAGED STATE OF LENS, UNABLE TO PERFORM LENS INSPECTION (DIMENSIONAL/ REFRACTIVE). CLAIM#(B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. IMPLANT/EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -02.00/+6.0/104 | OD | Physician | 09/01/2021 | 2021 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12544646&pc=QCB | 2023826-2021-03792 | 2021-01-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -11.0/2.5/86 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REMOVED AND REPLACED VERTICALLY WITH A SAME LENGTH LENS WITHIN THE SAME SURGERY, PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BROKEN. DIMENSIONAL INSPECTION COULD NOT BE PERFORMED DUE TO SIGNIFICANT LENS DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: PATIENT ID: (B)(6). WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE CODE (B)(4): OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT AND ACD < 3.0MM). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -11.0/2.5/86 | OD | Physician | 08/30/2021 | 2021 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12550770&pc=QCB | 2023826-2021-03816 | 2021-08-06 | Injury | STAAR SURGICAL COMPANY | 2021-09-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -10.0/1.0/5 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/1.0/5 | OD | Physician | 09/06/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12550771&pc=QCB | 2023826-2021-03817 | 2021-08-02 | Injury | STAAR SURGICAL COMPANY | 2021-09-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.0/3.0/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -11.0/3.0/92 | OD | Physician | 09/05/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12550772&pc=QCB | 2023826-2021-03815 | 2021-05-07 | Injury | STAAR SURGICAL COMPANY | 2021-09-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -9.0/2.0/178 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -9.0/2.0/178 | OS | Physician | 09/05/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12549952&pc=QCB | 2023826-2021-03702 | 2021-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -02.50/+5.5/081 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2021. THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS EXPLANTED. THE PROBLEM WAS RESOLVED. THE PATIENT IS USING HER SPECTACLES AND HER EYES ARE IN PERFECT CONDITION. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT AND WITH SEVERE HAPTIC DAMAGE. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION NOT PERFORMED DUE TO DAMAGED LENS STATE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM #: (B)(4). | -02.50/+5.5/081 | OS | Physician | 09/01/2021 | 2021 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12558351&pc=QCB | 2023826-2021-03831 | 2021-08-10 | Injury | STAAR SURGICAL COMPANY | 2021-09-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.5/1.5/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM#: (B)(4). | -11.5/1.5/087 | OS | Physician | 09/13/2021 | 2021 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12558389&pc=QCB | 2023826-2021-03821 | 2021-08-07 | Injury | STAAR SURGICAL COMPANY | 2021-09-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -7.0/1.5/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM#: (B)(4). | -7.0/1.5/088 | OD | Physician | 09/08/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12566293&pc=QCB | 2023826-2021-03838 |  | Injury | STAAR SURGICAL COMPANY | 2021-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -14.0/+2.5/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. LENS ROTATION WAS OBSERVED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION THE LENS WAS EXPLANTED ON (B)(6) 2021 AND LATER REPLACED WITH A LONGER LENGTH LENS. THE REPORTER STATED "AS OF LAST VISIT, THE PATIENT HAS UNAIDED VA RE 6/6-2 AND NO COMPLAINTS. VAULT 1 CT". H3: DEVICE EVALUATION - LENS WAS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. DATE OF EVENT: UNK. A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS FOUND. CLAIM # (B)(4). | -14.0/+2.5/109 | OD | Physician | 09/02/2021 | 2021 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12566297&pc=QCB | 2023826-2021-03839 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-10-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -13.0/+1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION WAS OBSERVED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: D 6A.- (B)(6) 2020 SHOULD BE CORRECTED TO (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. DATE OF EVENT: UNK. A LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINT FOUND. CLAIM # (B)(4). | -13.0/+1.5/093 | OS | Physician | 09/02/2021 | 2021 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12566501&pc=QCB | 2023826-2021-03840 | 2021-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.00/+1.0/159, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON 04 APR 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULT, LENS REMAINS IMPLANTED. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.00/+1.0/159, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT IS UNKNOWN. CLAIM (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS FOUND. CLAIM # (B)(4). | -7.00/+1.0/159 | OD | Physician | 09/02/2021 | 2021 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12568640&pc=QCB | 2023826-2021-03770 | 2021-07-10 | Injury | STAAR SURGICAL COMPANY | 2021-10-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/3.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/3.0/088 | OS | Physician | 09/13/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12568642&pc=QCB | 2023826-2021-03775 | 2021-08-23 | Injury | STAAR SURGICAL COMPANY | 2021-10-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -0850/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -08.50/1.0/091 | OS | Physician | 09/15/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12568643&pc=QCB | 2023826-2021-03767 | 2021-08-11 | Injury | STAAR SURGICAL COMPANY | 2021-10-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -12.50/1.5/092 | OD | Physician | 09/13/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12568644&pc=QCB | 2023826-2021-03777 | 2021-08-03 | Injury | STAAR SURGICAL COMPANY | 2021-10-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -9.50/3.50/098 (SPHERE /CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -9.50/3.50/098 | OD | Physician | 09/16/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12568657&pc=QCB | 2023826-2021-03778 | 2021-08-09 | Injury | STAAR SURGICAL COMPANY | 2021-10-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -10.00/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -10.00/2.0/088 | OS | Physician | 09/16/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12573145&pc=QCB | 2023826-2021-03868 | 2021-08-13 | Injury | STAAR SURGICAL COMPANY | 2021-10-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -6.0/1.5/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -6.0/1.5/91 | OS | Physician | 09/13/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12573241&pc=QCB | 2023826-2021-03877 | 2021-09-03 | Injury | STAAR SURGICAL COMPANY | 2021-10-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.5/2.0/88 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -11.5/2.0/88 | OS | Physician | 09/16/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12573245&pc=QCB | 2023826-2021-03871 | 2021-08-18 | Injury | STAAR SURGICAL COMPANY | 2021-10-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -10.0/1.0/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/1.0/088 | OD | Physician | 09/14/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12573260&pc=QCB | 2023826-2021-03874 | 2021-08-27 | Injury | STAAR SURGICAL COMPANY | 2021-10-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -6.5/1.5/165 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. MEDICAL DEVICE PROBLEM CODE: (B)(4): OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -6.5/1.5/165 | OS | Physician | 09/15/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12573314&pc=QCB | 2023826-2021-03875 | 2021-08-27 | Injury | STAAR SURGICAL COMPANY | 2021-10-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -9.5/1.5/006 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. MEDICAL DEVICE PROBLEM CODE: (B)(4): OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | -9.5/1.5/006 | OS | Physician | 09/15/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12573323&pc=QCB | 2023826-2021-03867 | 2021-08-25 | Injury | STAAR SURGICAL COMPANY | 2021-10-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.5/1.5/085 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -9.5/1.5/085 | OS | Physician | 09/13/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12573330&pc=QCB | 2023826-2021-03870 | 2021-08-05 | Injury | STAAR SURGICAL COMPANY | 2021-10-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -10.0/1.5/93 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -10.0/1.5/93 | OS | Physician | 09/13/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12573694&pc=QCB | 2023826-2021-03858 | 2021-08-07 | Injury | STAAR SURGICAL COMPANY | 2021-10-04 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.00/+3.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO THE LENS TORE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.00/+3.0/091 | OS | Physician | 09/05/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12566801&pc=QCB | 2023826-2021-03836 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -11.5/1.5/109, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021 AS A REPLACEMENT LENS TO CORRECT EXCESSIVE VAULT BUT IT DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. " IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PREVIOUS MFR REPORT IS NOT APPLICABLE, THE INFORMATION THAT WAS INITIALLY PROVIDED WAS INCORRECT. PER UPDATED INFORMATION THIS REPLACEMENT LENS DID RESOLVE THE PROBLEM. SEE MFR# 2023826-2021-03835 FOR CLAIM AGAINST THE INITIAL LENS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/1.5/109 | OD | Physician | 09/03/2021 | 2021 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12566805&pc=QCB | 2023826-2021-03835 | 2021-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.5/1.5/095, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE DID NOT RESOLVE THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -11.5/1.5/095, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE DID RESOLVE THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT - RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -11.5/1.5/095 | OD | Physician | 09/03/2021 | 2021 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12568637&pc=QCB | 2023826-2021-03766 | 2021-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -09.0/3.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - REPORTED LENS IMPLANTED IN PATIENTS RIGHT EYE(OD). - CORRECT TO PATIENTS LEFT EYE (OS). CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510 K- THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -09.0/3.0/087 | OS |  | 09/13/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12580318&pc=QCB | 2023826-2021-03878 | 2021-06-23 | Injury | STAAR SURGICAL COMPANY | 2021-10-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.5/2.5/88 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -12.5/2.5/88 | OD | Physician | 09/17/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12580319&pc=QCB | 2023826-2021-03887 | 2021-07-23 | Injury | STAAR SURGICAL COMPANY | 2021-10-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -11.50/+1.5/128 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. IT WAS REPORTED THE SURGEON USED THE WRONG AXIS FOR IMPLANTATION WITH NO ADVERSE EVENTS. LATER THE SURGEON EXCHANGED THE LENS WITH A SAME MODEL/LENGTH/POWER BUT DIFFERENT AXIS ON (B)(6) 2021 AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS USER ERROR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -11.50/+1.5/128 | OD | Physician | 10/05/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12580323&pc=QCB | 2023826-2021-03879 | 2021-08-04 | Injury | STAAR SURGICAL COMPANY | 2021-10-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -14.5/2.0/170 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM#: (B)(4). | -14.5/2.0/170 | OD | Physician | 09/20/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12580324&pc=QCB | 2023826-2021-03884 | 2021-07-26 | Injury | STAAR SURGICAL COMPANY | 2021-10-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -12.0/1.5/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -12.0/1.5/89 | OD | Physician | 09/20/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12580327&pc=QCB | 2023826-2021-03883 | 2021-07-17 | Injury | STAAR SURGICAL COMPANY | 2021-10-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -9.5/2.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -9.5/2.0/092 | OD | Physician | 09/20/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12580456&pc=QCB | 2023826-2021-03861 | 2021-03-30 | Injury | STAAR SURGICAL COMPANY | 2021-10-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -13.50/+1.0/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME LENGTH BUT DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT-RELATED FACTOR BUT THE DEVICE ALSO FAILED TO PERFORM. Manufacturer Narrative: AGE/DATE OF BIRTH: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM # (B)(4). | -13.50/+1.0/083 | OS | Physician | 09/06/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12580470&pc=QCB | 2023826-2021-03862 | 2021-07-17 | Injury | STAAR SURGICAL COMPANY | 2021-10-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -06.00/+3.5/042 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT-RELATED FACTOR BUT THE DEVICE ALSO FAILED TO PERFORM. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE EVALUATION NOT REQUIRED. (B)(4). CLAIM # (B)(4). | -06.00/+3.5/042 | OD | Physician | 09/06/2021 | 2021 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12580552&pc=QCB | 2023826-2021-03888 | 2021-08-23 | Malfunction | STAAR SURGICAL COMPANY | 2021-10-05 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -18.0/2.5/089, IMPLANTABLE COLLAMER LENS, TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE ON (B)(6) 2021 THE LENS WAS IMPLANTED AND REMOVED WITHIN THE SAME SURGERY AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT THERE WAS NO PATIENT INJURY. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. A LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINT WAS FOUND. CLAIM # (B)(4). | -18.0/2.5/089 | NA | Physician | 09/05/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12586890&pc=QCB | 2023826-2021-03902 | 2021-09-20 | Injury | STAAR SURGICAL COMPANY | 2021-10-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.0/3.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -9.0/3.0/090 | OS | Physician | 09/24/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12586891&pc=QCB | 2023826-2021-03885 | 2021-08-10 | Injury | STAAR SURGICAL COMPANY | 2021-10-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -10.0/1.5/82 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -10.0/1.5/82 | OS | Physician | 09/20/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12586892&pc=QCB | 2023826-2021-03904 | 2021-09-20 | Malfunction | STAAR SURGICAL COMPANY | 2021-10-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -7.0/1.0/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -7.0/1.0/92 | OS | Physician | 09/24/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12586896&pc=QCB | 2023826-2021-03899 | 2021-07-07 | Injury | STAAR SURGICAL COMPANY | 2021-10-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.5/2.0/86 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -9.5/2.0/86 | OD | Physician | 09/20/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12586897&pc=QCB | 2023826-2021-03903 | 2021-09-13 | Malfunction | STAAR SURGICAL COMPANY | 2021-10-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -10.5/2.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -10.5/2.0/091 | OD | Physician | 09/24/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12587296&pc=QCB | 2023826-2021-03891 | 2021-08-23 | Malfunction | STAAR SURGICAL COMPANY | 2021-10-06 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO_12.6, -18.00/2.5/90, IMPLANTABLE COLLAMER LENS, TORE/BROKE DURING INJECTION/DELIVERY INTO THE LEFT EYE (OS) ON (B)(6) 2021. THERE WAS PATIENT CONTACT BUT NO PATIENTS INJURY. THE LENS WAS IMPLANTED AND REMOVED WITHIN THE SAME SURGERY AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. A LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINT WAS FOUND. CLAIM # (B)(4). | -18.00/2.5/90 | OS | Physician | 09/06/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12587027&pc=QCB | 2023826-2021-03889 | 2021-04-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, --13.5/1.0/032, IMPLANTABLE COLLAMER LENS, TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE ON (B)(6) 2021. THERE WAS PATIENT CONTACT BUT NO PATIENTS INJURY. THE LENS WAS REPLACED WITHIN THE SAME SURGERY AS REMOVAL OF THE INITIAL LENS. IT WAS REPORTED THAT THE PROBLEM DID NOT RESOLVE. FOR ADDITIONAL INFORMATION THE REPORTER STATED " DOCTOR BROKE THE NEW LENS AGAIN." IN THE REPORTERS OPINION USER ERROR WAS THE CAUSE OF THIS EVENT. THE REPORTER ALSO STATED THAT IT WAS UNKNOWN IF THE DEVICE FAIL TO PERFORM AS INTENDED. FOR CAUSE DETAILS THE REPORTER STATED " DOCTOR DID NOT LOAD THE LENS WELL." Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE : UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -13.5/1.0/032 | NA | Physician | 09/06/2021 | 2021 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12587324&pc=QCB | 2023826-2021-03892 | 2021-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.00/+2.0/094, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) IN 2019. LOW VAULT WITH ROTATION HAS BEEN OBSERVED, LENS REMAINS IMPLANTED. THE REPORTER STATED THAT THE PATIENT CURRENTLY DOES NOT WISH TO EXCHANGE THE LENS. CAUSE OF EVENT IS UNKNOWN.Event Description: AL THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+2.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) IN 2019. LOW VAULT WITH LENS ROTATION WAS OBSERVED AND THE LENS WAS EXCHANGED AND REPLACED WITH A LONGER LENGTH LENS ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: B5: PER UPDATED INFORMATION: LENS EXCHANGED WITH A LONGER LENGTH LENS AND PROBLEM RESOLVED. H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. Manufacturer Narrative: H6- MEDICAL DEVICE PROBLEM CODE: ADDED CODE 2923. H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B1: ADVERSE EVENT. B2: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. B5: ADDITIONAL INFORMATION: AL THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+2.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) IN 2019. LOW VAULT WITH LENS ROTATION WAS OBSERVED AND THE LENS WAS EXCHANGED AND REPLACED WITH A LONGER LENGTH LENS ON (B)(6) /2022. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. H6: IMPACT CODE: 4629. CLAIM#(B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US.A LENS WORK ORDER SEARCH WAS PERFORMED. ONE SIMILAR COMPLAINT WAS FOUND. CLAIM # (B)(4). | -12.00/+2.0/094 | OD | Physician | 09/28/2022 | 2022 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12587411&pc=QCB | 2023826-2021-03893 | 2021-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICM513.2, -10/+2/164, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS, DIFFERENT POWER AND THE PROBLEM RESOLVED. "USER ERROR" WAS REPORTED TO BE THE CAUSE OF THIS EVENT. IT WAS REPORTED THAT THE PROBLEM DID NOT FAIL TO PERFORM AS INTENDED. CAUSE DETAILS " MANIFEST REFRACTION AXIS TYPED 75 INSTEAD OF 175 CAUSING THE AXIS OF THE ORDERED TICL TO BE WRONG." Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WERE FOUND. CLAIM # (B)(4). | -10/+2/164 | OS | Physician | 09/06/2021 | 2021 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12594834&pc=QCB | 2023826-2021-03925 | 2021-08-16 | Injury | STAAR SURGICAL COMPANY | 2021-10-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -14.5/2.0/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -14.5/2.0/088 | OD | Physician | 09/25/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12594922&pc=QCB | 2023826-2021-03930 | 2021-08-26 | Injury | STAAR SURGICAL COMPANY | 2021-10-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -14.5/1.5/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -14.5/1.5/91 | OS | Physician | 09/26/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12595354&pc=QCB | 2023826-2021-03923 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-10-07 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM, VTICMO13.7,-18.00/3.5/101, IMPLANTABLE COLLAMER LENS TORE DURING LOADING. THERE WAS NO PATIENT CONTACT. CAUSE OF EVENT WAS UNKNOWN, REPORTER STATED " LENS TEAR DURING MOVEMENT ON THE CARTRIDGE." Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE FOUND. CLAIM # (B)(4). | -18.00/3.5/101 | NA | Physician | 09/07/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12595191&pc=QCB | 2023826-2021-03922 | 2021-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-17.00/2.0/059, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS AND WAS IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, INSIDE A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE : UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE FOUND. CLAIM # (B)(4). | -17.00/2.0/059 | OD | Physician | 09/07/2021 | 2021 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12595355&pc=QCB | 2023826-2021-03921 | 2021-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM512.6,-5.00/+4.5/088, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. . CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.6MM, VTICM512.6,-5.00/+4.5/088, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE FOUND. CLAIM # (B)(4). | -5.00/+4.5/088 | OD | Physician | 09/07/2021 | 2021 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12599516&pc=QCB | 2023826-2021-03933 | 2021-09-07 | Injury | STAAR SURGICAL COMPANY | 2021-10-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -6.0/1.0/74 (SPHERE/C YLINDER/ AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. LATER ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM #: (B)(4). | -6.0/1.0/74 | OD | Physician | 09/27/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12599526&pc=QCB | 2023826-2021-03935 | 2021-09-04 | Injury | STAAR SURGICAL COMPANY | 2021-10-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -5.0/1.0/075 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.0/1.0/075 | OS | Physician | 09/27/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12599757&pc=QCB | 2023826-2021-03914 | 2021-09-01 | Malfunction | STAAR SURGICAL COMPANY | 2021-10-08 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON INJECTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.0/+2.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2021. THE LENS WAS REMOVED INTRA-OPERATIVELY. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AT A LATER DATE AND THE PROBLEM IS RESOLVED. NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -16.0/+2.0/095 | OD | Physician | 09/13/2021 | 2021 | 36
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12601875&pc=QCB | 2023826-2021-03938 | 2021-08-25 | Injury | STAAR SURGICAL COMPANY | 2021-10-08 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -10.0/2.5/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -10.0/2.5/093 | OD | Physician | 09/27/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12599746&pc=QCB | 2023826-2021-03918 | 2021-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -13.5/+2.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. EXCESSIVE VAULT IS REPORTED. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-UPDATED INFORMATION: ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.5/+2.0/091 | OS | Physician | 09/16/2021 | 2021 | 46 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12602572&pc=QCB | 2023826-2021-03942 | 2021-08-11 | Injury | STAAR SURGICAL COMPANY | 2021-10-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1,-11/3/25, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REMOVED AND A LENS OF LONGER LENGTH WAS LATER IMPLANTED DUE TO LOW VAULT WITH ROTATION. PROBLEM WAS RESOLVED. IN THE REPORTERS OPINION "DEVICE AND USER ERROR" WAS THE CAUSE OF THE EVENT. THE REPORTER STATED " PROBLEM BELONG TO INCORRECT MEASUREMENT AND USER ERROR." Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE FOUND. CLAIM # (B)(4). | -11/3/25 | OD | Physician | 09/09/2021 | 2021 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12602576&pc=QCB | 2023826-2021-03945 | 2021-09-09 | Malfunction | STAAR SURGICAL COMPANY | 2021-10-09 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -14.0/+2.0/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, TORE/BROKE DURING INJECTION/DELIVERY INTO THE LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS IMPLANTED, REMOVED AND REPLACED WITHIN THE SAME SURGERY. IT WAS REPORTED THAT THERE WAS NO PATIENT INJURY. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. (B)(4). CLAIM # (B)(4). | -14.0/+2.0/92 | OS | Physician | 09/09/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12602461&pc=QCB | 2023826-2021-03924 | 2021-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-11.5/1.5/015, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021 .EXCESSIVE VAULT (ARCH HEIGHT NOT APPROPRIATE) WAS OBSERVED. ON (B)(6) 2021 THE LENS WAS REPOSITIONED. THE LENS WAS LATER EXCHANGED FOR A SAME LENGTH LENS BUT IMPLANTED VERTICALLY. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY, INSIDE A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY & RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE FOUND. CLAIM #: (B)(4). | -11.5/1.5/015 | OS | Physician | 09/08/2021 | 2021 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12602574&pc=QCB | 2023826-2021-03943 | 2021-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1,-12.50/2.00/94, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6)2021 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: B5- PER UPDATED INFORMATION THE REPLACEMENT LENS WAS IMPLANTED VERTICALLY. CAUSE IS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE/DEBRIS. VISUAL INSPECTION FOUND ONE OF THE HAPTIC TORN AND STRETCHED OUT. DRY CLEAR RESIDUE AND DEBRIS THROUGHOUT THE LENS WAS ALSO OBSERVED. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE FOUND. CLAIM # (B)(4). | -12.50/2.00/94 | OD | Physician | 09/09/2021 | 2021 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12602577&pc=QCB | 2023826-2021-03944 | 2021-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -11.50/+2.00/81, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE/DEBRIS. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN AND DEBRIS AND RESIDUE THROUGHOUT THE LENS. CLAIM# (B)(4). Manufacturer Narrative: B5- PER UPDATED INFORMATION THE REPLACEMENT LENS WAS IMPLANTED VERTICALLY. CAUSE IS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE : UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE FOUND. CLAIM # (B)(4). | -11.50/+2.00/81 | OS | Physician | 09/09/2021 | 2021 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12606068&pc=QCB | 2023826-2021-03854 | 2021-09-13 | Malfunction | STAAR SURGICAL COMPANY | 2021-10-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -07.50/0.5/094 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -07.50/0.5/094 | OS | Physician | 09/24/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12606072&pc=QCB | 2023826-2021-03856 | 2021-07-29 | Injury | STAAR SURGICAL COMPANY | 2021-10-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.00/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -8.00/1.5/092 | OS | Physician | 09/24/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12606074&pc=QCB | 2023826-2021-03844 | 2021-08-05 | Injury | STAAR SURGICAL COMPANY | 2021-10-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.50/1.0/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER SPHERICAL LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -9.50/1.0/111 | OS | Physician | 09/20/2021 | 2021 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12611493&pc=QCB | 2023826-2021-03939 | 2021-08-25 | Injury | STAAR SURGICAL COMPANY | 2021-10-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -8.0/1.5/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -8.0/1.5/89 | OS | Physician | 09/27/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12611512&pc=QCB | 2023826-2021-03993 | 2020-08-17 | Injury | STAAR SURGICAL COMPANY | 2021-10-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.5/4.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/4.0/089 | OS | Physician | 09/29/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12611523&pc=QCB | 2023826-2021-03994 | 2020-08-17 | Injury | STAAR SURGICAL COMPANY | 2021-10-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.5/3.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM#: (B)(4). | -10.5/3.5/091 | OD | Physician | 09/29/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12611524&pc=QCB | 2023826-2021-03996 | 2020-02-26 | Injury | STAAR SURGICAL COMPANY | 2021-10-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED 12.6MM VTICMO12.6; -9.5/4.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -9.5/4.0/089 | OS | Physician | 09/29/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12611535&pc=QCB | 2023826-2021-03995 | 2020-01-20 | Injury | STAAR SURGICAL COMPANY | 2021-10-11 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.5/1.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/1.5/091 | OS | Physician | 09/29/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12603563&pc=QCB | 2023826-2021-03947 | 2021-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.50/+1.0/091, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LOW VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER UPDATED INFORMATION NO MEDICAL OR SURGICAL INTERVENTION IS NECESSARY. PROBLEM RESOLED. THE REPORTER STATED THAT THE PATIENT IS HAPPY. "CLOSED MONITORING IS PLANNED" WAS REPORTED. (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 -7.50/+1.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. LOW VAULT WAS OBSERVED. LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS AND PROBLEM WAS RESOLVED. ALL FINE AND PATIENT IS HAPPY. CLAIM#(B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US.WORK ORDER SEARCH: NO SIMILAR COMPLAINT WERE FOUND. CLAIM # (B)(4). | -7.50/+1.0/091 | OD | Physician | 09/10/2021 | 2021 | 50 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12606070&pc=QCB | 2023826-2021-03855 | 2021-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/1.0/024 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTION OF INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/1.0/024 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED VERTICALLY WITH A SAME LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.50/1.0/024 | OD | Physician | 09/24/2021 | 2021 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12619781&pc=QCB | 2023826-2021-03985 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-10-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -06.50/+3.5/145 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) IN 2011. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: AGE: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE : UNK. DATE OF EVENT: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -06.50/+3.5/145 | OD | Physician | 09/13/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12619784&pc=QCB | 2023826-2021-03998 | 2021-08-31 | Injury | STAAR SURGICAL COMPANY | 2021-10-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.0/0.5/172 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -9.0/0.5/172 | OS | Physician | 09/29/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12619785&pc=QCB | 2023826-2021-04001 | 2021-04-21 | Injury | STAAR SURGICAL COMPANY | 2021-10-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -7.0/4.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -7.0/4.0/092 | OS | Physician | 09/29/2021 | 2021 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12619789&pc=QCB | 2023826-2021-03997 | 2020-02-25 | Injury | STAAR SURGICAL COMPANY | 2021-10-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.5/4.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -9.5/4.0/093 | OD | Physician | 09/29/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12619790&pc=QCB | 2023826-2021-04000 | 2021-09-14 | Injury | STAAR SURGICAL COMPANY | 2021-10-12 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED 12.6MM VTICMO12.6; -9.0/3.5/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -9.0/3.5/89 | OD | Physician | 09/29/2021 | 2021 | 40
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12619770&pc=QCB | 2023826-2021-03982 | 2021-09-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AND REMOVED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -08.50/+2.5/111 (SPHERE/CYLINDER/AXIS), ON (B)(6) 2021 DUE TO THE LENS WAS IMPLANTED UPSIDE DOWN. THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -08.50/+2.5/111 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO THE LENS WAS IMPLANTED UPSIDE DOWN AND WAS DAMAGED. THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED DURING THE SAME SURGERY WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -08.50/+2.5/111 | NA | Other Health Care Professional | 09/15/2021 | 2021 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12628887&pc=QCB | 2023826-2021-04010 | 2021-08-27 | Injury | STAAR SURGICAL COMPANY | 2021-10-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.5/1.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM#: (B)(4). | -9.5/1.5/089 | OS | Physician | 09/29/2021 | 2021 | 46
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12628892&pc=QCB | 2023826-2021-04007 | 2021-08-07 | Injury | STAAR SURGICAL COMPANY | 2021-10-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -14.0/2.0/067 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND LATER ON THE SAME DATE REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -14.0/2.0/067 | OS | Physician | 09/29/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12628893&pc=QCB | 2023826-2021-04005 | 2020-11-03 | Injury | STAAR SURGICAL COMPANY | 2021-10-13 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.5/3.5/97 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -13.5/3.5/97 | OD | Physician | 09/29/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12628834&pc=QCB | 2023826-2021-03988 | 2021-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.50/+0.5/038 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED AND REPLACED ON (B)(6) 2021. THE PROBLEM WAS RESOLVED AND THE PATIENT IS HAPPY. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -08.50/+0.5/038 | OD | Physician | 09/14/2021 | 2021 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12637186&pc=QCB | 2023826-2021-03954 | 2021-06-28 | Malfunction | STAAR SURGICAL COMPANY | 2021-10-14 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.5/+1.0/105 (SPHERE/CYLINDER/AXIS) TORE DURING INJECTION INTO THE LEFT (OS) EYE. THERE WAS PATIENT CONTACT WITH THE LENS BUT NO PATIENT INJURY. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY ON (B)(6) 2021. AN ALTERNATE LENS WAS IMPLANTED AT A LATER DATE AND THE PROBLEM IS RESOLVED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -7.5/+1.0/105 | OS | Physician | 09/20/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12637280&pc=QCB | 2023826-2021-04022 | 2021-09-24 | Injury | STAAR SURGICAL COMPANY | 2021-10-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2; -13.5/2.5/114 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -13.5/2.5/114 | OD | Physician | 09/29/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12637283&pc=QCB | 2023826-2021-04021 | 2021-09-02 | Injury | STAAR SURGICAL COMPANY | 2021-10-14 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -15.0/2.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND LATER THE SAME DATE REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -15.0/2.0/089 | OS | Physician | 09/29/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12637140&pc=QCB | 2023826-2021-03957 | 2021-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.0/+1.5/109 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE SURGEON REPORTS THAT, "BECAUSE THE ASTIGMATISM SHEET ROTATES, IT IS REPLACED WITH A NON-ASTIGMATISM SHEET." Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISIBLE INSPECTION FOUND HAPTIC BENT/BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4) LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#:(B)(4). | -11.0/+1.5/109 | OD | Physician | 09/23/2021 | 2021 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12637153&pc=QCB | 2023826-2021-03952 | 2021-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -15.0/+1.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LOW VAULT WAS REPORTED. REPORTEDLY, THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5: ON (B)(6) 2021; THE LENS WAS EXCHANGED FOR A LONGER LENTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. IT IS REPORTED, "PATIENT IS HAPPY." CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -15.0/+1.5/086 | OD | Physician | 09/20/2021 | 2021 | 46 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12642783&pc=QCB | 2023826-2021-04027 | 2021-08-26 | Injury | STAAR SURGICAL COMPANY | 2021-10-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -12.00/+3.00/88 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -12.00/+3.00/88 | OS | Physician | 10/01/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12643536&pc=QCB | 2023826-2021-04013 | 2021-09-08 | Malfunction | STAAR SURGICAL COMPANY | 2021-10-15 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -13.0/0.5/049, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LOW VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. IT WAS REPORTED THAT MEDICAL OR SURGICAL INTERVENTION WAS NOT NECESSARY AND THAT THE PROBLEM DID NOT RESOLVE. THE REPORTER STATED "PATIENT DOES NOT WANT TO EXCHANGE THE LENS. IS INFORMED ABOUT CATARACT RISK." CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS FOUND. CLAIM # (B)(4). | -13.0/0.5/049 | OD | Physician | 09/15/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12643696&pc=QCB | 2023826-2021-04016 | 2021-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -14.00/3.50/90, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR THE SAME LENGTH, DIFFERENT MODEL AND POWER LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN A MICRO-CENTRIFUGE VIAL WITH DEBRIS. VISUAL INSPECTION FOUND ONE OF THE HAPTICS BENT AND DEBRIS ON LENS OPTIC. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -14.00/3.50/90 | OS | Physician | 09/16/2021 | 2021 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12651397&pc=QCB | 2023826-2021-04049 | 2021-07-29 | Injury | STAAR SURGICAL COMPANY | 2021-10-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.0/2.5/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -13.0/2.5/89 | OD | Physician | 10/10/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12651436&pc=QCB | 2023826-2021-04032 | 2021-09-02 | Injury | STAAR SURGICAL COMPANY | 2021-10-18 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -9.5/1.5/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. MEDICAL DEVICE PROBLEM CODE: (B)(4): OFF-LABEL USE (ACD < 3.0MM). CLAIM#: (B)(4). | -9.5/1.5/91 | OS | Physician | 10/06/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12662186&pc=QCB | 2023826-2021-04052 | 2021-06-10 | Injury | STAAR SURGICAL COMPANY | 2021-10-19 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -10.5/1.5/114 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -10.5/1.5/114 | OD | Physician | 10/11/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12665972&pc=QCB | 2023826-2021-03963 | 2021-08-19 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.50/1.0/113 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -7.50/1.0/113 | OD | Physician | 09/26/2021 | 2021 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12665986&pc=QCB | 2023826-2021-03980 | 2021-08-18 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -11.50/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.50/1.0/089 | OS | Physician | 09/29/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12665996&pc=QCB | 2023826-2021-03966 | 2021-08-11 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.50/2.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED VERTICALLY WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. REPORT SOURCE - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.50/2.0/086 | OD | Physician | 09/26/2021 | 2021 | 30
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12666002&pc=QCB | 2023826-2021-03970 | 2021-09-08 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.50/1.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. REPORT SOURCE - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.50/1.5/094 | OS | Physician | 09/27/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12666012&pc=QCB | 2023826-2021-03962 | 2021-08-28 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.50/0.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED IN A SECOND SURGERY ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. REPORT SOURCE - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -7.50/0.5/095 | OD | Physician | 09/27/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12666020&pc=QCB | 2023826-2021-03977 | 2020-03-30 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/2.0/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.50/2.0/072 | OD | Physician | 09/29/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12666026&pc=QCB | 2023826-2021-03969 | 2021-08-26 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -12.50/2.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.50/2.5/082 | OD | Physician | 09/27/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12666041&pc=QCB | 2023826-2021-03961 | 2021-09-01 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/3.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.00/3.5/091 | OD | Physician | 09/26/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12666046&pc=QCB | 2023826-2021-03968 | 2021-09-16 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -16.00/3.5/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2021 DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -16.00/3.5/078 | OS | Physician | 09/27/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12666052&pc=QCB | 2023826-2021-03967 | 2021-09-02 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -16.00/3.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM# (B)(4). | -16.00/3.5/083 | OD | Physician | 09/27/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12666055&pc=QCB | 2023826-2021-03965 | 2021-08-11 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.0/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -15.0/1.5/090 | OD | Physician | 09/26/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12666058&pc=QCB | 2023826-2021-03971 | 2021-09-08 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.50/1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. REPORT SOURCE - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.50/1.0/092 | OD | Physician | 09/27/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12666064&pc=QCB | 2023826-2021-03976 | 2020-04-27 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/4.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPLACED VERTICALLY WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. REPORT SOURCE - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.50/4.0/082 | OS | Physician | 09/21/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12669501&pc=QCB | 2023826-2021-04040 | 2021-09-16 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM # (B)(4). | -10.50/+1.5/089 | OS | Physician | 09/20/2021 | 2021 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12669507&pc=QCB | 2023826-2021-04034 | 2021-09-09 | Malfunction | STAAR SURGICAL COMPANY | 2021-10-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/010 (SPHERE/CYLINDER/AXIS), TORE DURING LOADING, AND THERE WAS NO PATIENT CONTACT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). CLAIM # (B)(4). | -10.50/+1.0/010 | NA | Physician | 09/28/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12669515&pc=QCB | 2023826-2021-04036 | 2021-07-20 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM TMICL12.1 IMPLANTABLE COLLAMER LENS, -10.00/+3.0/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. (B)(4). CLAIM # (B)(4). | -10.00/+3.0/091 | OS | Other Health Care Professional | 09/20/2021 | 2021 | 48
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12669516&pc=QCB | 2023826-2021-04037 | 2021-07-29 | Malfunction | STAAR SURGICAL COMPANY | 2021-10-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON WAS INSERTING A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.00/+3.5/115 (SPHERE/CYLINDER/AXIS) AND THE LENS TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD) EYE. THERE WAS PATIENT CONTACT BUT NO INJURY. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN (THE SURGEON THOUGHT THE LENS CAUGHT IN THE INJECTOR SYSTEM BUT THE ERROR WAS FROM THE SURGEON, THE LENS WAS NOT LOADED WELL). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -13.00/+3.5/115 | OD | Physician | 09/20/2021 | 2021 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12669518&pc=QCB | 2023826-2021-04039 | 2021-09-16 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.00/+0.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -08.00/+0.5/094 | OD | Physician | 09/20/2021 | 2021 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12669527&pc=QCB | 2023826-2021-04038 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-10-20 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON WAS INSERTING A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+3.0/045 (SPHERE/ CYLINDER/ AXIS) AND THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS) EYE. THERE WAS PATIENT CONTACT BUT NO INJURY. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN (THE SURGEON THOUGHT THE LENS CAUGHT IN THE INJECTOR SYSTEM BUT THE ERROR WAS FROM THE SURGEON, THE LENS WAS NOT LOADED WELL). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). | -18.00/+3.0/045 | OS | Physician | 09/20/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12666018&pc=QCB | 2023826-2021-03974 | 2021-06-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Optical Decentration; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -16.50/4.0/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTION: H6 - DEVICE PROBLEM CODE: 1360 (OPTICAL DECENTRATION) ADDED. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -16.50/4.0/097 | OS | Physician | 09/27/2021 | 2021 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12666023&pc=QCB | 2023826-2021-03964 | 2021-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - REPORTED IMPLANT DATE AS (B)(6) 2021 - CORRECT TO (B)(6) 2021. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. REPORT SOURCE - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.50/1.5/093 | OD | Physician | 09/26/2021 | 2021 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12669523&pc=QCB | 2023826-2021-04048 | 2021-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/165 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE INITIAL LENS WAS IMPLANTED VERTICALLY AND THE REPLACEMENT LENS WAS IMPLANTED HORIZONTALLY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISIBLE INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). CLAIM # (B)(4). | -10.00/+1.5/165 | OS | Physician | 09/29/2021 | 2021 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12669569&pc=QCB | 2023826-2021-04101 | 2021-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -5.5/1.0/088, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THIS EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. IT WAS REPORTED THAT AFTER THE ICL WAS REMOVED, THE PATIENT DID NOT WANT TO IMPLANT AGAIN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -5.5/1.0/088 | OD | Physician | 09/20/2021 | 2021 | 44 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12669570&pc=QCB | 2023826-2021-04099 | 2021-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -10.5/2.0/180 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. IT WAS REPORTED THAT THE FIRST ICL WAS IMPLANTED VERTICALLY AND THE SECOND ICL WAS IMPLANTED HORIZONTALLY. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTICS TORN,BROKEN AND STRETCHED OUT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, WEIGHT: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM #: (B)(4). | -10.5/2.0/180 | OD | Physician | 09/20/2021 | 2021 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12669571&pc=QCB | 2023826-2021-04102 | 2021-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -5.5/5.0/97, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. IT WAS REPORTED THAT THE REPLACEMENT LENS WAS SELECTED FOR VERTICAL IMPLANTATION. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AND THE HAPTICS BROKEN AND BENT. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO LENS DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.5/5.0/97 | OS | Physician | 09/20/2021 | 2021 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12669572&pc=QCB | 2023826-2021-04100 | 2021-02-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -5.0/1.0/064, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THIS EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. IT WAS REPORTED THAT AFTER THE ICL WAS REMOVED, THE PATIENT DID NOT WANT TO IMPLANT AGAIN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS BROKEN AND BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -5.0/1.0/064 | OS | Physician | 09/20/2021 | 2021 | 44 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12669629&pc=QCB | 2023826-2021-04103 | 2021-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -9.50/2.5/112, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LENS ROTATION WAS OBSERVED, UNKNOWN IF LENS WAS REPOSITIONED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER UPDATED INFORMATION A 13.2MM TMICL13.2 -9.50/2.50/112 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE PROBLEM DID NOT RESOLVE. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL AND DIOPTER LENS. THE REPORTER STATED "SO FAR THE ICL HAS NOT APPEARED TO HAVE ROTATED. THE NEW ICL THAT WAS PLACED IS SPHERICAL. PATIENT IS DOING WELL. " CLAIM#: (B)(4). Manufacturer Narrative: AGE, SEX, WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. (B)(4). CLAIM # (B)(4). | -9.50/2.5/112 | OD | Physician | 09/20/2021 | 2021 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12676111&pc=QCB | 2023826-2021-04096 | 2021-09-09 | Injury | STAAR SURGICAL COMPANY | 2021-10-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -10.0/1.5/062 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM #: (B)(4). | -10.0/1.5/062 | OS | Physician | 10/20/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12676124&pc=QCB | 2023826-2021-04098 | 2021-04-20 | Injury | STAAR SURGICAL COMPANY | 2021-10-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -7.5/1.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -7.5/1.0/090 | OS | Physician | 10/20/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12677879&pc=QCB | 2023826-2021-04043 | 2021-09-16 | Malfunction | STAAR SURGICAL COMPANY | 2021-10-21 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+1.5/110 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED DURING THE SAME SURGERY WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/501K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -12.50/+1.5/110 | OS | Physician | 09/24/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12676148&pc=QCB | 2023826-2021-04097 | 2021-08-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.0/1.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -11.0/1.0/089 | OD | Physician | 10/20/2021 | 2021 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685349&pc=QCB | 2023826-2021-04083 | 2021-08-10 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.0/+4.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "BECAUSE NO SAME DIOPTER LENS TO CHOSE WHEN RESERVING THE LENS FOR THE SECOND TIME, SO SELECT A SIMILAR TO THE EQUIVALENT SPHERICAL LENS." Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -18.0/+4.0/087 | OS | Physician | 09/29/2021 | 2021 | 38
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685353&pc=QCB | 2023826-2021-04074 | 2021-08-27 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/054 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.0/+1.5/054 | OS | Physician | 09/29/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685360&pc=QCB | 2023826-2021-04088 | 2021-08-27 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+1.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -9.5/+1.0/089 | OS | Physician | 09/29/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685362&pc=QCB | 2023826-2021-04069 | 2021-09-01 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.0/+2.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -18.0/+2.0/095 | OD | Physician | 09/29/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685378&pc=QCB | 2023826-2021-04073 | 2021-09-24 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -8.5/+3.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/+3.5/091 | OD | Physician | 09/29/2021 | 2021 | 19
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685392&pc=QCB | 2023826-2021-04078 | 2021-09-08 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -9.5/+1.5/081 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/+1.5/081 | OD | Physician | 09/29/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685394&pc=QCB | 2023826-2021-04072 | 2021-08-13 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+5.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+5.5/086 | OD | Physician | 09/29/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685402&pc=QCB | 2023826-2021-04067 | 2021-06-01 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -8.5/+1.5/091 | OD | Physician | 09/29/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685420&pc=QCB | 2023826-2021-04084 | 2021-09-27 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.5/+2.0/179 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(4) 2021. ON (B)(4) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE PROBLEM IS RESOLVED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -10.5/+2.0/179 | OS | Physician | 09/29/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685435&pc=QCB | 2023826-2021-04066 | 2021-07-17 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.5/+0.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -9.5/+0.5/087 | OD | Physician | 09/29/2021 | 2021 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685439&pc=QCB | 2023826-2021-04082 | 2021-07-23 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.5/+1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -10.5/+1.5/091 | OD | Physician | 09/29/2021 | 2021 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685442&pc=QCB | 2023826-2021-04075 | 2021-08-02 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -10.0/+1.0/074 | OS | Physician | 09/29/2021 | 2021 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685449&pc=QCB | 2023826-2021-04077 | 2021-09-09 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -10.5/+1.0/074 | OS | Physician | 09/29/2021 | 2021 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685451&pc=QCB | 2023826-2021-04071 | 2021-09-06 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.0/+1.5/007 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY THE FIRST IMPLANTED LENS WAS MARK AT 87 DEGREES. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+1.5/007 | OS | Physician | 09/29/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685342&pc=QCB | 2023826-2021-04065 | 2021-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -4.5/+1.0/108 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "THE PATIENT IS TOO HIGH AFTER THE TRANSISTOR IS IMPLANTED AND THE NEW CRYSTAL IS REQUIRED. THE PATIENT DOES NOT ACCEPT THE REPLACEMENT SURGERY, REQUIRING THE CRYSTAL, AND REMOVE THE CRYSTAL AFTER MULTIPLE COMMUNICATION." Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC BROKEN. DUE TO SIGNIFICANT LENS WIDTH DAMAGE, UNABLE TO CONFIRM CORRECT RE-HYDRATION OF THE LENS, HENCE TO SIZING. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -4.5/+1.0/108 | OS | Physician | 09/29/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685376&pc=QCB | 2023826-2021-04064 | 2021-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -3.0/+1.0/068 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "THE PATIENT IS TOO HIGH AFTER THE TRANSISTOR IS IMPLANTED AND THE NEW CRYSTAL IS REQUIRED. THE PATIENT DOES NOT ACCEPT THE REPLACEMENT SURGERY, REQUIRING THE CRYSTAL, AND REMOVE THE CRYSTAL AFTER MULTIPLE COMMUNICATION." Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -3.0/+1.0/068 | OD | Physician | 09/29/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12685415&pc=QCB | 2023826-2021-04090 | 2021-09-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -4.0/+0.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY THE PROBLEM IS NOT RESOLVED THOUGH THE LENS HAS BEEN EXPLANTED. NO FURTHER SURGICAL TREATMENT IS PLANNED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -4.0/+0.5/093 | OD | Physician | 09/30/2021 | 2021 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12687189&pc=QCB | 2023826-2021-04110 | 2021-09-01 | Malfunction | STAAR SURGICAL COMPANY | 2021-10-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.0/2.5/111, IMPLANTABLE COLLAMER LENS GOT STUCK IN CARTRIDGE OR INJECTION SYSTEM. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. CAUSE OF EVENT WAS UNKNOWN. PROBLEM RESOLVED WAS REPORTED. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. A LENS WORK ORDER SEARCH WAS PERFORMED. NO SIMILAR COMPLAINTS FOUND. CLAIM # (B)(4). | -13.0/2.5/111 | NA | Physician | 09/24/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12688694&pc=QCB | 2023826-2021-04093 | 2021-08-14 | Injury | STAAR SURGICAL COMPANY | 2021-10-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.00/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER SPHERICAL LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.00/1.0/091 | OS | Physician | 10/06/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12688703&pc=QCB | 2023826-2021-04094 | 2021-08-30 | Malfunction | STAAR SURGICAL COMPANY | 2021-10-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -3.00/2.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REPLACED INTRAOPERATIVELY WITH A LONGER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S LEFT EYE (OS) DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -3.00/2.0/081 | OS | Physician | 10/06/2021 | 2021 | 37
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12688704&pc=QCB | 2023826-2021-04092 | 2021-08-14 | Injury | STAAR SURGICAL COMPANY | 2021-10-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/1.0/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER SPHERICAL LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.50/1.0/073 | OD | Physician | 10/06/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12688709&pc=QCB | 2023826-2021-04095 | 2021-09-08 | Injury | STAAR SURGICAL COMPANY | 2021-10-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.5/1.0/08 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.5/1.0/08 | OS | Physician | 10/08/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12693735&pc=QCB | 2023826-2021-04127 | 2021-08-16 | Injury | STAAR SURGICAL COMPANY | 2021-10-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -7.0/4.5/94 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -7.0/4.5/94 | OS | Physician | 10/22/2021 | 2021 | 34
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12693854&pc=QCB | 2023826-2021-04124 | 2021-08-20 | Injury | STAAR SURGICAL COMPANY | 2021-10-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.0/3.5/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON 1(B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -9.0/3.5/89 | OD | Physician | 10/22/2021 | 2021 | 21
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12693913&pc=QCB | 2023826-2021-04125 | 2021-09-06 | Injury | STAAR SURGICAL COMPANY | 2021-10-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.0/1.0/34 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -10.0/1.0/34 | OD | Physician | 10/22/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12693945&pc=QCB | 2023826-2021-04126 | 2021-09-06 | Injury | STAAR SURGICAL COMPANY | 2021-10-25 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.0/1.0/148 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -10.0/1.0/148 | OS | Physician | 10/22/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12701037&pc=QCB | 2023826-2021-04144 | 2021-08-16 | Injury | STAAR SURGICAL COMPANY | 2021-10-26 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.5/1.0/93 (SPHERE/ CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.5/1.0/93 | OS | Physician | 10/23/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12707980&pc=QCB | 2023826-2021-04146 | 2021-08-24 | Injury | STAAR SURGICAL COMPANY | 2021-10-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.5/1.5/74 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -12.5/1.5/74 | OS | Physician | 10/24/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12708004&pc=QCB | 2023826-2021-04149 | 2021-10-20 | Malfunction | STAAR SURGICAL COMPANY | 2021-10-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.5/3.0/71 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -11.5/3.0/71 | OS | Physician | 10/25/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12708009&pc=QCB | 2023826-2021-04150 | 2021-09-04 | Injury | STAAR SURGICAL COMPANY | 2021-10-27 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -13.0/1.5/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -13.0/1.5/094 | OD | Physician | 10/25/2021 | 2021 | 45
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12716075&pc=QCB | 2023826-2021-04129 | 2021-09-08 | Malfunction | STAAR SURGICAL COMPANY | 2021-10-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+2.0/093 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -18.00/+2.0/093 | OD | Physician | 09/29/2021 | 2021 | 39
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12716098&pc=QCB | 2023826-2021-04137 | 2021-09-29 | Malfunction | STAAR SURGICAL COMPANY | 2021-10-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50/+0.5/046 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK RACE: UNK EXPLANT DATE: UNK PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -17.50/+0.5/046 | OD | Physician | 10/01/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12716156&pc=QCB | 2023826-2021-04138 | 2021-09-29 | Malfunction | STAAR SURGICAL COMPANY | 2021-10-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50/+0.5/153 (SPHERE / CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.50/+0.5/153 | OS | Physician | 10/01/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12716610&pc=QCB | 2023826-2021-04159 | 2021-09-08 | Injury | STAAR SURGICAL COMPANY | 2021-10-28 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -9.0/1.5/108 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -9.0/1.5/108 | OS | Physician | 10/27/2021 | 2021 | 25
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12716073&pc=QCB | 2023826-2021-04128 | 2021-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50/+4.5/174 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND LENS DISLOCATION/SUBLUXATION. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED SO THE LENS WAS EXPLANTED. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISIBLE INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.50/+4.5/174 | OD | Physician | 09/29/2021 | 2021 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12716087&pc=QCB | 2023826-2021-04131 | 2021-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -15.50/+1.5/122 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(4) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -15.50/+1.5/122 | OD | Physician | 09/29/2021 | 2021 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12724512&pc=QCB | 2023826-2021-04165 | 2021-09-01 | Injury | STAAR SURGICAL COMPANY | 2021-10-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.0/2.0/82 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S (B)(4). CLAIM#: (B)(4). | -11.0/2.0/82 | OD | Physician | 10/27/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12724514&pc=QCB | 2023826-2021-04163 | 2020-08-28 | Injury | STAAR SURGICAL COMPANY | 2021-10-29 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -17.0/1.5/84 (SPHERE/ CYLINDER/ AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.0/1.5/84 | OS | Physician | 10/27/2021 | 2021 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12724227&pc=QCB | 2023826-2021-04155 | 2021-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.0/+1.0/077 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-PREVIOUSLY STATED THE LENS WAS EXPLANTED AND THEN REPLACED WITH A SHORTER LENS. CORRECTED INFORMATION-THE LENS WAS EXCHANGED WITH A SHORTER LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. DATE OF EVENT - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -8.0/+1.0/077 | OS | Physician | 10/06/2021 | 2021 | 45 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725171&pc=QCB | 2023826-2021-04171 | 2021-08-19 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.5/1.0/155, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -8.5/1.0/155 | OD | Physician | 09/30/2021 | 2021 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725181&pc=QCB | 2023826-2021-04170 | 2021-08-19 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.0/1.5/112, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -10.0/1.5/112 | OS | Physician | 09/30/2021 | 2021 | 43
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725071&pc=QCB | 2023826-2021-04181 | 2021-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2 -9.5/3.5/073, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THIS EXPLANT RESOLVED THE PROBLEM. THE REPORTER STATED " THE PATIENTS ARCH HEIGHT WAS TOO HIGH AND THE LENS NEEDED TO BE REPLACED, BUT THE PATIENTS WAS CONCERNED ABOUT THE RISK OF A SECOND OPERATION AND DECIDED TO REMOVE THE LENS". CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -9.5/3.5/073 | OS | Physician | 09/29/2021 | 2021 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725076&pc=QCB | 2023826-2021-04169 | 2021-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -10.0/1.0/127, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME LENGTH DIFFERENT MODEL LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A BROKEN HAPTIC. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH:NO SIMILAR COMPLAINTS FOUND. CLAIM # (B)(4). | -10.0/1.0/127 | OS | Physician | 09/29/2021 | 2021 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12725252&pc=QCB | 2023826-2021-04180 | 2021-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.50/5.0/143, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. FOR STATUS OF THE EYE (SIGNS / SYMPTOMS) THE REPORTER STATED "STILL HAS EDEMA AND STITCH DUE TO WOUND LEAKAGE INTRA-OP. AC IS NOT SHALLOW ANYMORE. BUT IT WILL RECOVER SOON." CAUSE OF EVENT WAS REPORTED TO BE "PATIENT-RELATED FACTOR" AND WAS REPORTED THAT THE DEVICE FAILED TO PERFORM AS INTENDED. FOR CAUSE DETAILS THE REPORTER STATED "PATIENT HAS NARROW SULCUS. WTW IS 100% CORRECT BUT THE SULCUS IS NARROW." Manufacturer Narrative: D4: CORRECTED TO "SERIAL #: (B)(6). EXPIRATION DATE: 30 - JUN - 2024" IN INITAL MDR. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -7.50/5.0/143 | OD | Physician | 09/30/2021 | 2021 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12728117&pc=QCB | 2023826-2021-04114 | 2021-09-07 | Injury | STAAR SURGICAL COMPANY | 2021-11-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.50/2.0/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.50/2.0/070 | OS | Physician | 10/11/2021 | 2021 | 20
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12728122&pc=QCB | 2023826-2021-04120 |  | Injury | STAAR SURGICAL COMPANY | 2021-11-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2, -09.00/1.50/018 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. DATE OF EVENT - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -09.00/1.50/018 | OD | Physician | 10/13/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12728123&pc=QCB | 2023826-2021-04122 | 2021-09-04 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6, -9.00/4.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REPLACED INTRAOPERATIVELY WITH A LONGER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S RIGHT EYE (OD) DUE TO LOW VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K- THIS PRODUCT IS NOT MARKETED IN THE US CLAIM# (B)(4). | -9.00/4.0/092 | OD | Physician | 10/14/2021 | 2021 | 26
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12728124&pc=QCB | 2023826-2021-04121 | 2021-05-05 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6, -14.00/3.0/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REPLACED INTRAOPERATIVELY WITH A LONGER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S LEFT EYE (OS) DUE TO LOW VAULT.THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -14.00/3.0/103 | OS | Physician | 10/14/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12728129&pc=QCB | 2023826-2021-04119 |  | Injury | STAAR SURGICAL COMPANY | 2021-11-01 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/1.0/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, MEASUREMENTS ARE ON THE BORDERLINE BETWEEN 2 SIZES, POSSIBLE PATIENT-RELATED FACTOR. Manufacturer Narrative: WEIGHT - UNK, ETHNICITY- UNK, RACE- UNK, DATE OF EVENT- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.50/1.0/107 | OD | Physician | 10/12/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12741582&pc=QCB | 2023826-2021-04218 | 2020-03-14 | Injury | STAAR SURGICAL COMPANY | 2021-11-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -12.0/1.5/066 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -12.0/1.5/066 | OS | Physician | 10/27/2021 | 2021 | 42
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12741583&pc=QCB | 2023826-2021-04217 | 2021-09-01 | Injury | STAAR SURGICAL COMPANY | 2021-11-02 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -14.5/2.0/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -14.5/2.0/90 | OS | Physician | 10/27/2021 | 2021 | 24
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12749134&pc=QCB | 2023826-2021-04224 | 2021-07-20 | Injury | STAAR SURGICAL COMPANY | 2021-11-03 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO_13.2, -10.0/2.0/94, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LOW VAULT WITH ROTATION, PROBLEM RESOLVED. IT WAS ALSO REPORTED THAT ON THE SAME DAY THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM #: (B)(4). | -10.0/2.0/94 | OD | Physician | 10/04/2021 | 2021 | 33
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12749544&pc=QCB | 2023826-2021-04242 | 2021-08-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -12.0/1.0/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6 - TYPE OF INVESTIGATION- 4110 SHOULD BE REMOVED AS IT'S NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -12.0/1.0/087 | OD | Physician | 10/28/2021 | 2021 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12755751&pc=QCB | 2023826-2021-04205 | 2021-10-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS LOW VAULT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: ADDITIONAL INFORMATION: REPORTEDLY A 13.2MM VTICMO13.2 LENS WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2021. H6: INVESTIGATION CODE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN AND BLOOD STAINS ON LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -15.50/+1.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(4) 2021. LOW VAULT WAS OBSERVED. LENS WAS EXCHANGED ON 15 MAR 2022 FOR A LONGER LENGTH LENS AND PROBLEM WAS RESOLVED. ALL FINE AND PATIENT IS HAPPY. CLAIM# (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER, AGE, SEX, WEIGHT, ETHNICITY, RACE - UNK. DATE OF EVENT - UNK. PRODUCT CODE UNK; ESUBMITTER SOFTWARE DOES NOT ALLO FOR AN UNK OR BLANK ENTRY. IMPLANT DATE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE - UNK. CLAIM# (B)(4). | -15.50/+1.0/091 | OS | Physician | 10/12/2021 | 2021 | 53 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12755754&pc=QCB | 2023826-2021-04203 | 2021-09-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY ON (B)(6) 2021. AN ALTERNATE LENS WAS IMPLANTED AT A LATER DATE AND THE PROBLEM WAS RESOLVED. REPORTEDLY, THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -11.5/+2.5/091 | OS | Physician | 10/11/2021 | 2021 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12755757&pc=QCB | 2023826-2021-04204 | 2021-10-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS LOW VAULT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PATIENT IDENTIFIER, AGE, SEX, WEIGHT, ETHNICITY, RACE - UNK. DATE OF EVENT - UNK. PRODUCT CODE UNK; ESUBMITTER SOFTWARE DOES NOT ALLO FOR AN UNK OR BLANK ENTRY. IMPLANT DTE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE - UNK. CLAIM# (B)(4). Manufacturer Narrative: B5: ADDITIONAL INFORMATION: REPORTEDLY A 13.2MM VTICMO13.2, -14.5/+0.5/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2021. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. H6: INVESTIGATION CODE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -14.50/+0.5/102 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON 07 SEP 2021. LOW VAULT WAS OBSERVED. LENS WAS EXCHANGED ON 15 MAR 2022 FOR A LONGER LENGTH LENS AND PROBLEM WAS RESOLVED. ALL FINE AND PATIENT IS HAPPY. CLAIM# (B)(4). | -14.50/+0.5/102 | OD | Physician | 10/12/2021 | 2021 | 53 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12755760&pc=QCB | 2023826-2021-04207 | 2021-10-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/101 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE SURGEON REPORTS EXCESSIVE VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -8.0/+1.0/101 | OD | Physician | 10/13/2021 | 2021 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12755762&pc=QCB | 2023826-2021-04206 | 2021-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -13.0/+1.5/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN VIAL IN LIQUID. VISIBLE INSPECTION FOUND OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, & RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -13.0/+1.5/095 | OS | Physician | 10/13/2021 | 2021 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12756093&pc=QCB | 2023826-2021-04250 | 2021-04-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8.50/1.0/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6 -TYPE OF INVESTIGATION - 4110 SHOULD BE REMOVED AS IT'S NOT APPLICABLE. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -8.50/1.0/094 | OS | Physician | 10/29/2021 | 2021 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12756137&pc=QCB | 2023826-2021-04249 | 2021-06-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -7.00/1.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - TYPE OF INVESTIGATION: 4110 SHOULD BE REMOVED AS IT'S NOT APPLICABLE. CLAIM#: (B)(4). | -7.00/1.0/089 | OD | Physician | 10/28/2021 | 2021 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12756232&pc=QCB | 2023826-2021-04211 |  | Injury | STAAR SURGICAL COMPANY | 2021-11-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -10.50/+3.0/030 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS HAVING ROTATED, WAS SURGICALLY REPOSITIONED AND THEN HAD ROTATED AGAIN. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANT DATE: UNK. (B)(4). CLAIM # (B)(4). | -10.50/+3.0/030 | OD | Physician | 10/13/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12763555&pc=QCB | 2023826-2021-04255 | 2021-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.5/1.5/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -12.5/1.5/90 | OS | Physician | 11/01/2021 | 2021 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12763645&pc=QCB | 2023826-2021-04256 | 2021-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.0/1.5/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -11.0/1.5/092 | OD | Physician | 11/01/2021 | 2021 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12763658&pc=QCB | 2023826-2021-04254 | 2021-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.5/1.0/079 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -10.5/1.0/079 | OS | Physician | 11/01/2021 | 2021 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12763693&pc=QCB | 2023826-2021-04253 | 2021-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -14.0/1.0/109 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -14.0/1.0/109 | OD | Physician | 10/29/2021 | 2021 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12763750&pc=QCB | 2023826-2021-04234 | 2021-07-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.5/2.0/125, IMPLANTABLE COLLAMER LENS WAS FOUND DAMAGED DURING LOADING THE LENS INTO THE CARTRIDGE. THERE WAS NO PATIENT CONTACT. IT WAS REPORTED THAT A REPLACEMENT LENS WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -12.5/2.0/125 | NA | Physician | 10/06/2021 | 2021 | 49 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12763754&pc=QCB | 2023826-2021-04235 | 2021-07-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -12.0/1.5/70, IMPLANTABLE COLLAMER LENS FOUND THE LENS DAMAGE DURING LOADING THE LENS INTO THE CARTRIDGE. THERE WAS NO PATIENT CONTACT. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -12.0/1.5/70 | NA | Physician | 10/06/2021 | 2021 | 49 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12763775&pc=QCB | 2023826-2021-04233 | 2021-07-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.5/2.5/092, IMPLANTABLE COLLAMER LENS FOUND THE LENS DAMAGE DURING LOADING THE LENS INTO THE CARTRIDGE. THERE WAS NO PATIENT CONTACT. IT WAS REPORTED THAT A REPLACEMENT LENS WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -13.5/2.5/092 | NA | Physician | 10/06/2021 | 2021 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12764662&pc=QCB | 2023826-2021-04240 | 2021-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM, TMICL13.7, -13.5/3/85, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. FOREIGN BODY ATTACHED TO POSTERIOR INFERIOR CORNEA WAS OBSERVED. LENS REMAINS IMPLANTED. IT WAS REPORTED THAT MEDICAL OR SURGICAL INTERVENTION IS NOT NECESSARY. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED " SYMPTOMS NORMAL EXCEPT FOR FOREIGN BODY FOUND DURING SLIT LAMP EXAM." CAUSE OF EVENT WAS UNKNOWN. FOR CAUSE DETAILS THE REPORTER STATED " BELIEVED TO BE A FOREIGN BODY SOMEHOW LEFT IN THE EYE DURING SURGERY. AS THIS IS RELATED TO PROCEDURE, UNCERTAIN WHETHER THIS SHOULD BE USER ERROR OR OTHER." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: SEX, WEIGHT, ETHNICITY, RACE: UNK. (B)(4). CLAIM # (B)(4). | -13.5/3/85 | OS | Other Health Care Professional | 10/08/2021 | 2021 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12775932&pc=QCB | 2023826-2021-04277 | 2021-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -15.00/3.00/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(60 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -15.00/3.00/92 | OS | Physician | 11/01/2021 | 2021 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12775935&pc=QCB | 2023826-2021-04274 | 2021-02-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -13.50/+3.00/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -13.50/+3.00/95 | OS | Physician | 11/01/2021 | 2021 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12775936&pc=QCB | 2023826-2021-04275 | 2021-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -12.50/+3.50/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -12.50/+3.50/91 | OD | Physician | 11/01/2021 | 2021 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12775937&pc=QCB | 2023826-2021-04276 | 2021-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -12.00/+3.00/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -12.00/+3.00/91 | OS | Physician | 11/01/2021 | 2021 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12783571&pc=QCB | 2023826-2021-04283 | 2021-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -13.0/1.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -13.0/1.0/092 | OD | Physician | 11/04/2021 | 2021 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12783572&pc=QCB | 2023826-2021-04282 | 2021-05-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -5.0/3.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -5.0/3.0/091 | OS | Physician | 11/03/2021 | 2021 | 44 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12783573&pc=QCB | 2023826-2021-04281 | 2021-05-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -5.0/2.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -5.0/2.5/090 | OD | Physician | 11/03/2021 | 2021 | 44 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12783781&pc=QCB | 2023826-2021-04215 | 2021-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/025 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -09.50/+1.0/025 | OD | Physician | 10/13/2021 | 2021 | 44 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12790674&pc=QCB | 2023826-2021-04184 | 2021-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -11.00/1.50/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.00/1.50/092 | OS | Physician | 10/19/2021 | 2021 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12790678&pc=QCB | 2023826-2021-04189 | 2020-12-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.00/1.5/091 | OS | Physician | 10/21/2021 | 2021 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12790679&pc=QCB | 2023826-2021-04187 | 2021-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -12.00/2.0/096 (SPHERE/ CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -12.00/2.0/096 | OD | Physician | 10/19/2021 | 2021 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12790680&pc=QCB | 2023826-2021-04183 | 2021-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -8.50/1.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -8.50/1.5/080 | OS | Physician | 10/19/2021 | 2021 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12790681&pc=QCB | 2023826-2021-04192 | 2021-10-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -11.00/1.0/091 | OD | Physician | 10/22/2021 | 2021 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12790682&pc=QCB | 2023826-2021-04117 | 2021-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -10.50/2.50/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.50/2.50/092 | OD | Physician | 10/19/2021 | 2021 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12790683&pc=QCB | 2023826-2021-04185 | 2021-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -9.00/2.00/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.00/2.00/092 | OD | Physician | 10/19/2021 | 2021 | 43 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12790691&pc=QCB | 2023826-2021-04200 | 2021-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.00/2.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITION/CORRECTION INFORMATION H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -14.00/2.5/090 | OS | Physician | 10/25/2021 | 2021 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12790693&pc=QCB | 2023826-2021-04199 | 2021-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.5/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: ADDITION/CORRECTION INFORMATION: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE(S) WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -13.5/2.5/092 | OD | Physician | 10/25/2021 | 2021 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12790694&pc=QCB | 2023826-2021-04197 | 2021-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.00/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/2.0/088 | OS | Physician | 10/23/2021 | 2021 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12791231&pc=QCB | 2023826-2021-04284 | 2021-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -14.00/+1.0/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM # (B)(4). | -14.00/+1.0/092 | OD | Other Health Care Professional | 10/19/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12791923&pc=QCB | 2023826-2021-04320 | 2021-05-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event; Insufficient Information | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -13.0/+2.50/89, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) AND REMOVED ON (B)(6) 2021, REASON FOR REMOVAL IS UNKNOWN. IT WAS REPORTED THAT THERE WAS NO INJURY TO THE PATIENT. ON (B)(6) 2021 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -13.0/+2.50/89, IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENT RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REMOVED INTRAOPERATIVELY WITH NO INJURY TO THE PATIENT. ON (B)(6) 2021, A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -13.0/+2.50/89 | OD | Physician | 10/11/2021 | 2021 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12800657&pc=QCB | 2023826-2021-04297 | 2021-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -9.00/+1.50/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND LATER ON THE SAME DATE REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM AND IT WAS NOTED "CORNEAL AND MAIN INCISION NOT 100% RECOVER, BUT THE PATIENT IS VERY HAPPY WITH THE NEW LENS". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED "WTW AND ACD CALCULATE TO 13.2 SIZE BUT IT HAPPEN HIGH VAULT AFTER IMPLANTATION". Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -9.00/+1.50/090 | OD | Physician | 11/05/2021 | 2021 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12800720&pc=QCB | 2023826-2021-04324 | 2021-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -11.5/1.0/028 (SPHERE/ CYLINDER/ AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: REPLACEMENT LENS WAS IMPLANTED VERTICALLY. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM #: (B)(4). | -11.5/1.0/028 | OD | Physician | 11/06/2021 | 2021 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12800757&pc=QCB | 2023826-2021-04323 | 2021-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -7.5/0.5/96 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -7.5/0.5/96 | OD | Physician | 11/06/2021 | 2021 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12801638&pc=QCB | 2023826-2021-04327 | 2021-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-10.00/+2.0/159, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS THE DEVICE, "SIZING" WAS REPORTED FOR CAUSE DETAILS. Manufacturer Narrative: WEIGHT - RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -10.00/+2.0/159 | OD | Physician | 10/12/2021 | 2021 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12806812&pc=QCB | 2023826-2021-04329 | 2021-10-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -18.0/3.5/101 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/3.5/101 | OD | Physician | 11/06/2021 | 2021 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12807718&pc=QCB | 2023826-2021-04272 | 2021-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.5/+3.0/090 (SPHERE/ CYLINDER/ AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -14.5/+3.0/090 | OD | Physician | 10/22/2021 | 2021 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12807722&pc=QCB | 2023826-2021-04264 | 2021-10-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.5/+1.0/074 (SPHERE /CYLINDER/ AXIS) TORE DURING LOADING. THIS WAS DISCOVERED AFTER OPENING THE VIAL ON (B)(6) 2021. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: REPLACEMENT LENS WAS IMPLANTED AT A LATER DATE. THE PROBLEM IS RESOLVED. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -15.5/+1.0/074 | NA | Physician | 10/20/2021 | 2021 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12807725&pc=QCB | 2023826-2021-04261 | 2021-09-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.0/+2.0/081 (SPHERE/ CYLINDER/ AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LOW VAULT IS REPORTED. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE: "LENS TOO SMALL?". Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -8.0/+2.0/081 | OS | Physician | 10/18/2021 | 2021 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12807762&pc=QCB | 2023826-2021-04267 | 2021-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.0/+1.5/090 (SPHERE/ CYLINDER/ AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT MODEL LENS DUE TO LENS ROTATION. Manufacturer Narrative: H6 - " INVESTIGATION TYPE 4110:LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT" IS NOT APPLICABLE IN INITIAL MDR CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH DEBRIS ON LENS. VISUAL INSPECTION FOUND HAPTIC TORN AND DEBRIS THROUGHOUT THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.0/+1.5/090 | OD | Physician | 10/21/2021 | 2021 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12807793&pc=QCB | 2023826-2021-04268 | 2019-06-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.0/+2.5/088 (SPHERE/ CYLINDER/ AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2019 AND (B)(6) 2020 THE LENS WAS REPOSITIONED. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT MODEL LENS DUE TO LENS ROTATION AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6: " INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT" IS NOT APPLICABLE IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND THICK RESIDUE THROUGH THE LENS HAPTIC. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.0/+2.5/088 | OS | Physician | 10/21/2021 | 2021 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12807818&pc=QCB | 2023826-2021-04269 | 2021-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -11.5/+1.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZED DIFFERENT MODEL LENS DUE TO LENS ROTATION AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE "INITIAL TORIC LENS GOT ROTATED TWICE AND SUSPECTED ZONULAR WEAKNESS. SINCE THE PATIENT WAS UNHAPPY WITH IT, EXCHANGED THE SAME WITH SPHERICAL LENS LEAVING CYLINDER AS RESIDUAL." Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | -11.5/+1.5/089 | OD | Physician | 10/22/2021 | 2021 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12807860&pc=QCB | 2023826-2021-04257 | 2021-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -8.5/+3.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT MODEL LENS DUE TO LENS ROTATION. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS OTHER: "LENS ROTATED AND DR. DETERMINED ROTATING LENS WOULDN'T CORRECT RESIDUAL ASTIGMATISM." Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -8.5/+3.0/088 | OD | Physician | 10/18/2021 | 2021 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12807966&pc=QCB | 2023826-2021-04271 | 2021-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.5/+3.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: CORRECTED TO:" THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.5/+3.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS OTHER". CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -15.5/+3.0/089 | OD | Physician | 10/22/2021 | 2021 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12808185&pc=QCB | 2023826-2021-04265 | 2021-10-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -1.5/+2.5/091 (SPHERE/CYLINDER/AXIS) TORE DURING INJECTION INTO THE LEFT (OS) EYE. THERE WAS PATIENT CONTACT WITH THE LENS BUT NO PATIENT INJURY. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY ON (B)(6) 2021. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -1.5/+2.5/091 | OS | Physician | 10/20/2021 | 2021 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12808830&pc=QCB | 2023826-2021-04334 | 2021-10-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.0/1.5/090, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS THAT WAS IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PT INFO:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.0/1.5/090 | OD | Physician | 10/15/2021 | 2021 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12814415&pc=QCB | 2023826-2021-04301 | 2021-09-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -9.00/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REPLACED INTRAOPERATIVELY WITH A LONGER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S RIGHT EYE (OD) DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.00/1.5/090 | OD | Physician | 10/27/2021 | 2021 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12814418&pc=QCB | 2023826-2021-04300 | 2021-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/1.5/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/1.5/111 | OD | Physician | 10/26/2021 | 2021 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12814421&pc=QCB | 2023826-2021-04302 | 2021-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.50/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -11.50/2.0/090 | OS | Physician | 10/27/2021 | 2021 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12814422&pc=QCB | 2023826-2021-04307 | 2021-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -11.00/2.0/091 (SPHERE/ CYLINDER/ AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS IMPLANTED VERTICALLY DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -11.00/2.0/091 | OS | Physician | 10/27/2021 | 2021 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12814423&pc=QCB | 2023826-2021-04309 | 2021-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -2.50/1.00/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS IMPLANTED VERTICALLY DUE TO LOW VAULT ON 17/SEP/2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -2.50/1.00/093 | OS | Physician | 10/27/2021 | 2021 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12814424&pc=QCB | 2023826-2021-04308 | 2021-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -6.00/1.50/090 (SPHERE/ CYLINDER/ AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS IMPLANTED VERTICALLY DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -6.00/1.50/090 | OD | Physician | 10/27/2021 | 2021 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12814426&pc=QCB | 2023826-2021-04304 | 2021-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.50/3.0/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -11.50/3.0/111 | OD | Physician | 10/27/2021 | 2021 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12814428&pc=QCB | 2023826-2021-04306 | 2021-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.00/2.5/000 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.00/2.5/000 | OD | Physician | 10/27/2021 | 2021 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12814429&pc=QCB | 2023826-2021-04310 | 2021-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/2.0/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT : UNK. .ETHNICITY: UNK. RACE: UNK. DATE REC¿D BY MFR: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.50/2.0/071 | OD | Physician | 10/28/2021 | 2021 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12814430&pc=QCB | 2023826-2021-04314 | 2021-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/2.0/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). Manufacturer Narrative: B5 - REPORTED AS REPLACEMENT LENS BEING IMPLANTED VERTICALLY AND RESOLVED PROBLEM - CORRECTION, REPLACEMENT LEN IMPLANTED AND RESOLVED PROBLEM. CLAIM# (B)(4). | -12.50/2.0/112 | OS | Physician | 10/28/2021 | 2021 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12814431&pc=QCB | 2023826-2021-04316 | 2021-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/2.0/147 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS IMPLANTED DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -11.00/2.0/147 | OD | Physician | 10/29/2021 | 2021 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12814432&pc=QCB | 2023826-2021-04315 | 2021-08-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -7.00/1.5/089 (SPHERE/ CYLINDER/ AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -7.00/1.5/089 | OS | Physician | 10/29/2021 | 2021 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12814433&pc=QCB | 2023826-2021-04313 | 2021-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.00/1.5/095 | OD | Physician | 10/28/2021 | 2021 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12814435&pc=QCB | 2023826-2021-04317 | 2021-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/1.0/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS IMPLANTED DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -11.00/1.0/099 | OS | Physician | 10/29/2021 | 2021 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12814436&pc=QCB | 2023826-2021-04311 | 2021-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.50/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.50/1.5/092 | OS | Physician | 10/28/2021 | 2021 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12817108&pc=QCB | 2023826-2021-04362 | 2021-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -11.50/+2.00/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.50/+2.00/92 | OS | Physician | 11/10/2021 | 2021 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12817112&pc=QCB | 2023826-2021-04359 | 2021-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -18.0/2.0/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -18.0/2.0/095 | OD | Physician | 11/07/2021 | 2021 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12824603&pc=QCB | 2023826-2021-04402 | 2021-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -13.00/+1.50/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4). | -13.00/+1.50/95 | OS | Physician | 11/10/2021 | 2021 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12831547&pc=QCB | 2023826-2021-04412 | 2021-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.5/1.5/69 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. MEDICAL DEVICE PROBLEM CODE: (B)(4). MEDICAL DEVICE PROBLEM CODE: (B)(4). CLAIM#: (B)(4). | -10.5/1.5/69 | OS | Physician | 11/11/2021 | 2021 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12831572&pc=QCB | 2023826-2021-04405 | 2021-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -5.5/5.0/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.5/5.0/088 | OD | Physician | 11/10/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12831581&pc=QCB | 2023826-2021-04409 | 2021-10-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -16.0/4.5/104 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -16.0/4.5/104 | OD | Physician | 11/11/2021 | 2021 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12831614&pc=QCB | 2023826-2021-04406 | 2021-10-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -8.0/1.0/108 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/1.0/108 | OD | Physician | 11/10/2021 | 2021 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12831622&pc=QCB | 2023826-2021-04407 | 2021-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -8.0/1.0/070 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. MEDICAL DEVICE PROBLEM CODE: (B)(4). CLAIM#: (B)(4). | -8.0/1.0/070 | OS | Physician | 11/10/2021 | 2021 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12831623&pc=QCB | 2023826-2021-04410 | 2021-10-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -13.5/4.5/177 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -13.5/4.5/177 | OD | Physician | 11/11/2021 | 2021 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12833463&pc=QCB | 2023826-2021-04289 | 2021-10-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON WAS PREPARING A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/089 (SPHERE/CYLINDER/AXIS), FOR DELIVERY INTO THE PATIENTS LEFT EYE (OS) AND FOUND AN UNIDENTIFIED OBJECT WHILE FOLDING THE ICL IN THE CARTRIDGE. THE UNIDENTIFIED OBJECT WAS REMOVED FROM THE ICL AND THE ICL WAS IMPLANTED INTO THE EYE. THE ICL LENS REMAINED IMPLANTED AND THE PATIENT IS DOING WELL. THE EYE LOOKS CALM AND OK. Manufacturer Narrative: H3: DEVICE EVALUATION: THE OBJECT WAS RETURNED IN LIQUID, IN A VIAL, LOOSE IN THE LENS BOX. A THIN CIRCULAR OBJECT WAS NOTED FLOATING INSIDE THE VIAL. CLAIM # (B)(4). Manufacturer Narrative: H6: BASED ON THE RESULTS OF THE INVESTIGATION, THE PARTICULATE FOUND ON LENS IS LIKELY TO BE POLYETHYLENE A COMMON PLASTIC MATERIAL. DUE TO THE WIDESPREAD USE OF THIS MATERIAL WE ARE UNABLE TO DETERMINE ANY DEFINITIVE ROOT CAUSE FOR THIS EVENT. OPERATORS INVOLVED IN THE MANUFACTURING OF THIS NONCONFORMING LENS HAVE BEEN RETRAINED IN IDENTIFICATION OF COSMETIC DEFECTS SUCH AS PARTICULATES ON LENS. THERE IS NO INDICATION THAT SUGGESTS A DEFICIENCY IN THE MANUFACTURING PROCESS, FURTHER ACTION IS NOT REQUIRED AT THIS TIME. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.00/+1.5/089 | OS | Physician | 10/18/2021 | 2021 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12833490&pc=QCB | 2023826-2021-04294 | 2021-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+3.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -18.00/+3.5/095 | OS | Physician | 10/19/2021 | 2021 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12833757&pc=QCB | 2023826-2021-04415 | 2021-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1,-8.0/1.0/170, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. LOW VAULT WAS OBSERVED , LENS REMAINS IMPLANTED. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT IS THE DEVICE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER UPDATED INFORMATION ON 11-NOV-2021, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM RESOLVED. CLAIM #(B)(4). Manufacturer Narrative: B5: INITIAL REPORTED -8.00/1.0/170 (SPHERE/CYLINDER/AXIS) SHOULD BE CORRECTED TO -8.00/1.0/179 (SPHERE/CYLINDER/AXIS). H3: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC BROKEN. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT - RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM # (B)(4). | -8.0/1.0/170 | OS | Physician | 10/18/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12833765&pc=QCB | 2023826-2021-04414 | 2021-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1,-7.5/1.0/170, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LOW VAULT WAS OBSERVED , LENS REMAINS IMPLANTED. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT IS THE DEVICE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- INITIAL REPORTED -7.5/1.0/170 (SPHERE/CYLINDER/AXIS) SHOULD BE CORRECTED TO -7.50/1.0/176 (SPHERE/CYLINDER/AXIS). H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND LENS HAPTIC BROKEN. CLAIM # (B)(4). Manufacturer Narrative: B5- PER UPDATED INFORMATION ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.5/1.0/170 | OD | Physician | 10/18/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12833801&pc=QCB | 2023826-2021-04413 |  | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER STATED THAT A 13.2MM, TMICL13.2, -10.50/2.5/086, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. IT WAS REPORTED THAT THE PATIENTS POST OP CONDITION IS GREAT. Manufacturer Narrative: AGE - RACE: UNK. OUTCOMES ATTRIBUTED TO ADVERSE EVENT: UNK. CLAIM # (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -10.50/2.5/086 | OS | Physician | 10/18/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840480&pc=QCB | 2023826-2021-04394 | 2021-05-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -7.0/+0.5/101 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH DIFFERENT AXIS LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISIBLE INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -7.0/+0.5/101 | OD | Physician | 11/01/2021 | 2021 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840485&pc=QCB | 2023826-2021-04386 | 2021-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -8.5/+3.5/091 (SPHERE/ CYLINDER/ AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -8.5/+3.5/091 | OD | Physician | 10/29/2021 | 2021 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840488&pc=QCB | 2023826-2021-04379 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+3.0/111 (SPHERE/CYLINDER/AXIS) TORE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.0/+3.0/111 | NA | Physician | 10/28/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840515&pc=QCB | 2023826-2021-04423 | 2021-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -12.0/1.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -12.0/1.5/091 | OD | Physician | 11/12/2021 | 2021 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840528&pc=QCB | 2023826-2021-04375 | 2021-07-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.0/+2.0/092 (SPHERE/ CYLINDER/ AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. REPORTEDLY, THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.0/+2.0/092 | OS | Physician | 10/26/2021 | 2021 | 17 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840537&pc=QCB | 2023826-2021-04421 | 2021-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -10.5/1.0/109 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -10.5/1.0/109 | OS | Physician | 11/12/2021 | 2021 | 17 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840579&pc=QCB | 2023826-2021-04383 | 2021-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+3.5/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. THE PROBLEM IS RESOLVED. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -10.5/+3.5/082 | OS | Physician | 10/29/2021 | 2021 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840581&pc=QCB | 2023826-2021-04400 | 2021-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: HOLD FOR DD 2.11 THE REPORTER STATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS LOW VAULT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL.Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE SURGEON REPORTS LOW VAULT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-UPDATED INFORMATION: A 12.1MM VTICM5_12.1, -8.0/+1.0/104 (SPHERE/CYLINDER/AXIS) LENS WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH, DIFFERENT MODEL LENS. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT CODE UNK; ESUBMITTER SOFTWARE DOES NOT ALLO FOR AN UNK OR BLANK ENTRY THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT CODE UNK; ESUBMITTER SOFTWARE DOES NOT ALLO FOR AN UNK OR BLANK ENTRY THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -8.0/+1.0/104 | OD | Physician | 11/03/2021 | 2021 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840582&pc=QCB | 2023826-2021-04387 | 2021-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -8.5/+1.0/089 | OS | Physician | 10/29/2021 | 2021 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840583&pc=QCB | 2023826-2021-04374 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -15.5/+1.5/070 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE SURGEON REPORTS EXCESSIVE VAULT. REPORTEDLY, THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -15.5/+1.5/070 | OS | Physician | 10/26/2021 | 2021 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840584&pc=QCB | 2023826-2021-04354 | 2021-09-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.50/+2.0/090 (SPHERE/ CYLINDER/ AXIS), TORE DURING LOADING INTO THE CARTRIDGE AND THERE WAS NO PATIENT CONTACT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -08.50/+2.0/090 | NA | Physician | 10/21/2021 | 2021 | 44 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840590&pc=QCB | 2023826-2021-04396 | 2021-10-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON INJECTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/071 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2021. THERE WAS PATIENT CONTACT WITH THE LENS BUT THE LENS WAS NOT FULLY IMPLANTED. NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.5/+2.0/071 | OD | Physician | 11/01/2021 | 2021 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840591&pc=QCB | 2023826-2021-04371 | 2021-09-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/066 (SPHERE/ CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH DIFFERENT AXIS LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. (B)(4). | -12.5/+1.5/066 | OS | Physician | 10/23/2021 | 2021 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840633&pc=QCB | 2023826-2021-04395 | 2021-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+1.0/015 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "THE PATIENT WAS NERVOUS AND REFUSED TO TAKE THE SURGERY AGAIN." Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISIBLE INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: PT INFO-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.0/+1.0/015 | OS | Physician | 11/01/2021 | 2021 | 48 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840650&pc=QCB | 2023826-2021-04376 | 2021-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -7.5/+1.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT AXIS LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -7.5/+1.5/094 | OS | Physician | 10/27/2021 | 2021 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840654&pc=QCB | 2023826-2021-04377 | 2021-08-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRA-OPERATIVELY ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. REPORTEDLY, THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. SEE MFR REP# 2023826-2021-04378 FOR REPLACEMENT LENS.Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRA-OPERATIVELY ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. REPORTEDLY, THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. SEE MFR REP# 2023826-2021-04378 FOR REPLACEMENT LENS. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -13.5/+2.0/092 | OS | Physician | 10/27/2021 | 2021 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840918&pc=QCB | 2023826-2021-04416 | 2021-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-7.5/1.5/93, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -7.5/1.5/93 | OS | Physician | 11/15/2021 | 2021 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12840919&pc=QCB | 2023826-2021-04418 | 2021-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-6.5/4.0/81, (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.5/4.0/81 | OS | Physician | 11/15/2021 | 2021 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12841577&pc=QCB | 2023826-2021-04425 | 2021-09-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2021 A 13.2MM, VTICMO13.2, -17.50/1.0/168 IMPLANTABLE COLLAMER LENS TORE DURING LOADING IN THE CARTRIDGE, THERE WAS NO PATIENT CONTACT. ON (B)(6) 2021 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED TO BE THE DEVICE. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -17.50/1.0/168 | NA | Physician | 10/19/2021 | 2021 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12841583&pc=QCB | 2023826-2021-04424 | 2021-09-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -3.50/2.5/025 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. EXCESSIVE VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. IT WAS REPORTED THAT MEDICAL OR SURGICAL INTERVENTION WAS NOT NECESSARY AND THAT THE PROBLEM RESOLVED. THE REPORTER STATED THAT THE LENS WAS NOT REPLACED BECAUSE THE VAULT WAS REDUCING GRADUALLY, SO THE DOCTOR DIDNT WANT TO REPLACE THE LENS. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -3.50/2.5/025 | OS | Physician | 10/19/2021 | 2021 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12849374&pc=QCB | 2023826-2021-04451 | 2021-06-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.0/1.5/004 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULTING. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -9.0/1.5/004 | OD | Physician | 11/11/2021 | 2021 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12849379&pc=QCB | 2023826-2021-04452 | 2021-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -13.5/2.5/101 (SPHERE/CYLINDER/AXIS) IMPLANTABLE LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING AND ON (B)(6) 2021 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. THEN ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED "BIG STS". Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. CLAIM #(B)(4). | -13.5/2.5/101 | OD | Physician | 11/17/2021 | 2021 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12849409&pc=QCB | 2023826-2021-04453 | 2021-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -14/03/81 (SPHERE/CYLINDER/AXIS) IMPLANTABLE LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING AND ON (B)(6) 2021 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. THEN ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND NOTED "BIG STS". Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -14/03/81 | OS | Physician | 11/17/2021 | 2021 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12849502&pc=QCB | 2023826-2021-04441 | 2021-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -10.50/+2.0/79, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.50/+2.0/79 | OS | Physician | 10/20/2021 | 2021 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12849519&pc=QCB | 2023826-2021-04440 | 2021-09-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2021 A 12.6MM, VTICMO12.6, -12.5/3.5/70, IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS IMPLANTED, REMOVED AND REPLACED WITHIN THE SAME SURGERY AND THE PROBLEM RESOLVED. THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.5/3.5/70 | NA | Physician | 10/20/2021 | 2021 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12849826&pc=QCB | 2023826-2021-04445 | 2021-10-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: MISALIGNMENT OF THE AXIS AFTER IMPLANTABLE COLLAMER LENS WAS REPORTED, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION A 13.2MM VTICM5_13.2 -12.5/1.50/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6)2021. ON (B)(6)2021 LENS REPOSITIONING WAS PERFORMED DUE TO LOW VAULT NO ASSOCIATED TO LOW VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. THIS PRODUCT IS NOT MARKETED IN THE US. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM # (B)(4). | -12.5/1.50/92 | OS | Physician | 10/21/2021 | 2021 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12850284&pc=QCB | 2023826-2021-04446 | 2021-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -10.50/1.0/093, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THIS EXPLANT RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/1.0/093 | OD | Physician | 10/22/2021 | 2021 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12850303&pc=QCB | 2023826-2021-04448 | 2021-08-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2021 A 12.6MM, VTICMO12.6, -8.5/1.5/89, IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS IMPLANTED AND REMOVED WITHIN THE SAME SURGERY. THERE WAS NO PATIENT INJURY. ON (B)(6) 2021 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4) | -8.5/1.5/89 | NA | Physician | 10/22/2021 | 2021 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12852320&pc=QCB | 2023826-2021-04337 | 2021-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.50/1.5/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -15.50/1.5/103 | OD | Physician | 10/31/2021 | 2021 | 43 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12852324&pc=QCB | 2023826-2021-04338 | 2021-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.50/1.5/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -15.50/1.5/073 | OS | Physician | 10/31/2021 | 2021 | 43 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12852325&pc=QCB | 2023826-2021-04343 | 2021-10-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -5.0/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON 09/OCT/2021 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT.THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -5.0/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT.THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENTS INFORMATION - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: PATIENTS INFORMATION - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.0/1.0/087 | OS | Physician | 11/01/2021 | 2021 | 47 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12852327&pc=QCB | 2023826-2021-04342 | 2021-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED VERTICALLY A 13.2MM VTICMO13.2, -6.50/0.5/035 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -6.50/0.5/035 | OS | Physician | 10/30/2021 | 2021 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12852329&pc=QCB | 2023826-2021-04340 | 2021-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -17.00/1.0/116 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -17.00/1.0/116 | OD | Physician | 10/31/2021 | 2021 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12852330&pc=QCB | 2023826-2021-04346 | 2021-08-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/3.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -10.50/3.0/091 | OS | Physician | 11/05/2021 | 2021 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12852331&pc=QCB | 2023826-2021-04347 | 2021-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.0/3.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.0/3.5/097 | OD | Physician | 11/05/2021 | 2021 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12852333&pc=QCB | 2023826-2021-04336 | 2021-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.50/1.0/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -14.50/1.0/071 | OD | Physician | 10/31/2021 | 2021 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12852335&pc=QCB | 2023826-2021-04344 | 2021-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED VERTICALLY A 12.1MM VTICMO12.1, -10.50/1.0/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -10.50/1.0/073 | OS | Physician | 11/01/2021 | 2021 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12857856&pc=QCB | 2023826-2021-04505 | 2021-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -11.5/2.0/87 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. EXCESSIVE VAULTING WAS REPORTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. LENS REMAINS IMPLANTED. Manufacturer Narrative: B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR . B5 - "THE LENS WAS REPLACED WITH A SHORTER LENGTH/ SAME MODEL LENS ON (B)(6) 2022. THIS RESOLVED THE PROBLEM." SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - CORRECTED TO HEALTH EFFECT - IMPACT CODE: 4629 . CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -11.5/2.0/87 | OD | Physician | 11/16/2021 | 2021 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12857861&pc=QCB | 2023826-2021-04432 | 2021-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -9.5/2.0/082 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE LENS WAS PLACED VERTICALLY AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.5/2.0/082 | OS | Physician | 11/14/2021 | 2021 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12857889&pc=QCB | 2023826-2021-04506 | 2021-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -12.0/2.0/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. EXCESSIVE VAULTING WAS REPORTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: B1: CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2: CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR B5: "THE LENS WAS REPLACED WITH A SHORTER LENGTH/ SAME MODEL LENS ON 15/03/2022. THIS RESOLVED THE PROBLEM." SHOULD BE ADDED TO THE INITIAL MDR. D6B: EXPLANT DATE (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. H1: CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - CORRECTED TO HEALTH EFFECT - IMPACT CODE: 4629. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED TO CLAIM # (B)(4). CLAIM# (B)(4). | -12.0/2.0/92 | OS | Physician | 11/16/2021 | 2021 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12857969&pc=QCB | 2023826-2021-04426 | 2021-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-11.5/1.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND SIMILAR POWER DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B2- DISABILITY OR PERMANENT DAMAGE SELECTED IN INITIAL MDR IS NOT APPLICABLE. CLAIM# (B)(4). | -11.5/1.5/086 | OS | Physician | 11/15/2021 | 2021 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12858034&pc=QCB | 2023826-2021-04427 | 2021-08-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -14.0/2.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH AND SIMILAR POWER LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.0/2.0/090 | OS | Physician | 11/15/2021 | 2021 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12858503&pc=QCB | 2023826-2021-04485 | 2021-10-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50/+2.0/091 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THERE WAS NO PATIENT INJURY. ANOTHER ICL LENS WAS IMPLANTED AND THE PATIENT IS VERY HAPPY. THE PROBLEM WAS RESOLVED. CLAIM # 7(B)(4). | -08.50/+2.0/091 | OD | Physician | 10/25/2021 | 2021 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12858506&pc=QCB | 2023826-2021-04486 | 2020-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.00/+3.0/106 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR/DEVICE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -12.00/+3.0/106 | OD | Physician | 10/25/2021 | 2021 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12865977&pc=QCB | 2023826-2021-04508 | 2021-10-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -9.5/1.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND LATER THE SAME DATE REPLACED WITH A LONGER LENGTH LENS; DUE TO LENS DISLOCATION/SUBLUXATION AND NOTED THAT THE LENS HAD ROTATED 90 DEGREES. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -9.5/1.5/091 | OS | Other Health Care Professional | 11/18/2021 | 2021 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12871955&pc=QCB | 2023826-2021-04510 | 2021-10-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2021 A 12.6MM, VTICMO12.6, -18/06/56, IMPLANTABLE COLLAMER LENS, TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE / LENS STUCK IN CARTRIDGE OR INJECTION SYSTEM. THERE WAS NO PATIENT INJURY. IN THE REPORTERS OPINION THE DEVICE WAS THE CAUSE OF THIS EVENT, " BECAUSE OF CARTRIDGE" WAS REPORTED FOR CAUSE DETAILS. "NO SYMPTOMS" REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18/06/56 | NA | Physician | 10/25/2021 | 2021 | 46 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12872816&pc=QCB | 2023826-2021-04495 | 2021-10-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS RIGHT EYE (OD). THE LENS WAS REPORTED AS HAVING ZERO VAULTING AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: EXPIRATION DATE UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/143 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS REPOSITIONED ON 21-OCT-2021 BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS EXPLANTED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. CORRECTED DATA: D2: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB. E1: (B)(6). CLAIM # (B)(4). | -10.50/+1.0/143 | OD | Physician | 10/28/2021 | 2021 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12872825&pc=QCB | 2023826-2021-04504 | 2021-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.00/+3.5/098 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED SO THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5 - CORRECTED TO : THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.00/+3.5/098 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2018. THE LENS WAS EXCHANGED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED SO THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.00/+3.5/098 | OS | Physician | 10/29/2021 | 2021 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12872827&pc=QCB | 2023826-2021-04503 | 2021-10-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/077 (SPHERE/ CYLINDER/ AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS WAS TORN UPON DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/ LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.50/+1.0/077 | OD | Physician | 10/29/2021 | 2021 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12872829&pc=QCB | 2023826-2021-04496 | 2021-10-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS LEFT EYE (OS). THE LENS WAS REPORTED AS HAVING ZERO VAULTING AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.00/+2.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS EXPLANTED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. D2: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB. E1: (B)(6). CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -13.00/+2.0/088 | OS | Physician | 10/28/2021 | 2021 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12879718&pc=QCB | 2023826-2021-04523 | 2021-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -13.50/1.00/92, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS THAT WAS IMPLANTED VERTICALLY DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS EXCHANGED RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH DEBRIS ON PRODUCT, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A HAPTIC TORN. DEBRIS THROUGHOUT THE LENS WAS OBSERVED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | -13.50/1.00/92 | OD | Physician | 10/27/2021 | 2021 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12879719&pc=QCB | 2023826-2021-04525 | 2021-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2021 A 12.6MM, VTICMO_12.6, -13.0/1.5/071, IMPLANTABLE COLLAMER LENS TORE/ BROKE DURING INJECTION/ DELIVERY INTO THE EYE. THE LENS WAS LATER REMOVED. IT WAS REPORTED THAT THERE WAS NO PATIENT INJURY. ON (B)(6) 2021 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.0/1.5/071 | NA | Physician | 10/27/2021 | 2021 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12879843&pc=QCB | 2023826-2021-04528 | 2021-10-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -7.00/1.0/97, IMPLANTABLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED. THERE WAS NO PATIENT INJURY. IN THE REPORTER OPINION THE CAUSE OF THIS EVENT WAS USER ERROR. IT WAS REPORTER THAT IT IS UNKNOWN IF THE DEVICE FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: (B)(6). THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.00/1.0/97 | OD | Physician | 10/28/2021 | 2021 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12880226&pc=QCB | 2023826-2021-04529 | 2021-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2021 A 12.6MM, VTICMO12.6, -13.5/1.5/092, IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. ON (B)(6) 2021 THE LENS WAS REMOVED AND REPLACED AND THE PROBLEM RESOLVED. THERE WAS NO PATIENT INJURY. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.5/1.5/092 | NA | Physician | 10/29/2021 | 2021 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12880249&pc=QCB | 2023826-2021-04530 | 2021-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1,-13.5/1.5/089, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: B5- PER UPDATED INFORMATION THE REPORTER STATED "THE PATIENT WAS IMPLANTED WITH A HORIZONTAL LENS OF 12.1MM WITH A HIGH ARCH HEIGHT. SINCE SMALLER LENSES WERE NOT AVAILABLE, THE DOCTOR DECIDED TO IMPLANT A VERTICAL LENS OF 12.1MM TO OBTAIN THE IDEAL ARCH HEIGHT." H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). WEIGHT - RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM # (B)(4). | -13.5/1.5/089 | OD | Physician | 10/29/2021 | 2021 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12880265&pc=QCB | 2023826-2021-04531 | 2021-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-12.0/2.0/91, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). WEIGHT - RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM # (B)(4). Manufacturer Narrative: B5- PER UPDATED INFORMATION THE REPORTER STATED "THE PATIENT WAS IMPLANTED WITH 12.6MM LENS HORIZONTALLY, BUT THE ARCH HEIGHT WAS HIGH. BECAUSE THE PATIENT ORDERED BOTH HORIZONTAL AND VERTICAL LENSES, AFTER COMMUNICATION WITH THE PATIENT, THE DOCTOR DECIDED TO IMPLANT 12.6MM LENS VERTICALLY TO OBTAIN A RELATIVELY LOW ARCH HEIGHT." H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | -12.0/2.0/91 | OD | Physician | 10/29/2021 | 2021 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12880304&pc=QCB | 2023826-2021-04536 | 2021-10-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2,-13.0/+1.0/003, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A DIFFERENT POWER LENS DUE TO "SURGEON HAS CHOSEN WRONG LENS POWER". THIS EXCHANGE RESOLVED THE PROBLEM. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS "USER ERROR". Manufacturer Narrative: WEIGHT - RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US CLAIM # (B)(4). | -13.0/+1.0/003 | OD | Physician | 10/29/2021 | 2021 | 47 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12880350&pc=QCB | 2023826-2021-04534 | 2021-07-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6,-8.50/1.0/099, IMPLANTABLE COLLAMER LENS TORE/BROKE DURING LOADING. THE REPORTER STATED "THE LENS GOT CAUGHT IN THE CARTRIDGE AND WAS DAMAGED WHILE LOADING THE LENS INTO THE CARTRIDGE." THE LENS DID NOT TOUCH THE PATIENT. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: WEIGHT - RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -8.50/1.0/099 | NA | Physician | 10/29/2021 | 2021 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12880380&pc=QCB | 2023826-2021-04535 | 2021-07-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6,-7.50/1.5/091, IMPLANTABLE COLLAMER LENS TORE/BROKE DURING LOADING. THE REPORTER STATED "THE LENS GOT CAUGHT IN THE CARTRIDGE AND WAS DAMAGED WHILE LOADING THE LENS INTO THE CARTRIDGE." THE LENS DID NOT TOUCH THE PATIENT. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINT WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | -7.50/1.5/091 | NA | Physician | 10/29/2021 | 2021 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12880522&pc=QCB | 2023826-2021-04532 | 2021-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2,-11.0/1.5/098, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT'S WEIGHT, ETHNICITY AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DEBRIS/RESIDUE WAS OBSERVED THROUGHOUT THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | -11.0/1.5/098 | OS | Physician | 10/29/2021 | 2021 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12883935&pc=QCB | 2023826-2021-04435 | 2021-10-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -10.0/2.5/87 (SPHERE/ CYLINDER/ AXIS) LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND SIMILAR POWER. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS PATIENT RELATED FACTOR; SHORT ACD BUT LONG WTW. Manufacturer Narrative: ADDITIONAL INFORMATION: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/2.5/87 | OD | Physician | 11/17/2021 | 2021 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12883937&pc=QCB | 2023826-2021-04436 | 2021-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6,-18.0/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND SIMILAR POWER DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -18.0/2.5/093 | OS | Physician | 11/18/2021 | 2021 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12884051&pc=QCB | 2023826-2021-04434 | 2021-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2,-12.0/2.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND SIMILAR POWER DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -12.0/2.5/081 | OS | Physician | 11/16/2021 | 2021 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12884998&pc=QCB | 2023826-2021-04459 | 2021-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -6.00/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -6.00/1.5/089 | OS | Physician | 11/08/2021 | 2021 | 44 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12885020&pc=QCB | 2023826-2021-04454 | 2021-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED HORIZONTALLY A 13.2MM VTICMO13.2, -11.00/2.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -11.00/2.5/094 | OD | Physician | 11/08/2021 | 2021 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12885024&pc=QCB | 2023826-2021-04460 | 2021-08-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -15.0/2.5/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.0/2.5/108 | OD | Physician | 11/08/2021 | 2021 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12885027&pc=QCB | 2023826-2021-04467 | 2021-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT - UNK. PATIENT ETHNICITY - UNK. PATIENT RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.00/2.0/093 | OD | Physician | 11/10/2021 | 2021 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12885028&pc=QCB | 2023826-2021-04468 | 2021-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.00/1.5/063 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.00/1.5/063 | OS | Physician | 11/11/2021 | 2021 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12885029&pc=QCB | 2023826-2021-04473 | 2021-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.50/2.0/095 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -8.50/2.0/095 | OS | Physician | 11/11/2021 | 2021 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12885031&pc=QCB | 2023826-2021-04469 | 2021-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.00/2.0/116 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.00/2.0/116 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#: (B)(4). | -13.00/2.0/116 | OD | Physician | 11/11/2021 | 2021 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12885040&pc=QCB | 2023826-2021-04464 | 2021-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.50/2.5/012 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.50/2.5/012 | OS | Physician | 11/10/2021 | 2021 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12885042&pc=QCB | 2023826-2021-04479 | 2021-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.00/1.5/091 (SPHERE/ CYLINDER/ AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTION: THIS MDR (B)(4) IS NOT APPLICABLE AS IT IS A DUPLICATE OF MFR#: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/1.5/091 | OD | Physician | 11/12/2021 | 2021 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12885043&pc=QCB | 2023826-2021-04455 | 2021-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -14.00/1.5/173 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -14.00/1.5/173 | OD | Physician | 11/08/2021 | 2021 | 43 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12885050&pc=QCB | 2023826-2021-04458 | 2021-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/1.0/113 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.50/1.0/113 | OD | Physician | 11/08/2021 | 2021 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12885053&pc=QCB | 2023826-2021-04461 | 2021-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.00/3.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -7.00/3.5/094 | OS | Physician | 11/10/2021 | 2021 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12885056&pc=QCB | 2023826-2021-04462 | 2021-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.50/2.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US CLAIM# (B)(4). | -12.50/2.5/090 | OS | Physician | 11/10/2021 | 2021 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12885064&pc=QCB | 2023826-2021-04466 | 2021-05-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -12.0/2.5/094 (SPHERE/ CYLINDER/ AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -12.0/2.5/094 | OD | Physician | 11/10/2021 | 2021 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12885069&pc=QCB | 2023826-2021-04477 | 2021-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.00/1.0/167 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -14.00/1.0/167 | OS | Physician | 11/10/2021 | 2021 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12885074&pc=QCB | 2023826-2021-04476 | 2021-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.00/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.00/2.0/093 | OS | Physician | 11/10/2021 | 2021 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12888635&pc=QCB | 2023826-2021-04564 | 2021-11-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT WHILE ATTEMPTING TO IMPLANT A 12.6MM VTICM5_12.6; -8.5/1.5/110 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) IT HAD TORN DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT WITH NO PATIENT INJURY. LENS WAS NOT IMPLANTED. ON THE SAME DATE IN A SEPARATE SURGERY A REPLACEMENT LENS OF THE SAME MODEL AND LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.5/1.5/110 | OS | Physician | 11/17/2021 | 2021 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12888641&pc=QCB | 2023826-2021-04563 | 2021-10-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -12.5/1.5/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS IMPLANTED UPSIDE DOWN AND INTRAOPERATIVELY THE LENS WAS REMOVED WITH NO PATIENT INJURY. ON (B)(6) 2021 A REPLACEMENT LENS OF THE SAME MODEL/LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -12.5/1.5/087 | OD | Physician | 11/19/2021 | 2021 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12888643&pc=QCB | 2023826-2021-04614 | 2021-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.0/1.5/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - METHOD CODE 3331, DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -11.0/1.5/086 | OS | Physician | 11/25/2021 | 2021 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12888656&pc=QCB | 2023826-2021-04613 | 2021-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.0/1.5/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -11.0/1.5/92 | OD | Physician | 11/25/2021 | 2021 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12910541&pc=QCB | 2023826-2021-04439 | 2021-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -12.0/3.0/88 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.0/3.0/88 | OS | Physician | 11/18/2021 | 2021 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12910769&pc=QCB | 2023826-2021-04538 | 2021-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -13.0/3.5/109 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: D4: VTICMO12.1. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.0/3.5/109 | OD | Physician | 11/18/2021 | 2021 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12911036&pc=QCB | 2023826-2021-04539 | 2021-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -13.0/3.5/70 (SPHERE/ CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.0/2.5/094 | OS | Physician | 11/18/2021 | 2021 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12911286&pc=QCB | 2023826-2021-04621 | 2021-10-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON IMPLANTED 12.6MM VTICMO12.6; -15.0/2.0/072 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021 IT HAD TORE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SAME MODEL AND LENGTH LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.0/2.0/072 | OD | Physician | 11/29/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12911289&pc=QCB | 2023826-2021-04622 | 2021-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.5/3.5/94 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.5/3.5/94 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/3.5/94 | OD | Physician | 11/29/2021 | 2021 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12914036&pc=QCB | 2023826-2021-04624 | 2021-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -13.50/+3.00/86 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULTING. IT WAS REPORTED THERE IS NO PLAN TO IMPLANT ANOTHER LENS AND THE PROBLEM IS RESOLVED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -13.50/+3.00/86 | OD | Physician | 12/01/2021 | 2021 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12914050&pc=QCB | 2023826-2021-04581 | 2021-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -15.50/+3.5/139 (SPHERE/ CYLINDER/ AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING AND PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE INCISION WAS ENLARGED AND SUTURES WERE REQUIRED. THE CAUSE OF THE EVENT WAS DUE TO PATIENT RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTIVE VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM#: (B)(4). | -15.50/+3.5/139 | OD | Physician | 11/03/2021 | 2021 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12914055&pc=QCB | 2023826-2021-04583 | 2021-10-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -10.50/+2.5/094 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS WAS TORN DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH A SHORTER LENS ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -10.50/+2.5/094 | OS | Physician | 11/04/2021 | 2021 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12920200&pc=QCB | 2023826-2021-04628 | 2020-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1,-9.5/+1.5/068, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS EXPLANTED AND LATER REPLACED ON THE SAME DAY WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. IT WAS REPORTED THAT THE PROBLEM RESOLVED. " NO ROTATION OF THE TORIC ICL, BUT STILL NO VAULT" WAS REPORTED FOR STATUS OF THE EYE (SIGNS/SYMPTOMS). CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WERE FOUND. CLAIM # (B)(4). | -9.5/+1.5/068 | OS | Physician | 11/02/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12920238&pc=QCB | 2023826-2021-04629 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6,-9.5/+1.5/077, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020 AS A REPLACEMENT LENS TO CORRECT LOW VAULT WITH ROTATION. IT WAS REPORTED THAT THE ROTATION RESOLVED BUT LOW VAULT IS STILL NOTED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINTS WERE FOUND. CLAIM # (B)(4). | -9.5/+1.5/077 | OS | Physician | 11/02/2021 | 2021 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12920471&pc=QCB | 2023826-2021-04620 | 2021-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -8.5/+1.5/011 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT AND LENS TEAR/BREAK DURING LOADING. THE VAULT VALUE PRE-EXCHANGE IS REPORTED AS 1020UM; POST-EXCHANGE REPORTED VALUE IS 600UM. Manufacturer Narrative: B5-THE REPORTER STATES THE PROBLEM WAS RESOLVED AFTER EXCHANGE. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -8.5/+1.5/011 | OD | Physician | 11/04/2021 | 2021 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12921180&pc=QCB | 2023826-2021-04540 | 2021-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.0/1.5/110 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -11.0/1.5/110 | OD | Physician | 11/18/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12921485&pc=QCB | 2023826-2021-04631 | 2021-09-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -15.0/2.0/4 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. INTRAOPERATIVELY THE LENS WAS REMOVED DUE TO THE LENS HAS TORE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -15.0/2.0/4 | OD | Physician | 12/01/2021 | 2021 | 55 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12921555&pc=QCB | 2023826-2021-04568 | 2021-11-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS LEFT EYE (OS). AT ONE WEEK POST-OP, THERE WAS A REPORT OF LOW VAULTING AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5: THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/087 (SPHERE/CYLINDER/AXIS). THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT IS ASYMPTOMATIC AND HAPPY WITH VISION. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: D2: COMMON DEVICE NAME: PHAKIC TORIC INTRAOCULAR LENS, QCB. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. DATE OF EVENT: UNK. EXPIRATION DATE: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. CLAIM # (B)(4). | -09.50/+1.0/087 | OS | Physician | 11/03/2021 | 2021 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12930009&pc=QCB | 2023826-2021-04634 | 2021-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -12.5/1.0/171 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH LENS BUT DIFFERENT AXIS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/1.0/171 | OD | Physician | 11/15/2021 | 2021 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12942150&pc=QCB | 2023826-2021-04545 | 2021-09-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -8.5/3.0/095 (SPHERE/ CYLINDER/ AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/3.0/095 | OS | Physician | 11/21/2021 | 2021 | 49 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12942416&pc=QCB | 2023826-2021-04543 | 2021-08-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -13.5/2.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM) ADDED TO INITAL MDR. CLAIM#: (B)(4). | -13.5/2.0/094 | OD | Physician | 11/21/2021 | 2021 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12942420&pc=QCB | 2023826-2021-04637 | 2020-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+4.5/004 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) UPSIDE DOWN ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXPLANTED. REPORTEDLY, THE LENS GOT STUCK IN THE CARTRIDGE AND WAS TORN DURING AN ATTEMPT TO RE-IMPLANT THE LENS. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR AND THE LENS IS REPORTED TO HAVE FAILED TO PERFORM AS INTENDED. NO PATIENT INJURY WAS REPORTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -8.5/+4.5/004 | OS | Physician | 11/09/2021 | 2021 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12942451&pc=QCB | 2023826-2021-04542 | 2021-08-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: BS THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -13.0/3.0/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -13.0/3.0/087 | OS | Physician | 11/21/2021 | 2021 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12942458&pc=QCB | 2023826-2021-04546 | 2021-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2 VTICMO 13.2 IMPLANTABLE COLLAMER OF -11.0/1.5/92 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/1.5/92 | OS | Physician | 11/21/2021 | 2021 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12942477&pc=QCB | 2023826-2021-04544 | 2021-09-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -10.0/3.5/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.0/3.5/089 | OD | Physician | 11/21/2021 | 2021 | 49 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12942521&pc=QCB | 2023826-2021-04640 | 2021-10-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+0.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) UPSIDE DOWN. THIS OCCURRED ON (B)(6) 2021. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY AND THE PROBLEM IS RESOLVED. REPORTEDLY NO PATIENT INJURY OCCURRED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -15.0/+0.5/092 | OD | Physician | 11/09/2021 | 2021 | 44 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12943189&pc=QCB | 2023826-2021-04661 | 2021-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -10.50/+1.50/87 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT SPHERICAL LENS DUE TO LENS ROTATION. THIS EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND DEBRIS THROUGHOUT THE LENS AND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -10.50/+1.50/87 | OD | Physician | 12/02/2021 | 2021 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12943229&pc=QCB | 2023826-2021-04660 | 2021-11-15 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -6.5/1.5/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULTING. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.5/1.5/89 | OD | Physician | 11/15/2021 | 2021 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12951993&pc=QCB | 2023826-2021-04589 | 2021-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.00/3.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -10.00/3.0/088 | OS | Physician | 11/15/2021 | 2021 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12951998&pc=QCB | 2023826-2021-04586 | 2021-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -08.00/1.0/129 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -08.00/1.0/129 | OS | Physician | 11/15/2021 | 2021 | 46 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12952003&pc=QCB | 2023826-2021-04591 | 2021-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -11.00/2.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.00/2.0/081 | OS | Physician | 11/14/2021 | 2021 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12952005&pc=QCB | 2023826-2021-04590 | 2021-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/1.00/034 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4) | -11.00/1.00/034 | OS | Physician | 11/14/2021 | 2021 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12952007&pc=QCB | 2023826-2021-04595 | 2021-10-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.50/1.50/115 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.50/1.50/115 | OD | Physician | 11/15/2021 | 2021 | 46 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12952082&pc=QCB | 2023826-2021-04597 | 2021-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -10.00/2.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.00/2.0/092 | OS | Physician | 11/15/2021 | 2021 | 43 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12952091&pc=QCB | 2023826-2021-04602 | 2021-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -07.00/1.5/092 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -07.00/1.5/092 | OS | Physician | 11/18/2021 | 2021 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12952101&pc=QCB | 2023826-2021-04603 | 2021-09-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -08.00/1.5/091 , (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -08.00/1.5/091 | OD | Physician | 11/18/2021 | 2021 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12952106&pc=QCB | 2023826-2021-04605 | 2021-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.5/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -11.5/2.5/091 | OD | Physician | 11/18/2021 | 2021 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12952110&pc=QCB | 2023826-2021-04606 | 2021-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/1.0/090 , (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.50/1.0/090 | OS | Physician | 11/18/2021 | 2021 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12953753&pc=QCB | 2023826-2021-04596 | 2021-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -12.50/3.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.50/3.0/095 | OD | Physician | 11/15/2021 | 2021 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12957840&pc=QCB | 2023826-2021-04662 | 2021-10-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -9.0/2.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. INTRAOPERATIVELY THE LENS WAS REMOVED DUE TO THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE PATIENT'S EYE. THERE WAS NO PATIENT INJURY. ON (B)(6) 2021 A REPLACEMENT LENS WAS IMPLANTED OF THE SAME MODEL/LENGTH AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED "LENS STUCK IN PLUNGER AND GOT DAMAGED WHILE INJECTING INTO THE EYE". Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -9.0/2.0/090 | OS | Physician | 12/03/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12960473&pc=QCB | 2023826-2021-04550 | 2021-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -6.5/0.5/60 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6)2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.5/0.5/60 | OS | Physician | 11/21/2021 | 2021 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12960479&pc=QCB | 2023826-2021-04548 | 2021-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -10.5/1.0/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/1.0/91 | OD | Physician | 11/21/2021 | 2021 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12960508&pc=QCB | 2023826-2021-04549 | 2021-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -6.5/3.0/92 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)2021, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -6.5/3.0/92 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)2021, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -6.5/3.0/92 | OS | Physician | 11/21/2021 | 2021 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12965086&pc=QCB | 2023826-2021-04693 | 2021-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -8.0/1.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED ROTATION NOT ASSOCIATED WITH LOW VAULTING. THE SURGEON REPOSITIONED THE LENS ON (B)(6) 2021 BUT THIS DID NOT RESOLVE THE PROBLEM. THEN ON (B(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -8.0/1.5/091 | OD | Physician | 12/07/2021 | 2021 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12965113&pc=QCB | 2023826-2021-04694 | 2021-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.0/1.0/112 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE PATIENT EXPERIENCED ROTATION NOT ASSOCIATED WITH LOW VAULTING. ON (B)(6) 2021, THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021, THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.0/1.0/112 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE PATIENT EXPERIENCED ROTATION NOT ASSOCIATED WITH LOW VAULTING. ON (B)(6) 2021, THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021, THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -6.0/1.0/112 | OS | Physician | 12/07/2021 | 2021 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12965353&pc=QCB | 2023826-2021-04648 | 2021-10-31 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50/+2.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: AGE: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: NA. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -10.50/+2.5/092 | OS | Physician | 11/10/2021 | 2021 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12972660&pc=QCB | 2023826-2021-04551 | 2021-11-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -9.5/1.5/168 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE WAS IMPLANTED VERTICALLY AND THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -9.5/1.5/168 | OD | Physician | 11/21/2021 | 2021 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12972666&pc=QCB | 2023826-2021-04557 | 2021-02-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.0/1.5/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM #: (B)(4). | -11.0/1.5/086 | OS | Physician | 11/22/2021 | 2021 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12972725&pc=QCB | 2023826-2021-04554 | 2021-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -8.5/1.0/09 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.5/1.0/09 | OS | Physician | 11/21/2021 | 2021 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12972741&pc=QCB | 2023826-2021-04553 | 2021-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -10.5/2.0/075 (SPHERE/ CYLINDER/ AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -10.5/2.0/075 | OD | Physician | 11/21/2021 | 2021 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12972745&pc=QCB | 2023826-2021-04559 | 2021-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.7 VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF -10.0/1.5/89 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.0/1.5/89 | OS | Physician | 11/22/2021 | 2021 | 13 MO | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12972761&pc=QCB | 2023826-2021-04552 | 2021-11-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/011 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. THE REPORTER STATES THAT THE FIRST LENS WAS IMPLANTED VERTICALLY. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -9.0/1.0/011 | OS | Physician | 11/21/2021 | 2021 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12972763&pc=QCB | 2023826-2021-04558 | 2021-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.7 VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF -9.5/2.0/87 (SPHERE/ CYLINDER/ AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -9.5/2.0/87 | OD | Physician | 11/22/2021 | 2021 | 13 MO | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12972765&pc=QCB | 2023826-2021-04556 | 2021-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -11.0/2.5/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6)2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -11.0/2.5/93 | OD | Physician | 11/22/2021 | 2021 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12972931&pc=QCB | 2023826-2021-04655 | 2021-10-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/085 (SPHERE/ CYLINDER/ AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED DURING THE SAME SURGERY WITH ANOTHER SAME MODEL/ LENGTH LENS (THE LENS WAS IMPLANTED VERTICALLY) BUT THE PROBLEM WAS NOT RESOLVED. SEE MFR. REPORT # 2023826-2021-04656 FOR REPLACEMENT LENS. CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.00/+1.5/085 | NA | Physician | 11/11/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12972988&pc=QCB | 2023826-2021-04719 | 2021-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6 VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.5/4.0/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON 13-SEPT-2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM #: (B)(4). | -11.5/4.0/087 | OD | Physician | 11/23/2021 | 2021 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973014&pc=QCB | 2023826-2021-04654 | 2021-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+2.0/170 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+2.0/170 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -10.50/+2.0/170 | OS | Physician | 11/11/2021 | 2021 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973337&pc=QCB | 2023826-2021-04708 | 2021-03-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+5.0/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. INTRAOPERATIVELY THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.0/+5.0/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. INTRAOPERATIVELY THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -13.0/+5.0/073 | OS | Physician | 11/15/2021 | 2021 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973339&pc=QCB | 2023826-2021-04710 | 2021-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -11.5/+2.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD),IT WAS TORN. THIS OCCURRED ON (B)(6) 2021. REPORTEDLY THE LENS REMAINS IMPLANTED AND NO PATIENT INJURY OCCURRED.Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -11.5/+2.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD),IT WAS TORN. THIS OCCURRED ON (B)(6) 2021. REPORTEDLY THE LENS REMAINS IMPLANTED AND NO PATIENT INJURY OCCURRED. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: B3- DATE OF EVENT CORRECTED TO (B)(6) 2021 IN THE INITIAL MDR. B5- "ON (B)(6) 2021 THE LENS WAS EXPLANTED FOR A SAME MODEL AND LENGTH LENS. THIS RESOLVED THE PROBLEM." SHOULD BE ADDED TO THE INITIAL MDR. D6B - IF EXPLANNTED, GIVE DATE (B)(6) 2021 SHOULD BE ADDED IN INITIAL MDR. H1- CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6- HEALTH EFFECT IMPACT CODE 2199 CORRECTED TO 4627 IN INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -11.5/+2.0/087 | OD | Physician | 08/02/2021 | 2021 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973342&pc=QCB | 2023826-2021-04703 | 2021-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -13.5/+2.5/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -13.5/+2.5/092 | OS | Physician | 11/11/2021 | 2021 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973348&pc=QCB | 2023826-2021-04698 | 2021-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -13.0/+2.0/091 | OD | Physician | 11/10/2021 | 2021 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973361&pc=QCB | 2023826-2021-04700 | 2021-06-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -9.0/+1.5/173 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: B5 - THE LENS WAS REPLACED AT A LATER DATE AND PROBLEM IS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: B5 - PER UPDATED EMAIL, THE LENS WAS REMOVED AND PATIENT DECIDED NOT TO IMPLANT A NEW LENS ANYMORE. THEREFORE THERE IS NO EXCHANGED LENS INFORMATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE - UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -9.0/+1.5/173 | OS | Physician | 11/11/2021 | 2021 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12973362&pc=QCB | 2023826-2021-04696 | 2021-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -12.5/+2.0/072 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY:UNK. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN VIAL DRY WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -12.5/+2.0/072 | OS | Physician | 11/11/2021 | 2021 | 41 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12979005&pc=QCB | 2023826-2021-04656 | 2021-10-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED (VERTICALLY) A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/172 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD). THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2021-04655 AND THE PROBLEM WAS NOT RESOLVED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE REPORTER INDICATED THE PROBLEM OF HIGH VAULT HAS BEEN RESOLVED. THE VAULT WAS 550UM AT THE LAST VISIT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.00/+1.5/172 | OD | Physician | 11/11/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12980833&pc=QCB | 2023826-2021-04723 | 2021-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -8.5/2.0/106 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)2021, THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -8.5/2.0/106 | OS | Physician | 11/24/2021 | 2021 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12980868&pc=QCB | 2023826-2021-04720 | 2021-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/92 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6)2021 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -12.0/1.5/92 | OD | Physician | 11/24/2021 | 2021 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12980875&pc=QCB | 2023826-2021-04721 | 2021-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -7.5/0.5/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -7.5/0.5/93 | OD | Physician | 11/24/2021 | 2021 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12980888&pc=QCB | 2023826-2021-04722 | 2021-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -8.0/0.5/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US CLAIM# (B)(4). | -8.0/0.5/087 | OS | Physician | 11/24/2021 | 2021 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12980931&pc=QCB | 2023826-2021-04724 | 2021-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -10.0/2.0/010 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.0/2.0/010 | OD | Physician | 11/24/2021 | 2021 | 42 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12981002&pc=QCB | 2023826-2021-04801 | 2021-12-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -8.00/1.0/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULTING. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -8.00/1.0/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULTING. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- PER FINAL CQ RECEIVED ON 31-JAN-2022: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -8.00/1.0/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LEN WAS EXCHANGED FOR A SPHERICAL MODEL LENS LONGER IN LENGTH (13.2MM VICM5_13.2) ON (B)(6) 2021. THE PROBLEM WAS RESOLVED AND "PATIENT IS HAPPY!" CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -8.00/1.0/95 | OS | Physician | 12/08/2021 | 2021 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12981015&pc=QCB | 2023826-2021-04800 | 2021-12-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -7.50/1.0/71 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULTING. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- PER FINAL CQ RECEIVED ON 31-JAN-2022: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -7.50/1.0/71 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LEN WAS EXCHANGED FOR A SPHERICAL MODEL LENS LONGER IN LENGTH (13.2MM VICM5_13.2) ON (B)(6) 2021. THE PROBLEM WAS RESOLVED AND "PATIENT IS HAPPY!" CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH, NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -7.50/1.0/71 | OD | Physician | 12/08/2021 | 2021 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12993487&pc=QCB | 2023826-2021-04884 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICMO12.1; -13.50/1.5/071 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). IT WAS INDICATED THAT THE PATIENT EXPERIENCED LOW VAULTING. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -13.50/1.5/071 | OS | Physician | 11/19/2021 | 2021 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12993567&pc=QCB | 2023826-2021-04883 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICMO13.2; -13.00/1.5/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). IT WAS INDICATED THAT THE PATIENT EXPERIENCED LOW VAULTING. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -13.00/1.5/095 | OD | Physician | 11/19/2021 | 2021 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=12993612&pc=QCB | 2023826-2021-04880 | 2021-10-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -5.0/1.5/107 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. INTRAOPERATIVELY THE LENS WAS REMOVED DUE TO THE LENS WAS IMPLANTED UPSIDE DOWN. THERE WAS NO PATIENT INJURY. ON (B)(6) 2021 A REPLACEMENT LENS OF THE SAME MODEL/LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -5.0/1.5/107 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. INTRAOPERATIVELY THE LENS WAS REMOVED DUE TO THE LENS WAS IMPLANTED UPSIDE DOWN. THERE WAS NO PATIENT INJURY. ON (B)(6) 2021 A REPLACEMENT LENS OF THE SAME MODEL/LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -5.0/1.5/107 | OS | Physician | 11/16/2021 | 2021 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13001716&pc=QCB | 2023826-2021-04673 | 2021-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/2.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US CLAIM# (B)(4). | -11.00/2.0/096 | OD | Physician | 11/22/2021 | 2021 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13001724&pc=QCB | 2023826-2021-04678 | 2021-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -12.0/1.0/168 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON 13/SEP/2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -12.0/1.0/168 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -12.0/1.0/168 | OS | Physician | 11/24/2021 | 2021 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13001769&pc=QCB | 2023826-2021-04610 | 2021-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -5.00/1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.00/1.0/092 | OD | Physician | 11/22/2021 | 2021 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13001773&pc=QCB | 2023826-2021-04663 | 2021-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.50/1.0/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -11.50/1.0/070 | OS | Physician | 11/22/2021 | 2021 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13001779&pc=QCB | 2023826-2021-04670 | 2021-08-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -12.00/2.00/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REPLACED INTRAOPERATIVELY WITH A LONGER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S LEFT EYE (OS) DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -12.00/2.00/093 | OS | Physician | 11/22/2021 | 2021 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13001781&pc=QCB | 2023826-2021-04675 | 2021-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -16.00/0.5/177 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.00/0.5/177 | OS | Physician | 11/22/2021 | 2021 | 47 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13001783&pc=QCB | 2023826-2021-04612 | 2021-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -7.50/1.0/009 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.50/1.0/009 | OD | Physician | 11/21/2021 | 2021 | 17 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13001864&pc=QCB | 2023826-2021-04669 | 2021-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.50/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -10.50/1.0/091 | OS | Physician | 11/22/2021 | 2021 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13001869&pc=QCB | 2023826-2021-04677 | 2021-10-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.50/1.5/180 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, INITIAL AND EXCHANGE LENS WERE IMPLANTED VERTICALLY. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.5/180 | OD | Physician | 11/24/2021 | 2021 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13001882&pc=QCB | 2023826-2021-04608 | 2021-06-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.50/1.5/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -9.50/1.5/110 | OS | Physician | 11/21/2021 | 2021 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13001887&pc=QCB | 2023826-2021-04667 | 2021-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.00/2.5/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -8.00/2.5/108 | OD | Physician | 11/22/2021 | 2021 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13001889&pc=QCB | 2023826-2021-04674 | 2021-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -17.00/1.0/157 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -17.00/1.0/157 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -17.00/1.0/157 | OD | Physician | 11/22/2021 | 2021 | 47 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13001895&pc=QCB | 2023826-2021-04679 | 2021-10-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.00/1.0/132 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -10.00/1.0/132 | OD | Physician | 11/24/2021 | 2021 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13001904&pc=QCB | 2023826-2021-04607 | 2021-06-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.50/2.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#: (B)(4). | -9.50/2.0/092 | OD | Physician | 11/21/2021 | 2021 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13001911&pc=QCB | 2023826-2021-04668 | 2021-11-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.00/2.50/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.00/2.50/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK ETHNICITY: UNK PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK ETHNICITY: UNK PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -7.00/2.50/088 | OS | Physician | 11/22/2021 | 2021 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13001965&pc=QCB | 2023826-2021-04666 | 2021-11-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/2.0/091 | OS | Physician | 11/22/2021 | 2021 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13001967&pc=QCB | 2023826-2021-04672 | 2021-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.00/1.0/129 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -9.00/1.0/129 | OS | Physician | 11/22/2021 | 2021 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13001969&pc=QCB | 2023826-2021-04680 | 2021-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.50/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.50/2.5/092 | OS | Physician | 11/24/2021 | 2021 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13002083&pc=QCB | 2023826-2021-04804 | 2021-11-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -07.00/+2.0/095 (SPHERE/CYLINDER/AXIS), WAS SCRATCHED AND IT WAS UNKNOWN IF THERE WAS ANY PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE BACK UP LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN VIAL WITH MOISTURE. VISUAL INSPECTION FOUND OPTIC AND HAPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. CLAIM # (B)(4). | -07.00/+2.0/095 | NA | Other Health Care Professional | 11/22/2021 | 2021 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13009692&pc=QCB | 2023826-2021-04665 | 2021-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.50/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -14.50/1.5/091 | OS | Physician | 11/22/2021 | 2021 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13011401&pc=QCB | 2023826-2021-04824 | 2021-08-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Unintended Movement | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.50/+3.0/071 (SPHERE/ CYLINDER/ AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS WAS INSERTED UPSIDE DOWN INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITHIN THE SAME SURGERY WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.50/+3.0/071 | OS | Physician | 11/21/2021 | 2021 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13011415&pc=QCB | 2023826-2021-04827 | 2021-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT THE PROBLEM WAS NOT RESOLVED SO THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED (THE LENS WAS IMPLANTED VERTICALLY). THE CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT THE PROBLEM WAS NOT RESOLVED SO THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED (THE LENS WAS IMPLANTED VERTICALLY). THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -12.00/+2.0/087 | OD | Physician | 11/15/2021 | 2021 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13014885&pc=QCB | 2023826-2021-04806 | 2021-11-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.50/+3.5/071 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED ON (B)(6) 2021 WITH ANOTHER SAME MODEL/ LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.50/+3.5/071 | OD | Physician | 11/15/2021 | 2021 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13014896&pc=QCB | 2023826-2021-04807 | 2021-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/109 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/109 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PER ADDITIONAL INFORMATION RECEIVED ON 14-APR-2022. THE INITAL LENS (12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/109) WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. THE REPORTER STATED "THE EXPLANT RESOLVED THE PROBLEM. NO FURTHER SURGICAL INTERVENTION WILL BE PLANNED." CLAIM# (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -10.00/+1.5/109 | OD | Physician | 11/16/2021 | 2021 | 49 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13014899&pc=QCB | 2023826-2021-04826 | 2021-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -07.00/+1.0/004 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING AND THE PROBLEM WAS RESOLVED. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -07.00/+1.0/004 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING AND THE PROBLEM WAS RESOLVED. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -07.00/+1.0/004 | OS | Physician | 11/15/2021 | 2021 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13014905&pc=QCB | 2023826-2021-04808 | 2021-10-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/112 (SPHERE/CYLINDER/AXIS), TORE DURING LOADING AND THERE WAS NO PATIENT CONTACT. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. THE LENS WAS DISCARDED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.50/+1.5/112 | NA | Physician | 11/17/2021 | 2021 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13014959&pc=QCB | 2023826-2021-04811 | 2020-05-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00/+4.5/061 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TEARING DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00/+4.5/061 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THERE WAS A PROBLEM WITH THE LENS. THERE WAS NO PATIENT INJURY. THE SURGEON DID NOT IMPLANT ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.00/+4.5/061 | OS | Physician | 11/17/2021 | 2021 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13023011&pc=QCB | 2023826-2021-04904 | 2021-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -10.0/1.5/167 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH/MODEL LENS BUT DIFFERENT AXIS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS NOTED AS "OTHER". Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON THE LENS WITH FIBERS ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -10.0/1.5/167 | OD | Physician | 11/22/2021 | 2021 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13023220&pc=QCB | 2023826-2021-04901 | 2021-11-20 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Unintended Movement | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+4.0/091 (SPHERE/CYLINDER/AXIS) WAS "STICKY" AND WOULD NOT OPEN DURING INJECTION INTO THE LEFT (OS) EYE. LENS STUCK IN CARTRIDGE/INJECTION SYSTEM IS ALSO REPORTED. THERE WAS PATIENT CONTACT WITH THE LENS BUT NO PATIENT INJURY. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY ON (B)(4) 2021. AN ALTERNATE LENS WAS IMPLANTED IN A SEPARATE SURGERY. SEE MFR FILE# (B)(4) FOR REPLACEMENT LENS. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -11.0/+4.0/091 | OS | Physician | 11/26/2021 | 2021 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13023221&pc=QCB | 2023826-2021-04899 | 2021-09-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1, -11.0/+3.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXPLANTED AND THEN REPLACED WITH A LONGER LENS DUE TO LENS ROTATION AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1, -11.0/+3.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXPLANTED AND THEN REPLACED WITH A LONGER LENS DUE TO LENS ROTATION AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE AND RESIDUE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY AND RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY AND RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+3.0/091 | OD | Physician | 11/25/2021 | 2021 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13023223&pc=QCB | 2023826-2021-04896 | 2021-10-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/070 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY ON (B)(6) 2021. ON THE SAME DATE IN A SEPARATE SURGERY A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. REPORTEDLY, THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -10.0/+1.0/070 | OS | Physician | 11/24/2021 | 2021 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13023225&pc=QCB | 2023826-2021-04892 | 2021-11-18 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -5.5/+2.0/091 (SPHERE/ CYLINDER/ AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. EXCESSIVE VAULT IS REPORTED. REPORTEDLY, THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5 - "REPORTEDLY, THE PATIENT DOES NOT WANT A LENS EXCHANGE. PATIENT IS HAPPY. LENS REMAINS IMPLANTED." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). | -5.5/+2.0/091 | OS | Physician | 11/19/2021 | 2021 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13023226&pc=QCB | 2023826-2021-04891 | 2021-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.0/+1.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -15.0/+1.5/091 | OD | Physician | 11/19/2021 | 2021 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13023227&pc=QCB | 2023826-2021-04889 | 2021-05-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT AS THE SURGEON INJECTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/110 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2021. THE LENS WAS INJECTED, REMOVED, AND REPLACED INTRA-OPERATIVELY. THE PROBLEM IS RESOLVED. NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT AS THE SURGEON INJECTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+1.5/110 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD), THE LENS TORE/BROKE. THIS OCCURRED ON (B)(6) 2021. THE LENS WAS INJECTED, REMOVED, AND REPLACED INTRA-OPERATIVELY. THE PROBLEM IS RESOLVED. NO PATIENT INJURY OCCURRED AND THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: PREVIOUS G3 STATEMENT SHOULD HAVE READ 03/15/2022. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -11.0/+1.5/110 | OD | Physician | 11/18/2021 | 2021 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13023229&pc=QCB | 2023826-2021-04888 | 2021-02-09 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/072 (SPHERE/ CYLINDER/ AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/ BROKE. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRA-OPERATIVELY ON (B)(6) 2021 AND THE PROBLEM WAS RESOLVED. REPORTEDLY, THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -8.5/+1.5/072 | OS | Physician | 11/18/2021 | 2021 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13023232&pc=QCB | 2023826-2021-04893 | 2021-11-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -5.0/+1.0/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. LOW VAULT IS REPORTED. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -5.0/+1.0/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. LOW VAULT IS REPORTED. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -5.0/+1.0/094 | OD | Physician | 11/19/2021 | 2021 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13023238&pc=QCB | 2023826-2021-04900 | 2021-10-31 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Unintended Movement | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+3.5/086 (SPHERE/CYLINDER/AXIS) FLIPPED AFTER INJECTION INTO THE LEFT (OS) EYE. THERE WAS PATIENT CONTACT WITH THE LENS BUT NO PATIENT INJURY. REPORTEDLY THE LENS WAS THEN "TORN DURING THE HAND OVER HAND REMOVAL TECHNIQUE." THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY ON (B)(6) 2021. AN ALTERNATE LENS WAS IMPLANTED AT A LATER DATE. SEE MFR REPORT # 2023826-2021-04901 FOR REPLACEMENT LENS. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND THAT THE HAPTIC WAS BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.0/+3.5/086 | OS | Physician | 11/26/2021 | 2021 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13027732&pc=QCB | 2023826-2021-04731 | 2021-10-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -12.0/2.5/97 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE LENS WAS IMPLANTED VERTICALLY AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.0/2.5/97 | OD | Physician | 11/25/2021 | 2021 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13027742&pc=QCB | 2023826-2021-04728 | 2021-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT 13.2 VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -11.5/2.0/94 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: D4: 30-NOV-2023. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/2.0/94 | OD | Physician | 11/24/2021 | 2021 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13027753&pc=QCB | 2023826-2021-04734 | 2021-10-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -13.0/3.0/170 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE LENS WAS IMPLANTED VERTICALLY AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.0/3.0/170 | OS | Physician | 11/25/2021 | 2021 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13053929&pc=QCB | 2023826-2021-04833 | 2021-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -14.00/0.5/140 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN. ETHNICITY: UNKNOWN. RACE: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -14.00/0.5/140 | OD | Physician | 11/28/2021 | 2021 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13053931&pc=QCB | 2023826-2021-04835 | 2021-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.00/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.00/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -8.00/1.5/089 | OD | Physician | 11/28/2021 | 2021 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13053935&pc=QCB | 2023826-2021-04845 | 2021-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.00/1.0/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.00/1.0/110 | OD | Physician | 11/30/2021 | 2021 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13053938&pc=QCB | 2023826-2021-04836 | 2021-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.00/1.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.00/1.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -9.00/1.0/095 | OD | Physician | 11/28/2021 | 2021 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13053943&pc=QCB | 2023826-2021-04832 | 2021-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -13.50/0.5/036 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.50/0.5/036 | OS | Physician | 11/28/2021 | 2021 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13053944&pc=QCB | 2023826-2021-04838 | 2021-04-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.50/2.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.50/2.0/094 | OD | Physician | 11/29/2021 | 2021 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13053946&pc=QCB | 2023826-2021-04834 | 2021-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/1.5/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.00/1.5/072 | OS | Physician | 11/28/2021 | 2021 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13053950&pc=QCB | 2023826-2021-04861 | 2021-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.00/2.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/2.0/084 | OD | Physician | 12/01/2021 | 2021 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13053951&pc=QCB | 2023826-2021-04866 | 2021-04-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/2.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED IN A SECOND SURGERY WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -11.50/2.5/087 | OS | Physician | 12/01/2021 | 2021 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13053954&pc=QCB | 2023826-2021-04849 | 2021-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -8.50/1.5/154 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -8.50/1.5/154 | OS | Physician | 12/01/2021 | 2021 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13053960&pc=QCB | 2023826-2021-04855 | 2021-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.50/3.5/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.50/3.5/078 | OS | Physician | 12/01/2021 | 2021 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13053962&pc=QCB | 2023826-2021-04859 | 2021-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.50/4.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -13.50/4.0/090 | OS | Physician | 12/01/2021 | 2021 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13053963&pc=QCB | 2023826-2021-04868 | 2021-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -5.50/1.5/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -5.50/1.5/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -5.50/1.5/074 | OD | Physician | 12/01/2021 | 2021 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13053964&pc=QCB | 2023826-2021-04847 | 2021-11-22 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1, -10.00/1.50/088 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REPLACED INTRAOPERATIVELY WITH A LONGER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S LEFT EYE (OS) DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1, -10.00/1.50/088 (SPHERE/CYLINDER/AXIS) , IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REPLACED INTRAOPERATIVELY WITH A LONGER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S LEFT EYE (OS) DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK; ETHNICITY: UNK; RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK; ETHNICITY: UNK; RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -10.00/1.50/088 | OS | Physician | 11/30/2021 | 2021 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13053965&pc=QCB | 2023826-2021-04870 | 2021-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.00/2.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.00/2.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -8.00/2.5/089 | OS | Physician | 12/01/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13053968&pc=QCB | 2023826-2021-04852 | 2021-10-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.00/1.0/169 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -7.00/1.0/169 | OS | Physician | 12/01/2021 | 2021 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13053969&pc=QCB | 2023826-2021-04862 | 2021-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.00/0.5/160 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.00/0.5/160 | OD | Physician | 12/01/2021 | 2021 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13053974&pc=QCB | 2023826-2021-04857 | 2021-10-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.00/1.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -9.00/1.0/088 | OS | Physician | 12/01/2021 | 2021 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13053976&pc=QCB | 2023826-2021-04865 | 2021-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.00/1.0/003 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -10.00/1.0/003 | OD | Physician | 12/01/2021 | 2021 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13053978&pc=QCB | 2023826-2021-04875 | 2021-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/1.5/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED IN A SECOND SURGERY WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -11.00/1.5/072 | OS | Physician | 12/03/2021 | 2021 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13053980&pc=QCB | 2023826-2021-04878 | 2021-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -9.00/1.50/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2021 DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -9.00/1.50/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2021 DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.00/1.50/081 | OS | Physician | 12/03/2021 | 2021 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13053987&pc=QCB | 2023826-2021-04837 | 2021-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -10.50/1.0/044 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). Manufacturer Narrative: CORRECTION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -10.50/1.0/044 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULT. IN A SECOND SURGERY ON (B)(6)2021 A LONGER LENGTH REPLACEMENT LENS WAS IMPLANTED . THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR. D6B REPORTED AS ON (B)(6) 2021 / CORRECT TO ON (B)(6) 2021. CLAIM#: (B)(4). | -10.50/1.0/044 | OD | Physician | 11/29/2021 | 2021 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13053991&pc=QCB | 2023826-2021-04877 | 2021-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6, -11.50/2.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.50/2.0/092 | OD | Physician | 12/03/2021 | 2021 | 43 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13053996&pc=QCB | 2023826-2021-04848 | 2021-11-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/2.5/117 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -10.50/2.5/117 | OD | Physician | 11/30/2021 | 2021 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13053998&pc=QCB | 2023826-2021-04846 | 2021-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.00/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K- THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.00/1.0/087 | OS | Physician | 11/30/2021 | 2021 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13053999&pc=QCB | 2023826-2021-04854 | 2021-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.00/2.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.00/2.0/085 | OD | Physician | 12/01/2021 | 2021 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13054001&pc=QCB | 2023826-2021-04858 | 2021-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.00/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -9.00/1.0/089 | OD | Physician | 12/01/2021 | 2021 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13054010&pc=QCB | 2023826-2021-04860 | 2021-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.00/3.5/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.00/3.5/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -15.00/3.5/112 | OD | Physician | 12/01/2021 | 2021 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13054012&pc=QCB | 2023826-2021-04869 | 2021-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -6.50/2.0/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -6.50/2.0/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -6.50/2.0/077 | OD | Physician | 12/01/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13054019&pc=QCB | 2023826-2021-04876 | 2021-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.50/2.0/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.50/2.0/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -11.50/2.0/105 | OS | Physician | 12/03/2021 | 2021 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13054020&pc=QCB | 2023826-2021-04867 | 2021-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -5.50/1.0/106 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM# (B)(4). | -5.50/1.0/106 | OS | Physician | 12/01/2021 | 2021 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13054282&pc=QCB | 2023826-2021-04737 | 2021-10-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT 13.2 VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -12.0/0.5/89 (SPHERE/ CYLINDER/ AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.0/0.5/89 | OD | Physician | 11/29/2021 | 2021 | 45 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13057157&pc=QCB | 2023826-2021-04914 | 2021-09-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50/+3.0/090 (SPHERE/CYLINDER/AXIS),WITHIN THE SAME SURGERY DUE TO THE LENS WAS IMPLANTED INTO THE WRONG EYE (OS LENS IMPLANTED IN OD). THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. THE EYE IS QUIET AND STABLE. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: B5- UPDATED: THE REPORTER INDICATED THE SURGEON INSERTED THE LEFT (OS) LENS (12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -14.50/+3.0/090) INTO THE RIGHT (OD) EYE BY MISTAKE ON (B)(6) 2021. THE LENS WAS REMOVED INTRAOPERATIVELY. LATER ON (B)(6) 2021 A LENS OF THE SAME PARAMETERS WAS IMPLANTED INTO THE RIGHT (OD) EYE AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. CLAIM# (B)(4). | -14.50/+3.0/090 | OD | Physician | 11/26/2021 | 2021 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13057197&pc=QCB | 2023826-2021-04831 | 2021-10-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/112 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/112 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON PRODUCT. VISUAL INSPECTION FOUND ONE HAPTIC TORN AND FIBERS ON THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -12.00/+1.5/112 | OD | Physician | 11/21/2021 | 2021 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13057207&pc=QCB | 2023826-2021-04830 | 2021-10-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/+3.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS EXPLANTED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. (B)(4) Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.00/+3.0/090 | OD | Physician | 11/21/2021 | 2021 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13057231&pc=QCB | 2023826-2021-04909 | 2021-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+4.0/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (THE LENS WAS IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WERE FIBERS ON THE LENS SURFACE. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). | -09.00/+4.0/096 | OS | Physician | 11/23/2021 | 2021 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13057240&pc=QCB | 2023826-2021-04818 | 2021-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -08.50/+3.0/160 (SPHERE/CYLINDER/AXIS), (IMPLANTED VERTICALLY) INTO THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED HORIZONTALLY) AND THE PROBLEM WAS RESOLVED.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -08.50/+3.0/160 (SPHERE/CYLINDER/AXIS), (IMPLANTED VERTICALLY) INTO THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED HORIZONTALLY) AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -08.50/+3.0/160 | OS | Physician | 11/22/2021 | 2021 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13057244&pc=QCB | 2023826-2021-04817 | 2021-10-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -09.50/+2.0/088 (SPHERE/CYLINDER/AXIS), TORE DURING DELIVERY INTO THE PATIENTS EYE. THERE WAS NO PATIENT CONTACT OR INJURY. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED.Event Description: THE REPORTER INDICATED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -09.50/+2.0/088 (SPHERE/CYLINDER/AXIS), TORE DURING DELIVERY INTO THE PATIENTS EYE. THERE WAS NO PATIENT CONTACT OR INJURY. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -09.50/+2.0/088 | NA | Physician | 11/22/2021 | 2021 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13064425&pc=QCB | 2023826-2021-04982 | 2021-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED EXCESSIVE VAULTING. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.00/2.50/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) IN (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULT. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: B5-THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.00/2.50/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) IN (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULT. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -9.00/2.50/095 | OS | Physician | 10/26/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13064515&pc=QCB | 2023826-2021-04981 | 2021-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5 13.7, -9.00/4.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) IN (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULT. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5 13.7, -9.00/4.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) IN (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULT. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -9.00/4.0/089 | OD | Physician | 10/26/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13065338&pc=QCB | 2023826-2021-04744 | 2021-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -6.0/1.5/70 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE WAS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -6.0/1.5/70 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -6.0/1.5/70 | OD | Physician | 11/29/2021 | 2021 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13065378&pc=QCB | 2023826-2021-04741 | 2021-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT 13.2 VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -6.0/1.0/73 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.0/1.0/73 | OS | Physician | 11/29/2021 | 2021 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082028&pc=QCB | 2023826-2021-04989 | 2021-12-06 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT WHILE ATTEMPTING TO LOAD A 13.2MM TMICL13.2; -8.00/1.00/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS; THE LENS HAD TORE LOADING INTO THE CARTRIDGE. THIS OCCURRED ON (B)(6) 2021. A BACK UP LENS OF THE SAME MODEL/LENGTH AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS NOTED AS "OTHER" AND CLARIFIED THERE WAS NO PRODUCT ISSUE.Event Description: THE REPORTER INDICATED THAT WHILE ATTEMPTING TO LOAD A 13.2MM TMICL13.2; -8.00/1.00/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS; THE LENS HAD TORE LOADING INTO THE CARTRIDGE. THIS OCCURRED ON (B)(6) 2021. A BACK UP LENS OF THE SAME MODEL/LENGTH AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS NOTED AS "OTHER" AND CLARIFIED THERE WAS NO PRODUCT ISSUE. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -8.00/1.00/092 | NA | Other Health Care Professional | 12/13/2021 | 2021 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082048&pc=QCB | 2023826-2021-04986 | 2021-12-07 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -5.0/+4.0/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: B5- PER UPDATE RECEIVED ON 09-MAR-2022 "MONITORING IS STILL GOING ON. NO REPLACEMENT WILL BE PLANNED. PATIENT IS HAPPY." CLAIM# (B)(4). | -5.0/+4.0/90 | OD | Physician | 12/08/2021 | 2021 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082436&pc=QCB | 2023826-2021-04961 | 2021-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -17.0/+3.5/083 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: PT. INFO-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -17.0/+3.5/083 | OD | Physician | 11/29/2021 | 2021 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082437&pc=QCB | 2023826-2021-04960 | 2021-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: REPORTEDLY AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED ON (B)(6) 2021 INTO THE PATIENT'S LEFT (OS) EYE. ON (B)(6) 2021 THE LENS WAS THEN EXPLANTED AND THEN REPLACED WITH AN ALTERNATE LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: A 13.7MM TMICL13.7, -14.5/+2.5/114 (SPHERE/CYLINDER/AXIS) LENS WAS IMPLANTED. THE LENS WAS EXCHANGED WITH A SHORTER LENS ON (B)(6) 2021 DUE TO "ICL INVERTED." THE PROBLEM WAS RESOLVED. THE CAUSE IS REPORTED AS UNKNOWN. H6-INVESTIGATION CODE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: PRODUCT CODE: UNK; ESUBMITTER SOFTWARE DOES NOT ALLOW FOR A BLANK/UNK ENTRY. (B)(4). | -14.5/+2.5/114 | OS | Physician | 12/02/2021 | 2021 | 41 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082441&pc=QCB | 2023826-2021-04963 | 2021-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.5/+2.5/062 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THEN REPLACED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(6). THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -13.5/+2.5/062 | OS | Physician | 11/29/2021 | 2021 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082448&pc=QCB | 2023826-2021-04964 | 2021-08-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/102 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXPLANTED AND THEN REPLACED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: (B)(6). WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -11.0/+1.0/102 | OD | Physician | 11/29/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082450&pc=QCB | 2023826-2021-04973 | 2021-11-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+3.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED WITH A SAME VERTICAL DIFFERENT AXIS LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.5/+3.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED WITH A SAME VERTICAL DIFFERENT AXIS LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6: LENS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A HAPTIC TORN AND BROKEN LENS. DIMENSIONAL INSPECTION FOUND LENS TO BE WITHIN SPECFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.5/+3.5/091 | OS | Physician | 12/01/2021 | 2021 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082704&pc=QCB | 2023826-2021-04926 | 2021-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.50/+4.5/180 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.50/+4.5/180 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -15.50/+4.5/180 | OD | Physician | 11/30/2021 | 2021 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082710&pc=QCB | 2023826-2021-04913 | 2021-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.50/+1.0/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.50/+1.0/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -12.50/+1.0/103 | OS | Physician | 11/26/2021 | 2021 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13082713&pc=QCB | 2023826-2021-04912 | 2021-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+1.0/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+1.0/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -13.00/+1.0/074 | OD | Physician | 11/26/2021 | 2021 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13085408&pc=QCB | 2023826-2021-04934 | 2021-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+2.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT THE PROBLEM WAS NOT RESOLVED SO THE LENS WAS EXPLANTED. THE LENS WAS REPLACED WITH A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+2.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2015. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT THE PROBLEM WAS NOT RESOLVED SO THE LENS WAS EXPLANTED. THE LENS WAS REPLACED WITH A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -12.50/+2.5/094 | OD | Physician | 12/03/2021 | 2021 | 46 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13090499&pc=QCB | 2023826-2021-04943 | 2021-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.00/2.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/2.0/084 | OD | Physician | 12/06/2021 | 2021 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13090508&pc=QCB | 2023826-2021-04950 | 2021-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.00/2.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/2.0/092 | OD | Physician | 12/08/2021 | 2021 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13090511&pc=QCB | 2023826-2021-04940 | 2021-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.00/1.5/012 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.00/1.5/012 | OS | Physician | 12/06/2021 | 2021 | 46 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13090516&pc=QCB | 2023826-2021-04946 | 2021-02-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.50/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -8.50/1.5/092 | OD | Physician | 12/08/2021 | 2021 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13090536&pc=QCB | 2023826-2021-04939 | 2021-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.00/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.00/2.0/095 | OD | Physician | 12/05/2021 | 2021 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13090542&pc=QCB | 2023826-2021-04951 | 2021-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -11.00/1.5/091 | OS | Physician | 12/08/2021 | 2021 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13090552&pc=QCB | 2023826-2021-04955 | 2021-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -08.50/1.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -08.50/1.5/082 | OD | Physician | 12/09/2021 | 2021 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13090557&pc=QCB | 2023826-2021-04937 | 2021-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -8.50/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.50/2.0/091 | OD | Physician | 12/05/2021 | 2021 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13090558&pc=QCB | 2023826-2021-04942 | 2021-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/1.5/092 (PHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT:UNKNOWN, ETHINICITY:UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.00/1.5/092 | OD | Physician | 12/06/2021 | 2021 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13090579&pc=QCB | 2023826-2021-04931 | 2021-11-30 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/075 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS WAS TORN DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -12.00/+2.0/075 | OS | Physician | 12/01/2021 | 2021 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13090626&pc=QCB | 2023826-2021-04990 | 2021-12-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -09.00/+2.0/090 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS WAS INSERTED UPSIDE DOWN INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/ LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.00/+2.0/090 | OD | Other Health Care Professional | 12/06/2021 | 2021 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13090645&pc=QCB | 2023826-2021-04935 | 2021-10-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.5/050 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (THE LENS WAS IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -12.00/+1.5/050 | OD | Physician | 12/03/2021 | 2021 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13090653&pc=QCB | 2023826-2021-04907 |  | Malfunction | STAAR SURGICAL COMPANY | 2021-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -7.5/2.5/110 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021 AS A REPLACEMENT LENS TO A PREVIOUS REFRACTIVE SURPRISE (MYOPIA OVER CORRECTION). THE SURGEON REPORTS THE PATIENT NOW EXPERIENCES MYOPIA UNDER CORRECTION. LENS REMAINS IMPLANTED. SEE MDR# 2023826- 2021-04906 FOR INITIAL LENS IMPLANT.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -7.5/2.5/110 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021 AS A REPLACEMENT LENS TO A PREVIOUS REFRACTIVE SURPRISE (MYOPIA OVER CORRECTION). THE SURGEON REPORTS THE PATIENT NOW EXPERIENCES MYOPIA UNDER CORRECTION. LENS REMAINS IMPLANTED. SEE MDR# 2023826- 2021-04906 FOR INITIAL LENS IMPLANT. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -7.5/2.5/110 | OS | Physician | 12/12/2021 | 2021 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13090655&pc=QCB | 2023826-2021-04933 |  | Injury | STAAR SURGICAL COMPANY | 2021-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.50/+0.5/026 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT HAS BECOME HYPEROPIC AND HAS MAINLY READING PROBLEMS NOW. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.50/+0.5/026 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT HAS BECOME HYPEROPIC AND HAS MAINLY READING PROBLEMS NOW. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. DATE OF EVENT: UNK. EXPLANT DATE: NA. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: THE SURGEON DECIDED TO CORRECT THE REFRACTIVE SURPRISE WITH A LASER AND THE PATIENT IS HAPPY NOW WITH THE OUTCOME. H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - LASER SURGERY. (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. DATE OF EVENT: UNK. EXPLANT DATE: NA. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.50/+0.5/026 | OS | Physician | 12/02/2021 | 2021 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13091132&pc=QCB | 2023826-2021-04747 | 2021-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -5.5/5.5/088 (SPHERE/ CYLINDER/ AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -5.5/5.5/088 | OD | Physician | 11/30/2021 | 2021 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13091157&pc=QCB | 2023826-2021-04748 | 2021-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -12.0/5.0/035 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -12.0/5.0/035 | OS | Physician | 11/30/2021 | 2021 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13091189&pc=QCB | 2023826-2021-04756 | 2021-11-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -9.5/1.0/068 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.5/1.0/068 | OD | Physician | 12/01/2021 | 2021 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13091343&pc=QCB | 2023826-2021-04753 | 2020-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -7.00/+1.50/114 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS OF LONGER LENGTH AND DIFFERENT POWER DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT - RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -7.00/+1.50/114 | OD | Physician | 11/30/2021 | 2021 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13097272&pc=QCB | 2023826-2021-04759 | 2021-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS (ICL), -5.0/1.0/168 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -5.0/1.0/168 | OD | Physician | 12/01/2021 | 2021 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13097288&pc=QCB | 2023826-2021-04760 | 2021-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS (ICL), -10.5/1.0/084 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.5/1.0/084 | OD | Physician | 11/30/2021 | 2021 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13102084&pc=QCB | 2023826-2021-05013 | 2021-12-21 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2021-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICM5_13.2; -6.0/+6.0/87 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. IT WAS INDICATED THAT THE PATIENT EXPERIENCED EXCESSIVE VAULTING; BUT NO SURGICAL INTERVENTION IS PLANNED DUE TO THE PATIENT FEELS GOOD AND THE SURGEON WILL MONITOR THE PATIENT OVER THE NEXT MONTHS. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -6.0/+6.0/87 | OS | Physician | 12/22/2021 | 2021 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13102472&pc=QCB | 2023826-2021-04763 | 2021-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1 VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -12.50/4.0/85 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.50/4.0/85 | OS | Physician | 12/01/2021 | 2021 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13102473&pc=QCB | 2023826-2021-04762 | 2021-10-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6 VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -14.50/1.50/104 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -14.50/1.50/104 | OD | Physician | 12/01/2021 | 2021 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13113526&pc=QCB | 2023826-2021-04774 | 2021-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -10.5/3.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -10.5/3.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT IDENTIFIER:(B)(6). PATIENT WEIGHT, ETHNICITY AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER:(B)(6). PATIENT WEIGHT, ETHNICITY AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/3.0/091 | OS | Physician | 12/01/2021 | 2021 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13113544&pc=QCB | 2023826-2021-04773 | 2021-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -10.0/1.0/105 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -10.0/1.0/105 | OD | Physician | 12/01/2021 | 2021 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13113566&pc=QCB | 2023826-2021-04772 | 2021-10-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -6.0/2.5/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -6.0/2.5/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -6.0/2.5/086 | OS | Physician | 12/01/2021 | 2021 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13113568&pc=QCB | 2023826-2021-04776 | 2021-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -11.0/2.5/89 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -11.0/2.5/89 | OD | Physician | 12/01/2021 | 2021 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13115041&pc=QCB | 2023826-2021-05017 | 2021-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -9.0/1.0/9 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT DIFFERENT AXIS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -9.0/1.0/9 | OD | Physician | 12/23/2021 | 2021 | 43 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13115046&pc=QCB | 2023826-2021-05020 | 2021-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -14.0/4.5/94 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED WITH LOW VAULTING. ON (B)(6) 2021 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVED THE PROBLEM. THEN ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -14.0/4.5/94 | OS | Physician | 12/24/2021 | 2021 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13115356&pc=QCB | 2023826-2021-04769 | 2021-11-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -9.5/2.0/88 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PM/510K: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE CODE: (B)(4). CLAIM# (B)(4). | -9.5/2.0/88 | OS | Physician | 12/01/2021 | 2021 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13122734&pc=QCB | 2023826-2021-05004 | 2021-09-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.0/+1.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD). ON (B)(6) 2021 THE LENS WAS EXPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE INITIAL LENS WAS IMPLANTED (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED AND REPLACED WITH A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. AFTER THE REPLACEMENT LENS WAS IMPLANTED, UNAIDED VA IS 6/6 AND VAULT IS 200. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR, "LENS TOO SHORT, EXCHANGE FOR DIFFERENT DIOPTER, SIZING ISSUES." CLAIM# (B)(4). | -11.0/+1.0/091 | OD | Physician | 12/06/2021 | 2021 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13122735&pc=QCB | 2023826-2021-05005 | 2021-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.5/+2.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE. ON (B)(6) 2021 THE LENS WAS EXPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE INITIAL LENS WAS IMPLANTED INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2021. THE LENS WAS EXPLANTED AND REPLACED WITH A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. AFTER REPLACEMENT, THE UNAIDED VA IS 6/7.5-1. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR; THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. CLAIM# (B)(4) Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE- UNK. DATE OF EVENT- UNK. IMPLANT DATE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -11.5/+2.0/091 | OS | Physician | 12/06/2021 | 2021 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13122736&pc=QCB | 2023826-2021-05002 | 2020-10-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2021-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/069 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). ON (B)(6) 2021 THE LENS WAS EXPLANTED AND THEN REPLACED WITH AN ALTERNATE LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: B5-ADDITIONAL INFORMATION: THE INITIAL LENS WAS IMPLANTED ON (B)(6) 2020. THE LENS WAS REPLACED WITH A LONGER LENS DUE TO LOW VAULT. UNAIDED VA IS 6/6 AFTER REPLACEMENT. REPORTEDLY THE VAULT IS STILL LOW. REFERENCE MFR FILE# (B)(4) FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT IS REPORTED AS USER ERROR; THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE CODE: (B)(4). CLAIM#: (B)(4). | -9.5/+2.5/069 | OS | Physician | 12/06/2021 | 2021 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13146122&pc=QCB | 2023826-2022-00019 | 2020-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -14.50/+3.00/85 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2020 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. THEN ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS; ALONG WITH AN INDICATION OF "RESIDUAL ASTIGMATISM WAS CORRECTED BY LASER" AND THIS RESOLVED THE PROBLEM.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -14.50/+3.00/85 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2020 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. THEN ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS; ALONG WITH AN INDICATION OF "RESIDUAL ASTIGMATISM WAS CORRECTED BY LASER" AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID . VISUAL INSPECTION FOUND NO VISABLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -14.50/+3.00/85 | OS | Physician | 12/27/2021 | 2021 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13146198&pc=QCB | 2023826-2022-00036 | 2021-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -11.0/+2.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULTING WITH ROTATION. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER CLARIFIED THE CORRECT LENS USED WAS A VTICM5_12.1; -12.00/2.5/073 (SPHERE/CLYINDER/AXIS) IMPLANTABLE COLLAMER LENS. REPORTER INITIALLY PROVIDED THE INCORRECT SERIAL NUMBER. H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -11.0/+2.0/092 | OD | Physician | 12/28/2021 | 2021 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13146350&pc=QCB | 2023826-2022-00037 | 2021-04-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -9.0/2.5/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.0/2.5/090 | OS | Physician | 12/06/2021 | 2021 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13147241&pc=QCB | 2023826-2022-00039 | 2021-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -9.0/1.5/88 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.0/1.5/88 | OS | Physician | 12/05/2021 | 2021 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13147355&pc=QCB | 2023826-2022-00041 | 2021-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -14.50/3.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON THE SAME DAY DIFFERENT SURGERY THE LENS WAS EXCHANGED FOR A LONGER LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -14.50/3.0/091 | OD | Physician | 12/06/2021 | 2021 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13161086&pc=QCB | 2023826-2022-00050 | 2021-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -8.5/1.5/077 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6)2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -8.5/1.5/077 | OS | Physician | 12/08/2021 | 2021 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13161151&pc=QCB | 2023826-2022-00045 | 2021-07-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-01-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -5.5/1.0/079 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD). ON (B)(6) 2021 THE LENS WAS IMPLANTED; REMOVED AND REPLACED INTRA-OPERATIVELY FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -5.5/1.0/079 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD). ON (B)(6) 2021 THE LENS WAS IMPLANTED; REMOVED AND REPLACED INTRA-OPERATIVELY FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -5.5/1.0/079 | OD | Physician | 12/07/2021 | 2021 | 49 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13161246&pc=QCB | 2023826-2022-00051 | 2021-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -7.0/0.5/000 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM# (B)(4). | -7.0/0.5/000 | OS | Physician | 12/08/2021 | 2021 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13167750&pc=QCB | 2023826-2022-00004 | 2021-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.50/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM# (B)(4). | -9.50/2.0/093 | OD | Physician | 12/14/2021 | 2021 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13167752&pc=QCB | 2023826-2022-00007 | 2021-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.50/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -10.50/2.5/093 | OS | Physician | 12/15/2021 | 2021 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13167753&pc=QCB | 2023826-2022-00005 | 2021-12-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.50/2.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. REPORTEDLY, EYE QUIET, VISION 6/6, PATIENT ACCEPTING 0.5 CYLINDER, PATIENT HAPPY. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.50/2.0/092 | OS | Physician | 12/15/2021 | 2021 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13167754&pc=QCB | 2023826-2022-00009 | 2021-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -8.50/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(6). | -8.50/1.0/093 | OS | Physician | 12/15/2021 | 2021 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13167755&pc=QCB | 2023826-2022-00008 | 2021-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.00/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.00/2.0/093 | OD | Physician | 12/15/2021 | 2021 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13167758&pc=QCB | 2023826-2022-00012 | 2021-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/1.5/05 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -13.00/1.5/05 | OD | Physician | 12/16/2021 | 2021 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13167759&pc=QCB | 2023826-2022-00006 | 2021-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.50/1.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -9.50/1.5/081 | OD | Physician | 12/15/2021 | 2021 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13167761&pc=QCB | 2023826-2022-00011 | 2021-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.50/1.5/114 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHINICITY: UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -8.50/1.5/114 | OS | Physician | 12/15/2021 | 2021 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13167763&pc=QCB | 2023826-2022-00015 | 2021-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -4.00/1.0/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). UNDER 21 YRS OF AGE AT DATE OF IMPLANT. CLAIM# (B)(4). | -4.00/1.0/102 | OD | Physician | 12/17/2021 | 2021 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13167764&pc=QCB | 2023826-2022-00010 | 2021-12-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.50/2.0/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -14.50/2.0/103 | OD | Physician | 12/15/2021 | 2021 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13172842&pc=QCB | 2023826-2022-00094 | 2021-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -8.0/2.0/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6)2021, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -8.0/2.0/93 | OD | Physician | 12/10/2021 | 2021 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13172843&pc=QCB | 2023826-2022-00093 | 2021-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -6.0/1.5/89 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -6.0/1.5/89 | OS | Physician | 12/10/2021 | 2021 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13172845&pc=QCB | 2023826-2022-00095 | 2021-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF 12.0/1.5/78 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REMOVED DUE TO LOW VAULT. A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED AT A LATER DATE AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/78 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED LONGER LENGTH DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. (B)(4). Manufacturer Narrative: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.0/1.5/78 | OS | Physician | 12/10/2021 | 2021 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13172867&pc=QCB | 2023826-2022-00096 | 2021-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -10.5/1.5/90 (SPHERE/ CYLINDER/ AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -10.5/1.5/90 | OS | Physician | 12/10/2021 | 2021 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13173033&pc=QCB | 2023826-2022-00076 | 2021-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -7.0/1.0/172 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND A DIFFERENT AXIS DUE TO LOW VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.0/1.0/172 | OS | Physician | 12/30/2021 | 2021 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13173075&pc=QCB | 2023826-2022-00058 |  | Injury | STAAR SURGICAL COMPANY | 2022-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -02.50/+3.0/076 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS). THE LENS WAS REPORTED AS LENS ROTATION AND WAS REPOSITIONED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. (B)(4). | -02.50/+3.0/076 | OS | Physician | 12/10/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13173107&pc=QCB | 2023826-2022-00055 | 2021-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.50/+1.0/126 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD). THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.50/+1.0/126 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD). THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.0/115 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD) . THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. D4: SERIAL #: (B)(6). H4: DEVICE MANUFACTURE DATE: 01-SEP-2021. CLAIM # (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00/+1.0/115 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.50/+1.0/126 | OD | Physician | 12/09/2021 | 2021 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13182532&pc=QCB | 2023826-2022-00031 | 2021-12-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.5/+1.5/089 (SPHERE/CYLINDER/AXIS) GOT STUCK IN THE CARTRIDGE/INJECTION SYSTEM.. THIS OCCURRED ON 14-(B)(6) 2021. THERE WAS PATIENT CONTACT WITH THE LENS BUT NO PATIENT INJURY AND THE LENS WAS NOT IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AND THE PROBLEM WAS RESOLVED. REPORTEDLY, "FIRST LENS GOT STUCK IN INJECTOR AND GOT DAMAGED. SECOND LENS IMPLANTATION WAS UNEVENTFUL."Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.5/+1.5/089 (SPHERE/CYLINDER/AXIS) GOT STUCK IN THE CARTRIDGE/INJECTION SYSTEM.. THIS OCCURRED ON (B)(6) 2021. THERE WAS PATIENT CONTACT WITH THE LENS BUT NO PATIENT INJURY AND THE LENS WAS NOT IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS DEVICE. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AND THE PROBLEM WAS RESOLVED. REPORTEDLY, "FIRST LENS GOT STUCK IN INJECTOR AND GOT DAMAGED. SECOND LENS IMPLANTATION WAS UNEVENTFUL." Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. . CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -7.5/+1.5/089 | NA | Physician | 12/17/2021 | 2021 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13182564&pc=QCB | 2023826-2022-00032 | 2021-11-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.0/+1.5/089 (SPHERE/CYLINDER/AXIS) TORE DURING INJECTION INTO THE PATIENT'S RIGHT (OD) EYE. THIS OCCURRED ON (B)(6) 2021. THERE WAS PATIENT CONTACT WITH THE LENS BUT THE LENS WAS NOT FULLY IMPLANTED; NO PATIENT INJURY REPORTED. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AND THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "INITIAL LENS GOT DAMAGED DURING INJECTION INTO THE EYE. SECOND LENS IMPLANTATION WAS SMOOTH AND PATIENT IS DOING FINE."Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -7.0/+1.5/089 (SPHERE/CYLINDER/AXIS) TORE DURING INJECTION INTO THE PATIENT'S RIGHT (OD) EYE. THIS OCCURRED ON (B)(6) 2021. THERE WAS PATIENT CONTACT WITH THE LENS BUT THE LENS WAS NOT FULLY IMPLANTED; NO PATIENT INJURY REPORTED. AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AND THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, "INITIAL LENS GOT DAMAGED DURING INJECTION INTO THE EYE. SECOND LENS IMPLANTATION WAS SMOOTH AND PATIENT IS DOING FINE." Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -7.0/+1.5/089 | OD | Physician | 12/17/2021 | 2021 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13182946&pc=QCB | 2023826-2022-00100 | 2021-11-18 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -16.00/+1.50/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULTING. IT WAS ALSO REPORTED THAT THE PATIENT IS PREGNANT AND WILL WOULD LIKE TO DO A REPLACEMENT SURGERY NEXT YEAR. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -16.00/+1.50/95 | OD | Physician | 12/22/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13182948&pc=QCB | 2023826-2022-00101 | 2021-11-18 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -16.00/+2.00/73 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULTING. IT WAS ALSO REPORTED THAT THE PATIENT IS PREGNANT AND WILL WOULD LIKE TO DO A REPLACEMENT SURGERY NEXT YEAR. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -16.00/+2.00/73 | OS | Physician | 12/22/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13183004&pc=QCB | 2023826-2022-00102 | 2021-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE WHILE IMPLANTING A 12.6MM VTICMO12.6; -13.5/1.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021; IT HAD TORE/BROKE DURING INJECTION/DELIVERY. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A BACK UP LENS. THERE WAS PATIENT CONTACT WITH NO INJURY. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -13.5/1.5/091 | OS | Physician | 01/05/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13183168&pc=QCB | 2023826-2022-00056 | 2021-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.50/+0.5/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS). THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+0.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS). THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. D4: SERIAL #: (B)(6). H4: DEVICE MANUFACTURE DATE: 27-AUG-2021. CLAIM # (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+0.5/092 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021 . THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). | -11.50/+0.5/085 | OS | Physician | 12/09/2021 | 2021 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13192750&pc=QCB | 2023826-2022-00104 | 2021-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -9.5/1.0/144 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -9.5/1.0/144 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.5/1.0/144 | OD | Physician | 12/14/2021 | 2021 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13192992&pc=QCB | 2023826-2022-00105 | 2021-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -9.5/1.0/168 (SPHERE/ CYLINDER/A XIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -9.5/1.0/168 | OD | Physician | 12/14/2021 | 2021 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13193203&pc=QCB | 2023826-2022-00106 | 2021-11-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -14.5/2.5/067 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION :UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -14.5/2.5/067 | OS | Physician | 12/15/2021 | 2021 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13206130&pc=QCB | 2023826-2022-00084 | 2021-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.50/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.50/1.0/091 | OS | Physician | 12/22/2021 | 2021 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13206175&pc=QCB | 2023826-2022-00082 | 2021-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/4.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.50/4.0/093 | OD | Physician | 12/22/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13206176&pc=QCB | 2023826-2022-00090 | 2021-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -5.50/1.50/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULT. A LONGER LENGTH LENS WAS IMPLANTED ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PT INFO: UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -5.50/1.50/098 | OS | Physician | 12/23/2021 | 2021 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13206200&pc=QCB | 2023826-2022-00088 | 2021-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.00/1.50/091 (SPHERE/ CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -10.00/1.50/091 | OD | Physician | 12/23/2021 | 2021 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13206204&pc=QCB | 2023826-2022-00083 | 2021-12-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.50/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -14.50/2.5/092 | OD | Physician | 12/22/2021 | 2021 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13206217&pc=QCB | 2023826-2022-00081 | 2021-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/3.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -10.50/3.0/086 | OS | Physician | 12/22/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13206237&pc=QCB | 2023826-2022-00086 | 2021-10-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.0/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. ON (B)(6) 2021, A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.0/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. ON (B)(6) 2021, A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -7.0/2.5/091 | OS | Physician | 12/22/2021 | 2021 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13211571&pc=QCB | 2023826-2022-00126 | 2021-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11/+2/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED LENS ROTATION; WHICH THE SURGEON REPOSITIONED ON (B)(6) 2021 BUT THIS DID NOT RESOLVE THE PROBLEM. THEN ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -11/+2/095 | OD | Physician | 01/05/2022 | 2022 | 46 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13211622&pc=QCB | 2023826-2022-00125 | 2021-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -16/+3.5/40 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LENS ROTATION; WHICH THE SURGEON REPOSITIONED ON (B)(6) 2021 BUT THIS DID NOT RESOLVE THE PROBLEM. THEN ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND IT WAS NOTED THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -16/+3.5/40 | OS | Physician | 01/05/2022 | 2022 | 46 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13226416&pc=QCB | 2023826-2022-00107 | 2021-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -17.0/2.5/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -17.0/2.5/90 | OD | Physician | 12/15/2021 | 2021 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13227086&pc=QCB | 2023826-2022-00108 | 2021-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -12.5/1.0/118 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.5/1.0/118 | OS | Physician | 12/15/2021 | 2021 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13227139&pc=QCB | 2023826-2022-00109 | 2021-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -10.5/1.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.5/1.0/090 | OS | Physician | 12/15/2021 | 2021 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13231233&pc=QCB | 2023826-2022-00110 | 2021-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.5/72 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON THE SAME DAY DIFFERENT SURGERY THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -9.0/1.5/72 | OD | Physician | 12/15/2021 | 2021 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13231887&pc=QCB | 2023826-2022-00112 | 2021-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -15.50/2.0/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS PATIENT- RELATED FACTOR DUE TO "ASIAN PATIENT, ANATOMICALLY INDISPOSED." Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -15.50/2.0/90 | OS | Physician | 12/15/2021 | 2021 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13232196&pc=QCB | 2023826-2022-00113 | 2021-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -8.0/1.5/113 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -8.0/1.5/113 | OS | Physician | 12/17/2021 | 2021 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13233194&pc=QCB | 2023826-2022-00065 |  | Injury | STAAR SURGICAL COMPANY | 2022-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -06.00/+4.5/072 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. A REFRACTIVE TREATMENT WAS PERFORMED. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. HEALTH IMPACT - ADDITIONAL SURGERY: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -06.00/+4.5/072 | OS | Physician | 12/14/2021 | 2021 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13233202&pc=QCB | 2023826-2022-00063 | 2021-11-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+3.0/008 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.00/+3.0/008 | OS | Physician | 12/13/2021 | 2021 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13233208&pc=QCB | 2023826-2022-00064 | 2021-08-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00/+1.5/091 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS WAS INSERTED UPSIDE DOWN INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -08.00/+1.5/091 | OS | Physician | 12/13/2021 | 2021 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13255769&pc=QCB | 2023826-2022-00071 | 2021-12-15 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Difficult to Fold, Unfold or Collapse; Unintended Movement | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.00/+1.5/109 (SPHERE/ CYLINDER/ AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS REMAINED IN THE FOLDED POSITION (WOULD NOT UNFOLD) WHEN DELIVERED INTO THE PATIENTS RIGHT EYE (OD). THE LENS WAS INSERTED AGAIN DURING THE SAME SURGERY AND AGAIN REMAINED IN THE FOLDED POSITION. THE SURGEON MANAGED TO OPEN THE LENS BUT IT WENT INTO THE EYE UPSIDE DOWN. THE LENS WAS REMOVED WITH NO PATIENT INJURY. THE SURGEON DID NOT IMPLANT ANOTHER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5 - "THE SURGERY WAS REARRANGED FOR (B)(6) HOWEVER HAD TO BE CANCELLED AGAIN DUE TO SURGEON ILLNESS. SUBSEQUENTLY THE PATIENT DECIDED AGAINST FURTHER SURGERY." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.00/+1.5/109 | OD | Physician | 12/16/2021 | 2021 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13255778&pc=QCB | 2023826-2022-00066 |  | Injury | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00/+1.0/151 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. PHACO EMULSIFICATION SURGERY WAS PERFORMED AND AN IOL LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. THE PATIENT IS ASYMPTOMATIC AND IS HAPPY WITH THE RESULTS. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND ONE HAPTIC BROKEN. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. REFRACTIVE VERIFICATION WAS PERFORMED AND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.00/+1.0/151 | OS | Physician | 12/15/2021 | 2021 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13255795&pc=QCB | 2023826-2022-00069 | 2021-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -03.00/+1.0/143 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FIBER ON THE LENS SURFACE. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -03.00/+1.0/143 | OS | Physician | 12/15/2021 | 2021 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13255815&pc=QCB | 2023826-2022-00073 | 2021-12-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -13.50/+0.5/147 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6: TYPE OF INVESTIGATION - 4117. B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -13.50/+0.5/147 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6)2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. H6: HEALTH IMPACT - 2199. CLAIM # (B)(4). Manufacturer Narrative: B5 - REPORTEDLY, "THE PATIENT IS HAPPY WITH THE WAY IT IS. HE DO NOT WANT AN EXCHANGE." PROBLEM RESOLVED. CLAIM# (B)(4). | -13.50/+0.5/147 | OS | Physician | 12/17/2021 | 2021 | 50 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13255816&pc=QCB | 2023826-2022-00072 | 2021-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -15.00/+1.0/153 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: H6: TYPE OF INVESTIGATION - 4117. B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -15.00/+1.0/153 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON 10-DEC-2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. H6: HEALTH IMPACT - 2199. CLAIM # (B)(4). Manufacturer Narrative: B5 - THE LENS WAS REPLACED ON (B)(6) 2022 WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. PATIENT IS HAPPY. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -15.00/+1.0/153 | OD | Physician | 12/17/2021 | 2021 | 50 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13256337&pc=QCB | 2023826-2022-00150 | 2021-04-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -18.0/+3.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -18.0/+3.0/094 | OS | Physician | 12/23/2021 | 2021 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13256356&pc=QCB | 2023826-2022-00149 | 2021-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATES THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.5/+4.5/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -10.5/+4.5/086 | OD | Physician | 12/23/2021 | 2021 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13256373&pc=QCB | 2023826-2022-00146 | 2019-04-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/106 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY DUE TO EXCESSIVE VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B3- (B)(6) 2019. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. DATE OF EVENT: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -8.5/+1.5/106 | OS | Physician | 12/22/2021 | 2021 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13256531&pc=QCB | 2023826-2022-00158 | 2021-12-15 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -11.00/+3.5/103 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THERE WAS PATIENT CONTACT; IT WAS REPORTED THE LENS WAS IMPLANTED UPSIDE DOWN AND WOULDN'T OPEN UP; "HAPPENED ON 3 OCCASIONS AND BACK UP LENS USED BUT HAD TO BE MANIPULATED IN ORDER TO OPEN UP". ON THE SAME DATE THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH LENS BUT THE PROBLEM WAS NOT RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: D4 - UNIQUE IDENTIFIER (UDI) # (B)(4) SHOULD BE REMOVED FROM INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: B1 - CORRECTED TO PRODUCT PROBLEM (E.G., DEFECTS/MALFUNCTIONS) IN INITIAL MDR. B2 - "REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE" SHOULD BE REMOVED FROM INITAL MDR. H1 - CORRECTED TO MALFUNCTION IN INITIAL MDR. CLAIM #(B)(4). | -11.00/+3.5/103 | OS | Physician | 12/17/2021 | 2021 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13262008&pc=QCB | 2023826-2022-00062 |  | Injury | STAAR SURGICAL COMPANY | 2022-01-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON PLANS TO EXPLANT AND EXCHANGE AN ICL IMPLANTABLE COLLAMER LENS. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -13.00/+3.0/095, INTO THE PATIENTS LEFT EYE (OS). THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | -13.00/+3.0/095 | OS | Other Health Care Professional | 12/17/2021 | 2021 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13266332&pc=QCB | 2023826-2022-00134 | 2021-12-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.50/2.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -13.50/2.5/090 | OD | Physician | 12/30/2021 | 2021 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13266336&pc=QCB | 2023826-2022-00132 | 2021-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -16.50/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -16.50/2.5/093 | OD | Physician | 12/30/2021 | 2021 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13266340&pc=QCB | 2023826-2022-00135 | 2021-12-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.50/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -14.50/1.5/093 | OD | Physician | 12/30/2021 | 2021 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13266533&pc=QCB | 2023826-2022-00131 | 2021-11-24 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 12.6MM VTICMO 12.6, -9.00/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS FROM THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. A SHORTER LENGTH LENS WAS IMPLANTED IN A SECOND SURGERY ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.00/1.0/085 | OS | Physician | 12/29/2021 | 2021 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13286481&pc=QCB | 2023826-2022-00181 |  | Injury | STAAR SURGICAL COMPANY | 2022-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS, -13.00/+4.5/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS). THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REPOSITIONED AT 7 DAYS POST-OP. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -13.00/+4.5/089 | OS | Physician | 12/20/2021 | 2021 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13286577&pc=QCB | 2023826-2022-00195 |  | Injury | STAAR SURGICAL COMPANY | 2022-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: IT WAS REPORTED THAT A 12.6MM TMICL12.6; -12.50/3.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO A PATIENT'S EYE ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE THE "LENS TOO LARGE". ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). | -12.50/3.0/093 | NA | Other Health Care Professional | 12/27/2021 | 2021
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13286579&pc=QCB | 2023826-2022-00237 | 2021-12-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -6.5/+1.0/105 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. IT WAS NOTED THE LENS HAD TORE DURING INJECTION INTO THE EYE. THERE WAS PATIENT CONTACT WITH NO INJURY BUT IT WAS NOTED THERE IS MINOR DAMAGE TO THE HAPTICS OF THE LENS. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. THERE IS A PLANNED LENS EXCHANGED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -6.5/+1.0/105 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. IT WAS NOTED THE LENS HAD TORE DURING INJECTION INTO THE EYE. THERE WAS PATIENT CONTACT WITH NO INJURY BUT IT WAS NOTED THERE IS MINOR DAMAGE TO THE HAPTICS OF THE LENS. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. THERE IS A PLANNED LENS EXCHANGED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED TO ADVERSE EVENT IN INITIAL MDR. ADDITIONAL INFORMATION: " LENS WAS EXCHANGED ON (B)(6) 2022 FOR ANOTHER SAMEMODEL/ LENGTH LENS. PROBLEM WAS RESOLVED. REPORTEDLY, "ALL GOOD. PATIENT IS HAPPY." (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. CORRECTED TO SERIOUS INJURY IN INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -6.5/+1.0/105 | OS | Physician | 12/20/2021 | 2021 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13286580&pc=QCB | 2023826-2022-00178 | 2021-12-21 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -6.0/+6.0/99 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE PATIENT EXPERIENCED EXCESSIVE VAULTING BUT IT WAS REPORTED THAT THE PATIENT FEELS GOOD AND THE SURGEON DECIDED NO SURGICAL INTERVENTION IS PLANNED AND WILL BE MONITORED OVER THE NEXT MONTHS. LENS REMAINS IMPLANTED.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -6.0/+6.0/99 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE PATIENT EXPERIENCED EXCESSIVE VAULTING BUT IT WAS REPORTED THAT THE PATIENT FEELS GOOD AND THE SURGEON DECIDED NO SURGICAL INTERVENTION IS PLANNED AND WILL BE MONITORED OVER THE NEXT MONTHS. LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -6.0/+6.0/99 | OD | Physician | 12/22/2021 | 2021 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13286582&pc=QCB | 2023826-2022-00208 | 2021-11-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -11/+1.5/78 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING WITH LENS ROTATION. THIS RESOLVED THE PROBLEM.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -11/+1.5/78 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULTING WITH LENS ROTATION. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -11/+1.5/78 | OS | Physician | 01/13/2022 | 2022 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13290513&pc=QCB | 2023826-2022-00209 | 2021-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -14.5/2.5/087(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: E1- INITIAL REPORTER NAME AND ADDRESS CORRECTED TO CHINA. G2- REPORT SOURCE CORRECTED TO CHINA. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -14.5/2.5/087 | OD | Physician | 12/23/2021 | 2021 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13295448&pc=QCB | 2023826-2022-00211 | 2021-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -9.5/1.0/055 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6)2021, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.5/1.0/055 | OS | Physician | 12/23/2021 | 2021 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13295567&pc=QCB | 2023826-2022-00212 | 2021-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -5.0/0.5/105 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -5.0/0.5/105 | OD | Physician | 12/23/2021 | 2021 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13296109&pc=QCB | 2023826-2022-00207 | 2022-01-03 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Unintended Movement | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -11.50/0.50/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS IMPLANTED UPSIDE DOWN. INTRAOPERATIVELY THE LENS WAS REMOVED AND DURING REMOVAL THE LENS HAD TORE. ALSO INTRAOPERATIVELY A REPLACEMENT LENS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THERE WAS PATIENT CONTACT WITH INJURY OF CORNEAL ENDOTHELIAL TOUCH. LATER THE DOCTOR REPORTED THE PATIENT IS RECOVERING WELL; CORNEA IS CLEARING UP AND VISION IS 6/6. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND A PATIENT RELATED FACTOR; THE PATIENT KEPT ROLLING THE EYE UPWARD SO IT WOULD NOT COVER THE POSITION OF THE LENS BEFORE IT FULLY FLIPPED OVER.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -11.50/0.50/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS IMPLANTED UPSIDE DOWN. INTRAOPERATIVELY THE LENS WAS REMOVED AND DURING REMOVAL THE LENS HAD TORE. ALSO INTRAOPERATIVELY A REPLACEMENT LENS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THERE WAS PATIENT CONTACT WITH INJURY OF CORNEAL ENDOTHELIAL TOUCH. LATER THE DOCTOR REPORTED THE PATIENT IS RECOVERING WELL; CORNEA IS CLEARING UP AND VISION IS 6/6. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND A PATIENT RELATED FACTOR; THE PATIENT KEPT ROLLING THE EYE UPWARD SO IT WOULD NOT COVER THE POSITION OF THE LENS BEFORE IT FULLY FLIPPED OVER. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -11.50/0.50/092 | OS | Physician | 01/02/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13296133&pc=QCB | 2023826-2022-00240 | 2022-01-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -14.50/3.0/109 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. INTRAOPERATIVELY THE LENS WAS REMOVED DUE TO THE LENS WAS IMPLANTED UPSIDE DOWN. THERE WAS NO PATIENT INJURY. ON (B)(6) 2022 A REPLACEMENT LENS OF THE SAME MODEL / LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4). Manufacturer Narrative: G3: CORRECTED TO 10-JAN-2022 IN INITIAL MDR. H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | -14.50/3.0/109 | OD | Physician | 01/10/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13297476&pc=QCB | 2023826-2022-00187 | 2021-12-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.50/+1.5/004 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.50/+1.5/004 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR B5 - "THE LENS WAS REPLACED WITH A LONGER LENGTH/ SAME MODEL LENS ON (B)(6) 2022. THIS RESOLVED THE PROBLEM." SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - CORRECTED TO HEALTH EFFECT - IMPACT CODE: 4629. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: NA. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: NA. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -08.50/+1.5/004 | OS | Physician | 12/22/2021 | 2021 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13297561&pc=QCB | 2023826-2022-00180 |  | Injury | STAAR SURGICAL COMPANY | 2022-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -14.50/+2.0/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS). THE LENS WAS EXPLANTED DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE (VA WAS 20/400). ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND ONE HAPTIC TORN. THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -14.50/+2.0/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS). THE PATIENT HAD NORMAL EYE PRESSURE, GOOD VAULT BUT IS +5.50 AND IS SEEING 20/400. THE LENS WAS EXPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLATED: UNK. EXPLANTED: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -14.50/+2.0/096 | OS | Other Health Care Professional | 12/20/2021 | 2021 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13308090&pc=QCB | 2023826-2022-00214 | 2021-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -5.5/1.0/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON THE SAME DAY DIFFERENT SURGERY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -5.5/1.0/90 | OS | Physician | 12/23/2021 | 2021 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13308091&pc=QCB | 2023826-2022-00216 | 2021-09-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -5.5/1.5/94 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -5.5/1.5/94 | OD | Physician | 12/23/2021 | 2021 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13308893&pc=QCB | 2023826-2022-00218 | 2021-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.0/1.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON THE SAME DAY DIFFERENT SURGERY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.0/1.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON THE SAME DAY DIFFERENT SURGERY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: B3- DATE OF EVENT: (B)(6) 2021. CLAIM# (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. DATE OF EVENT: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.0/1.0/093 | OS | Physician | 12/23/2021 | 2021 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13311395&pc=QCB | 2023826-2022-00243 | 2022-01-11 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT WHILE ATTEMPTING TO LOAD A 12.6MM TMICL12.6; -12.50/2.5/105 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS; THE LENS HAD TORE AFTER LOADING INTO THE CARTRIDGE. THIS OCCURRED ON (B)(6) 2022. THERE WAS NO PATIENT CONTACT. A BACK UP LENS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -12.50/2.5/105 | NA | Other Health Care Professional | 01/13/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13311414&pc=QCB | 2023826-2022-00193 | 2017-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+2.5/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+2.5/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2017 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5- PER EMAIL COMMUNICATION ON (B)(6)2022 IT WAS REPORTED THAT THE LENS (T253011) WAS NOT EXPLANTED AND THERE WAS NO ISSUE WITH THIS SN. THIS EVENT WAS REPORTED IN ERROR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4). | -11.50/+2.5/091 | OS | Physician | 12/23/2021 | 2021 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13315661&pc=QCB | 2023826-2022-00215 | 2021-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -4.0/0.5/121 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -4.0/0.5/121 | OD | Physician | 12/23/2021 | 2021 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13317680&pc=QCB | 2023826-2022-00206 | 2021-11-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS -15.0/+2.0/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY THE REPLACEMENT LENS WAS IMPLANTED VERTICALLY. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -15.0/+2.0/085 | OD | Physician | 12/31/2021 | 2021 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13317686&pc=QCB | 2023826-2022-00199 | 2021-12-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -9.5/+2.0/007 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE VERTICALLY IMPLANTED LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6- WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.5/+2.0/007 | OD | Physician | 12/27/2021 | 2021 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13317687&pc=QCB | 2023826-2022-00201 | 2021-11-11 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT, DURING LOADING A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.0/+2.5/101 (SPHERE/ CYLINDER/AXIS), IT TORE. THIS WAS DISCOVERED DURING LOADING AND THERE WAS NO PATIENT CONTACT WITH THIS LENS. THIS OCCURRED ON (B)(6) 2021. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN, THE SURGEON STATES, "I FELT RESISTANCE WHILE LOADING THE INJECTOR." AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AT A LATER DATE. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.0/+2.5/101 | NA | Physician | 12/28/2021 | 2021 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13318275&pc=QCB | 2023826-2022-00224 | 2021-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -8.0/1.0/163 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -8.0/1.0/163 | OD | Physician | 12/30/2021 | 2021 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13318428&pc=QCB | 2023826-2022-00192 | 2021-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.50/+2.0/101 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS REPLACED WITH A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE RESIDUAL ASTIGMATISM WAS CORRECTED BY LRI. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -12.50/+2.0/101 | OD | Physician | 12/23/2021 | 2021 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13319437&pc=QCB | 2023826-2022-00227 | 2021-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.0/2.0/92 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.0/2.0/92 | OD | Physician | 12/31/2021 | 2021 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13319621&pc=QCB | 2023826-2022-00228 | 2021-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.0/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.0/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.0/2.0/089 | OS | Physician | 12/31/2021 | 2021 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13320020&pc=QCB | 2023826-2022-00229 | 2021-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -11.5/0.5/110 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH SPHERICAL LENS DUE TO EXCESSIVE VAULT. "THE PATIENT'S ASTIGMATISM IS RELATIVELY LOW, AND DUE TO INVENTORY AND TIME REASONS, THE SURGEON DECIDED TO CHANGE TICL TO ICL." THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/0.5/110 | OS | Physician | 01/04/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13321362&pc=QCB | 2023826-2022-00231 | 2021-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -12.5/1.5/128 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.5/1.5/128 | OD | Physician | 01/05/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13334846&pc=QCB | 2023826-2022-00163 | 2021-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/1.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/1.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -10.50/1.0/088 | OD | Physician | 01/04/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13334853&pc=QCB | 2023826-2022-00164 | 2021-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.00/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -9.00/1.0/090 | OS | Physician | 01/04/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13334854&pc=QCB | 2023826-2022-00167 | 2021-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.5/1.5/172 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -17.5/1.5/172 | OD | Physician | 01/04/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13334855&pc=QCB | 2023826-2022-00162 | 2021-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.5/1.0/068 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/1.0/068 | OD | Physician | 01/04/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13334856&pc=QCB | 2023826-2022-00168 | 2020-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -11.00/2.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.00/2.5/090 | OS | Physician | 01/07/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13334858&pc=QCB | 2023826-2022-00166 | 2021-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -6.00/1.5/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -6.00/1.5/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -6.00/1.5/084 | OS | Physician | 01/04/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13350363&pc=QCB | 2023826-2022-00278 | 2021-12-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -9.00/+3.0/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULTING WITH ROTATION. THE LENS WAS EXPLANTED ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -9.00/+3.0/91 | OD | Physician | 01/12/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13350401&pc=QCB | 2023826-2022-00299 | 2021-12-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -11.5/1.0/94 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT IMPLANTED VERTICALLY ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -11.5/1.0/94 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT IMPLANTED VERTICALLY ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -11.5/1.0/94 | OD | Physician | 01/24/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13350415&pc=QCB | 2023826-2022-00279 | 2021-12-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -8.50/+2.0/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULTING WITH ROTATION. THE LENS WAS EXPLANTED ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -8.50/+2.0/95 | OS | Physician | 01/12/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13350448&pc=QCB | 2023826-2022-00297 | 2021-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6; -11.0/2.5/070 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING AND ON (B)(6) 2021 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. THEN ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6; -11.0/2.5/070 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING AND ON (B)(6) 2021 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. THEN ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -11.0/2.5/070 | OS | Physician | 01/21/2022 | 2022 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13350519&pc=QCB | 2023826-2022-00296 | 2021-12-03 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-01-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT AFTER LOADING A 13.2MM VTICMO13.2; -7.5/1.5/179 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS; IT WAS NOTICED THE LENS HAD A CRACK ON THE LENS. THIS OCCURRED ON (B)(6) 2021. THE LENS WAS NOT IMPLANTED OR USED. ON (B)(6) 2022 A REPLACEMENT LENS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND DEBRIS THROUGHOUT THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# : (B)(4). | -7.5/1.5/179 | NA | Physician | 01/18/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13358321&pc=QCB | 2023826-2022-00234 | 2021-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -18.0/2.5/084 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -18.0/2.5/084 | OS | Physician | 01/05/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13358949&pc=QCB | 2023826-2022-00236 | 2021-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -4.50/2.5/126 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS OTHER. THE LENS DID NOT FAIL TO PERFORM AS INTENDED. "WRONG DATA TYPED BY THE USER RESULTING IN WRONG REFRACTION AND WRONG ICL SIZE, WE WERE SURE AFTER EXPLANTING THE ICL AND REDOING ALL THE TESTS FOR THE PATIENT." Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -4.50/2.5/126 | OD | Physician | 01/06/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13385983&pc=QCB | 2023826-2022-00350 | 2021-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.5/+1.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY THERE IS NO PLAN TO IMPLANT A REPLACEMENT LENS. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -10.5/+1.0/087 | OD | Physician | 12/28/2021 | 2021 | 48 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13388188&pc=QCB | 2023826-2022-00286 | 2021-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -06.00/+1.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS EXPLANTED. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (LENS WAS IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -06.00/+1.0/087 | OS | Physician | 12/30/2021 | 2021 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13389440&pc=QCB | 2023826-2022-00354 | 2021-12-15 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-01-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -6.0/2.0/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. DURING INJECTION INTO THE EYE THE LENS HAD TORE. THERE WAS PATIENT CONTACT WITH NO INJURY. INTRAOPERATIVELY THE LENS WAS REMOVED. ON (B)(6) 2022 AT REPLACEMENT LENS OF THE SAME MODEL/LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -6.0/2.0/95 | OS | Physician | 01/24/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13390295&pc=QCB | 2023826-2022-00288 | 2021-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50/+2.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -08.50/+2.0/093 | OS | Physician | 12/31/2021 | 2021 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13398729&pc=QCB | 2023826-2022-00302 | 2021-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -6.5/2.0/071 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -6.5/2.0/071 | OS | Physician | 01/09/2022 | 2022 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13401432&pc=QCB | 2023826-2022-00303 | 2021-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -14.0/1.0/100 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE LENS WAS IMPLANTED VERTICALLY AND THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENTS INFORMATION: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -14.0/1.0/100 | OD | Physician | 01/09/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13404364&pc=QCB | 2023826-2022-00304 | 2021-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -10.0/1.0/085 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -10.0/1.0/085 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -10.0/1.0/085 | OS | Physician | 01/10/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13404707&pc=QCB | 2023826-2022-00305 | 2021-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -13.5/3.0/111 (SPHERE/ CYLINDER/ AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.5/3.0/111 | OS | Physician | 01/10/2022 | 2022 | 51 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13419057&pc=QCB | 2023826-2022-00357 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-02-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT WHILE ATTEMPTING TO LOAD A 12.6MM TMICL12.6; -16.00/3.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS IT HAD TORE. IT WAS NOTED "LENS BROKE PULLING THROUGH CARTRIDGE". THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. LENS WAS NOT IMPLANTED. A BACK UP LENS WAS USED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A USER ERROR AND NOTED "FORCEPS PULLED TEAR IN LENS AS IT CAME THROUGH CARTRIDGE". Manufacturer Narrative: CLAIM#: (B)(4). | -16.00/3.5/090 | NA | Other Health Care Professional | 01/24/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13423343&pc=QCB | 2023826-2022-00249 | 2021-12-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -7.50/1.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#:(B)(4). | -7.50/1.0/095 | OD | Physician | 01/09/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13423344&pc=QCB | 2023826-2022-00248 | 2021-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/1.5/069 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.00/1.5/069 | OD | Physician | 01/09/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13423346&pc=QCB | 2023826-2022-00255 | 2021-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.50/1.0/114 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -8.50/1.0/114 | OD | Physician | 01/10/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13423347&pc=QCB | 2023826-2022-00252 | 2021-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.00/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.00/2.0/089 | OD | Physician | 01/09/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13423348&pc=QCB | 2023826-2022-00254 | 2021-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -5.5/1.5/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: (B)(4). CLAIM#: (B)(4). | -5.5/1.5/079 | OS | Physician | 01/10/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13423350&pc=QCB | 2023826-2022-00251 | 2021-12-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/2.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED VERTICALLY WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.50/2.0/081 | OS | Physician | 01/09/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13423351&pc=QCB | 2023826-2022-00250 | 2021-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/2.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -11.00/2.0/085 | OS | Physician | 01/09/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13423353&pc=QCB | 2023826-2022-00262 | 2021-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/2.0/091 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM# (B)(4). | -10.50/2.0/091 | OS | Physician | 01/11/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13423355&pc=QCB | 2023826-2022-00258 | 2021-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -07.50/1.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#:(B)(4). | -07.50/1.0/094 | OS | Physician | 01/10/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13423357&pc=QCB | 2023826-2022-00256 | 2021-11-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.50/2.5/066 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.50/2.5/066 | OS | Physician | 01/10/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13423358&pc=QCB | 2023826-2022-00260 | 2021-10-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.50/2.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#:(B)(4). | -9.50/2.0/086 | OD | Physician | 01/11/2022 | 2022 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13423359&pc=QCB | 2023826-2022-00259 | 2021-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.00/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#:(B)(4). | -7.00/2.0/091 | OD | Physician | 01/10/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13423360&pc=QCB | 2023826-2022-00263 | 2021-11-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -10.50/1.0/090 | OS | Physician | 01/11/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13423364&pc=QCB | 2023826-2022-00264 | 2021-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.50/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#:(B)(4). | -8.50/1.5/092 | OS | Physician | 01/11/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13423366&pc=QCB | 2023826-2022-00267 | 2019-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -13.00/3.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# :(B)(4). | -13.00/3.0/095 | OD | Physician | 01/12/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13423369&pc=QCB | 2023826-2022-00272 | 2021-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.50/2.5/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#:(B)(4). | -12.50/2.5/098 | OD | Physician | 01/12/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13423371&pc=QCB | 2023826-2022-00273 | 2021-12-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.00/1.5/151 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#:(B)(4). | -12.00/1.5/151 | OS | Physician | 01/12/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13423373&pc=QCB | 2023826-2022-00270 | 2021-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -13.00/2.0/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) /2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -13.00/2.0/071 | OD | Physician | 01/12/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13423376&pc=QCB | 2023826-2022-00275 | 2021-12-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/4.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -11.0/4.5/088 | OS | Physician | 01/12/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13423377&pc=QCB | 2023826-2022-00271 | 2021-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -15.50/1.50/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -15.50/1.50/090 | OS | Physician | 01/12/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13423382&pc=QCB | 2023826-2022-00276 | 2021-12-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.00/1.0/064 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -17.00/1.0/064 | OD | Physician | 01/12/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13426913&pc=QCB | 2023826-2022-00376 | 2021-11-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8.0/+2.5/93 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULTING WITH ROTATION. ON (B)(6) 2021 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: IT WAS LATER INDICATED THAT ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. "PATIENT IS HAPPY". CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. B5 - IT WAS LATER REPORTED THAT THE PROBLEM WAS NOT RESOLVED. PATIENT NOW HAS REFRACTIVE CHANGE OVERTIME WITH THE REPLACEMENT LENS. SEE MFR# 2023826-2022-02627 FOR REPLACEMENT LENS CLAIM. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -8.0/+2.5/93 | OS | Physician | 01/18/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13430072&pc=QCB | 2023826-2022-00289 | 2021-12-29 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00/+0.5/140 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. THE LENS WILL BE MONITORED OVER THE NEXT MONTHS. THE CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00/+0.5/140 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. THE LENS WILL BE MONITORED OVER THE NEXT MONTHS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5 - "THE LENS WILL REMAIN IMPLANTED. SURGEON DECIDED NO TO DO ANY FURTHER INTERVENTIONS. PATIENT IS HAPPY." SHOULD BE ADDED TO INITAL MDR. CLAIM #: (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: NA. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: NA. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -09.00/+0.5/140 | OD | Physician | 01/03/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13430078&pc=QCB | 2023826-2022-00290 | 2021-12-29 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00/+1.0/042 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. THE LENS WILL BE MONITORED OVER THE NEXT MONTHS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. EXPLANT DATE: NA. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: B5 - "THE LENS WILL REMAIN IMPLANTED. SURGEON DECIDED NO TO DO ANY FURTHER INTERVENTIONS. PATIENT IS HAPPY." SHOULD BE ADDED TO INITAL MDR. CLAIM #: (B)(4). | -09.00/+1.0/042 | OS | Physician | 01/03/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13447136&pc=QCB | 2023826-2022-00394 | 2022-01-09 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -12/+1.0/94 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS IMPLANTED UPSIDE DOWN; AND INTRAOPERATIVELY THE LENS WAS REMOVED BUT IT HAD TORE. THERE WAS NO PATIENT INJURY. ON (B)(6) 2022 A REPLACEMENT LENS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "UPSIDE OPENING OF ICL".Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -12/+1.0/94 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS IMPLANTED UPSIDE DOWN; AND INTRAOPERATIVELY THE LENS WAS REMOVED BUT IT HAD TORE. THERE WAS NO PATIENT INJURY. ON (B)(6) 2022 A REPLACEMENT LENS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "UPSIDE OPENING OF ICL". Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -12/+1.0/94 | OD | Physician | 01/18/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13447169&pc=QCB | 2023826-2022-00291 | 2021-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -09.00/+2.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (LENS WAS IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -09.00/+2.0/088 | OS | Physician | 01/05/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13447212&pc=QCB | 2023826-2022-00292 | 2021-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -06.00/+3.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (LENS WAS IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL . VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -06.00/+3.0/093 | OD | Physician | 01/05/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13459221&pc=QCB | 2023826-2022-00391 | 2021-12-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT, DURING LOADING A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.5/+1.5/093 (SPHERE/CYLINDER/AXIS), IT TORE. THIS WAS DISCOVERED DURING LOADING AND THERE WAS NO PATIENT CONTACT WITH THE LENS. THIS OCCURRED ON (B)(6) 2021. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN: "THE LENS GOT CAUGHT IN THE CARTRIDGE AND WAS DAMAGED WHILE LOADING THE LENS INTO THE CARTRIDGE." AN ALTERNATE LENS WAS SUCCESSFULLY IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -12.5/+1.5/093 | NA | Physician | 01/12/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13459223&pc=QCB | 2023826-2022-00388 | 2021-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.0/+1.0/111 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT AXIS LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | -10.0/+1.0/111 | OS | Physician | 01/11/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13459225&pc=QCB | 2023826-2022-00383 | 2021-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -13.5/+3.5/084 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH DIFFERENT MODEL LENS DUE TO LOW VAULT AND LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -13.5/+3.5/084 | OD | Physician | 01/11/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13459234&pc=QCB | 2023826-2022-00386 | 2021-12-23 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE OLLAMER LENS, -12.5/+2.0/077 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REMOVED INTRAOPERATIVELY DUE TO EXCESSIVE VAULT AND THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN. ETHNICITY: UNKNOWN. RACE: UNKNOWN. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -12.5/+2.0/077 | OS | Physician | 01/11/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13460822&pc=QCB | 2023826-2022-00295 | 2021-12-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.50/+3.0/090 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO THE LENS TORE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -13.50/+3.0/090 | OD | Physician | 01/07/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13471830&pc=QCB | 2023826-2022-00364 | 2021-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED VERTICALLY A 12.6MM VTICMO12.6, -13.5/2.0/124 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED VERTICALLY A 12.6MM VTICMO12.6, -13.5/2.0/124 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.5/2.0/124 | OS | Physician | 01/24/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13471837&pc=QCB | 2023826-2022-00361 | 2022-01-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -6.00/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.00/1.0/091 | OD | Physician | 01/20/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13471850&pc=QCB | 2023826-2022-00360 | 2021-04-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -14.00/2.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. IN A SECOND SURGERY ON (B)(6) 2021 A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR, DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -14.00/2.5/088 | OS | Physician | 01/18/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13471853&pc=QCB | 2023826-2022-00362 | 2021-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.50/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -09.50/1.0/087 | OD | Physician | 01/20/2022 | 2022 | 17 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13476690&pc=QCB | 2023826-2022-00423 | 2021-12-29 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -16.0/1.5/074 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULTING AND INTRAOPERATIVELY THE LENS WAS REMOVED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN AND BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -16.0/1.5/074 | OS | Physician | 01/11/2022 | 2022 | 43 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13483819&pc=QCB | 2023826-2022-00322 | 2021-11-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6 VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -10.0/3.0/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT, ETHNICITY AND RACK: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.0/3.0/90 | OD | Physician | 01/11/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13488903&pc=QCB | 2023826-2022-00401 | 2021-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -06.00/+2.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO SIZING ISSUE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: SEX: UNK. WEIGHT: UNK. EHNICITY: UNK. RACE: UNK. DATE OF EVENT:UNK. CLAIM (B)(4). Manufacturer Narrative: B5: THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE EYE WAS STABLE WITH NEW IMPLANT. THE CAUSE OF THE EVENT WAS UNKNOWN. CLAIM # (B)(4). | -06.00/+2.5/094 | OS | Other Health Care Professional | 01/11/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13488916&pc=QCB | 2023826-2022-00400 | 2021-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -11.00/+3.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THERE WAS NO PAIN BUT MILD INFLAMMATION. SEE MFR. REPORT # 2023826-2022-00428 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -11.00/+3.5/084 (SPHERE/CYLINDER/AXIS), IN THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS REPLACED WIH A LONGER LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THERE WAS NO PAIN BUT MILD INFLAMMATION. THE INFLAMMATION WAS MILD CELL AND FLARE, WHICH WAS MILD AND NORMAL FOR THE SURGERY, AND HAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. CLAIM # (B)(4). | -11.00/+3.5/084 | OS | Other Health Care Professional | 01/10/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13488951&pc=QCB | 2023826-2022-00404 | 2021-12-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.00/+1.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -12.00/+1.0/095 | OD | Physician | 01/09/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13502202&pc=QCB | 2023826-2022-00324 | 2021-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6 VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -8.5/3.5/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT:UNKNOWN, ETHNICITY: UNKNOWN, RACE:UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM# (B)(4). | -8.5/3.5/095 | OS | Physician | 01/11/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13502203&pc=QCB | 2023826-2022-00325 | 2021-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6 VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -9.5/3.0/073 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -9.5/3.0/073 | OD | Physician | 01/11/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13502205&pc=QCB | 2023826-2022-00326 | 2021-12-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6 VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -9.0/0.5/71 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM# (B)(4). | -9.0/0.5/71 | OD | Physician | 01/11/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13502303&pc=QCB | 2023826-2022-00396 | 2022-01-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -7.0/1.0/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -7.0/1.0/095 | OD | Physician | 01/28/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13502321&pc=QCB | 2023826-2022-00395 | 2021-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -16.5/2.5/106 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -16.5/2.5/106 | OD | Physician | 01/28/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13502324&pc=QCB | 2023826-2022-00398 | 2021-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -16.5/1.5/110 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -16.5/1.5/110 | OS | Physician | 01/31/2022 | 2022 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13502326&pc=QCB | 2023826-2022-00397 | 2021-12-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/130 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/130 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.0/1.0/130 | OD | Physician | 01/31/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13502328&pc=QCB | 2023826-2022-00399 | 2022-01-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS -12.5/+1.5/101 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. CAUSE REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS -12.5/+1.5/101 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -12.5/+1.5/101 | OS | Physician | 02/02/2022 | 2022 | 45 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13502403&pc=QCB | 2023826-2022-00402 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -11.50/+2.5/109 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD). THE SURGEON PLANS TO EXPLANT/EXCHANGE THE LENS FOR A LONGER LENS. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: SEX: UNK. WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. DATE OF EVENT: UNK. MODEL #: EXPIRATION DATE UNK. IMPLANT DATE : UNK. EXPLANT DATE: UNK. DEVICE MANUFACTURE DATE: UNK. CLAIM # (B)(4). | -11.50/+2.5/109 | OD | Other Health Care Professional | 01/12/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13502423&pc=QCB | 2023826-2022-00418 | 2022-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.50/+4.0/102 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2021. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR B5 - "THE LENS WAS LATER EXCHANGED ON 26/07/2022 FOR A SAME LENGTH LENS DUE TO REFRACTIVE SURPRISE. THIS RESOLVED THE PROBLEM." SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE (B)(6) 2022 SHOULD BE ADDED TO INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4629 SHOULD BE ADDED TO THE INITAL MDR. CLAIM #(B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | -12.50/+4.0/102 | OD | Physician | 01/14/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13509875&pc=QCB | 2023826-2022-00329 | 2021-12-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.6 MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -9.5/2.0/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). ON (B)(6) 2021 THE LENS WAS IMPLANTED; REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -9.5/2.0/91 | OS | Physician | 01/12/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13509877&pc=QCB | 2023826-2022-00328 | 2021-12-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.6 MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -10.5/1.5/92 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD). ON (B)(6) 2021 THE LENS WAS IMPLANTED; REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -10.5/1.5/92 | OD | Physician | 01/12/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13509879&pc=QCB | 2023826-2022-00334 | 2021-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO 12.1 IMPLANTABLE COLLAMER LENS OF -10.0/0.5/085 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.0/0.5/085 | OS | Physician | 01/12/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13509880&pc=QCB | 2023826-2022-00335 | 2021-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VICMO 12.1 IMPLANTABLE COLLAMER LENS OF -12.5/2.5/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.5/2.5/086 | OS | Physician | 01/12/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13510201&pc=QCB | 2023826-2022-00336 | 2021-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -8.0/1.5/141 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -8.0/1.5/141 | OD | Physician | 01/17/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13510354&pc=QCB | 2023826-2022-00403 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -13.00/+2.5/069 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE SURGEON PLANS TO EXPLANT/EXCHANGE THE LENS FOR A LONGER LENS. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: SEX: UNK. WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. DATE OF EVENT: UNK. OTHER #: EXPIRATION DATE UNK. EXPLANT DATE: UNK. DEVICE MANUFACTURE DATE: UNK. CLAIM # (B)(4). | -13.00/+2.5/069 | OS | Other Health Care Professional | 01/12/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13510358&pc=QCB | 2023826-2022-00419 | 2022-01-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -09.00/+4.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2021. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -09.00/+4.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2021. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(6). CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.00/+4.0/089 | OS | Physician | 01/14/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13515881&pc=QCB | 2023826-2022-00414 | 2021-11-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: HOLD FOR DD 2.11 THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -10.00/+2.0/088 | OS | Physician | 01/12/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13515912&pc=QCB | 2023826-2022-00415 | 2021-12-23 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50/+1.0/081 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED ON (B)(6) 2021 WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50/+1.0/081 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED ON (B)(6) 2021 WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -08.50/+1.0/081 | OD | Physician | 01/12/2022 | 2022 | 43 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13529746&pc=QCB | 2023826-2022-00482 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-02-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -9.00/+6.0/058 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). IT WAS REPORTED THAT THE PATIENT EXPERIENCED EXCESSIVE VAULTING. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.00/+6.0/058 | OS | Physician | 01/18/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13529778&pc=QCB | 2023826-2022-00442 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-02-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -8.00/+6.0/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). IT WAS REPORTED THAT THE PATIENT EXPERIENCED EXCESSIVE VAULTING. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -8.00/+6.0/086 | OD | Physician | 01/18/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13534851&pc=QCB | 2023826-2022-00370 | 2021-11-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.00/4.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -14.00/4.0/079 | OS | Physician | 01/26/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13534853&pc=QCB | 2023826-2022-00377 | 2021-12-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.50/3.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US . CLAIM# (B)(4). | -7.50/3.5/089 | OS | Physician | 01/25/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13534863&pc=QCB | 2023826-2022-00366 | 2022-01-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -7.00/1.0/020 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2022 DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -7.00/1.0/020 | OD | Physician | 01/26/2022 | 2022 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13534868&pc=QCB | 2023826-2022-00372 | 2021-12-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED VERTICALLY A 12.6MM VTICM5_12.6, -7.50/1.0/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: DATA CORRECTION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -7.50/1.0/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON 19/NOV/2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. IN A SECOND SURGERY ON (B)(6) 2022 A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. H6 HEALTH EFFECT - IMPACT CODE REPORTED AS 4629. ADDING 4627 EXPLANT. CLAIM# (B)(4). | -7.50/1.0/073 | OS | Physician | 01/27/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13534873&pc=QCB | 2023826-2022-00365 | 2021-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.50/1.5/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.50/1.5/107 | OD | Physician | 01/24/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13534885&pc=QCB | 2023826-2022-00374 | 2021-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -11.50/2.50/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPLACED VERTICALLY WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.50/2.50/092 | OS | Physician | 01/28/2022 | 2022 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13538842&pc=QCB | 2023826-2022-00434 | 2021-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.00/+2.0/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "LENS TOO BIG". Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: B3: 10-NOV-2021. CLAIM#: (B)(4). | -6.00/+2.0/92 | OD | Physician | 01/17/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13542043&pc=QCB | 2023826-2022-00441 | 2021-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -17.00/3.0/062 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED (ASSESSED ON (B)(6) 2021), THE LENS WAS REPOSITIONED ON (B)(6) 2021 AND ROTATED (B)(6) 2021, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS REPLACED LENS WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT: UNK. ETHNIITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -17.00/3.0/062 | OS | Physician | 01/21/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13542052&pc=QCB | 2023826-2022-00440 | 2021-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -18.00/2.5/119 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED (ASSESSED ON (B)(6) 2021), THE LENS WAS REPOSITIONED ON 06/DEC/2021 AND ROTATED 18/DEC/2021, BUT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS REPLACED LENS WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.00/2.5/119 | OD | Physician | 01/21/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13545934&pc=QCB | 2023826-2022-00490 | 2022-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -14.00/+1.5/151 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -14.00/+1.5/151 | OD | Physician | 01/17/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13546060&pc=QCB | 2023826-2022-00436 | 2021-12-20 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.50/1.50/74 (SPHERE/CYLINDER/AXIS), TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED AND THE PROBLEM RESOLVED.Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.50/1.50/74 (SPHERE/CYLINDER/AXIS), TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT CONTACT. THE BACKUP LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -13.50/1.50/74 | NA | Physician | 01/18/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13546403&pc=QCB | 2023826-2022-00532 | 2021-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICTED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.5/+3.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO SURGEON PREFERENCE CHANGE AND THEN NOTED "SURGEON PREFERENCE CHANGE IS REFERRING TO ICL SIZE CHANGE NOTED FROM POSTOPERATIVE ICL VAULT". THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -11.5/+3.5/089 | OD | Physician | 01/25/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13546405&pc=QCB | 2023826-2022-00533 | 2021-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICTED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.5/+3.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO SURGEON PREFERENCE CHANGE AND THEN NOTED "SURGEON PREFERENCE CHANGE IS REFERRING TO ICL SIZE CHANGE NOTED FROM POSTOPERATIVE ICL VAULT". THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -11.5/+3.5/089 | OD | Physician | 01/25/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13551157&pc=QCB | 2023826-2022-00445 | 2021-12-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -14.0/1.0/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -14.0/1.0/087 | OD | Physician | 02/07/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13551162&pc=QCB | 2023826-2022-00446 | 2022-01-20 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -5.5/1.0/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). Manufacturer Narrative: G3 DATE IN INITIAL MDR SHOULD BE 09-FEB-2022 NOT 30-JUN-2023. CLAIM # (B)(4). | -5.5/1.0/086 | OD | Physician | 02/09/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13551164&pc=QCB | 2023826-2022-00447 | 2022-01-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -12.5/1.0/073 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.5/1.0/073 | OD | Physician | 02/09/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13551175&pc=QCB | 2023826-2022-00443 | 2022-01-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -14.0/2.5/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -14.0/2.5/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -14.0/2.5/094 | OD | Physician | 02/06/2022 | 2022 | 45 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13552601&pc=QCB | 2023826-2022-00483 | 2021-12-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -10.00/+3.5/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS ROTATED AT POST-OP DAY 1 AND WAS EXPLANTED ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENS BUT THE PROBLEM WAS NOT RESOLVED, THE LENS ROTATED AGAIN. SEE MFR. REPORT # 2023826-2022-00484 FOR REPLACEMENT LENS. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A SPECIMEN CUP IN LIQUID. VISUAL INSPECTION FOUND NO VISBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.00/+3.5/091 | OS | Physician | 01/19/2022 | 2022 | 41 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13552604&pc=QCB | 2023826-2022-00484 | 2022-01-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -09.50/+3.5/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS ROTATED AT POST-OP DAY 1 AND WAS EXPLANTED ON (B)(6) 2022. THE LENS WAS REPLACED WITH A NON TORIC LENS. SEE MFR. REPORT # 2023826-2022-00483 FOR THE INITIAL LENS. THE PATIENT WILL EVENTUALLY REQUIRE PRK SURGERY FOR HIS ASTIGMATISM BECAUSE HIS EYE ANATOMY IS NOT CONDUCIVE TO AN ICL REMAINING ON AXIS. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM # (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A SPECIMEN CUP IN LIQUID. VISUAL INSPECTION FOUND NO DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -09.50/+3.5/085 | OS | Physician | 01/19/2022 | 2022 | 41 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13552650&pc=QCB | 2023826-2022-00494 | 2022-01-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -04.50/+4.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. THE EYE IS CLINICALLY WELL AND THE PATIENT WILL BE MONITORED. THE CAUSE OF THE EVENT WAS DUE TO A LONG ICL LENS.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -04.50/+4.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. THE EYE IS CLINICALLY WELL AND THE PATIENT WILL BE MONITORED. THE CAUSE OF THE EVENT WAS DUE TO A LONG ICL LENS. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: NA. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: NA. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -04.50/+4.5/090 | OS | Physician | 01/18/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13552678&pc=QCB | 2023826-2022-00493 | 2022-01-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -05.50/+5.5/117 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. THE EYE HAS NO SYMPTOMS AND THE PATIENT WILL BE MONITORED. THE CAUSE OF THE EVENT WAS DUE TO A LONG ICL LENS.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -05.50/+5.5/117 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. THE EYE HAS NO SYMPTOMS AND THE PATIENT WILL BE MONITORED. THE CAUSE OF THE EVENT WAS DUE TO A LONG ICL LENS. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: NA. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: NA. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -05.50/+5.5/117 | OD | Physician | 01/18/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13552689&pc=QCB | 2023826-2022-00495 | 2021-11-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -09.00/+1.0/172 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021 (IMPLANTED IN A VERTICAL POSITION). THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED IN A HORIZONTAL POSITION) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE CODE: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.00/+1.0/172 | OD | Physician | 01/19/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13552692&pc=QCB | 2023826-2022-00492 | 2022-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00/+2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WILL BE EXPLANTED/EXCHANGED DUE TO WRONG DATA FOR LENS CALCULATION WAS USED (CALCULATION ERROR AND MISTAKE IN THE SUBJECTIVE REFRACTION BY CLINIC). THE LENS REMAINED IMPLANTED. CAUSE OF THE EVENT WAS DUE TO USER ERROR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00/+2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WILL BE EXPLANTED/EXCHANGED DUE TO WRONG DATA FOR LENS CALCULATION WAS USED (CALCULATION ERROR AND MISTAKE IN THE SUBJECTIVE REFRACTION BY CLINIC). THE LENS REMAINED IMPLANTED. CAUSE OF THE EVENT WAS DUE TO USER ERROR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5- PER ADDITIONAL INFORMATION RECEIVED ON 10- MAY-2022: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00/+2.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED FOR A LENS OF THE SAME MODEL/LENGTH BUT DIFFERENT POWER ON (B)(6) 2022 DUE TO USING THE WRONG DATA FOR LENS CALCULATION (CALCULATION ERROR AND MISTAKE IN THE SUBJECTIVE REFRACTION BY CLINIC). THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS DUE TO USER ERROR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -09.00/+2.0/092 | OS | Physician | 01/18/2022 | 2022 | 48 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13552724&pc=QCB | 2023826-2022-00497 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS RIGHT EYE (OD). THE LENS WAS REPORTED AS LENS ROTATION AND REMAINED IMPLANTED. CURRENTLY THERE IS NO TREATMENT PLANNED DUE TO PATIENT IS PREGNANT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. DATE OF EVENT: UNK. EXPIRATION DATE: UNK. IMPLANT DATE: UNK. EXPLANT DATE: NA. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE MANUFACTURE DATE: UNK. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50/+2.0/119 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD). THE LENS WAS REPORTED AS LENS ROTATION AND REMAINED IMPLANTED. CURRENTLY THERE IS NO TREATMENT PLANNED DUE TO PATIENT IS PREGNANT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. D2: COMMON DEVICE NAME: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB. E1: NAME: (B)(6). CLAIM # (B)(4). | -07.50/+2.0/119 | OD | Physician | 01/20/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13552907&pc=QCB | 2023826-2022-00480 | 2021-12-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.5/2.5/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.5/2.5/094 | OD | Physician | 02/11/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13553237&pc=QCB | 2023826-2022-00541 | 2022-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -4.0/+0.5/108 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS IMPLANTED UPSIDE DOWN; AND ON (B)(6) 2022 THE LENS WAS EXPLANTED. THERE WAS NO PATIENT INJURY. IT WAS INDICATED A NEW LENS IS BEING ORDERED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -4.0/+0.5/108 | OD | Physician | 02/11/2022 | 2022 | 46 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13565840&pc=QCB | 2023826-2022-00454 | 2021-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -11.0/2.0/104 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -11.0/2.0/104 | OS | Physician | 01/21/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13577872&pc=QCB | 2023826-2022-00515 | 2022-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.0/4.0/121 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.0/4.0/121 | OD | Physician | 01/28/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13577877&pc=QCB | 2023826-2022-00516 | 2022-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -12.0/2.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK, ETHNICITY- UNK, RACE- UNK, PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.0/2.5/090 | OS | Physician | 01/28/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13580869&pc=QCB | 2023826-2022-00571 |  | Injury | STAAR SURGICAL COMPANY | 2022-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -9.50/1.0/080 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. IT WAS REPORTED THERE IS LOW VAULT. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -9.50/1.0/080 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. IT WAS REPORTED THERE IS LOW VAULT. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP, IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. B5 - THE LENS WAS EXPLANTED AT A LATER DATE. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -9.50/1.0/080 | OS | Other Health Care Professional | 01/20/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13580886&pc=QCB | 2023826-2022-00570 |  | Injury | STAAR SURGICAL COMPANY | 2022-02-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -11.00/2.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. IT WAS REPORTED THERE IS LOW VAULT. LENS REMAINS IMPLANTED. . ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -11.00/2.0/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. IT WAS REPORTED THERE IS LOW VAULT. LENS REMAINS IMPLANTED. . ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN A SPECIMEN CUP, IN LIQUID WITH DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. B5: THE LENS WAS EXPLANTED AT A LATER DATE. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -11.00/2.0/089 | OD | Other Health Care Professional | 01/20/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13583013&pc=QCB | 2023826-2022-00597 | 2021-12-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -12.0/1.0/92 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. EXCESSIVE VAULT WAS NOTED AND ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME MODEL; SAME LENGTH; DIFFERENT AXIS LENS WHICH WAS IMPLANTED VERTICALLY, AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: PATIENT INFORMATION:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -12.0/1.0/92 | OS | Physician | 01/24/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13586579&pc=QCB | 2023826-2022-00458 | 2021-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.0/1.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/1.5/091 | OD | Physician | 01/24/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13586599&pc=QCB | 2023826-2022-00457 | 2021-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -11.5/1.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) AT AN UNKNOWN DATE. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -11.5/1.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. (B)(6) 2021. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -11.5/1.5/091 | OD | Physician | 01/24/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13588585&pc=QCB | 2023826-2022-00525 | 2021-12-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -12.50/5.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS OTHER. THE LENS DID NOT FAIL TO PERFORM AS INTENDED. CAUSE OF EVENT IS PATIENT-RELATED FACTOR - NARROW SULCUS. Manufacturer Narrative: AGE, WEIGHT-ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -12.50/5.0/090 | OD | Physician | 02/14/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13588594&pc=QCB | 2023826-2022-00538 | 2021-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.0/0.5/111 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(4) 2021. ON (B)(4) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.0/0.5/111 | OD | Physician | 02/17/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13588644&pc=QCB | 2023826-2022-00531 | 2021-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -6.50/1.50/079 (SPHERE/CYLINDER/AXIS) DIOPTER WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -6.50/1.50/079 | OD | Physician | 02/16/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13588735&pc=QCB | 2023826-2022-00574 | 2021-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -03.00/+4.5/176 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING AND DUE TO SURGEON PREFERENCE CHANGE (SURGEON PREFERENCE CHANGE IS REFERRING TO ICL SIZE CHANGE NOTED FROM POSTOPERATIVE ICL VAULT). THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS GOOD. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -03.00/+4.5/176 | OS | Physician | 01/25/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13588759&pc=QCB | 2023826-2022-00572 |  | Injury | STAAR SURGICAL COMPANY | 2022-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -08.00/+6.0/094 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD). THE LENS WAS REPORTED AS A REFRACTIVE MISS AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS SUCCESSFULLY REMOVED AND REPLACED WITH A PIOL LENS. THE LENS WAS DISCARDED. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -08.00/+6.0/094 | OD | Physician | 01/27/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13588913&pc=QCB | 2023826-2022-00600 | 2021-12-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -9.5/3.0/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO THE INITIAL LENS HAD SHIFTED IN FRONT OF THE IRIS (DISPLACEMENT). THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -9.5/3.0/095 | OD | Physician | 02/14/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13588949&pc=QCB | 2023826-2022-00603 | 2022-01-24 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -8.50/+2.50/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS HAD TORE DURING INJECTION INTO THE PATIENT'S EYE. INTRAOPERATIVELY THE LENS WAS REMOVED AND THE PROBLEM RESOLVED. THERE WAS NO PATIENT INJURY. IT WAS INDICATED THERE IS NO PLANS TO REPLACE THE LENS. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -8.50/+2.50/89 | OD | Physician | 02/18/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13593899&pc=QCB | 2023826-2022-00460 | 2021-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -12.0/2.5/173 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICTY:UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.0/2.5/173 | OS | Physician | 01/26/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13593915&pc=QCB | 2023826-2022-00464 | 2021-12-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -12.5/2.5/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.5/2.5/089 | OD | Physician | 01/26/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13593916&pc=QCB | 2023826-2022-00463 | 2021-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -12.5/1.0/88 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6)2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.5/1.0/88 | OD | Physician | 01/26/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13593918&pc=QCB | 2023826-2022-00459 | 2021-11-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -11.0/3.0/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6)2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO:UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/3.0/095 | OD | Physician | 01/26/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13593930&pc=QCB | 2023826-2022-00462 | 2021-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -7.5/5.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6)2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -7.5/5.5/093 | OD | Physician | 01/26/2022 | 2022 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13595952&pc=QCB | 2023826-2022-00581 | 2020-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00/+3.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2020 BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (LENS WAS IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00/+3.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2020 BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (LENS WAS IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHINICITY: UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN, ETHINICITY: UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -18.00/+3.5/093 | OS | Physician | 01/26/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13595989&pc=QCB | 2023826-2022-00580 | 2022-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/148 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: B5: THE REPORTER INDICATED THE PATIENTS RE TURNED OUT BEAUTIFUL FOLLOWING THE ROTATION. REFRACTION IS +0.25 AND SHE IS SEEING 6/4.5 UNAIDED. CLAIM # (B)(4). Manufacturer Narrative: PT INFO. EXPLANT DATE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -09.50/+1.0/148 | OD | Physician | 01/25/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13596000&pc=QCB | 2023826-2022-00573 |  | Injury | STAAR SURGICAL COMPANY | 2022-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -04.00/+6.0/059 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS REPORTED AS A REFRACTIVE MISS AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PT INFO: UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: B5: THE LENS WAS SUCCESSFULLY REMOVED AND REPLACED WITH A PIOL LENS. THE LENS WAS DISCARDED. CLAIM # (B)(4). | -04.00/+6.0/059 | OS | Physician | 01/27/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13596007&pc=QCB | 2023826-2022-00605 | 2022-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -7.00/+2.00/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED LENS ROTATION WHICH WAS NOTED ON (B)(6) 2022. THEN ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS/MODEL BUT DIFFERENT AXIS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -7.00/+2.00/92 | OD | Physician | 02/18/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13596008&pc=QCB | 2023826-2022-00583 | 2021-08-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.0/091 (SPHERE/CYLINDER/AXIS), TORE DURING LOADING INTO THE CARTRIDGE AND THERE WAS NO PATIENT CONTACT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE CODE: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -11.00/+1.0/091 | NA | Physician | 01/28/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13596023&pc=QCB | 2023826-2022-00582 | 2022-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -04.50/+2.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS LOW VAULTING WITH LENS ROTATION. THE LENS WAS REPOSITIONED ON (B)(6)2021 BUT THE PROBLEM WAS NOT RESOLVED. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -04.50/+2.0/093 | OS | Physician | 01/27/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13596055&pc=QCB | 2023826-2022-00606 | 2022-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -8.50/+1.50/88 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LENS ROTATION WHICH WAS NOTED ON (B)(6) 2022. THEN ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS/MODEL BUT DIFFERENT AXIS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS OPTIC TORN/SPLIT AND THE LENS HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). UNDER 21YRS OF AGE AT DATE OF IMPLANT. CLAIM# (B)(4). | -8.50/+1.50/88 | OS | Physician | 02/18/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13596355&pc=QCB | 2023826-2022-00467 | 2021-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11.50/2.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS DEVICE "LARGE LENS/HIGH VAULT". Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -11.50/2.0/094 | OS | Physician | 01/27/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13596356&pc=QCB | 2023826-2022-00466 | 2021-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -9.50/2.0/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON(B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS DEVICE "LARGE LENS/HIGH VAULT".Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -9.50/2.0/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS DEVICE "LARGE LENS/HIGH VAULT". Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#:(B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#:(B)(4). | -9.50/2.0/93 | OD | Physician | 01/27/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13599453&pc=QCB | 2023826-2022-00616 | 2021-12-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -15.5/1.5/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. LOW VAULT WITH LENS ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2021. THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -15.5/1.5/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. LOW VAULT WITH LENS ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2021. THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -15.5/1.5/094 | OS | Physician | 01/28/2022 | 2022 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13599706&pc=QCB | 2023826-2022-00469 | 2021-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM0 12.6 IMPLANTABLE COLLAMER LENS OF -14.50/3.0/70 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS DEVICE "LARGE LENS/HIGH VAULT". Manufacturer Narrative: PATIENT INFORMATION:UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -14.50/3.0/70 | OS | Physician | 01/28/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13599741&pc=QCB | 2023826-2022-00468 | 2021-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM0 12.6 IMPLANTABLE COLLAMER LENS OF -18.0/3.0/132 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS DEVICE "LARGE LENS/HIGH VAULT". Manufacturer Narrative: PT INFO: UNKNOWN. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -18.0/3.0/132 | OD | Physician | 01/28/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13603042&pc=QCB | 2023826-2022-00584 | 2021-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-02-25 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/081 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED.Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/081 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. THERE WAS FIBER ON THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. REFRACTIVE (FUNCTIONAL) VERIFICATION WAS PERFORMED AND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER, INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.50/+1.0/081 | OS | Physician | 01/28/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13641135&pc=QCB | 2023826-2022-00655 | 2021-10-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50/+2.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING WITH LENS ROTATION. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED SO THE LENS WAS EXPLANTED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -08.50/+2.0/095 | OD | Physician | 01/31/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13641167&pc=QCB | 2023826-2022-00643 | 2022-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, IN THE PATIENTS EYE ON (B)(6) 2022. THE SURGEON REPORTED THE LENS WAS OVER VAULTED AND PLANNED TO REPLACE THE LENS THE NEXT DAY FOR A SMALLER LENS. THE PATIENT RETURNED THE NEXT DAY AND THE SURGEON INDICATED HE BELIEVED THERE WAS RETAINED VISCOELASTIC BEHIND THE LENS. THE SURGEON DECIDED TO MONITOR THE PATIENT AND VAULT. THE LENS REMAINED IMPLANTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM TMICL13.7 IMPLANTABLE COLLAMER LENS, -14.00/+3.0/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. D2: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK EXPIRATION DATE UNK IMPLANTED: UNK MANUFACTURED: UNK. (B)(4) | -14.00/+3.0/085 | OD | Physician | 02/08/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13641289&pc=QCB | 2023826-2022-00653 | 2022-01-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -06.50/+1.0/070 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (LENS WAS IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H6: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM #(B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTICS BENT AND DEFORMED. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | -06.50/+1.0/070 | OD | Physician | 01/31/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13654076&pc=QCB | 2023826-2022-00470 | 2021-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -8.0/1.50/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6)2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE DID NOT RESOLVE THE PROBLEM. CAUSE IS REPORTED AS DEVICE "LOW VAULT." SEE MFR# 2022-00471 FOR EXCHANGE LENS. Manufacturer Narrative: WEIGHT:UNKNOWN,ETHINICITY:UNKNOWN: RACE:UNKNOWN. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. THE RETURNED LENS DID NOT MEET THE ORIGINAL VALUE MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). | -8.0/1.50/91 | OD | Physician | 02/04/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13654276&pc=QCB | 2023826-2022-00471 | 2021-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 13.2 MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -8.50/2.0/110 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. STATUS OF EYE WNL (WITHIN LIMITS), SHALLOW VAULT BUT ACCEPTABLE. CAUSE IS REPORTED AS DEVICE "LOW VAULT." SEE MFR# 2022-00470 FOR INITIAL LENS. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL, DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. THE RETURN LENS DID NOT MEET THE ORIGINAL VALUE MEASURED AT THE TIME OF MANUFACTURING. THE LENS RETURNED IS NOT THE SIZE STATED IN THE CLAIM. UNABLE TO IDENTIFY WHICH LENS PERTAINS TO WHICH CLAIM AS THEY WERE BOTH RETURNED IN THE SAME BAG. CLAIM# (B)(4). | -8.50/2.0/110 | OD | Physician | 02/04/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13654283&pc=QCB | 2023826-2022-00472 | 2021-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.6 MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -9.0/2.50/94 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE DID NOT RESOLVE THE PROBLEM. STATUS OF EYE "LOW VAULT AFTER EXCHANGE." CAUSE IS REPORTED AS DEVICE "LOW VAULT." SEE MFR # 2022-00473 FOR EXCHANGE LENS. Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. THE RETURNED LENS DID NOT MEET THE ORIGINAL VALUE MEASURED AT THE TIME OF MANUFACTURING. THE LENS RETURNED IS NOT THE MODEL/SIZE STATED IN THE CLAIM. UNABLE TO IDENTIFY WHICH LENS PERTAINS TO WHICH CLAIM AS THEY WERE BOTH RETURNED IN THE SAME BAG. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/2.50/94 | OS | Physician | 02/04/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13654519&pc=QCB | 2023826-2022-00473 | 2021-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 13.2 MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -9.0/2.50/110 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) AS AN EXCHANGE LENS ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE DID NOT RESOLVE THE PROBLEM. STATUS OF EYE "LOW VAULT AFTER EXCHANGE." CAUSE IS REPORTED AS DEVICE "LOW VAULT." SEE MFR# 2022-00474 FOR EXCHANGE LENS. Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BENT. THE RETURN LENS DID NOT MEET THE ORIGINAL VALUE MEASURED AT THE TIME OF MANUFACTURING. THE LENS RETURNED IS NOT THE THE SIZE STATED IN THE CLAIM. UNABLE TO IDENTIFY WHICH LENS PERTAINS TO WHICH CLAIM AS THEY WERE BOTH RETURNED IN THE SAME BAG. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/2.50/110 | OS | Physician | 02/04/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13654601&pc=QCB | 2023826-2022-00474 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 13.7 MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF -8.50/2.0/92 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) AS A SECOND EXCHANGE LENS ON (B)(6) 2021 TO CORRECT LOW VAULT. THE PATIENT IS STILL EXPERIENCING A LOW VAULT OF 70UM. THE LENS REMAINS IMPLANTED. PER REPORTER, NO INTERVENTION IS PLANNED. CAUSE IS REPORTED AS DEVICE "LOW VAULT." SEE MFR# 2022-00473 FOR FIRST EXCHANGE LENS. Manufacturer Narrative: PT INFO. DATE OF EVENT: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -8.50/2.0/92 | OS | Physician | 02/04/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13656869&pc=QCB | 2023826-2022-00647 | 2022-01-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.50/+2.5/094 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. THE EYE IS QUIET AND VAULT IS IDEAL. PATIENT IS COMPLAINING ABOUT VISION. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE SURGEON DECIDED TO TOP UP WITH PRK FOR THE RESIDUAL POWER. AS OF NOW HAS DECIDED NOT TO UNDERGO AN EXCHANGE. H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - PRK SURGERY. CLAIM #(B)(4). Manufacturer Narrative: PT INFO: UNK. EXPLANT DATE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). | -17.50/+2.5/094 | OD | Physician | 02/01/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13656871&pc=QCB | 2023826-2022-00656 | 2021-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -17.00/+1.0/113 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH A SAME LENGTH BUT DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -17.00/+1.0/113 | OD | Physician | 02/07/2022 | 2022 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13656872&pc=QCB | 2023826-2022-00648 | 2022-01-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/090 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. THE EYE IS QUIET AND VAULT IS IDEAL. PATIENT IS COMPLAINING ABOUT VISION. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: B5: THE SURGEON DECIDED TO TOP UP WITH PRK FOR THE RESIDUAL POWER. AS OF NOW HAS DECIDED NOT TO UNDERGO AN EXCHANGE. H6: HEALTH IMPACT - ADDITIONAL SURGERY: 4625 - PRK SURGERY. CLAIM #(B)(4). | -11.50/+2.0/090 | OS | Physician | 02/01/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13664213&pc=QCB | 2023826-2022-00713 | 2022-01-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -13.0/2.0/94 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.0/2.0/94 | OD | Physician | 03/02/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13664263&pc=QCB | 2023826-2022-00712 | 2022-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.0/2.0/75 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/2.0/75 | OS | Physician | 03/02/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13664270&pc=QCB | 2023826-2022-00714 | 2022-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.5/1.5/77 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO.:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -9.5/1.5/77 | OS | Physician | 03/02/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13664368&pc=QCB | 2023826-2022-00715 | 2021-12-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -13.5/2.0/113 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.5/2.0/113 | OD | Physician | 03/02/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13664484&pc=QCB | 2023826-2022-00716 | 2022-02-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -9.5/1.5/2 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH/MODEL LENS BUT IMPLANTED HORIZONTALLY; WHICH IT WAS NOTED THE INITIAL IMPLANT WAS IMPLANTED VERTICALLY. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | -9.5/1.5/2 | OS | Physician | 03/02/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13670593&pc=QCB | 2023826-2022-00590 | 2021-12-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -18.0/3.0/157 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -18.0/3.0/157 | OS | Physician | 02/21/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13670652&pc=QCB | 2023826-2022-00589 | 2021-12-05 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -14.5/2.5/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY:UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -14.5/2.5/089 | OD | Physician | 02/21/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13670702&pc=QCB | 2023826-2022-00609 | 2021-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.5/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -10.5/2.0/089 | OD | Physician | 12/01/2021 | 2021 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13670724&pc=QCB | 2023826-2022-00611 | 2021-12-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -15.0/2.0/077 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. THE EXCHANGE LENS WAS IMPLANTED VERTICALLY. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -15.0/2.0/077 | OD | Physician | 02/23/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13670751&pc=QCB | 2023826-2022-00607 | 2022-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.5/1.0/074 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -12.5/1.0/074 | OS | Physician | 02/21/2022 | 2022 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13672905&pc=QCB | 2023826-2022-00724 | 2022-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Patient Device Interaction Problem | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -15.5/2/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING WITH ROTATION WHICH A LENS REPOSITIONING WAS PERFORMED ON (B)(6) 2022 BUT THIS DID NOT RESOLVE THE PROBLEM. THEN ON (B)(6) 2022 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH NO VISIBLE DAMAGE TO LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -15.5/2/92 | OD | Physician | 09/23/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13674374&pc=QCB | 2023826-2022-00739 | 2020-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -11.5/2.5/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6)2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. GMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). Manufacturer Narrative: B5 - "REPORTEDLY, 'T561847/12.1MM LENS WAS IMPLANTED IN A VERTICAL POSITION, WITH A LOW ARCH AFTER IMPLANTATION. SECONDARY SURGERY: T858877/12.6MM LENS WAS IMPLANTED IN A HORIZONTAL POSITION.'" SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). | -11.5/2.5/93 | OD | Physician | 03/04/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13675210&pc=QCB | 2023826-2022-00738 | 2022-02-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.5/1.0/074 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO:UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -9.5/1.0/074 | OS | Physician | 03/03/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13684326&pc=QCB | 2023826-2022-00557 | 2021-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -14.50/2.0/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -14.50/2.0/073 | OS | Physician | 02/17/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13684353&pc=QCB | 2023826-2022-00550 | 2021-12-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.00/2.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -12.00/2.0/084 | OS | Physician | 02/14/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13684356&pc=QCB | 2023826-2022-00558 | 2021-12-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -7.50/2.0/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -7.50/2.0/098 | OD | Physician | 02/17/2022 | 2022 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13684363&pc=QCB | 2023826-2022-00556 | 2021-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -13.00/1.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.00/1.0/088 | OD | Physician | 02/17/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13684366&pc=QCB | 2023826-2022-00564 | 2022-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.00/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.00/1.5/093 | OD | Physician | 02/18/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13684386&pc=QCB | 2023826-2022-00559 | 2021-12-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -18.0/1.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -18.0/1.5/088 | OS | Physician | 02/17/2022 | 2022 | 44 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13684406&pc=QCB | 2023826-2022-00554 | 2022-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -12.50/1.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2016. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.50/1.0/082 | OS | Physician | 02/16/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13684407&pc=QCB | 2023826-2022-00555 | 2021-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.5/2.5/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -14.5/2.5/071 | OS | Physician | 02/17/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13684408&pc=QCB | 2023826-2022-00565 | 2022-01-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.00/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.00/2.0/088 | OS | Physician | 02/18/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13684412&pc=QCB | 2023826-2022-00560 | 2022-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -13.50/1.50/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.50/1.50/089 | OD | Physician | 02/18/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13684415&pc=QCB | 2023826-2022-00568 | 2021-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -9.0/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.0/1.5/093 | OS | Physician | 02/18/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13684419&pc=QCB | 2023826-2022-00566 | 2021-12-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 -9.50/2.0/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US.(B)(4). | -9.50/2.0/072 | OS | Physician | 02/18/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13690336&pc=QCB | 2023826-2022-00668 | 2021-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -15.0/1.5/166 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. THE EXCHANGE LENS WAS IMPLANTED VERTICALLY. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT:UNKNOWN, ETHINICITY:UNKNOWN, RACE:UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -15.0/1.5/166 | OS | Physician | 02/28/2022 | 2022 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13690436&pc=QCB | 2023826-2022-00673 | 2022-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.0/1.5/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE U.S. DEVICE CODE: (B)(4). CLAIM# (B)(4). | -10.0/1.5/091 | OS | Physician | 02/28/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13690447&pc=QCB | 2023826-2022-00670 | 2022-01-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -13.0/1.5/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -13.0/1.5/095 | OD | Physician | 02/28/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13690455&pc=QCB | 2023826-2022-00675 | 2022-01-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -7.0/1.5/045 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6)2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.0/1.5/045 | OD | Physician | 02/28/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13690522&pc=QCB | 2023826-2022-00672 | 2022-01-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -10.0/3.0/006 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE U.S. DEVICE CODE: (B)(4). CLAIM# (B)(4). | -10.0/3.0/006 | OS | Physician | 02/28/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13690548&pc=QCB | 2023826-2022-00671 | 2022-01-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -11.5/3.0/177 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -11.5/3.0/177 | OD | Physician | 02/28/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13690630&pc=QCB | 2023826-2022-00638 | 2022-01-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -9.0/2.0/080 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6)2022. ON (B)(6)2022 THE LENS WAS EXPLANTED. ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -9.0/2.0/080 | OD | Physician | 02/25/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13690634&pc=QCB | 2023826-2022-00639 | 2021-12-31 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.0/1.0/070 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS INTRAOPERATIVELY IMPLANTED AND REMOVED. ON THE SAME DATE IN A SEPARATE SURGERY IT WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO:UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -10.0/1.0/070 | OS | Physician | 12/31/2021 | 2021 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13690660&pc=QCB | 2023826-2022-00674 | 2022-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.0/1.5/088 | OD | Physician | 02/28/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13699598&pc=QCB | 2023826-2022-00749 | 2022-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; Unintended Movement | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.00/+3.0/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO THE LENS WAS IMPLANTED UPSIDE DOWN. THERE WAS NO PATIENT INJURY. THE LENS WAS DAMAGED WHEN EXPLANTED FROM THE EYE. ANOTHER LENS WAS NOT IMPLANTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.00/+3.0/082 | OS | Physician | 02/09/2022 | 2022 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13699604&pc=QCB | 2023826-2022-00748 | 2022-01-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -14.00/+1.5/059 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. VISION OK FOR NOW. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PT INFO: UNK. EXPLANT DATE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: B5: THE LENS WAS EXPLANTED AND REPLACED FOR A SHORTER LENS ON (B)(6) 2022. THE PROBLEM WAS RESOLVED. EVERYTHING IS OK. CLAIM # (B)(4). | -14.00/+1.5/059 | OS | Physician | 02/11/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13699610&pc=QCB | 2023826-2022-00751 | 2021-12-31 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/094 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY. THIS WAS DUE TO THE LENS WAS STUCK WITH PLUNGER AND TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS CLEAR AND PATIENT IS DOING VERY WELL. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.00/+2.0/094 | OS | Physician | 02/10/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13700079&pc=QCB | 2023826-2022-00743 | 2021-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -10.0/1.0/029 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS. THE REPLACEMENT LENS WAS PLACED IN THE HORIZONTAL POSITION. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON THE LENS. DIMENSIONAL FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -10.0/1.0/029 | OS | Physician | 02/07/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13700123&pc=QCB | 2023826-2022-00750 | 2021-12-22 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/003 (SPHERE/CYLINDER/AXIS) WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED ON (B)(6) 2022 WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING VERY WELL. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.00/+1.5/003 | OD | Physician | 02/10/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13713558&pc=QCB | 2023826-2022-00679 | 2022-01-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.0/3.0/096 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -9.0/3.0/096 | OD | Physician | 03/01/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13713568&pc=QCB | 2023826-2022-00727 | 2022-01-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.0/1.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -12.0/1.0/091 | OD | Physician | 03/02/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13713629&pc=QCB | 2023826-2022-00728 | 2022-02-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.0/2.5/179 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6)2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.0/2.5/179 | OS | Physician | 03/02/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13713652&pc=QCB | 2023826-2022-00734 | 2022-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -8.5/1.0/180 (CYLINDER/SPHERE/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6)2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#(B)(4). | -8.5/1.0/180 | OD | Physician | 03/02/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13713696&pc=QCB | 2023826-2022-00729 | 2021-11-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -8.0/1.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -8.0/1.0/092 | OS | Physician | 03/02/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13713759&pc=QCB | 2023826-2022-00735 | 2022-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.0/2.5/084 (CYLINDER/SPHERE/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/2.5/084 | OD | Physician | 03/02/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13713865&pc=QCB | 2023826-2022-00725 | 2022-01-19 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 -16.5/4.0/062 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#:(B)(4). | -16.5/4.0/062 | OS | Physician | 03/02/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13713897&pc=QCB | 2023826-2022-00736 | 2022-01-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -8.0/4.0/077 (CYLINDER/SPHERE/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#: (B)(4). | -8.0/4.0/077 | OS | Physician | 03/02/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13713929&pc=QCB | 2023826-2022-00730 | 2022-01-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.5/2.5/083 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.5/2.5/083 | OD | Physician | 03/02/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13725473&pc=QCB | 2023826-2022-00771 | 2022-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -8.0/2.0/003 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT DIFFERENT AXIS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3:DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/2.0/003 | OS | Physician | 09/23/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13725499&pc=QCB | 2023826-2022-00774 | 2022-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -13.5/+1.0/87 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. LENS REMAINS IMPLANTED. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE LENS WAS EXPLANTED ON (B)(6) 2022. LATER THE ON THE SAME DATE A REPLACEMENT LENS OF A SHORTER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(6). | -13.5/+1.0/87 | OS | Physician | 02/18/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13725647&pc=QCB | 2023826-2022-00773 | 2022-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -13.5/+1.5/94 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. LENS REMAINS IMPLANTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: IT WAS LATER REPORTED THAT THE LENS WAS EXPLANTED ON (B)(6) 2022. LATER THE ON THE SAME DATE A REPLACEMENT LENS OF A SHORTER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). | -13.5/+1.5/94 | OD | Physician | 02/18/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13725845&pc=QCB | 2023826-2022-00745 | 2022-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -1.0/1.5/123 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT. THE LENS REMAINS IMPLANTED. CAUSE REPORTED AS DEVICE. "LENS TOO BIG". Manufacturer Narrative: ADDITIONAL INFORMATION: B1-ADVERSE EVENT. B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). B5-THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -10.0/1.5/123 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED. A REPLACEMENT LENS OF A DIFFERENT TORIC MODEL AND SHORTER LENGTH LENS WAS IMPLANTED DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS DEVICE. "LENS TOO BIG". D6B- (B)(6) 2022. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -1.0/1.5/123 | OS | Physician | 02/08/2022 | 2022 | 44 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13726407&pc=QCB | 2023826-2022-00746 |  | Injury | STAAR SURGICAL COMPANY | 2022-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 13.2MM TMICL 13.2 IMPLANTABLE COLLAMER LENS OF -7.50/4.0/099 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DATE OF EVENT: UNK. EXPLAINED: UNK. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4) | -7.50/4.0/099 | OD | Other Health Care Professional | 02/10/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13726455&pc=QCB | 2023826-2022-00753 | 2022-02-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50/+3.0/093 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS QUIET AND THE PATIENT IS HAPPY. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -17.50/+3.0/093 | OD | Physician | 02/11/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13726547&pc=QCB | 2023826-2022-00754 | 2022-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.50/+0.5/045 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS LOW VAULTING AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6)2022 AND WAS REPLACED WITH A LONGER LENS. THE PROBLEM WAS RESOLVED AND PATIENT IS HAPPY. CAUSE OF THE EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -08.50/+0.5/045 | OD | Physician | 02/11/2022 | 2022 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13726649&pc=QCB | 2023826-2022-00756 | 2022-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+5.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2022 BUT THE PROBLEM WAS NOT RESOLVED. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED WITH A VICMO13.2 IMPLANTABLE COLLAMER LENS, -10.00 DIOPTER. CLAIM # (B)(4). Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2022. THE PROBLEM WAS NOT RESOLVED WITH THE IMPLANT OF THE REPLACEMENT LENS. THE PATIENT HAD AN ISCHEMIC OPTIC NEUROPATHY DUE TO VISCO. AT ONE POINT THE PATIENT ONLY HAD LIGHT PERCEPTION BUT IS NOW RECOVERING. PATIENT HAS MYDRIASIS DUE TO IRIS ISCHEMIA AND HAS PIGMENT DISPERSION. THE DOCTOR WANTS TO FOCUS ON RECOVERY AND STABILIZATION OF THE EYE AND THEN LOOK FOR CORRECTION OF THE RESIDUAL ASTIGMATISM WITH MAYBE A PRK. SEE MFR. REPORT # 2023826-2022-01776 FOR REPLACEMENT LENS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: NA. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM #: (B)(4). | -12.50/+5.0/093 | OD | Physician | 02/14/2022 | 2022 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13737265&pc=QCB | 2023826-2022-00776 | 2022-02-11 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/090 (SPHERE/CYLINDER/AXIS) LENS TEAR/BREAK DURING LOADING. THE LENS WAS NOT IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THE SURGERY WAS POSTPONED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -12.0/1.5/090 | OS | Physician | 02/11/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13739918&pc=QCB | 2023826-2022-00775 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-03-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -9.5/+2.5/062 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) . LOW VAULT IS REPORTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: WEIGHT, ETHNICITY, WEIGHT-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -9.5/+2.5/062 | OS | Physician | 01/25/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13771539&pc=QCB | 2023826-2022-00621 | 2022-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.00/2.0/069 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY:UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.00/2.0/069 | OS | Physician | 02/22/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13771541&pc=QCB | 2023826-2022-00627 | 2021-12-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -09.50/2.0/093 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6- HEALTH EFFECT- IMPACT CODE CORRECTION: CHANGED CODE 4629 TO 2199. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE U.S. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -09.50/2.0/093 | OS | Physician | 02/23/2022 | 2022 | 46 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13771542&pc=QCB | 2023826-2022-00623 | 2021-09-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1, -10.00/2.5/095 DIOPTER, IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REPLACED INTRAOPERATIVELY WITH A LONGER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S LEFT EYE (OS) DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.00/2.5/095 | OS | Physician | 02/21/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13771544&pc=QCB | 2023826-2022-00628 | 2022-01-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6, -9.50/1.00/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: B3- CORRECTED . CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -9.50/1.00/082 | OS | Physician | 02/24/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13771545&pc=QCB | 2023826-2022-00625 | 2022-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.00/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK .ETHNICITY - UNK .RACE - UNK PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US . DEVICE PROBLEM CODE: 1494 - OFF-LABEL USE, ACD<3.0MM. CLAIM# (B)(4). | -9.00/1.5/090 | OD | Physician | 02/21/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13771546&pc=QCB | 2023826-2022-00633 | 2022-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/0.5/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULT. IN A SECOND SURGERY ON (B)(6) 2022 A LONGER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE U.S. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -10.50/0.5/070 | OD | Physician | 02/25/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13771547&pc=QCB | 2023826-2022-00624 | 2022-01-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.50/2.0/116 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT:UNKNOWN, ETHINICITY:UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM# (B)(4). | -11.50/2.0/116 | OD | Physician | 02/21/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13771549&pc=QCB | 2023826-2022-00635 | 2021-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.50/1.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. A SHORTER LENGTH LENS WAS IMPLANTED IN A SECOND SURGERY ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -8.50/1.0/088 | OD | Physician | 02/25/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13771551&pc=QCB | 2023826-2022-00631 | 2021-04-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1, -14.50/3.50/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -14.50/3.50/092 | OD | Physician | 02/24/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13771552&pc=QCB | 2023826-2022-00630 | 2021-11-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -9.00/2.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY- UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -9.00/2.5/081 | OS | Physician | 02/24/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13771553&pc=QCB | 2023826-2022-00632 | 2021-12-31 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -11.0/1.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT.THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN, ETINICITY: UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: CORRECTION DATA: B5 - REPORTED AS IMPLANTED AND REPLACED INTRAOPERATIVELY. - CORRECT TO IMPLANTED AND REMOVED INTRAOPERATIVELY ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. IN A SECOND SURGERY ON (B)(6) 2021 A SHORTER LENGTH LENS WAS IMPLANTED ON (B)(6) 2021 TO RESOLVE THE PROBLEM. H1 REPORTED AS SERIOUS INJURY - CORRECT TO MALFUNCTION. CLAIM# (B)(4). | -11.0/1.5/080 | OD | Physician | 02/25/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13771560&pc=QCB | 2023826-2022-00636 | 2021-05-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.50/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. A SHORTER LENGTH LENS WAS IMPLANTED IN A SECOND SURGERY ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -8.50/1.0/090 | OS | Physician | 02/25/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13771563&pc=QCB | 2023826-2022-00634 | 2022-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.50/0.5/146 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: HEALTH EFFECT - IMPACT CODE: CORRECTED TO 4629. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.50/0.5/146 | OS | Physician | 02/25/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13778396&pc=QCB | 2023826-2022-00828 | 2021-12-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -16.5/3.0/082 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE LENS HAD ROTATED 5 MOTHS AFTER IMPLANT AND THE SURGEON THEN PERFORMED A LENS ROTATION ON (B)(6)2021 BUT THIS DID NOT RESOLVE THE PROBLEM. THEN ON (B)(6)2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT SPHERICAL LENS DUE TO FEAR OF ROTATION; AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -16.5/3.0/082 | OS | Physician | 03/11/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13778398&pc=QCB | 2023826-2022-00827 | 2021-12-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -17.0/3.5/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS HAD ROTATED 6 WEEKS AFTER IMPLANT AND THE SURGEON THEN PERFORMED A LENS ROTATION ON (B)(6) 2021 BUT THIS DID NOT RESOLVE THE PROBLEM. THEN ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT SPHERICAL LENS DUE TO FEAR OF ROTATION; AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -17.0/3.5/92 | OD | Physician | 03/11/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13793510&pc=QCB | 2023826-2022-00777 | 2021-12-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -12.0/3.50/104 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LENS ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2021 AND THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT, ETHNICITY AND RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH RESIDUE/ DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC DEFORMED. HAPTIC DEFORMED/ STRETCHED OUT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -12.0/3.50/104 | OS | Physician | 02/14/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13794224&pc=QCB | 2023826-2022-00784 | 2021-10-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.7 VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF -17.0/2.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS AN EXCHANGE LENS IN THE PATIENTS RIGHT EYE (OD). THE LENS IS STILL SLIGHTLY OFFSET DOWNWARD. REPORTEDLY, THE PATIENT IS HAPPY WITH THE OUTCOME, THEREFORE THE DOCTOR DOES NOT CONSIDER THIS AS AN ISSUE. THE LENS REMAINS IMPLANTED. SEE MFR # 2023826-2021-00835 FOR INITIAL LENS. Manufacturer Narrative: WEIGHT:UNKNOWN: ETHINICITY:UNKNOWN: RACE:UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -17.0/2.0/090 | OD | Other | 02/14/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13795102&pc=QCB | 2023826-2022-00791 | 2022-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -05.00/+1.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED DUE TO AN OVER VAULT. THE LENS WAS REPLACED WITH A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. CLAIM # (B)(4). Manufacturer Narrative: SEX: UNK, WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, DATE OF EVENT: UNK, EXPLANT DATE: UNK, WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -05.00/+1.5/090 | OS | Physician | 02/14/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13795110&pc=QCB | 2023826-2022-00790 | 2022-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM TMICL13.7 IMPLANTABLE COLLAMER LENS, -14.50/+4.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED DUE TO AN OVER VAULT. THE LENS WAS REPLACED WITH A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK RACE: UNK DATE OF EVENT: UNK EXPLANT DATE: UNK WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS DUE TO DEVICE. CLAIM # (B)(4). | -14.50/+4.0/092 | OS | Physician | 02/14/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13795124&pc=QCB | 2023826-2022-00789 | 2022-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, --5.00/+1.5/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED DUE TO AN OVER VAULT. THE LENS WAS REPLACED WITH A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED AND REPLACED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. CLAIM # (B)(4). Manufacturer Narrative: H6 - WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: SEX: UNK, WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, DATE OF EVENT: UNK, AND EXPLANT DATE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -5.00/+1.5/087 | OD | Physician | 02/14/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13795127&pc=QCB | 2023826-2022-00799 | 2022-01-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00/+1.5/014 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.00/+1.5/014 | OD | Physician | 02/18/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13795666&pc=QCB | 2023826-2022-00829 | 2022-01-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.5/3.5/137 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE PATIENT EXPERIENCED LOW VAULT WITH ROTATION. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH LENS BUT IMPLANTED HORIZONTALLY INSTEAD OF VERTICAL AS THE INITIAL IMPLANT. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -12.5/3.5/137 | OS | Physician | 03/14/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13795669&pc=QCB | 2023826-2022-00830 |  | Injury | STAAR SURGICAL COMPANY | 2022-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -10.0/3.5/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED WITH LOW VAULTING. ON (B)(6)2022 THE LENS WAS REPOSITIONED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -10.0/3.5/094 | OS | Physician | 03/15/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13795877&pc=QCB | 2023826-2022-00781 | 2022-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF -11.5/3.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED DUE TO THE LENS IMPLANTED UPSIDE DOWN. A REPLACEMENT LENS OF THE SAME MODEL AND SIZE WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. "IMPROVEMENTS IN OS AFTER REPLACEMENT/EXCHANGE. SLIGHT RESIDUAL ASTIGMATISM. CONTINUE MURO 128, PROLENSA, INVELTYS WITH CHANGE TO FLAREY WHEN OUT OF INVELTYS. CAUSE REPORTED AS UNKNOWN. LENS DID NOT FAIL TO PERFORM AS INTENDED. FURTHER CLARIFICATION HAS BEEN REQUESTED, BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF -11.5/3.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED AND THE BACKUP LENS OF SAME MODEL AND SIZE WAS IMPLANTED DUE TO THE LENS IMPLANTED UPSIDE DOWN. REPORTEDLY, THE PROBLEM WAS RESOLVED. "IMPROVEMENTS IN OS AFTER REPLACEMENT/EXCHANGE. SLIGHT RESIDUAL ASTIGMATISM DUE TO "HIGH AMOUNTS OF CYL PRE OP. CONTINUE MURO 128, PROLENSA, INVELTYS WITH CHANGE TO FLAREY WHEN OUT OF INVELTYS. CAUSE REPORTED AS UNKNOWN. LENS DID NOT FAIL TO PERFORM AS INTENDED. D6B- (B)(6) 2022. CLAIM# (B)(4). Manufacturer Narrative: EXPLANT DATE:UNKNOWN. CLAIM#: (B)(4). | -11.5/3.0/091 | OS | Other Health Care Professional | 02/14/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13795924&pc=QCB | 2023826-2022-00782 |  | Injury | STAAR SURGICAL COMPANY | 2022-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.1MM TMICL 12.1 IMPLANTABLE COLLAMER LENS OF -7.50/2.0/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6)2022. ON (B)(6) 2022 THE LENS WAS REMOVED AND REPLACED FOR SPHERICAL LENS DUE TO LENS SHIFTED IN THE EYE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN A VIAL WITH LIQUID. VISUAL INPSECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: DATE OF EVENT-UNK. NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#:(B)(4). | -7.50/2.0/088 | NA | Other Health Care Professional | 02/14/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13795938&pc=QCB | 2023826-2022-00783 | 2022-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Unintended Movement | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.1MM TMICL 12.1 IMPLANTABLE COLLAMER LENS OF -7.00/2.0/080 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED AND REPLACED FOR SPHERICAL LENS DUE TO LENS SHIFTED IN THE EYE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: THE REPORTER INDICATED THAT A 12.1MM TMICL 12.1 IMPLANTABLE COLLAMER LENS OF -7.00/2.0/080 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED AND REPLACED FOR SPHERICAL LENS DUE TO LENS SHIFTED IN THE EYE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT A 12.1MM TMICL 12.1 IMPLANTABLE COLLAMER LENS OF -7.00/2.0/080 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2022. ON 10-9B)(6) 2022 THE LENS WAS EXPLANTED AND ON THE SAME DAY DIFFERENT SURGERY A REPLACEMENT SPHERICAL LENS WAS IMPLANTED. THIS WAS DUE TO LENS ROTATION AND REFRACTIVE CHANGE OVER TIME. THE PROBLEM WAS RESOLVED. IMPROVED VISION WITH RESIDUAL ASTIGMATISM. CAUSE REPORTED AS A PATIENT RELATED FACTOR. ADDITIONAL INFORMATION: H6-DEVICE EVALUATION: LENS RETURNED IN A VIAL, DRY WITH RESIDUE ON PRDUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. RESIDUE THROUGHOUT THE WHOLE LENS. LENS HAS AN ORANGE TINT TO IT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: DATE OF EVENT:UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: DATE OF EVENT:UNKNOWN. CLAIM#: (B)(4). | -7.00/2.0/080 | NA | Other Health Care Professional | 02/14/2022 | 2022 | 32
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13802862&pc=QCB | 2023826-2022-00762 | 2022-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -13.0/2.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -13.0/2.0/092 | OS | Physician | 03/06/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13802865&pc=QCB | 2023826-2022-00764 | 2021-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -13.0/2.0/076 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#: (B)(4). | -13.0/2.0/076 | OD | Physician | 03/06/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13802996&pc=QCB | 2023826-2022-00808 | 2022-01-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -8.0/0.5/109 WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#: (B)(4). | -8.0/0.5/109 | OD | Physician | 03/10/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13803125&pc=QCB | 2023826-2022-00741 |  | Injury | STAAR SURGICAL COMPANY | 2022-03-18 04:00 | QCB | PHAKIC TORIC INTRAOCULAR LENS | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 13.2 MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -9.0/4.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED DUE TO SIZING ISSUE. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PATIENT IDENTIFIER-RACE: UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -9.0/4.0/091 | NA | Other | 02/07/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13804288&pc=QCB | 2023826-2022-00809 | 2022-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.0/2.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -12.0/2.5/092 | OS | Physician | 03/11/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13804293&pc=QCB | 2023826-2022-00810 | 2022-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.0/2.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE U.S. DEVICE CODE: (B)(4). CLAIM# (B)(4). | -10.0/2.0/092 | OS | Physician | 03/11/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13804296&pc=QCB | 2023826-2022-00813 | 2022-01-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -12.5/2.5/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE U.S. DEVICE CODE: (B)(4). CLAIM# (B)(4). | -12.5/2.5/094 | OD | Physician | 03/11/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13804300&pc=QCB | 2023826-2022-00814 | 2021-12-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -8.50/1.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE U.S. DEVICE CODE: (B)(4). CLAIM# (B)(4). | -8.50/1.0/093 | OD | Physician | 03/11/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13804307&pc=QCB | 2023826-2022-00815 | 2021-12-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.5/0.5/072 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE U.S. DEVICE CODE: (B)(4). CLAIM# (B)(4). | -9.5/0.5/072 | OS | Physician | 03/11/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13804315&pc=QCB | 2023826-2022-00817 | 2022-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -7.50/0.5/085 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -7.50/0.5/085 | OS | Physician | 03/11/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13808888&pc=QCB | 2023826-2022-00788 | 2021-12-07 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -8.0/1.0/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). ON (B)(6) 2021 THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. NO NEW LENS WAS IMPLANTED. THE PROBLEM WAS RESOLVED. CASE WAS REPORTED AS OTHER. Manufacturer Narrative: H6: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -8.0/1.0/088 | OS | Physician | 02/17/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13810329&pc=QCB | 2023826-2022-00802 | 2022-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/170 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHINICITY: UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -11.50/+2.0/170 | OS | Physician | 02/18/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13810338&pc=QCB | 2023826-2022-00803 | 2021-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50/+1.5/172 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS REMOVED LATER THAT SAME DAY DUE TO THE CRYSTAL SANK AFTER SURGERY, WHICH WAS CAUSED BY THE RELAXATION OF THE SUSPENSORY LIGAMENT. AFTER COMMUNICATION WITH THE PATIENT, THE CRYSTAL WAS REMOVED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | -07.50/+1.5/172 | OS | Physician | 02/18/2022 | 2022 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13810428&pc=QCB | 2023826-2022-00858 | 2021-11-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -9.00/+1.00/83 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2018. IT WAS NOTED THAT THE PATIENT HAD LENS ROTATION ON (B)(6) 2021. THE LENS WAS THEN EXCHANGED ON (B)(6) 2022 WITH A LONGER LENGTH/SPHERICAL LENS AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -9.00/+1.00/83 | OS | Physician | 03/17/2022 | 2022 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13829151&pc=QCB | 2023826-2022-00821 | 2022-01-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -13.5/2.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE U.S. DEVICE CODE: (B)(4). CLAIM# (B)(4). | -13.5/2.0/093 | OD | Physician | 03/13/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13830732&pc=QCB | 2023826-2022-00881 | 2021-05-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -8.00/+4.00/79 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED LENS ROTATION. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT AXIS LENS AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -8.00/+4.00/79 | OS | Physician | 03/18/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13831113&pc=QCB | 2023826-2022-00833 | 2022-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -8.50/1.0/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -8.50/1.0/086 | OS | Physician | 03/16/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13831114&pc=QCB | 2023826-2022-00834 | 2022-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -8.50/1.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -8.50/1.5/092 | OD | Physician | 03/16/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13831116&pc=QCB | 2023826-2022-00835 | 2022-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -7.0/1.0/099 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -7.0/1.0/099 | OD | Physician | 03/16/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13831146&pc=QCB | 2023826-2022-00844 | 2022-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -13.0/3.0/169 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#: (B)(4). | -13.0/3.0/169 | OD | Physician | 03/17/2022 | 2022 | 44 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13831147&pc=QCB | 2023826-2022-00845 | 2022-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.0/3.5/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -9.0/3.5/089 | OS | Physician | 03/17/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13831148&pc=QCB | 2023826-2022-00846 | 2022-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5/3.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM #: (B)(4). | -9.5/3.0/093 | OD | Physician | 03/17/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13831605&pc=QCB | 2023826-2022-00851 | 2022-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 13.2MM TMICL 13.2 IMPLANTABLE COLLAMER LENS OF -5.50/2.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED AND A REPLACEMENT LENS WAS IMPLANTED OF SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: WEIGHT-RACE: UNK. CLAIM#: (B)(4). | -5.50/2.0/092 | OS | Other Health Care Professional | 02/17/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13847457&pc=QCB | 2023826-2022-00899 | 2022-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -15.5/1.5/111 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON THE SAME DAY IN A SEPARATE SURGERY; THE LENS WAS EXPLANTED DUE TO LOW VAULT WITHOUT ROTATION. IT WAS NOTED THE PROBLEM AS NOT RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE PROBLEM WAS RESOLVED AND THERE IS NO PLANS TO REPLACE THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | -15.5/1.5/111 | OD | Physician | 02/20/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13847511&pc=QCB | 2023826-2022-00904 | 2021-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8/2.5/96 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED LENS ROTATION. ON (B)(6) 2022 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. AND NOTED "UNDERSIZING AND ROTATIONAL INSTABILITY EVEN IF VAULT IS ACCEPTABLE". Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2022; THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -8/2.5/96 | OD | Physician | 02/23/2022 | 2022 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13847553&pc=QCB | 2023826-2022-00900 | 2022-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -15.5/2.0/062 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON THE SAME DAY IN A SEPARATE SURGERY; THE LENS WAS EXPLANTED DUE TO LOW VAULT WITHOUT ROTATION. IT WAS NOTED THE PROBLEM AS NOT RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT THE PROBLEM WAS RESOLVED AND THERE IS NO PLANS TO REPLACE THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/ DEBRIS ON LENS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.5/2.0/062 | OS | Physician | 02/20/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13847564&pc=QCB | 2023826-2022-00905 | 2021-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8.5/2.5/51 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LENS ROTATION. ON (B)(6) 2022 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR. AND NOTED "UNDERSIZING AND ROTATIONAL INSTABILITY EVEN IF VAULT IS ACCEPTABLE". Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2022; THE LENS WAS EXPLANTED AND LATER THE SAME DATE REPLACED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN, DEFORMED AND STRETCHED OUT. DIMENSIONAL INSPECTION FOUND THE LENS WIDTH WITHIN SPECIFICATIONS. OVERALL LENGTH (DIAMETER) MEASUREMENTS COULD NOT BE PERFORMED DUE TO HAPTIC FOOTPLATES DAMAGE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -8.5/2.5/51 | OS | Physician | 02/23/2022 | 2022 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13865568&pc=QCB | 2023826-2022-00855 | 2022-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.1MM TMICL 12.1 IMPLANTABLE COLLAMER LENS OF -8.50/4.0/080 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS REMOVED AND A REPLACEMENT LENS WAS IMPLANTED OF LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CASE REPORTED AS OTHER- ICL TOO SHORT. Manufacturer Narrative: (B)(4). | -8.50/4.0/080 | OS | Other Health Care Professional | 02/21/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13865721&pc=QCB | 2023826-2022-00854 | 2022-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.1MM TMICL 12.1 IMPLANTABLE COLLAMER LENS OF -8.50/4.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS REMOVED AND A REPLACEMENT LENS WAS IMPLANTED OF LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CASE REPORTED AS OTHER- ICL TOO SHORT. Manufacturer Narrative: (B)(4). | -8.50/4.0/091 | OD | Other Health Care Professional | 02/21/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869473&pc=QCB | 2023826-2022-00687 | 2021-12-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7, -8.0/0.5/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2021 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM #(B)(4). | -8.0/0.5/077 | OS | Physician | 02/28/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869476&pc=QCB | 2023826-2022-00692 | 2022-02-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.50/1.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -9.50/1.0/086 | OS | Physician | 03/01/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869481&pc=QCB | 2023826-2022-00717 | 2021-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.00/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -11.00/1.5/089 | OD | Physician | 03/03/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869487&pc=QCB | 2023826-2022-00723 | 2021-11-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.50/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.50/1.5/092 | OD | Physician | 03/04/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869554&pc=QCB | 2023826-2022-00682 | 2022-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -10.00/1.5/092 | OD | Physician | 02/28/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869564&pc=QCB | 2023826-2022-00703 | 2021-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -6.50/1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: 723766. | -6.50/1.0/092 | OS | Physician | 03/02/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869567&pc=QCB | 2023826-2022-00721 | 2001-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED VERTICALLY A 12.1MM VTICMO12.1, -10.50/2.5/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS HORIZONTALLY DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#: (B)(4). | -10.50/2.5/074 | OD | Physician | 03/04/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869602&pc=QCB | 2023826-2022-00686 | 2021-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.00/2.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -9.00/2.5/088 | OS | Physician | 02/28/2022 | 2022 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869604&pc=QCB | 2023826-2022-00693 | 2022-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -13.50/4.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -13.50/4.0/087 | OS | Physician | 03/02/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869608&pc=QCB | 2023826-2022-00700 | 2022-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.5/1.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -7.5/1.0/075 | OS | Physician | 03/02/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869613&pc=QCB | 2023826-2022-00689 | 2022-02-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.50/3.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -9.50/3.0/090 | OD | Physician | 03/01/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869616&pc=QCB | 2023826-2022-00698 | 2022-01-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -16.5/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -16.5/2.5/091 | OS | Physician | 03/02/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869620&pc=QCB | 2023826-2022-00718 | 2021-08-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.50/1.0/090 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.50/1.0/090 | OS | Physician | 03/03/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869626&pc=QCB | 2023826-2022-00709 | 2022-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.00/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -12.00/1.5/095 | OS | Physician | 03/02/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869656&pc=QCB | 2023826-2022-00685 | 2021-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.50/2.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -9.50/2.5/094 | OD | Physician | 02/28/2022 | 2022 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869659&pc=QCB | 2023826-2022-00694 | 2021-12-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.50/2.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.50/2.5/089 | OD | Physician | 03/02/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869661&pc=QCB | 2023826-2022-00699 | 2020-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.50/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK RACE: UNK PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM#: (B)(4). | -15.50/2.0/091 | OD | Physician | 03/02/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869665&pc=QCB | 2023826-2022-00707 | 2021-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.50/1.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#: (B)(4). | -9.50/1.0/095 | OD | Physician | 03/02/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869666&pc=QCB | 2023826-2022-00688 | 2022-01-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -8.0/1.5/018 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -8.0/1.5/018 | OD | Physician | 03/01/2022 | 2022 | 44 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869668&pc=QCB | 2023826-2022-00696 | 2022-01-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.00/2.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#: (B)(4). | -13.00/2.0/084 | OS | Physician | 03/02/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869676&pc=QCB | 2023826-2022-00719 | 2022-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -13.50/2.5/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT :UNKNOWN, ETHNICITY: UNKNOWN, SEX: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -13.50/2.5/075 | OS | Physician | 03/03/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869762&pc=QCB | 2023826-2022-00683 | 2022-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.50/2.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RASE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -09.50/2.5/089 | OS | Physician | 02/28/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869771&pc=QCB | 2023826-2022-00708 | 2022-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED VERTICALLY A 12.1MM VTICMO12.1, -15.00/3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (N)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: 723771. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -15.00/3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS IMPLANTED VERTICALLY DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. CLAIM# (B)(4). | -15.00/3.0/093 | OD | Physician | 03/02/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13869773&pc=QCB | 2023826-2022-00720 | 2021-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -09.50/1.0/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -09.50/1.0/070 | OD | Physician | 03/04/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13877965&pc=QCB | 2023826-2022-00893 | 2022-01-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50/+6.0/042 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE DEBRIS ON PRODUCT IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM #: (B)(4). | -08.50/+6.0/042 | NA | Physician | 03/02/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13878666&pc=QCB | 2023826-2022-00973 | 2022-01-19 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -7.0/1.5/067 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. INTRAOPERATIVELY THE LENS WAS REMOVED. THE LENS WAS NOT REPLACED BUT INSTEAD THE SURGEON NOTED THE OPERATION WAS CHANGED TO EXCIMER LASER OPERATION. IT WAS INDICATED THE PROBLEM WAS NOT RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER CLARIFIED THAT THERE IS NO PLANS TO IMPLANT ANOTHER LENS, AFTER THE EVENTS OF THE SURGERY EVERYTHING IS GOOD AND THERE IS NOT ANY FURTHER DECISION DECIDED TO RESOLVE THE ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -7.0/1.5/067 | OS | Physician | 03/21/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13878798&pc=QCB | 2023826-2022-00942 | 2022-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -12.5/3.5/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -12.5/3.5/093 | OD | Physician | 02/27/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13878893&pc=QCB | 2023826-2022-00955 | 2021-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.5/2.0/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH LENS BUT IMPLANTED 48 DEGREES CLOCKWISE. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: PRODUCT EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -10.5/2.0/087 | OS | Physician | 03/18/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13894525&pc=QCB | 2023826-2022-00946 | 2022-02-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -14.5/1.0/91 (SPHERE/CYLINDER/AXIS) LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY FROM THE PATIENTS LEFT EYE (OS) ON (B)(4) 2022. THIS WAS DUE TO A LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTING FOUND THE HAPATIC TORN AND BROKEN. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO LENS HAPTIC BEING RETURNED BROKEN. CLAIM# (B)(4). | -14.5/1.0/91 | OS | Physician | 03/02/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13895589&pc=QCB | 2023826-2022-00896 | 2022-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.50/+1.0/101 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT THE PROBLEM WAS NOT RESOLVED SO THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN AND BROKEN. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO STATE OF DAMAGE LENS RETURNED. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -15.50/+1.0/101 | OD | Physician | 02/28/2022 | 2022 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13895715&pc=QCB | 2023826-2022-00894 | 2022-01-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+1.5/127 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS ROTATED 42 DEGREES CW AFTER HORIZONTAL IMPLANTATION (DURING SAME SURGERY). THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4). | -15.00/+1.5/127 | NA | Physician | 02/28/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13895756&pc=QCB | 2023826-2022-00897 | 2021-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.50/+2.0/106 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5: THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT THE PROBLEM WAS NOT RESOLVED SO THE LENS WAS EXPLANTED. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN AND BROKEN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -12.50/+2.0/106 | OD | Physician | 03/01/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13895760&pc=QCB | 2023826-2022-00898 | 2021-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+1.5/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. DEVICE CODE: (B)(4). CLAIM # (B)(4). | -09.00/+1.5/093 | OD | Physician | 03/02/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13895826&pc=QCB | 2023826-2022-00895 | 2021-09-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/110 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (LENS IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM #: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -10.50/+1.0/110 | NA | Physician | 03/03/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13911544&pc=QCB | 2023826-2022-00960 | 2021-10-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -9.0/+1.0/009 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY THE INITIAL LENS WAS IMPLANTED VERTICALLY. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC BROKEN, BENT AND DEFORMED TO THE LENS. AFTER REHYDRATION IT WAS VISIBLY CLEAR HAPTIC FOOTPLATE WAS BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/+1.0/009 | OD | Physician | 02/28/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13911561&pc=QCB | 2023826-2022-00966 | 2021-07-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT, WHILE ATTEMPTING TO IMPLANT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/+2.0/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS BECAME STUCK IN THE CARTRIDGE/INJECTION SYSTEM. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRA-OPERATIVELY ON (B)(6) 2021. REPORTEDLY, THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. SEE MFR REP# 2023826-2022-00967 FOR REPLACEMENT LENS. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -13.0/+2.0/085 | OS | Physician | 03/03/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13911593&pc=QCB | 2023826-2022-00969 | 2022-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -6.5/+0.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4). | -6.5/+0.5/090 | OS | Physician | 03/04/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13911596&pc=QCB | 2023826-2022-00958 | 2022-01-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -12.5/+3.0/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN AND BENT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+3.0/092 | OD | Physician | 02/28/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13911609&pc=QCB | 2023826-2022-00967 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -13.0/+2.0/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. REPORTEDLY, THE PROBLEM FROM THE PREVIOUS IMPLANT WAS NOT RESOLVED. SEE MFR REP# 2023826-2022-00966 FOR INITIAL LENS. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: B5 - IT WAS LATER REPORTED THAT THE PROBLEM WAS RESOLVED. CLAIM#: (B)(4). | -13.0/+2.0/086 | OS | Physician | 03/03/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13911612&pc=QCB | 2023826-2022-00959 | 2022-02-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.0/+5.0/103 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM #: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC TORN/SPLIT ON THE LENS. LENS RETURNED IN 2 HALVES, UNABLE TO PERFORM DIMENSIONAL INSPECTION. CLAIM# (B)(4). | -12.0/+5.0/103 | OS | Physician | 02/28/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13937894&pc=QCB | 2023826-2022-00871 | 2021-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.0/1.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.0/1.0/088 | OS | Physician | 03/10/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13937895&pc=QCB | 2023826-2022-00869 | 2022-01-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -11.00/3.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.00/3.0/087 | OS | Physician | 03/11/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13937898&pc=QCB | 2023826-2022-00875 | 2022-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.0/2.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -12.0/2.0/086 | OD | Physician | 03/11/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13937899&pc=QCB | 2023826-2022-00876 | 2022-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.0/2.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -12.0/2.5/088 | OS | Physician | 03/11/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13937900&pc=QCB | 2023826-2022-00862 | 2022-01-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.5/2.0/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.5/2.0/112 | OD | Physician | 03/06/2022 | 2022 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13937902&pc=QCB | 2023826-2022-00859 | 2021-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.5/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT :UNKNOWN, ETHINICITY:UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -14.5/1.5/089 | OS | Physician | 03/05/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13937907&pc=QCB | 2023826-2022-00873 | 2022-02-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.5/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -13.5/1.0/085 | OD | Physician | 03/11/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13937913&pc=QCB | 2023826-2022-00878 | 2022-02-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.0/3.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -13.0/3.0/092 | OS | Physician | 03/11/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13937943&pc=QCB | 2023826-2022-00879 | 2022-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -9.0/2.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -9.0/2.5/095 | OS | Physician | 03/11/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13937974&pc=QCB | 2023826-2022-00626 | 2022-01-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.0/1.5/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.0/1.5/079 | OS | Physician | 03/05/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13937987&pc=QCB | 2023826-2022-00872 | 2021-12-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.0/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK, ETHNICITY- UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.0/1.0/093 | OD | Physician | 03/10/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13938002&pc=QCB | 2023826-2022-00877 | 2022-02-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -10.5/2.0/091 | OD | Physician | 03/11/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13944315&pc=QCB | 2023826-2022-01023 | 2022-03-01 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2; -14.50/1.0/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE ON (B)(6) 2022. IT WAS NOTICED AFTER IMPLANTATION THE LENS HAD A NOTCH NEAR THE HAPTIC. THE LENS WAS REMOVED AND REPLACED WITH A BACK UP LENS INTRAOPERATIVELY WITH NO COMPLICATIONS. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL IN LIQUID. VISUAL INSPECTION FOUND THE LENS OPTIC AND HAPTIC TORN. ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | -14.50/1.0/088 | NA | Other Health Care Professional | 03/01/2022 | 2022 | 18
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13967160&pc=QCB | 2023826-2022-00954 | 2022-03-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF -9.50/1.0/090 (SPHERE/CYLINDER/AXIS) WAS BEING IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022 WHEN THE LENS WAS STUCK IN CARTRIDGE OR INJECTION SYSTEM. THE LENS WAS IMPLANTED; REMOVED AND REPLACED INTRA-OPERATIVELY ON (B)(6) 2022 FOR A SAME MODEL; SIZE AND POWER LENS. THE REPLACEMENT LENS RESOLVED THE PROBLEM. VISION IS "GOOD" GRADUALLY IMPROVING. CAUSE REPORTED AS USER ERROR. "LOADING ERROR". Manufacturer Narrative: (B)(4). | -9.50/1.0/090 | OS | Other Health Care Professional | 03/08/2022 | 2022 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13967795&pc=QCB | 2023826-2022-01022 | 2022-01-25 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-03-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -10.5/1.0/108 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). ON (B)(6) 2022 THE LENS WAS IMPLANTED; REMOVED AND REPLACED INTRA-OPERATIVELY FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE DID NOT RESOLVE THE PROBLEM. CAUSE REPORTED AS UNKNOWN. SEE MFR # 2023826-2022-00952 FOR EXCHANGE LENS. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -10.5/1.0/108 | OS | Physician | 03/11/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13967852&pc=QCB | 2023826-2022-00952 | 2022-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-03-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -10.5/1.5/96 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 AS AN EXCHANGE LENS TO CORRECT LOW VAULT. LOW VAULT IS STILL BEING OBSERVED. THE LENS REMAINS IMPLANTED. "THE DOCTOR CONSIDER CLOSE FOLLOW UP OF THIS PATIENT AND DO NOT CONSIDER CHANGING LENS FOR THE TIME BEING." CAUSE REPORTED AS UNKNOWN. SEE MFR# 2023826-2022-01022 FOR INITIAL LENS. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -10.5/1.5/96 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LOW VAULT WAS OBSERVED. ON (B)(6) 2022, LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -10.5/1.5/96 | OS | Physician | 10/07/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13979951&pc=QCB | 2023826-2022-01045 | 2021-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.0/2.0/079 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS OF DIFFERENT POWER DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -10.0/2.0/079 | OS | Physician | 03/03/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13981216&pc=QCB | 2023826-2022-01055 | 2022-02-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -4.0/1.0/013 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND LATER ON THE SAME DATE REPLACED WITH A SAME LENGTH/MODEL LENS BUT IMPLANTED VERTICALLY. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION-THE LENS WAS RETURNED DRY IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC DEFORMED, HAPTIC BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -4.0/1.0/013 | OS | Physician | 10/11/2022 | 2022 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13981507&pc=QCB | 2023826-2022-01044 | 2022-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.0/2.5/103 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING WITH ROTATION. ON (B)(6) 2022 THE SURGEON REPOSITIONED THE LENS BUT THIS DID NOT RESOLVE THE PROBLEM. THE SURGEON ALSO EXPLANTED THE LENS ON (B)(6) 2022 AND THEN REPLACED WITH LONGER LENGTH LENS LATER THAT. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -6.0/2.5/103 | OD | Physician | 04/01/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13983430&pc=QCB | 2023826-2022-01039 | 2021-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.0/+4.0/085 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. THE PROBLEM WAS NOT RESOLVED. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT AXIS LENS AND THE PROBLEM WAS RESOLVED. REPORTEDLY THE REPLACEMENT LENS WAS IMPLANTED VERTICALLY. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND ONE OF THE HAPTICS BROKEN AND FIBRE ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -11.0/+4.0/085 | OS | Physician | 03/11/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13983675&pc=QCB | 2023826-2022-01036 | 2021-11-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -14.5/+1.5/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. EXCESSIVE VAULT IS REPORTED. THE LENS REMAINS IMPLANTED. THE CAUSE IS REPORTED AS UNKNOWN; THE DOCTOR STATED THAT THE PATIENT DID NOT HAVE ANY PROBLEMS. Manufacturer Narrative: WEIGHT. ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -14.5/+1.5/094 | OD | Physician | 03/10/2022 | 2022 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13983728&pc=QCB | 2023826-2022-01040 | 2021-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+4.0/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER SPHERICAL MODEL LENS DUE TO LENS ROTATION AND THE PROBLEM WAS RESOLVED. REPORTEDLY THE RESIDUAL ASTIGMATISM WILL BE TREATED BY LASER TREATMENT. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -15.0/+4.0/088 | OS | Physician | 03/11/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13983744&pc=QCB | 2023826-2022-01037 | 2021-11-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -13.5/+1.0/133 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. EXCESSIVE VAULT IS REPORTED. THE LENS REMAINS IMPLANTED. THE CAUSE IS REPORTED AS UNKNOWN; THE DOCTOR STATED THAT THE PATIENT DID NOT HAVE ANY PROBLEMS. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -13.5/+1.0/133 | OS | Physician | 03/10/2022 | 2022 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13983752&pc=QCB | 2023826-2022-01038 | 2021-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.0/+3.5/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. THE PROBLEM WAS NOT RESOLVED. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT AXIS LENS AND THE PROBLEM WAS RESOLVED. REPORTEDLY THE REPLACEMENT LENS WAS IMPLANTED VERTICALLY. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -10.0/+3.5/089 | OD | Physician | 03/11/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13983787&pc=QCB | 2023826-2022-01041 | 2021-12-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -8.0/+1.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT SPHERE/AXIS LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, AFTER THE FIRST HORIZONTAL IMPLANTATION THE LENS WAS ROTATED 42 DEGREES CW. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT AND DEFORMED. DIMENSIONAL ANALYSIS PERFORMED FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -8.0/+1.5/090 | OS | Physician | 09/21/2022 | 2022 | 43 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13995849&pc=QCB | 2023826-2022-00911 | 2022-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/2.5/180 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO.: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.5/2.5/180 | OD | Physician | 03/14/2022 | 2022 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13995850&pc=QCB | 2023826-2022-00916 | 2022-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6, -11.50/1.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS PATIENT RELATED FACTOR, RELATED PATIENT EYE ANATOMY. Manufacturer Narrative: AGE OR DATE OF BIRTH:UNKNOWN, WEIGHT: UNKNOWN, ETHINICITY: UNKNOWN, SEX: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.50/1.0/082 | OS | Physician | 03/16/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13995853&pc=QCB | 2023826-2022-00919 | 2022-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -8.0/3.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO.: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -8.0/3.5/094 | OD | Physician | 03/17/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13995857&pc=QCB | 2023826-2022-00923 | 2022-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.0/1.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -14.0/1.5/082 | OD | Physician | 03/17/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13995858&pc=QCB | 2023826-2022-00918 | 2022-03-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.5/2.5/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED VERTICALLY WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -14.5/2.5/084 | OD | Physician | 03/16/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13995861&pc=QCB | 2023826-2022-00914 | 2022-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -17.0/1.0/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). OVER 45 YRS OF AGE AT DATE OF IMPLANT. CLAIM# (B)(4). | -17.0/1.0/074 | OD | Physician | 03/16/2022 | 2022 | 50 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13995863&pc=QCB | 2023826-2022-00921 | 2021-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.0/2.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -12.0/2.0/092 | OD | Physician | 03/17/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13995864&pc=QCB | 2023826-2022-00926 | 2022-03-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -07.50/1.0/050 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY - UNK, RACE- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -07.50/1.0/050 | OS | Physician | 03/17/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13995865&pc=QCB | 2023826-2022-00922 | 2022-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.0/4.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -14.0/4.0/092 | OD | Physician | 03/17/2021 | 2021 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13995868&pc=QCB | 2023826-2022-00930 | 2022-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5 12.1, -11.50/3.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -11.50/3.0/086 | OS | Physician | 03/17/2022 | 2022 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13995869&pc=QCB | 2023826-2022-00934 | 2022-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -14.0/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -14.0/2.0/095 | OS | Physician | 03/18/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13995871&pc=QCB | 2023826-2022-00920 | 2022-03-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -6.0/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -6.0/2.0/091 | OD | Physician | 03/17/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13995872&pc=QCB | 2023826-2022-00913 | 2021-12-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.0/1.0/068 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.0/1.0/068 | OD | Physician | 03/14/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13995876&pc=QCB | 2023826-2022-00935 | 2022-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.0/4.0/003 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.0/4.0/003 | OD | Physician | 03/18/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13995884&pc=QCB | 2023826-2022-00929 | 2022-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.00/3.50/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -15.00/3.50/071 | OD | Physician | 03/17/2022 | 2022 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13995885&pc=QCB | 2023826-2022-00931 | 2022-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -13.50/2.50/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -13.50/2.50/083 | OS | Physician | 03/17/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13995894&pc=QCB | 2023826-2022-00936 | 2022-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -7.5/3.5/180 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT :UNKNOWN, ETHINICITY:UNKNOWN, SEX:UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -7.5/3.5/180 | OS | Physician | 03/18/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13997991&pc=QCB | 2023826-2022-00823 | 2022-01-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.5/2.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS REMOVED DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. REPORTEDLY THE SHAPE OF CILIARY SULCUS WAS NOT GOOD, SO DECISION WAS NOT TO IMPLANT AGAIN. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -12.5/2.0/091 | OS | Physician | 03/11/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13998018&pc=QCB | 2023826-2022-00822 | 2022-01-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.5/1.5/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS REMOVED DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. REPORTEDLY THE SHAPE OF CILIARY SULCUS WAS NOT GOOD, SO DECISION WAS NOT TO IMPLANT AGAIN. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE CODE 1494: OFF-LABEL USE (ACD<3MM). CLAIM# (B)(4). | -10.5/1.5/089 | OD | Physician | 03/11/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13998033&pc=QCB | 2023826-2022-00974 | 2021-02-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -15.5/1.0/107 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED WITH A LOW VAULT OBSERVED. REPOSITIONING WAS PERFORMED BUT DID NOT RESOLVED. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS OF DIFFERENT AXIS WHICH WAS IMPLANTED VERTICALLY. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -15.5/1.0/107 | OD | Physician | 03/19/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13998213&pc=QCB | 2023826-2022-00976 | 2022-02-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -17.5/5.5/079 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. MISALIGNMENT FOUND AFTER LENS IMPLANTATION WAS REPORTED. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS OF DIFFERENT CYLINDER/AXIS. SEE MFR. REP#: 2023826-2022-00977 FOR REPLACEMENT LENS. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -17.5/5.5/079 | OS | Physician | 03/19/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=13998454&pc=QCB | 2023826-2022-00979 | 2021-11-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 -10.0/2.5/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REMOVED WITHIN THE SAME SURGERY AND REPLACED WITH AN IDENTICAL LENS DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. NO PATIENT INJURY. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#: (B)(4). | -10.0/2.5/088 | OD | Physician | 03/22/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14004330&pc=QCB | 2023826-2022-01053 | 2021-12-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5/1.5/108 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED AND REPLACED WITH A LONGER LENGTH LENS THAT WAS IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN AND BROKEN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #: (B)(4). | -9.5/1.5/108 | OS | Physician | 03/11/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14004890&pc=QCB | 2023826-2022-01051 | 2021-12-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -11.0/1.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS THAT WAS IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -11.0/1.0/092 | OD | Physician | 03/10/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14015269&pc=QCB | 2023826-2022-00986 | 2021-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -13.5/3.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022; THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -13.5/3.0/090 | OS | Physician | 03/14/2022 | 2022 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14015276&pc=QCB | 2023826-2022-00987 | 2021-11-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -7.5/1.0/105 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)2021 THE LENS WAS EXCHANGED FOR A SAME SIZE LENS DUE TO EXCESSIVE VAULT. A HORIZONTAL IMPLANTATION LENS WAS REPLACED WITH A VERTICAL IMPLANTED ONE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). (B)(4). | -7.5/1.0/105 | OS | Physician | 03/14/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14015336&pc=QCB | 2023826-2022-00988 | 2022-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6 VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.0/2.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). | -13.0/2.0/094 | OD | Physician | 03/14/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14015338&pc=QCB | 2023826-2022-00989 | 2021-12-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2 VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.5/3.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -11.5/3.0/089 | OS | Physician | 03/14/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14022356&pc=QCB | 2023826-2022-01106 | 2021-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -13.0/2.5/101 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. THE REPOSITIONING DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS REMOVED AND A REPLACEMENT LENS OF LONGER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION:UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: H6-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -13.0/2.5/101 | OS | Physician | 03/11/2022 | 2022 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14052181&pc=QCB | 2023826-2022-01058 | 2022-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -12.5/1.50/87 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6)2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.5/1.50/87 | OD | Physician | 03/18/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14052244&pc=QCB | 2023826-2022-01059 | 2022-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -11.5/1.50/88 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PATIENT INFORMATION:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.5/1.50/88 | OS | Physician | 03/18/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14061331&pc=QCB | 2023826-2022-01133 | 2022-01-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.5/+3.0/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. THE PROBLEM WAS NOT RESOLVED. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS. SEE MFR REP# 2023826-2022-01134 FOR REPLACEMENT LENS. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND OPTIC TORN AND HAPTIC BROKEN, BENT AND DEFORMED. DIMENSIONAL INSPECTION OF THE DIAMETER FOUND THE LENS TO BE WITHIN SPECIFICATION. UNABLE TO COMPLETE DIMENSIONAL INSPECTION OF THE WIDTH DUE TO LENS DAMAGE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -12.5/+3.0/080 | OD | Physician | 03/16/2022 | 2022 | 41 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14061337&pc=QCB | 2023826-2022-01140 | 2022-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -11.0/+2.0/095 | OD | Physician | 03/18/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14061340&pc=QCB | 2023826-2022-01141 | 2022-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.0/+2.0/084 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM # (B)(4). | -11.0/+2.0/084 | OS | Physician | 03/18/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14061341&pc=QCB | 2023826-2022-01138 | 2022-02-21 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+3.0/179 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED WITH AN ALTERNATE LENS INTRA-OPERATIVELY ON (B)(6) 2022.TTHE PROBLEM WAS RESOLVED. REPORTEDLY, THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -12.5/+3.0/179 | OS | Physician | 03/17/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14061342&pc=QCB | 2023826-2022-01125 | 2022-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS -10.0/+1.0/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT (OS) EYE ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO LOW VAULT. IT WAS REPLACED WITH A LONGER LENS. Manufacturer Narrative: B5-PREVIOUS STATEMENT: THE LENS WAS EXPLANTED DUE TO LOW VAULT AND WAS REPLACED WITH A LONGER LENS. CORRECTED STATEMENT: THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM IS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH, WEIGHT, ETHNICITY, RACE- UNK. DATE OF EVENT- UNK. CLAIM# (B)(4). | -10.0/+1.0/082 | OS | Physician | 03/17/2022 | 2022 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14061345&pc=QCB | 2023826-2022-01139 | 2022-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+1.5/075 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -13.5/+1.5/075 | OS | Physician | 03/18/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14061346&pc=QCB | 2023826-2022-01134 | 2022-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -12.0/+3.0/081 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO LENS ROTATION. REPORTEDLY, THERE ARE NO PLANS TO IMPLANT A REPLACEMENT LENS. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -12.0/+3.0/081 | OD | Physician | 03/16/2022 | 2022 | 41 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14061347&pc=QCB | 2023826-2022-01137 | 2022-02-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -8.0/+0.5/038 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT AXIS LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, AFTER THE FIRST HORIZONTAL IMPLANTATION THE LENS WAS ROTATED 42 DEGREES CW. THE FIRST LENS WAS IMPLANTED HORIZONTALLY AND THE REPLACEMENT LENS WAS IMPLANTED VERTICALLY. Manufacturer Narrative: DEVICE EVALUATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -8.0/+0.5/038 | OS | Physician | 03/17/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14061350&pc=QCB | 2023826-2022-01132 | 2021-12-27 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/084 (SPHERE/CYLINDER/AXIS) TORE DURING INSERTION INTO THE EYE. THIS OCCURRED ON (B)(6) 2021. THE LENS WAS REMOVED AND REPLACED INTRAOPERATIVELY WITH AN ALTERNATE LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -8.5/+1.5/084 | NA | Physician | 03/16/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14061352&pc=QCB | 2023826-2022-01136 | 2022-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATES THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.5/+3.0/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION BUT THE PROBLEM WAS NOT RESOLVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY THE REPLACEMENT LENS WAS VERTICALLY IMPLANTED. Manufacturer Narrative: H3: DEVICE EVALUATION LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -12.5/+3.0/093 | OD | Physician | 03/17/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14061356&pc=QCB | 2023826-2022-01124 | 2022-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS -10.5/+1.5/099 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT (OD) EYE ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO LOW VAULT. IT WAS REPLACED WITH A LONGER LENS. Manufacturer Narrative: B5: PREVIOUS STATEMENT: LENS WAS EXPLANTED DUE TO LOW VAULT AND REPLACED WITH A LONGER LENS. CORRECTED STATEMENT: LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM IS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH, WEIGHT, ETHNICITY, RACE- UNK. DATE OF EVENT- UNK. CLAIM# (B)(4). | -10.5/+1.5/099 | OD | Physician | 03/17/2022 | 2022 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14061359&pc=QCB | 2023826-2022-01126 | 2021-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -13.0/+2.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT AXIS LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, AFTER THE FIRST HORIZONTAL IMPLANTATION WAS REPLACED WITH A VERTICALLY IMPLANTED LENS. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -13.0/+2.0/089 | OS | Physician | 03/14/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14061360&pc=QCB | 2023826-2022-01127 | 2021-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -13.5/+2.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT AXIS LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY A HORIZONTALLY IMPLANTED LENS WAS REPLACED WITH A VERTICALLY IMPLANTED ONE. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE- UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). Manufacturer Narrative: PRODUCT EVALUATION: "H3- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS." CORRECTED DATA: H6- INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM: (B)(4). | -13.5/+2.5/094 | OS | Physician | 03/14/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14078041&pc=QCB | 2023826-2022-00994 | 2022-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -16.0/3.5/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE U.S. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -16.0/3.5/077 | OS | Physician | 03/19/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14078044&pc=QCB | 2023826-2022-00997 | 2021-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -09.5/1.5/117 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE U.S.(B)(4). CLAIM# (B)(4). | -09.5/1.5/117 | OD | Physician | 03/21/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14078200&pc=QCB | 2023826-2022-00999 | 2021-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.5/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.5/2.0/090 | OD | Physician | 03/21/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14078206&pc=QCB | 2023826-2022-01004 | 2022-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.5/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -11.5/1.5/092 | OS | Physician | 03/19/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14078214&pc=QCB | 2023826-2022-01008 | 2022-01-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.0/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE U.S. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -12.0/2.0/091 | OD | Physician | 03/19/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14078215&pc=QCB | 2023826-2022-01007 | 2022-01-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.0/0.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -10.0/0.5/092 | OD | Physician | 03/19/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14078218&pc=QCB | 2023826-2022-01016 | 2022-02-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.0/3.0/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -09.0/3.0/072 | OS | Physician | 03/22/2022 | 2022 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14078223&pc=QCB | 2023826-2022-01019 | 2021-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.5/5.0/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT:UNKNOWN, ETHINICITY:UNKNOWN, RACE:UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.5/5.0/098 | OD | Physician | 03/23/2022 | 2022 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14078365&pc=QCB | 2023826-2022-00912 | 2022-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -15.5/2.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE U.S. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -15.5/2.5/089 | OD | Physician | 03/19/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14078401&pc=QCB | 2023826-2022-01014 | 2021-08-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -09.50/2.0/125 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE U.S. (B)(4). CLAIM# (B)(4). | -09.50/2.0/125 | OD | Physician | 03/20/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14078408&pc=QCB | 2023826-2022-01017 | 2021-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.0/3.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. PMA/510K - THIS PRODUCT IS NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -9.0/3.5/086 | OS | Physician | 03/22/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14089671&pc=QCB | 2023826-2022-01110 | 2022-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -9.50/3.5/103 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED AND ON THE SAME DAY DIFFERENT SURGERY A REPLACEMENT LENS OF LONGER LENGTH LENS WAS IMPLANTED DUE TO LOW VAULT WITH ROTATION. THE REPLACEMENT RESOLVED THE PROBLEM. CAUSE REPORTED AS DEVICE. "LENS ROTATED EVEN THOUGH VAULT WAS 300 WHICH WOULD NORMALLY HAVE BEEN OK, BUT TOO SMALL FOR A TORIC PERHAPS." Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -9.50/3.5/103 | OS | Physician | 03/14/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14092563&pc=QCB | 2023826-2022-01196 | 2021-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 13.7MM TMICL13.7; -10.50/2.0/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS; WAS NOTICED TO BE DEFECTIVE AND INDICATED "SCRATCH OR CHUNK OUT OF LENS". IT WAS ALSO NOTED THERE WAS PATIENT CONTACT AND A BACK UP LENS WAS USED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT A 13.7MM TMICL13.7; -10.50/2.0/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS; WAS NOTICED TO BE DEFECTIVE AND INDICATED "SCRATCH OR CHUNK OUT OF LENS". LENS IMPLANTED THEN EXPLANTED. IT WAS ALSO NOTED THERE WAS PATIENT CONTACT AND A BACKUP LENS WAS USED. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A VIAL, DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN WITH THE LENS COVERED IN THICK RESIDUE. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). | -10.50/2.0/086 | NA | Other Health Care Professional | 03/18/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14107514&pc=QCB | 2023826-2022-01173 | 2021-03-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -12.00/+4.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM # (B)(4). | -12.00/+4.0/090 | OD | Other Health Care Professional | 03/22/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14107605&pc=QCB | 2023826-2022-01189 | 2022-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -13.0/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS LATER EXCHANGED FOR A SAME LENGTH LENS IMPLANTED VERTICALLY AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN, BENT, AND DEFORMED TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -13.0/2.0/091 | OD | Physician | 03/14/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14108109&pc=QCB | 2023826-2022-01198 | 2022-02-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.5/1.5/71 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH LENS BUT DIFFERENT AXIS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISIBLE INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -12.5/1.5/71 | OS | Physician | 09/26/2022 | 2022 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14112900&pc=QCB | 2023826-2022-00977 | 2022-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -17.5/5.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 AS A REPLACEMENT LENS. REPORTEDLY THE LENS REMAINED DEVIATED AFTER REPLACEMENT. DEFLECTION IS STILL OBVIOUS AND THE VAULT IS ONLY 219UM. DOCTOR CONSIDERS PATIENT'S SUSPENSORY LIGAMENT DYSPLASIA. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: PT INFO:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -17.5/5.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 AS A REPLACEMENT LENS. REPORTEDLY THE LENS REMAINED DEVIATED AFTER REPLACEMENT. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH RESIDE/DEBRIS ON PRODUCT IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -17.5/5.0/091 | OS | Physician | 04/14/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14123272&pc=QCB | 2023826-2022-01211 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-04-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF -5.5/1.5/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) AS AN EXCHANGE LENS. REPORTEDLY THE PROBLEM OF EXCESSIVE VAULTING IS RESOLVED. "BUT NOW WITH MOD RE LIKELY NEEDING ROTATION 1.25+1.50X88. MAY PERFORM LRI OR PRK, DILATED EXAM NEXT WEEK TO DETERMINE ICL POSITION. WE WANT TO AVOID INTRAOCULAR MANIPULATION, IF POSSIBLE. THE 2.00 TORIC LENS WAS NOT AVAIL IN TIME FOR EXCHANGE, NEEDED TO USE LOVER TORIC.'' THE LENS REMAINS IMPLANTED AND NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR# 2023826-2022-00853 FOR INITIAL IMPLANT. Manufacturer Narrative: WEIGHT-RACE: UNK. LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -5.5/1.5/089 | OS | Other Health Care Professional | 03/18/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14123275&pc=QCB | 2023826-2022-01209 | 2022-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -14.0/2.0/008 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS DUE TO LOW VAULT. THE REPLACEMENT LENS WAS IMPLANTED VERTICALLY AND THE PROBLEM WAS RESOLVED. THE CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEFORMED WITH FOREIGN MATERIAL ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -14.0/2.0/008 | OD | Physician | 03/17/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14123276&pc=QCB | 2023826-2022-01210 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-04-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF -8.5/1.0/082 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) AS AN EXCHANGE LENS. REPORTEDLY THE PROBLEM OF EXCESSIVE VAULTING IS RESOLVED. "BUT NOW WITH MOD RE LIKELY NEEDING ROTATION -0.50+1.00X81. MAY PERFORM LRI DILATED EXAM NEXT WEEK TO DETERMINE ICL POSITION. WE WANT TO AVOID INTRAOCULAR MANIPULATION, IF POSSIBLE." THE LENS REMAINS IMPLANTED AND NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR # 2023826-2022-00850 FOR INITIAL IMPLANT. Manufacturer Narrative: WEIGHT-RACE: UNK. CLAIM#: (B)(4). | -8.5/1.0/082 | OD | Other Health Care Professional | 03/18/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14138531&pc=QCB | 2023826-2022-01230 | 2022-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -9.0/2.0/083 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH/MODEL LENS BUT IMPLANTED VERTICALLY. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -9.0/2.0/083 | OS | Physician | 04/12/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14138532&pc=QCB | 2023826-2022-01229 | 2022-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -9.5/1.5/94 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH LENS BUT IMPLANTED VERTICALLY AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN/BENT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -9.5/1.5/94 | OD | Physician | 04/12/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14138551&pc=QCB | 2023826-2022-01176 | 2022-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.50/+1.0/012 (SPHERE/CYLINDER/AXIS), (IMPLANTED VERTICALLY) IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED TRANVERSALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -08.50/+1.0/012 | OD | Physician | 03/21/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14139052&pc=QCB | 2023826-2022-01062 | 2022-03-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT.. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). CLAIM# (B)(4). | -9.0/1.0/093 | OS | Physician | 03/21/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14139056&pc=QCB | 2023826-2022-01064 | 2022-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MMVTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -13.5/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. IN A SUBSEQUENT SURGERY LATER THAT DAY, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.5/2.0/089 | OD | Physician | 03/21/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14139057&pc=QCB | 2023826-2022-01065 | 2022-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -13.0/1.5/110 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6)2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.0/1.5/110 | OD | Physician | 03/21/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14139059&pc=QCB | 2023826-2022-01066 | 2021-12-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.5/2.5/116 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. IN A SUBSEQUENT SURGERY LATER THAT DAY, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -10.5/2.5/116 | OS | Physician | 03/21/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14139060&pc=QCB | 2023826-2022-01067 | 2021-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MMVTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -14.5/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. IN A SUBSEQUENT SURGERY LATER THAT DAY, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -14.5/2.0/089 | OS | Physician | 03/21/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14141490&pc=QCB | 2023826-2022-01115 | 2022-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MMVTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -12.0/3.5/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6)2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -12.0/3.5/095 | OD | Physician | 03/19/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14142335&pc=QCB | 2023826-2022-01233 | 2021-12-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.5/1.5/115 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6)2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). Manufacturer Narrative: D6B: DATE CORRECTED TO SAY 22-DEC-2021. CLAIM# (B)(4). | -11.5/1.5/115 | OD | Physician | 03/20/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14142793&pc=QCB | 2023826-2022-01234 | 2021-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -8.0/1.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -8.0/1.0/093 | OS | Physician | 03/21/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14145748&pc=QCB | 2023826-2022-01199 | 2021-12-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.5/+4.0/101 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. CLAIM # (B)(4). | -10.5/+4.0/101 | OS | Physician | 03/20/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14145774&pc=QCB | 2023826-2022-01235 | 2021-10-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -8.5/1.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -8.5/1.0/093 | OD | Physician | 03/21/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14147207&pc=QCB | 2023826-2022-01061 | 2022-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MMVTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -3.5/1.5/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED AND A SHORTER LENGTH LENS WAS IMPLANTED DUE TO EXCESSIVE VAULT. THE REPLACEMENT RESOLVED THE PROBLEM. IT IS UNCLEAR IF THE LENS WAS EXCHANGED WITHIN THE SAME SURGERY OR SUBSEQUENT SURGERIES WITHIN THE SAME DAY.. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -3.5/1.5/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS; OPPOSITE TORIC MEDIAN AND THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN." CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: A2- DOB: (B)(6) 1975 CORRECTED TO (B)(6) 2003. H6- DEVICE PROBLEM: OFF LABEL- AGE<21 YEARS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -3.5/1.5/90 | OS | Physician | 03/21/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14148168&pc=QCB | 2023826-2022-01236 | 2022-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS LATER EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.0/1.5/095 | OD | Physician | 03/21/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14149657&pc=QCB | 2023826-2022-01178 | 2022-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.50/+3.0/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PRODUCT EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM #: (B)(4). | -16.50/+3.0/074 | OS | Physician | 03/21/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14149754&pc=QCB | 2023826-2022-01177 | 2021-02-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -18.00/+4.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -18.00/+4.0/090 | OD | Physician | 03/21/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14151078&pc=QCB | 2023826-2022-01116 | 2022-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.5/3.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. OFF LABEL USE: AGE <21 YEARS. CLAIM# (B)(4). | -9.5/3.0/093 | OS | Physician | 03/19/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14151752&pc=QCB | 2023826-2022-01117 | 2021-08-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -9.5/3.0/103 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. OFF LABEL USE: AGE <21 YEARS. CLAIM# (B)(4). | -9.5/3.0/103 | OS | Physician | 03/21/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14161427&pc=QCB | 2023826-2022-01071 | 2022-03-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/1.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -11.0/1.0/083 | OS | Physician | 03/27/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14161430&pc=QCB | 2023826-2022-01077 | 2022-03-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -6.5/1.0/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: B5 - CORRECTED TO "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -6.5/1.0/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN." IN INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: B5 - CORRECTED TO "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -6.5/1.0/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN." IN INITIAL MDR. CLAIM #: (B)(4). | -6.5/1.0/070 | OD | Physician | 03/30/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14161434&pc=QCB | 2023826-2022-01081 | 2020-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.5/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2020. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.5/2.5/092 | OD | Physician | 03/30/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14161435&pc=QCB | 2023826-2022-01082 | 2021-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.50/1.5/151 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN2021. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -14.50/1.5/151 | OD | Physician | 03/30/2022 | 2022 | 46 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14161438&pc=QCB | 2023826-2022-01080 | 2022-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.0/3.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.0/3.0/095 | OD | Physician | 03/30/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14161439&pc=QCB | 2023826-2022-01088 | 2021-04-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -13.0/1.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.0/1.5/094 | OS | Physician | 04/01/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14161441&pc=QCB | 2023826-2022-01087 | 2021-05-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -6.5/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER SPHERICAL LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -6.5/1.0/089 | OS | Physician | 04/01/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14161443&pc=QCB | 2023826-2022-01086 | 2022-02-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.5/3.0/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED IN A SECOND SURGERY ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.5/3.0/109 | OD | Physician | 04/01/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14161444&pc=QCB | 2023826-2022-01078 | 2022-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.0/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -15.0/1.0/085 | OD | Physician | 03/30/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14161449&pc=QCB | 2023826-2022-01068 | 2022-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -10.5/4.5/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.5/4.5/073 | OD | Physician | 03/27/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14162329&pc=QCB | 2023826-2022-01219 |  | Injury | STAAR SURGICAL COMPANY | 2022-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A IMPLANTABLE COLLAMER LENS OF UNKNOWN MODEL AND SIZE WAS IMPLANTED INTO THE PATIENT'S EYE ON AN UNKNOWN DATE. THE REPORTER INDICATED THAT A SURGERY IS SCHEDULED TO REPLACE THE LENS. (EXTRACTION+REPLACEMENT). THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: (B)(4). Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -15.0/3.0/078 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. LENS ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2021 AND THIS DID NOT RESOLVE THE PROBLEM. REFRACTIVE TREATMENT WAS ALSO PERFORMED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. PRE-EXISTING CONDITION OF LATTICE DEGENERATION. CAUSE REPORTED AS A PATIENT RELATED FACTOR. H6-IMPACT CODE 4625- REFRACTIVE TREATMENT. CORRECTED DATA: B2- IN PREVIOUS MDR IS NOT APPLICABLE. THIS IS NOW CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMRNT/DAMAGE CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -15.0/3.0/078 | OS | Other Health Care Professional | 03/21/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14162874&pc=QCB | 2023826-2022-01251 | 2022-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5/4.5/160 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. IN A SUBSEQUENT SURGERY LATER THAT DAY, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. OFF LABEL USE: ACD <3.0. CLAIM # (B)(4). | -9.5/4.5/160 | OS | Physician | 03/23/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14171823&pc=QCB | 2023826-2022-01253 | 2021-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MMVTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -8.00/1.5/84 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6)2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -8.00/1.5/84 | OD | Physician | 03/24/2022 | 2022 | 44 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14172035&pc=QCB | 2023826-2022-01273 | 2022-03-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -6.0/+0.5/064, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -6.0/+0.5/064 | OD | Physician | 03/22/2022 | 2022 | 44 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14176756&pc=QCB | 2023826-2022-01254 | 2021-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MMVTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -7.5/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) ON (B)(6) 2021. IN A SUBSEQUENT SURGERY LATER THAT DAY, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-ETHNICITY: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -7.5/2.0/089 | OD | Physician | 03/28/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14179251&pc=QCB | 2023826-2022-01225 | 2021-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -6.5/1.0/92 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON 31-JUL-2021. ON (B)(6)2021 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS AND DIFFERENT POWER. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC TORN, BENT, DEFORMED AND STRETCHED OUT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT:UNKNOWN, ETHINICTY:UNKNOWN, RACE:UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -6.5/1.0/92 | OD | Physician | 03/23/2022 | 2022 | 46 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14187363&pc=QCB | 2023826-2022-01276 | 2021-11-22 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-04-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.0/3.5/1 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED. THIS RESOLVED THE PROBLEM. IT WAS ALSO NOTED "NEXT OPERATION DATE UNDECIDED". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -10.0/3.5/1 | OS | Physician | 04/14/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14187365&pc=QCB | 2023826-2022-01275 | 2021-11-22 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-04-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.5/3.0/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED. THIS RESOLVED THE PROBLEM. IT WAS ALSO NOTED "NEXT OPERATION DATE UNDECIDED". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON THE LENS.VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND STRETCHED OUT. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -10.5/3.0/91 | OD | Physician | 04/14/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14189870&pc=QCB | 2023826-2022-01282 | 2022-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5 12.6, -6.5/1.00/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: B5: PER UPDATED INFORMATION, THE LENS WAS EXPLANTED ON (B)(6) 2022. AND THE PROBLEM RESOLVED. CLAIM (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -6.5/1.00/083 | OD | Physician | 03/23/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14189897&pc=QCB | 2023826-2022-01283 | 2022-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5 12.6, -5.5/0.50/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). Manufacturer Narrative: B5: PER UPDATED INFORMATION THE LENS WAS EXPLANTED ON (B)(6) 2022 AND THE PROBLEM RESOLVED. CLAIM #: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -5.5/0.50/085 | OS | Physician | 03/23/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14197861&pc=QCB | 2023826-2022-01308 | 2022-03-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-04-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT WHILE ATTEMPTING TO IMPLANTE A 12.6MM TMICL12.6; -12.00/1.0/96 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE; THE LENS TORE. THIS OCCURRED ON (B)(6) 2022. THERE WAS NO PATIENT CONTACT. A BACK UP LENS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND USER ERROR AND NOTED "TEAR IN ICL NOTED UPON DELIVERY OF ICL INTO THE EYE". Manufacturer Narrative: CLAIM#: (B)(4). | -12.00/1.0/96 | NA | Other Health Care Professional | 04/11/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14206091&pc=QCB | 2023826-2022-01160 | 2021-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.5/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -8.5/1.5/091 | OS | Physician | 04/04/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14206103&pc=QCB | 2023826-2022-01169 | 2022-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6, -4.50/0.5/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -4.50/0.5/101 | OS | Physician | 04/08/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14206104&pc=QCB | 2023826-2022-01164 | 2021-12-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.0/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.0/2.0/088 | OS | Physician | 04/06/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14206105&pc=QCB | 2023826-2022-01171 | 2022-03-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.00/3.00/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.00/3.00/090 | OD | Physician | 04/08/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14206107&pc=QCB | 2023826-2022-01167 | 2022-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.5/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -15.5/2.5/092 | OD | Physician | 04/07/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14206108&pc=QCB | 2023826-2022-01166 | 2021-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.0/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -10.0/1.5/095 | OD | Physician | 04/06/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14206110&pc=QCB | 2023826-2022-01168 | 2022-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.5/1.0/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK, RACE- UNK, ETHNICITY- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.5/1.0/072 | OD | Physician | 04/07/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14206113&pc=QCB | 2023826-2022-01170 | 2022-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6, -8.50/1.50/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -8.50/1.50/099 | OD | Physician | 04/08/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14215065&pc=QCB | 2023826-2022-01301 | 2022-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM TMICL12.1 IMPLANTABLE COLLAMER LENS, -10.00/+3.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WILL BE REMOVED AND REPLACED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: SEX: UNK, WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, EVENT DATE: UNK, EXPLANT DATE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS BUT THE PROBLEM WAS NOT RESOLVED. SEE MFR. REPORT # 2023826-2022-01669 FOR REPLACEMENT LENS. H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CLAIM # (B)(4). | -10.00/+3.0/089 | OS | Other Health Care Professional | 04/04/2022 | 2022 | 44 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14215079&pc=QCB | 2023826-2022-01279 | 2022-03-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. PMA/510K: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.00/+1.5/088 | OS | Physician | 03/30/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14215119&pc=QCB | 2023826-2022-01227 | 2021-12-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -2.5/1.0/014 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME TORIC LENS OF THE SAME SIZE BUT DIFFERENT POWER -4.5/2.5/000 (SPHERE/CYLINDER/AXIS) LENS DUE TO REFRACTIVE SURPRISE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND LENS BENT/DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -2.5/1.0/014 | OS | Physician | 03/27/2022 | 2022 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14215167&pc=QCB | 2023826-2022-01337 | 2021-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -14.0/1.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS BUT DIFFERENT POWER. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHINICITY:UNKNOWN, RACE: UNKNOWN. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6 - INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). | -14.0/1.5/092 | OD | Physician | 03/27/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14215321&pc=QCB | 2023826-2022-01348 | 2021-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -12.00/+2.50/79 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER". Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -12.00/+2.50/79 | OS | Physician | 04/04/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14218810&pc=QCB | 2023826-2022-01350 | 2021-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5/0.5/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6)2021, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.5/0.5/088 | OD | Physician | 03/30/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14219331&pc=QCB | 2023826-2022-01351 | 2021-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5/0.5/074 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)2021, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -9.5/0.5/074 | OS | Physician | 03/30/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14220086&pc=QCB | 2023826-2022-01352 | 2022-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -10.5/1.0/096 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. IN A SUBSEQUENT SURGERY, LATER THAT DAY, THE LENS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.5/1.0/096 | OS | Physician | 03/30/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14220420&pc=QCB | 2023826-2022-01353 | 2022-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.0/2.0/100 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.0/2.0/100 | OS | Physician | 03/30/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14220818&pc=QCB | 2023826-2022-01104 | 2022-01-27 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 -9.5/2.5/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED WITHIN THE SAME SURGERY DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. NO PATIENT INJURY. ON (B)(6) 2022 AN IDENTICAL LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE IS UNKNOWN. Manufacturer Narrative: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. INVESTIGATION TYPE (B)(4): LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE CODE 1494: OFF-LABEL USE (AGE (B)(6)). CLAIM# (B)(4). | -9.5/2.5/092 | OD | Physician | 04/04/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14221172&pc=QCB | 2023826-2022-01354 | 2022-03-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.0/2.0/101 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2022, THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. ON (B)(6) 2022, A SHORTER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.0/2.0/101 | OS | Physician | 03/30/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14225692&pc=QCB | 2023826-2022-01355 | 2022-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -11.5/1.5/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHOTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE IS REPORTED AS UNKNOWN. REPORTEDLY. "FIRST LENS WAS REMOVED DUE TO HIGH VAULT AND IMPLANTED 12.6MM LENS. CHAMBER FORMED WELL AND PATIENT IS EXTREMELY HAPPY." Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -11.5/1.5/091 | OD | Physician | 03/31/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14228330&pc=QCB | 2023826-2022-01281 | 2021-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+0.5/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN/BENT. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.00/+0.5/087 | OD | Physician | 03/31/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14242468&pc=QCB | 2023826-2022-01356 | 2022-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -11.5/1.5/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHOTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE PROBLEM WAS RESOLVED. THE CAUSE IS REPORTED AS UNKNOWN. REPORTEDLY. "FIRST LENS WAS REMOVED DUE TO HIGH VAULT AND IMPLANTED 12.6MM LENS. CHAMBER FORMED WELL AND PATIENT IS EXTREMELY HAPPY." Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -11.5/1.5/089 | OS | Physician | 03/31/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14246467&pc=QCB | 2023826-2022-01357 | 2021-11-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MMVTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -3.00/+1.50/115 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). Manufacturer Narrative: D4 - UNIQUE IDENTIFIER (UDI) # (B)(4) SHOULD BE REMOVED FROM INITIAL MDR. CLAIM #(B)(4). | -3.00/+1.50/115 | OD | Physician | 03/30/2022 | 2022 | 41 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14246746&pc=QCB | 2023826-2022-01383 |  | Injury | STAAR SURGICAL COMPANY | 2022-04-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO A PATIENT'S EYE AND THERE MAY HAVE EXPERIENCED LENS ROTATION WHICH MAY REQUIRE A TICL REPOSITIONING. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THAT A 13.2MM TIMCL13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.00/2.0/101 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE. THE PATIENT EXPERIENCED LENS ROTATION; REPOSITIONING AND PRK WAS PERFORMED. ON (B)(6) 2022, THE LENS WAS REPOSITIONED BUT THAT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022, PRK WAS PERFORMED. REPORTEDLY "PRK FOR RESIDUAL REFRACTIVE ERROR. P.O ON (B)(6) 2023 OD UCVA 20/25". THE LENS REMAINS IMPLANTED. H6: INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: D4- MODEL#: TIMCL13.2 SHOULD BE CORRECTED TO TMICL13.2. B5- "TIMCL 13.2" SHOULD BE CORRECTED TO "TMICL 13.2". CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -8.00/2.0/101 | NA | Other Health Care Professional | 03/29/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14246772&pc=QCB | 2023826-2022-01385 | 2022-04-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-04-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO_13.2; -7.50/+5.0/104 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS HAD TORE/BREAK DURING INJECTION/DELIVERY INTO THE EYE; INTRAOPERATIVELY THE LENS WAS REMOVED. ON (B)(6) 2022 A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS NOTED AS "LENS GOT STUCK IN THE PLUNGER AND GOT TORN". Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). CLAIM#: (B)(4). | -7.50/+5.0/104 | OD | Physician | 04/21/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14246802&pc=QCB | 2023826-2022-01384 |  | Injury | STAAR SURGICAL COMPANY | 2022-04-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -11.5/+1.5/072 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND LATER ON THE SAME DATE REPLACED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -11.5/+1.5/072 | OS | Physician | 04/19/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14247031&pc=QCB | 2023826-2022-01359 | 2022-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF A -9.5/1.0/091 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -9.5/1.0/091 | OD | Physician | 03/31/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14247467&pc=QCB | 2023826-2022-01338 | 2022-02-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICH 12.6 IMPLANTABLE COLLAMER LENS OF -7.50/1.0/13 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. EXCESSIVE VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE REPORTED AS PATIENT RELATED FACTOR. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. D6B - EXPLANTED (B)(6) 2022. H6 - HEALTH EFFECT - HEALTH IMPACT CODE: 4629. CLAIM # (B)(4). | -7.50/1.0/13 | OD | Physician | 03/29/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14253219&pc=QCB | 2023826-2022-01340 |  | Injury | STAAR SURGICAL COMPANY | 2022-04-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 13.2MM TMICL 13.2 IMPLANTABLE COLLAMER LENS OF -8.50/4.0/125 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022. REPORTEDLY THIS WAS DUE TO THE "LENS ROTATED TO THE POINT WHERE IT WAS AFFECTING HIS VISION." POST-OP 1 DAY AND THE DOCTOR DOES NOT SEE ANYTHING OUT OF THE ORDINARY IN THEIR CHART. Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -8.50/4.0/125 | OD | Other Health Care Professional | 03/30/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14253801&pc=QCB | 2023826-2022-01315 | 2022-01-28 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-04-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -12.50/1.50/90 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS STUCK IN CARTRIDGE OR INJECTION SYSTEM AND TORE DURING INJECTION/DELIVERY INTO THE EYE ON (B)(6) 2022. LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. THERE WAS NO PATIENT INJURY. A REPLACEMENT LENS WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS UNKNOWN. FOR CAUSE DETAILS THE REPORTER STATED "DURING IMPLANTATION OF THE ICL PATIENT MOVED HIS HEAD AND PART OF THE PROXIMAL HAPTIC WAS INJECTED OUTSIDE THE WOUND WHICH WAS NOTED TO BE TORN SO DECISION WAS MADE TO EXPLANT THE ICL." Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.50/1.50/90 | NA | Physician | 03/31/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14254540&pc=QCB | 2023826-2022-01360 | 2022-02-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF A -12.0/2.5/090 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULTING WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -12.0/2.5/090 | OS | Physician | 03/31/2022 | 2022 | 44 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14254544&pc=QCB | 2023826-2022-01361 | 2021-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF A -12.0/3.0/66 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A LONGER LENGTH, SPHERICAL LENS DUE TO LOW VAULTING. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -12.0/3.0/66 | OD | Physician | 04/01/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14254592&pc=QCB | 2023826-2022-01343 |  | Injury | STAAR SURGICAL COMPANY | 2022-04-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 13.7MM TMICL 13.2 IMPLANTABLE COLLAMER LENS OF -9.50/1.0/099 (SPHERE/CYLINDER/AXIS) BROKE WHILE HANDLING DURING SURGERY ON THE PATIENT. THE LENS WAS NOT IMPLANTED AND THE DOCTOR HAD TO USE AN ADDITIONAL LENS. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: D4 - CORRECTED TO: MODEL# TMCIL13.7 IN INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4627 SHOULD BE ADDED TO THE INITAL MDR. CLAIM # (B)(4). | -9.50/1.0/099 | NA | Other Health Care Professional | 03/31/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14254615&pc=QCB | 2023826-2022-01362 | 2022-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF A -13.5/1.5/093 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.5/1.5/093 | OS | Physician | 03/31/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14254617&pc=QCB | 2023826-2022-01316 | 2022-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -15.0/2.0/082 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. FOR CAUSE DETAILS THE REPORTER STATED " ICL IS THOUGHT TO BE TOO LONG."Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -15.0/2.0/082 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. A REPLACEMENT LENS OF SHORTER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. FOR CAUSE DETAILS THE REPORTER STATED " ICL IS THOUGHT TO BE TOO LONG." ACD <3.00 MM, PRE-EXISTING PROGRESSIVE MYOPIA. Manufacturer Narrative: B5: ADDED ACD<3.00MM AND PROGRESSIVE MYOPIA. H6-INVESTIGATION CONCLUSION CODE 4319 ADDED TO INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -15.0/2.0/082 | OD | Physician | 03/31/2022 | 2022 | 44 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14255864&pc=QCB | 2023826-2022-01389 | 2022-02-03 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-04-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2; -18.00/+1.00/47 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE DURING INJECTION/DELIVERY INTO THE PATIENT'S LEFT EYE (OS). THIS OCCURRED ON (B)(6) 2022. THERE WAS PATIENT CONTACT WITH NO INJURY. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM#: (B)(4). | -18.00/+1.00/47 | OS | Physician | 03/31/2022 | 2022 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14255886&pc=QCB | 2023826-2022-01364 | 2022-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-04-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF A -11.0/1.5/169 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON 26-(B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULTING WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.0/1.5/169 | OD | Physician | 04/04/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14268750&pc=QCB | 2023826-2022-01270 | 2022-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -07.00/4.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US (B)(4). | -07.00/4.5/105 | OD | Physician | 04/12/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14268776&pc=QCB | 2023826-2022-01266 | 2021-11-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/4.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -11.0/4.0/091 | OD | Physician | 04/12/2022 | 2022 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14268800&pc=QCB | 2023826-2022-01265 | 2021-12-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/4.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -11.0/4.5/090 | OD | Physician | 04/09/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14268803&pc=QCB | 2023826-2022-01262 | 2022-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED VERTICALLY A 13.7MM VTICMO13.7, -15.0/2.0/173 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK, THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -15.0/2.0/173 | OD | Physician | 04/09/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14273614&pc=QCB | 2023826-2022-01391 |  | Injury | STAAR SURGICAL COMPANY | 2022-05-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -9.0/+1.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -9.0/+1.0/092 | OS | Physician | 04/04/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14281045&pc=QCB | 2023826-2022-01344 | 2022-03-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS OF UNKNOWN MODEL AND POWER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON AN UNKNOWN DATE. REPORTEDLY LOW VAULT WAS OBSERVED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -6.0/6.0/105 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. THE PROBLEM IS RESOLVED AND THE "PATIENT IS HAPPY AND WILL CHECK IN A YEAR." CAUSE REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM# (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -6.0/6.0/105 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE REPORTED AS UNKNOWN.IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM# (B)(4). | -6.0/6.0/105 | OS | Physician | 04/03/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14282386&pc=QCB | 2023826-2022-01214 | 2022-02-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -16.5/2.5/070 (SPHERE/CYLINDER/AXIS) WAS BEING IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS TEAR/ BREAK DURING INJECTION/DELIVERY INTO THE EYE WITH PATIENT CONTACT AND NO PATIENT INJURY. ON (B)(6) 2022 THE LENS WAS IMPLANTED;REMOVED INTRA-OPERATIVELY. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -16.5/2.5/070 | OS | Physician | 03/18/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14283686&pc=QCB | 2023826-2022-01366 | 2021-05-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO_12.6 IMPLANTABLE COLLAMER LENS OF -13.5/2.5/079 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. WITHOUT ROTATON. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -13.5/2.5/079 | OS | Physician | 04/05/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14288725&pc=QCB | 2023826-2022-01216 | 2022-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -13.0/0.5/76 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED DUE TO LOW VAULT. A REPLACEMENT LENS OF LONGER LENGTH WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -13.0/0.5/76 | OS | Physician | 03/18/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14288726&pc=QCB | 2023826-2022-01215 | 2022-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -13.5/0.5/08 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6)2022 THE LENS WAS REMOVED DUE TO LOW VAULT. A REPLACEMENT LENS OF LONGER LENGTH WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -13.5/0.5/08 | OD | Physician | 03/18/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14290579&pc=QCB | 2023826-2022-01367 | 2021-01-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO_12.6 IMPLANTABLE COLLAMER LENS OF -11.5/2.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. WITHOUT ROTATON. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. OFF LABEL USE - ACD 2.96. CLAIM#: (B)(4). | -11.5/2.0/090 | OD | Physician | 04/05/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14291604&pc=QCB | 2023826-2022-01403 | 2021-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO_13.2 IMPLANTABLE COLLAMER LENS OF -12.5/3.5/004 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IN A SUBSEQUENT SURGERY LATER THAT DAY, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -12.5/3.5/004 | OD | Physician | 04/06/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14296769&pc=QCB | 2023826-2022-01404 | 2021-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO_12.6 IMPLANTABLE COLLAMER LENS OF -11.0/1.0/74 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. IN A SUBSEQUENT SURGERY LATER THAT DAY, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT..THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -11.0/1.0/74 | OD | Physician | 04/06/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14307139&pc=QCB | 2023826-2022-01405 | 2022-03-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO_13.2 IMPLANTABLE COLLAMER LENS OF -10.5/1.5/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6)2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -10.5/1.5/091 | OD | Physician | 04/06/2022 | 2022 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14311528&pc=QCB | 2023826-2022-01156 |  | Injury | STAAR SURGICAL COMPANY | 2022-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. Manufacturer Narrative: DATE OF BIRTH: (B)(6). WEIGHT, ETHNICITY, RACE:UNK. EVENT DATE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.0/1.5/095 | OD | Physician | 04/06/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14311555&pc=QCB | 2023826-2022-01157 |  | Injury | STAAR SURGICAL COMPANY | 2022-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -10.0/2.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. Manufacturer Narrative: DATE OF BIRTH: (B)(6). WEIGHT, ETHNICITY, RACE:UNK. EVENT DATE: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.0/2.0/091 | OS | Physician | 04/06/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14311828&pc=QCB | 2023826-2022-01243 | 2021-10-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO INSERT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS -08.00/+1.0/002 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S RIGHT EYE (OD) BUT THE LENS TORE DURING LOADING INTO INJECTOR. THERE WAS NO PATIENT CONTACT. ANOTHER MODEL/DIOPTER LENS WAS INTRAOPERATIVELY IMPLANTED ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE IS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -08.00/+1.0/002 | OD | Physician | 04/14/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14311855&pc=QCB | 2023826-2022-01241 | 2022-04-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 -12.5/2.0/045 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED WITHIN THE SAME SURGERY AND REPLACED WITH AN IDENTICAL LENS DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. NO PATIENT INJURY. CAUSE IS UNKNOWN. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 -12.5/2.0/045 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED WITHIN THE SAME SURGERY AND REPLACED WITH AN IDENTICAL LENS DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. NO PATIENT INJURY. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6) (UNABLE TO ENTER IN FULL). THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -12.5/2.0/045 | OS | Physician | 08/05/2022 | 2022 | 52 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14322617&pc=QCB | 2023826-2022-01427 | 2022-04-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF -8.50/1.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED AND ON SAME DAY DIFFERENT SURGERY A LONGER LENGTH LENS WAS IMPLANTED AND THAT RESOLVED THE PROBLEM. CAUSE REPORTED OTHER AND DEVICE DID NOT FAIL TO PERFORM AS INTENDED. "LENGTH TOO SHORT." Manufacturer Narrative: ADDITIONAL DATA: B3- 05-APR-2022. CORRECTED DATA: B5- THE REPORTER INDICATED THAT A 12.6MM TIMCL 12.6 IMPLANTABLE COLLAMER LENS OF -8.5/1.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. IT WAS REPORTED LENS ROTATION DUE TO LOW VAULT. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON THE SAME DAY SEPARATE SURGERY THE LENS WAS REPLACED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS OTHER AND DEVICE DID NOT FAIL TO PERFORM AS INTENDED. REPORTEDLY "LENGTH TOO SHORT". CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -8.50/1.0/093 | OD | Other Health Care Professional | 04/07/2022 | 2022 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14324480&pc=QCB | 2023826-2022-01428 | 2022-02-11 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: IT WAS REPORTED THAT A 13.7MM VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF -12.5/2.0/013 (SPHERE/CYLINDER/AXIS) WAS OBSERVED TO TEAR/BREAK DURING INJECTION/DELIVERY INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS IMPLANTED; REMOVED INTRA-OPERATIVELY ON (B)(6) 2022. ON THE SAME DAY DIFFERENT SURGERY A REPLACEMENT LENS OF DIFFERENT AXIS WAS IMPLANTED VERTICALLY AND THE PROBLEM RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4): OFF-LABEL USE (PATIENT UNDER 21 YEARS OLD.) CLAIM#: (B)(4). | -12.5/2.0/013 | OS | Physician | 04/07/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14332610&pc=QCB | 2023826-2022-01397 | 2022-03-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, TMICL13.2, -4.5/1.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHAGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -4.5/1.0/091 | OD | Physician | 03/07/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14335675&pc=QCB | 2023826-2022-01323 | 2021-11-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -4.5/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: .THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -4.5/1.0/087 | OS | Physician | 04/20/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14335676&pc=QCB | 2023826-2022-01324 | 2022-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -6.5/1.0/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -6.5/1.0/108 | OS | Physician | 04/20/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14335678&pc=QCB | 2023826-2022-01321 | 2022-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.0/3.5/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS REPLACED IN A SECOND SURGERY WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -14.0/3.5/109 | OD | Physician | 04/19/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14335681&pc=QCB | 2023826-2022-01328 | 2022-03-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -09.50/1.5/113 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -09.50/1.5/113 | OS | Physician | 04/21/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14335685&pc=QCB | 2023826-2022-01325 | 2022-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.0/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -8.0/2.0/090 | OS | Physician | 04/20/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14335687&pc=QCB | 2023826-2022-01329 | 2022-01-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.00/1.5/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED VERTICALLY WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.00/1.5/072 | OD | Physician | 04/21/2022 | 2022 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14335690&pc=QCB | 2023826-2022-01333 | 2022-03-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.0/1.0/166 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -10.0/1.0/166 | OD | Physician | 04/22/2022 | 2022 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14335696&pc=QCB | 2023826-2022-01334 | 2022-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -07.50/0.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM#: (B)(4). | -07.50/0.5/088 | OS | Physician | 04/22/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14339447&pc=QCB | 2023826-2022-01345 | 2022-03-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -6.0/6.0/066 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE REPORTED AS UNKNOWN.IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT A12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -6.0/6.0/066 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. THE PROBLEM IS RESOLVED AND THE "PATIENT IS HAPPY AND WILL CHECK IN A YEAR." CAUSE REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. CLAIM # (B)(4). | -6.0/6.0/066 | OD | Physician | 04/03/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14340520&pc=QCB | 2023826-2022-01455 | 2022-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1, -6.0/+1.0/002 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LOW VAULT IS REPORTED. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR; IT IS UNKNOWN WHETHER THE DEVICE FAILED TO PERFORM AS EXPECTED. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S.. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDI TIONAL DATA: B5: THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO REFRACTIVE SURPRISE AND REPLACED WITH A LONGER LENS. POST-OP, THE PATIENT COMPLAINED OF PAIN AND POOR ABILITY TO READ. THE PROBLEM OF LOW VAULTING REMAINED UNCHANGED DESPITE SIZING UP. SEE MFR. REPORT # 2023826-2022-01794 FOR REPLACEMENT LENS. B4 & G3: (B)(6) 2022. CLAIM # (B)(4). | -6.0/+1.0/002 | OD | Physician | 04/04/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14340609&pc=QCB | 2023826-2022-01431 | 2022-03-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -14.5/5.0/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. LENS ROTATION WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. EXCHANGE LENS IMPLANTED VERTICALLY. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE OPTIC TORN AND HAPTIC TORN/BROKEN/BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -14.5/5.0/088 | OS | Physician | 04/09/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14342573&pc=QCB | 2023826-2022-01462 | 2021-01-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.5/+2.0/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B)(4). | -12.5/+2.0/087 | OS | Physician | 04/12/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14342575&pc=QCB | 2023826-2022-01464 | 2022-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -10.5/+2.0/079 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: AGE, WEIGHT, RACE-UNK. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -10.5/+2.0/079 | OD | Physician | 04/13/2022 | 2022 | 44 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14342600&pc=QCB | 2023826-2022-01458 | 2022-02-15 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT, WHILE IMPLANTING A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.5/+1.5/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS), THE LENS TORE/BROKE. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED WITH AN ALTERNATE LENS INTRA-OPERATIVELY ON (B)(6) 2022. THE PROBLEM WAS RESOLVED. REPORTEDLY, THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. OFF-LABEL USE (UNDER (B)(6) AT DATE OF IMPLANT). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -14.5/+1.5/082 | OS | Physician | 04/12/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14342649&pc=QCB | 2023826-2022-01465 | 2022-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -10.0/+2.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -10.0/+2.0/090 | OS | Physician | 04/13/2022 | 2022 | 44 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14342651&pc=QCB | 2023826-2022-01466 | 2021-12-27 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -8.5/+1.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. INTRAOPERATIVELY THE LENS WAS IMPLANTED, REMOVED, AND REPLACED WITH A LONGER LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT:UNKNOWN, ETHNICITY:UNKNOWN, RACE:UNKNOWN. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. (B(4). CLAIM# (B)(4). | -8.5/+1.5/090 | OD | Physician | 04/13/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14342664&pc=QCB | 2023826-2022-01467 | 2022-02-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.5/+2.0/100 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS EXCHANGED WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -13.5/+2.0/100 | OD | Physician | 04/13/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14342667&pc=QCB | 2023826-2022-01457 | 2021-12-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -6.5/+0.5/113 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SAME SIZE DIFFERENT AXIS LENS DUE TO LENS ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. OFF-LABEL USE (ACD < 3.0MM) CLAIM# (B)(4). | -6.5/+0.5/113 | OS | Physician | 04/12/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14342682&pc=QCB | 2023826-2022-01463 | 2022-01-18 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -6.5/+2.0/097 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENS DUE TO EXCESSIVE VAULT. THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S.. BUT NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -6.5/+2.0/097 | OD | Physician | 04/13/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14342686&pc=QCB | 2023826-2022-01460 | 2021-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -15.0/+2.5/074 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A LONGER LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT:UNKNOWN, ETHINICITY:UNKNOWN, RACE:UNKNOWN. PMA/510(K): THIS PRODUCT IS MANUFACTURED IN THE U.S. BUT NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -15.0/+2.5/074 | OS | Physician | 04/13/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14348443&pc=QCB | 2023826-2022-01239 | 2022-04-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -5.00/1.5/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LOW VAULT WITH ROTATION. THIS DID NOT RESOLVE THE PROBLEM. APPLYING FOR UPSIZING. Manufacturer Narrative: WEIGHT-RACE: UNK. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: PRODUCT EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND RESIDUE/DEBRIS ON THE LENS WITH NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -5.00/1.5/088 | OS | Physician | 09/28/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14349575&pc=QCB | 2023826-2022-01490 | 2022-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.5/1.0/099 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -11.5/1.0/099 | OS | Physician | 04/13/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14349577&pc=QCB | 2023826-2022-01492 | 2022-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.5/2.5/118 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6)2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.5/2.5/118 | OD | Physician | 04/13/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14349578&pc=QCB | 2023826-2022-01491 | 2022-01-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.0/1.5/052 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/1.5/052 | OS | Physician | 04/13/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14349582&pc=QCB | 2023826-2022-01494 | 2021-12-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -6.0/1.0/27 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)2021, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. (B)(4). | -6.0/1.0/27 | OS | Physician | 04/13/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14349955&pc=QCB | 2023826-2022-01429 | 2022-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -15.5/4.5/125 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO LOW VAULT WITH ROTATION AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. CLAIM #: (B)(4). | -15.5/4.5/125 | OS | Physician | 04/08/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14350128&pc=QCB | 2023826-2022-01433 |  | Injury | STAAR SURGICAL COMPANY | 2022-05-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 13.2 TMICL 13.2 IMPLANTABLE COLLAMER LENS OF -11.50/1.0/075 (SPHERE/CYLINDER/AXIS) LENS WAS IMPLANTED INTO THE PATIENTS EYE ON AN UNKNOWN DATE. THE LENS WAS REPORTEDLY REMOVED AND REPLACED FOR A SHORTER LENGTH DUE TO THE LENS BEING OVERSIZED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE EVALUATION: LENS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND LENS HAPTIC TORN. CLAIM#: (B)(4). | -11.50/1.0/075 | NA | Other Health Care Professional | 04/08/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14364432&pc=QCB | 2023826-2022-01517 | 2022-02-25 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -9.0/2.0/84 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). IT IS REPORTED THAT THE LENS HAD A TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS IMPLANTED; REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SAME MODEL; SIZE AND POWER LENS. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -9.0/2.0/84 | OD | Physician | 04/12/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14364773&pc=QCB | 2023826-2022-01478 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+2.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2012. THE LENS WAS REPORTED AS LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANT DATE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.00/+2.5/093 | OD | Physician | 04/11/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14365182&pc=QCB | 2023826-2022-01481 | 2022-04-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.00/+2.0/068, DURING THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED ON (B)(6) 2022 WITH ANOTHER SAME LENGTH BUT DIFFERENT MODEL LENS BUT THE PROBLEM WAS NOT RESOLVED (THE LENS HAD A LOW VAULT). SEE MFR. REPORT # 2023826-2022-01483 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: AT THE POST-OP VISIT ON (B)(6) 2022, THE PROBLEM OF LOW VAULT WITH THE REPLACEMENT LENS HAD RESOLVED AND THE PROBLEM WAS RESOLVED. THE EYE WAS NORMAL. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -14.00/+2.0/068 | OD | Physician | 04/14/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14365323&pc=QCB | 2023826-2022-01511 | 2022-04-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.00/+2.5/082 (SPHERE/CYLINDER/AXIS) WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS ON (B)(6) 2022 AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING WELL. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM #: (B)(4). | -10.00/+2.5/082 | OD | Physician | 04/13/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14365633&pc=QCB | 2023826-2022-01519 | 2022-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -7.50/1.50/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2022 AND THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS OF THE SAME LENGTH. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | -7.50/1.50/91 | OD | Physician | 04/12/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14365787&pc=QCB | 2023826-2022-01520 | 2022-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -6.50/1.0/83 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2022 AND THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS OF THE SAME LENGTH. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -6.50/1.0/83 | OS | Physician | 04/12/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14366136&pc=QCB | 2023826-2022-01512 | 2022-03-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.50/+3.5/114 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO THE LENS TORE DURING DELIVERY. THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. PATIENT IS HAPPY WITH VISUAL OUTCOME. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHINICITY: UNKNOWN, RACE: UNKNOWN. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.50/+3.5/114 | OD | Physician | 04/13/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14366140&pc=QCB | 2023826-2022-01515 | 2021-10-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.00/+1.0/092 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO LOW VAULTING. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC TORN/BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.00/+1.0/092 | NA | Physician | 04/15/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14366218&pc=QCB | 2023826-2022-01513 | 2022-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -03.50/+1.0/014 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM #: (B)(4). | -03.50/+1.0/014 | OD | Physician | 04/15/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14366322&pc=QCB | 2023826-2022-01527 | 2022-05-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2; -15.00/+6.0/097 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WHICH WAS INTENDED FOR A PATIENT'S RIGHT EYE (OD); HAD TORE DURING LOADING ON (B)(6) 2022. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: (B)(4) LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -15.00/+6.0/097 | OD | Physician | 05/11/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14366558&pc=QCB | 2023826-2022-01500 | 2022-03-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -11.5/2.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -11.5/2.0/091 | OS | Physician | 04/14/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14366626&pc=QCB | 2023826-2022-01499 | 2022-02-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -13.0/2.0/083 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. WITHIN THE SAME SURGERY, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: B2:MALFUNCTION. B5:THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -13.0/2.0/083 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. WITHIN THE SAME SURGERY, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. H6:IMPACT CODE 2199. CLAIM #(B)(4). | -13.0/2.0/083 | OD | Physician | 08/04/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14387943&pc=QCB | 2023826-2022-01407 | 2022-01-29 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -12.5/3.5/77 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. WITHIN THE SAME SURGERY, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -12.5/3.5/77 | OS | Physician | 04/14/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14387944&pc=QCB | 2023826-2022-01503 | 2022-03-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -5.5/0.5/178 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO:UNK. CLAIM# (B)(4). | -5.5/0.5/178 | OS | Physician | 04/14/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14387947&pc=QCB | 2023826-2022-01406 | 2022-01-29 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -14.0/4.0/106 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. WITHIN THE SAME SURGERY, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -14.0/4.0/106 | OD | Physician | 04/14/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14390054&pc=QCB | 2023826-2022-01410 | 2021-12-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -7.5/1.5/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. REPORTEDLY, "THE POSTOPERATIVE CONDITION IS GOOD." Manufacturer Narrative: PT INFO:UNK. CLAIM# (B)(4). | -7.5/1.5/91 | OS | Physician | 04/15/2022 | 2022 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14395947&pc=QCB | 2023826-2022-01489 | 2022-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -7.0/1.5/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION:UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -7.0/1.5/088 | OS | Physician | 04/15/2022 | 2022 | 22
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14399379&pc=QCB | 2023826-2022-01482 | 2022-04-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.00/+2.5/093, DURING THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REPLACED ON (B)(6) 2022 WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE EYE WAS NORMAL AND THE PATIENT WAS HAPPY. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -16.00/+2.5/093 | OS | Physician | 04/14/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14399473&pc=QCB | 2023826-2022-01483 | 2022-04-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -14.00/+2.0/095, INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPORTED AS HAVING A LOW VAULT AND REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2022-01481. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: B5: AT THE POST-OP VISIT ON (B)(6) 2022, THE PROBLEM OF LOW VAULT HAD RESOLVED AND THE PROBLEM WAS RESOLVED. THE EYE WAS NORMAL. CLAIM # (B)(4). | -14.00/+2.0/095 | OD | Physician | 04/27/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14399521&pc=QCB | 2023826-2022-01514 | 2022-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -02.50/+0.5/004 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM # (B)(4). | -02.50/+0.5/004 | OS | Physician | 04/15/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14399535&pc=QCB | 2023826-2022-01516 | 2021-10-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/094 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO LOW VAULTING. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.50/+1.5/094 | NA | Physician | 04/15/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14399923&pc=QCB | 2023826-2022-01319 | 2021-10-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2021 A 13.2MM, VTICM5_13.2, -10.0/2.5/94 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING LOADING, THERE WAS NO PATIENT CONTACT. "LENS TEAR DUE TO SNAGGING WHILE LOADING INJECTOR" WAS REPORTED. A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE U.S. CLAIM # (B)(4). | -10.0/2.5/94 | NA | Physician | 04/13/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14400461&pc=QCB | 2023826-2022-01504 | 2022-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.0/3.5/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | -12.0/3.5/090 | OS | Physician | 04/15/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14400523&pc=QCB | 2023826-2022-01505 | 2022-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5/3.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | -11.5/3.0/090 | OD | Physician | 04/15/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14402582&pc=QCB | 2023826-2022-01507 | 2022-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.5/1.5/082 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | -11.5/1.5/082 | OS | Physician | 04/15/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14402602&pc=QCB | 2023826-2022-01506 | 2022-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.0/1.0/070 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(4) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND ELEVATED IOP WITHOUT PUPIL BLOCK AND ANGLE CLOSURE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | -11.0/1.0/070 | OD | Physician | 04/15/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14402679&pc=QCB | 2023826-2022-01540 | 2022-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1, -7.5/+0.5/179 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENS DUE TO LOW VAULT. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR; IT IS UNKNOWN WHETHER THE DEVICE FAILED TO PERFORM AS EXPECTED. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2022 AND REPLACED WITH A LONGER LENS. POST-OP, THE PATIENT COMPLAINED OF PAIN AND POOR ABILITY TO READ. THE PROBLEM OF LOW VAULTING REMAINED UNCHANGED DESPITE SIZING UP. SEE MFR. REPORT # 2023826-2022-01795 FOR REPLACEMENT LENS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. EXPLANT DATE: UNK. CLAIM#: (B)(4). | -7.5/+0.5/179 | OS | Physician | 04/21/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14402691&pc=QCB | 2023826-2022-01541 |  | Injury | STAAR SURGICAL COMPANY | 2022-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER STATES THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.5/+5.0/131 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. ON (B)(6) 2022 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION BUT THE PROBLEM WAS NOT RESOLVED. ON (B)(6) 2022 THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS A PATIENT-RELATED FACTOR. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. CLAIM#: (B)(4). Manufacturer Narrative: ADDITONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4) | -15.5/+5.0/131 | OD | Physician | 04/22/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14402704&pc=QCB | 2023826-2022-01534 | 2022-03-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -8.0/+1.0/080 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY DUE TO EXCESSIVE VAULT AND THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.0/+1.0/080 | OS | Physician | 04/18/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14406592&pc=QCB | 2023826-2022-01420 | 2022-03-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.0/2.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK, (B)(4). CLAIM# (B)(4). | -11.0/2.0/094 | OD | Physician | 04/26/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14406597&pc=QCB | 2023826-2022-01418 | 2022-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.5/3.0/060 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE LENS WAS REPLACED IN A SECOND SURGERY ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM# (B)(4). | -11.5/3.0/060 | OS | Physician | 04/25/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14406598&pc=QCB | 2023826-2022-01416 | 2022-03-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -10.5/2.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REPLACED INTRAOPERATIVELY WITH A LONGER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S LEFT EYE (OS) DUE TO LOW VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND THE HAPTIC TORN, BENT, DEFORMED, AND STRETCHED OUT. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO DAMAGE OF LENS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -10.5/2.0/092 | OS | Physician | 04/25/2022 | 2022 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14406602&pc=QCB | 2023826-2022-01423 | 2022-02-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.0/2.0/087 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED VERTICALLY WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/2.0/087 | OS | Physician | 04/26/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14431901&pc=QCB | 2023826-2022-01587 | 2022-03-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50/+1.5/072 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A DIFFERENT MODEL BUT SAME LENGTH LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN.. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -07.50/+1.5/072 | OS | Physician | 04/18/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14431914&pc=QCB | 2023826-2022-01577 | 2021-03-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -10.50/+4.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -10.50/+4.0/091 | OS | Other Health Care Professional | 04/29/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14431984&pc=QCB | 2023826-2022-01576 | 2022-04-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -11.00/+3.0/136 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM #(B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). | -11.00/+3.0/136 | OD | Other Health Care Professional | 04/27/2022 | 2022 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14432212&pc=QCB | 2023826-2022-01566 | 2022-03-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -.6.5/1.5/87 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. CLAIM# (B)(4). | -6.5/1.5/87 | OS | Physician | 04/20/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14432224&pc=QCB | 2023826-2022-01592 | 2022-05-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -3.5/1.5/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION-THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | -3.5/1.5/90 | OD | Physician | 10/11/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14439502&pc=QCB | 2023826-2022-01595 |  | Injury | STAAR SURGICAL COMPANY | 2022-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS -8.0/+1.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE ON (B)(6) 2021. REPORTEDLY THE LENS WAS EXPLANTED DUE TO A VAULT ISSUE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. Manufacturer Narrative: SEX: UNK, WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. DATE OF EVENT: UNK. EXPLANT DATE: UNK. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -8.0/+1.0/094 | NA | Physician | 04/27/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14439564&pc=QCB | 2023826-2022-01371 | 2022-01-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -5.0/5.5/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS ROTATION NOT ASSOCIATED WITH A LOW VAULT WAS OBSERVED. REPOSITIONING WAS PERFORMED BUT DID NOT RESOLVED. ON (B)(6) 2022 THE LENS WAS EXPLANTED. NO MAJOR PROBLEMS AFTER EXPLANTATION. CAUSE OF EVENT DUE TO THE DEVICE AND PATIENT RELATED FACTOR. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -5.0/5.5/091 | OD | Physician | 04/19/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14439698&pc=QCB | 2023826-2022-01585 | 2022-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.50/+5.0/044 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS WAS IMPLANTED UPSIDE DOWN. THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -14.50/+5.0/044 | OS | Physician | 04/19/2022 | 2022 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14440445&pc=QCB | 2023826-2022-01619 | 2022-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -14.5/2.0/174 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) -2021 THE LENS WAS EXCHANGED FOR A SAME LENGTH, DIFFERENT MODEL AND POWER LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - RACE:UNK. THIS PRODUCT IS NOT MARKETED IN THE US TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN/BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). | -14.5/2.0/174 | OS | Physician | 04/18/2022 | 2022 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14440516&pc=QCB | 2023826-2022-01604 | 2021-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -13.0/4.5/071 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE ON (B)(6) 2021. LENS ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED AND THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - RACE: UNK. CLAIM# (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -13.0/4.5/071 | OS | Physician | 08/26/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14440684&pc=QCB | 2023826-2022-01617 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6; -9.00/2.0/085 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS DAMAGED DURING LOADING. THERE WAS NO PATIENT CONTACT. A BACK UP LENS WAS USED. THIS OCCURRED ON (B)(6) 2022. Manufacturer Narrative: (B)(4): LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -9.00/2.0/085 | NA | Physician | 05/06/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14440881&pc=QCB | 2023826-2022-01605 | 2021-10-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -13.0/4.0/105 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON 16-SEPT-2021. LENS ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED AND THIS DID NOT RESOLVE THE PROBLEM. ON 17-FEB-2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. ADDITIONAL INFORMATION PROVIDED BY THE REPORTER STATES; "ABNORMAL POSITION OF CILIARY SULCUS LEADS TO CRYSTAL ROTATION." CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT AND DEFORMED.. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -13.0/4.0/105 | OD | Physician | 04/20/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14441537&pc=QCB | 2023826-2022-01568 | 2022-03-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.0/0.5/8 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH SPHERICAL LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. OFF LABEL USE- ACD <3.0. CLAIM#: (B)(4). | -10.0/0.5/8 | OS | Physician | 04/20/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14453220&pc=QCB | 2023826-2022-01624 |  | Injury | STAAR SURGICAL COMPANY | 2022-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -8.50/0.5/106 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON 08-OCT-2020. ON 18-JAN-2022 THE LENS WAS EXPLANTED DUE TO LOW VAULT. A REPLACEMENT LENS OF LONGER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: PATIENT ID: (B)(6). PT INFO: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -8.50/0.5/106 | OD | Physician | 04/19/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14453325&pc=QCB | 2023826-2022-01623 |  | Injury | STAAR SURGICAL COMPANY | 2022-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -9.50/1.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON 16-JAN-2020. ON 09-DEC-2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT. A REPLACEMENT LENS OF LONGER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT ID: (B)(6). WEIGHT, RACE, ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -9.50/1.5/093 | OD | Physician | 04/19/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14455611&pc=QCB | 2023826-2022-01586 | 2022-04-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+2.0/084 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -11.00/+2.0/084 | NA | Physician | 04/25/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14455722&pc=QCB | 2023826-2022-01588 | 2022-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+0.5/178 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -09.50/+0.5/178 | OS | Physician | 04/22/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14456083&pc=QCB | 2023826-2022-01579 | 2022-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00/+6.0/086 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -12.00/+6.0/086 | OS | Physician | 04/29/2022 | 2022 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14457848&pc=QCB | 2023826-2022-01620 | 2021-12-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.0/2.0/072 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6)2021 LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN AND BROKEN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT-RACE:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). CLAIM # (B)(4). | -10.0/2.0/072 | OS | Physician | 04/21/2022 | 2022 | 41 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14459214&pc=QCB | 2023826-2022-01606 | 2022-02-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -8.5/2.0/080 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (ODO ON (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -8.5/2.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. D2-UDI # IS NOT APPLICABLE. CLAIM# (B)(4). | -8.5/2.0/080 | OD | Physician | 04/20/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14459228&pc=QCB | 2023826-2022-01375 | 2022-03-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 OF -11.5/2.5/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS OF DIFFERENT AXIS. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION:UNK. (B)(4). LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.5/2.5/088 | OS | Physician | 04/20/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14459241&pc=QCB | 2023826-2022-01572 | 2022-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -14.5/3.5/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. REPORTEDLY, "THE SECOND IS VERTICAL IMPLANTATION. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -14.5/3.5/089 | OS | Physician | 04/20/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14459442&pc=QCB | 2023826-2022-01573 | 2022-03-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.0/2.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. WITHIN THE SAME SURGERY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT.. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.0/2.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. WITHIN THE SAME SURGERY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT.. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. H6: IMPACT CODE 2199. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT-RACE: UNK. CLAIM#: (B)(4). | -12.0/2.0/094 | OS | Physician | 08/11/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14459608&pc=QCB | 2023826-2022-01575 | 2022-04-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -4.0/1.0/80 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022; THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. CLAIM#: (B)(4). | -4.0/1.0/80 | OS | Physician | 04/22/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14459631&pc=QCB | 2023826-2022-01574 | 2022-04-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -4.0/1.0/92 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022; THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. CLAIM#: (B)(4). | -4.0/1.0/92 | OD | Physician | 04/22/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14459636&pc=QCB | 2023826-2022-01598 | 2021-10-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -17.0/1.5/050 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021; THE LENS WAS REMOVED. IN A SUBSEQUENT SURGERY LATER THAT DAY, A SHORTER LENGTH LENS WAS IMPLANTED DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. CLAIM#: (B)(4). | -17.0/1.5/050 | OS | Physician | 04/22/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14459675&pc=QCB | 2023826-2022-01628 | 2022-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -3.00/6.0/099 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. CAUSE IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT-RACE: UNK. PMA/510(K): THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# 724405. Manufacturer Narrative: CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. B5 - "ON (B)(6) 2022 THE LENS WAS REPOSITIONE DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE PROBLEM WAS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4628 SHOULD BE ADDED TO THE INITAL MDR. CLAIM # (B)(4). | -3.00/6.0/099 | OD | Physician | 04/21/2022 | 2022 | 47 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14459732&pc=QCB | 2023826-2022-01629 | 2022-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -3.50/5.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. CAUSE IS REPORTED AS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR B5 - "THE SURGEON DECIDED TO PERFORM REFRACTIVE TREATMENT (PRK) DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE PROBLEM WAS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4625 SHOULD BE ADDED TO THE INITAL MDR. CLAIM # (B)(4). | -3.50/5.0/089 | OS | Physician | 04/21/2022 | 2022 | 47 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14459798&pc=QCB | 2023826-2022-01630 | 2022-03-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 12.6MM, VTICMO12.6, -13.0/2.0/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE LEFT EYE (OS). THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. THERE WAS NO PATIENT INJURY. A REPLACEMENT LENS WAS IMPLANTED AT A LATER DATE AND THE PROBLEM RESOLVED. Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.0/2.0/081 | OS | Physician | 04/21/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14459803&pc=QCB | 2023826-2022-01627 | 2022-04-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -6.00/2.0/108 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS IMPLANTED, REMOVED AND REPLACED WITH A SHORTER LENGTH LENS INTRAOPERATIVELY DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS PATIENT-RELATED FACTOR. Manufacturer Narrative: WEIGHT-RACE:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM # (B)(4). | -6.00/2.0/108 | OS | Physician | 04/20/2022 | 2022 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14466408&pc=QCB | 2023826-2022-01377 | 2022-04-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 -12.5/4.0/082 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED WITHIN THE SAME SURGERY AND REPLACED WITH A SHORTER LENGTH LENS OF SIMILAR DIOPTER DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. NO PATIENT INJURY. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4). | -12.5/4.0/082 | OS | Physician | 04/21/2022 | 2022 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14466726&pc=QCB | 2023826-2022-01607 | 2022-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -10.0/1.0/152 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE REPORTED AS UNKNOWN.IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 -10/+1/152 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT WAS OBSERVED. LENS EXPLANTED ON (B)(6) 2022. ON (B)(6) 2022 LENS WAS REPLACED WITH A LONGER LENGTH LENS. PROBLEM NOT RESOLVED. THE CAUSE IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: H3: PRODUCT EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | -10.0/1.0/152 | OD | Physician | 10/04/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14466949&pc=QCB | 2023826-2022-01610 | 2022-04-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11.0/1.50/92 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5: "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.50/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR THE SAME MODEL/SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS REPORTED AS "ALL FINE" AND "PATIENT IS HAPPY". THE CAUSE OF THE EVENT IS UNKNOWN." H6- HEALTH EFFECT- IMPACT CODE: 2199 SHOULD BE REMOVED. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | -11.0/1.50/92 | OD | Physician | 04/22/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14466950&pc=QCB | 2023826-2022-01611 | 2022-04-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11.0/2.0/110 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: PT INFO: UNK. LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). Manufacturer Narrative: "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.0/110 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR THE SAME MODEL/SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS REPORTED AS "ALL FINE" AND "PATIENT IS HAPPY". THE CAUSE OF THE EVENT IS UNKNOWN." H6- HEALTH EFFECT- IMPACT CODE: 2199 SHOULD BE REMOVED. CLAIM# (B)(4). | -11.0/2.0/110 | OS | Physician | 04/22/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14473879&pc=QCB | 2023826-2022-01445 | 2022-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1, -8.5/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REPLACED INTRAOPERATIVELY WITH A LONGER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S RIGHT EYE (OD) DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.5/1.0/091 | OD | Physician | 05/02/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14473890&pc=QCB | 2023826-2022-01446 | 2022-01-24 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -12.0/5.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -12.0/5.0/085 | OS | Physician | 05/02/2022 | 2022 | 17 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14473892&pc=QCB | 2023826-2022-01450 | 2022-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -6.0/2.0/001 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. CLAIM# (B)(4). | -6.0/2.0/001 | OS | Physician | 05/06/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14473910&pc=QCB | 2023826-2022-01449 | 2022-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.0/2.0/176 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -7.0/2.0/176 | OD | Physician | 05/06/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14474063&pc=QCB | 2023826-2022-01447 | 2022-02-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -5.5/4.5/115 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -5.5/4.5/115 | OS | Physician | 05/02/2022 | 2022 | 41 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14478004&pc=QCB | 2023826-2022-01631 |  | Injury | STAAR SURGICAL COMPANY | 2022-05-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM512.1, -3.50/3.5/057 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. MULTIPLE REPOSITIONINGS WERE PERFORMED ON THE FOLLOWING DAYS (B)(6) 2022 DUE TO REFRACTIVE SURPRISE BUT IT DID NOT RESOLVE THE PROBLEM. CAUSE OF EVENT WAS REPORTED TO BE THE DEVICE. Manufacturer Narrative: B5- PER UPDATED INFORMATION LENS REPOSITIONING WAS PERFORMED. IT WAS REPORTED THAT REFRACTION IS BETTER, AND THE VISUAL ACUITY IS BETTER. PATIENT IS HAPPY. CLAIM # (B)(4). Manufacturer Narrative: H4: CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -3.50/3.5/057 | OD | Physician | 04/23/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14492633&pc=QCB | 2023826-2022-01600 | 1997-05-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -6.0/3.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED DUE TO LOW VAULT WITHOUT ROTATION. IN A SUBSEQUENT SURGERY LATER THAT DAY, A LONGER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: CLAIM #: (B)(4). | -6.0/3.0/089 | OD | Physician | 04/25/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14492635&pc=QCB | 2023826-2022-01602 | 2022-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.5/2.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. IN A SUBSEQUENT SUGERY LATER THAT DAY, A SHORTER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -12.5/2.0/094 | OD | Physician | 04/25/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14492637&pc=QCB | 2023826-2022-01599 | 2022-03-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -6.0/3.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED DUE TO LOW VAULT. IN A SUBSEQUENT SURGERY LATER THAT DAY, A LONGER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: CLAIM #: (B)(4). | -6.0/3.0/093 | OD | Physician | 04/25/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14493174&pc=QCB | 2023826-2022-01613 | 2022-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -18.0/3.0/122 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. LOW VAULT WITH LENS ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2022. THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS REMOVED AND A REPLACEMENT LENS OF A LONGER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - RACE: UNK. CLAIM# (B)(4). | -18.0/3.0/122 | OD | Physician | 04/26/2022 | 2022 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14504607&pc=QCB | 2023826-2022-01670 | 2022-04-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 12.6MM, VTICM5_12.6, -13.5/2/0/062 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS IMPLANTED,REMOVED AND REPLACED INTRAOPERATIVELY AND THE PROBLEM RESOLVED. THERE WAS NO PATIENT INJURY. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS THE DEVICE. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.5/2/0/062 | NA | Physician | 04/26/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14504787&pc=QCB | 2023826-2022-01691 | 2022-05-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -11.5/+2.0/83 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.5/+2.0/83 | OS | Physician | 05/24/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14504798&pc=QCB | 2023826-2022-01690 | 2022-05-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -11.5/+2.0/112 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE SURGEON REPORTS EXCESSIVE VAULTING. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.5/+2.0/112 | OD | Physician | 05/24/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14517631&pc=QCB | 2023826-2022-01672 | 2022-02-03 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -10.50/+2.50/95 (SPHERE/CYLINDER/AXIS) STUCK IN CARTRIDGE AND LENS WAS FOUND TORN ON (B)(6)2022. THE REPORTER STATED "FELT A FRICTION FORCE WHEN THE ICL PASSED THROUGH THE CARTRIDGE NOZZLE. IT SEEMS THAT HE PUSHED THE LENS FORWARD DESPITE OF RESISTANCE. SINCE THE LENS WAS STUCK IN CARTRIDGE, HE DISCONTINUED TO DELIVER THE LENS INTO THE EYE. FINALLY, HE FOUND THE LENS TORN." THERE WAS NO PATIENT CONTACT OR INJURY. A REPLACEMENT LENS WAS IMPLANTED, AND THE PROBLEM RESOLVED. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE:UNK. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.50/+2.50/95 | NA | Physician | 04/28/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14519714&pc=QCB | 2023826-2022-01675 | 2022-02-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5/2.5/174 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED DUE TO LOW VAULT WITHOUT ROTATION. IN A SUBSEQUENT SUGERY LATER THAT DAY, A LONGER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. CLAIM#: (B)(4). | -9.5/2.5/174 | OD | Physician | 04/26/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14519807&pc=QCB | 2023826-2022-01676 | 2022-03-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -11.5/4.0/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: PATIENT INFORMATION: UNK. CLAIM# (B)(4). | -11.5/4.0/93 | OD | Physician | 04/28/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14523577&pc=QCB | 2023826-2022-01669 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM TMICL12.1 IMPLANTABLE COLLAMER LENS, -09.00/+3.0/099 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPORTED AS UNDESIRABLE REFRACTIVE OUTCOME AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4). CLAIM #: (B)(4). | -09.00/+3.0/099 | OS | Other Health Care Professional | 05/02/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14525766&pc=QCB | 2023826-2022-01678 | 2022-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -8.0/2.0/003 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6)2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - "ON (B)(6) 2022, THE LENS WAS EXCHANGED AND IMPLANTED VERTICALLY FOR A SHORTER LENGTH LENS DUE TO EXCCESIVE VAULT." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: PT INFO: UNK. (B)(4). CLAIM # (B)(4). | -8.0/2.0/003 | OD | Physician | 04/29/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14525820&pc=QCB | 2023826-2022-01544 | 2022-04-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -11.50/3.0/108 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS OF DIFFERENT AXIS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE:UNK. (B)(4). | -11.50/3.0/108 | OD | Physician | 04/28/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14525897&pc=QCB | 2023826-2022-01677 | 2022-01-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.0/2.5/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | -9.0/2.5/089 | OS | Physician | 04/29/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14526754&pc=QCB | 2023826-2022-01681 | 2022-01-27 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.5/1.0/084 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. WITHIN THE SAME SURGERY, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. CLAIM# : (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -9.5/1.0/084 | OS | Physician | 05/02/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14538461&pc=QCB | 2023826-2022-01616 | 2022-04-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: IT WAS REPORTED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.5/085 (SPHERE/CYLINDER/AXIS) WAS DAMAGED (TEAR/BREAK) WHILE INJECTING INTO THE RIGHT EYE (OD) AND ABORTED THE SURGERY WITHOUT IMPLANTING THE LENS. THERE WAS PATIENT CONTACT WITH NO PATIENT INJURY. A REPLACEMENT LENS OF THE SAME MODEL AND SIZE AND POWER WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE:UNK. (B)(4). CLAIM# (B)(4). | -9.0/1.5/085 | OD | Physician | 04/27/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14538524&pc=QCB | 2023826-2022-01702 | 2022-01-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -15.5/3.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). ON (B)(6) 2022 THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. A REPLACEMENT LENS OF LONGER LENGTH LENS AND POWER WAS IMPLANTED. THE PROBLEM WAS RESOLVED. CASE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE:UNK. CLAIM# (B)(4). | -15.5/3.0/089 | OS | Physician | 04/29/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14538891&pc=QCB | 2023826-2022-01663 | 2022-04-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -05.50/+1.0/095 (SPHERE/CYLINDER/AXIS), TORE DURING LOADING AND THERE WAS NO PATIENT CONTACT. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM #: (B)(4). | -05.50/+1.0/095 | NA | Physician | 05/01/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14552185&pc=QCB | 2023826-2022-01704 | 2022-03-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(4) 2022 A 13.7MM, VTICMO13.7, -9.00/1.5/118 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING LOADING, THERE WAS NO PATIENT CONTACT. ON (B)(4) 2022 A REPLACEMENT LENS WAS IMPLANTED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM #: (B)(4). | -9.00/1.5/118 | NA | Physician | 05/04/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14552992&pc=QCB | 2023826-2022-01706 | 2021-07-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-05-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2021 A 12.6MM, VTICMO12.6, -11.5/3.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. THERE WAS NO PATIENT INJURY. A REPLACEMENT LENS WAS IMPLANTED AT A LATER DATE AND THE PROBLEM RESOLVED. Manufacturer Narrative: WEIGH-RACE:UNK. CLAIM #: (B)(6). | -11.5/3.0/088 | NA | Physician | 05/06/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14554192&pc=QCB | 2023826-2022-01682 | 2019-12-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 112.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -18.0/2.0/104 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE:UNK. CLAIM# (B)(4). | -18.0/2.0/104 | OD | Physician | 05/04/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14554572&pc=QCB | 2023826-2022-01684 | 2022-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-05-31 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -12.50/3/009 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. CLAIM# (B)(4). | -12.50/3/009 | OD | Physician | 05/06/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14566164&pc=QCB | 2023826-2022-01713 | 2022-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -4.0/2.0/003 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. IN A SUBSEQUENT SURGERY LATER THAT DAY, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. CLAIM # (B)(4). | -4.0/2.0/003 | OD | Physician | 05/06/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14566405&pc=QCB | 2023826-2022-01714 | 2022-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -3.5/2.0/167 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. IN A SUBSEQUENT SUGERY LATER THAT DAY, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE:UNK. CLAIM# (B)(4). | -3.5/2.0/167 | OS | Physician | 05/06/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14567813&pc=QCB | 2023826-2022-01549 | 2021-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.5/1.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED WITH A LOW VAULT WAS OBSERVED. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS OF DIFFERENT AXIS. THIS RESOLVED THE PROBLEM. CAUSE OF UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN/BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK. CLAIM# (B)(4). | -12.5/1.5/093 | OS | Physician | 05/02/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14570032&pc=QCB | 2023826-2022-01738 | 2022-04-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -17.00/2.0/115 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. IT WAS REPORTED THE WRONG LENS WAS IMPLANTED DUE TO THE ASSISTANT IN THE OPERATING ROOM GRABBED A LENS INTENDED FOR A DIFFERENT PATIENT; ONCE THE MISTAKE WAS REALIZED; INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH THE CORRECT LENS AND THE PROBLEM WAS RESOLVED BUT NOTES "BLURRY WITH SOME GHOSTING BUT IS RESOLVING WITH MR". SEE MFR# 2023826-2022-01738 FOR REPLACEMENT LENS. Manufacturer Narrative: CLAIM#: (B)(4). | -17.00/2.0/115 | OD | Physician | 05/10/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14570524&pc=QCB | 2023826-2022-01111 | 2022-02-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -13.5/1.5/111 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS DUE TO EXCESSIVE VAULT. THE REPLACEMENT LENS WAS IMPLANTED VERTICALLY AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND OPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE:UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. (B)(4). | -13.5/1.5/111 | OD | Physician | 03/13/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14573309&pc=QCB | 2023826-2022-01557 | 2021-12-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.0/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. (B)(4). | -12.0/1.0/089 | OD | Physician | 05/13/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14573311&pc=QCB | 2023826-2022-01264 | 2022-04-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -8.5/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.5/1.0/093 | OS | Physician | 05/10/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14595249&pc=QCB | 2023826-2022-01705 | 2021-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.0/3.0/073 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGE FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -10.0/+3.0/073 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. IT WAS REPORTED LENS ROTATION NOT ASSOCIATED WITH LOW VAULT. LENS REPOSITIONING WAS ATTEMPTED. REPORTEDLY "LENS ROTATED; REDIALLING TRIED BUT FAILED". ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL; LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION:UNK. NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.0/3.0/073 | OS | Physician | 05/05/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14595528&pc=QCB | 2023826-2022-01778 | 2021-03-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT 12.6MM, VTICMO12.6, -12.00/0.5/117 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS' RIGHT EYE (OD) AND THE PATIENT EXPERIENCED LOW VAULT. LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "THE REPORTER INDICATED THAT 12.6MM, VTICMO12.6, -12.00/0.5/117 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULT. CAUSE OF THE EVENT WAS UNKNOWN. LENS REMAINS IMPLANTED." CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS WITH DIOPTER -12.00/0.5/117 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED A LOW VAULT. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON THE SAME DAY SEPARATE SURGERY THE LENS WAS REPLACED WITH A DIFFERENT MODEL, LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: " VAULT PERFECT AND NO SYMPTOMS". THE CAUSE OF THE EVENT IS UNKNOWN. D6B: "ON (B)(6) 2022" SHOULD BE ADDED. H6: HEALTH EFFECT- IMPACT CODE: "2199" SHOULD BE REMOVED AND CODE "4627" SHOULD BE ADDED. H6: MEDICAL DEVICE CODE:"2993" SHOULD BE ADDED. CLAIM#: (B)(4). Manufacturer Narrative: DATE OF EVENT:UNK. IMPLANT DATE: UNK. CLAIM#: (B)(4). | -12.00/0.5/117 | OD | Physician | 05/06/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14595636&pc=QCB | 2023826-2022-01779 | 2021-03-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT 12.6MM, VTICMO12.6, -10.50/1.0/079 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) AND THE PATIENT EXPERIENCED LOW VAULT. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "THE REPORTER INDICATED THAT 12.6MM, VTICMO12.6, -10.50/1.0/079 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULT WITHOUT ROTATION. LENS REMAINS IMPLANTED." CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON (B)(6) 2022, THE LENS WAS EXPLANTED. ON THE SAME DAY SEPARATE SURGERY THE LENS WAS REPLACED WITH A DIFFERENT MODEL, LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: " VAULT OK AND NO SYMPTOMS". D6B: "ON (B)(6) 2022" SHOULD BE ADDED. H6: HEALTH EFFECT- IMPACT CODE: "2199" SHOULD BE REMOVED AND CODE "4627" SHOULD BE ADDED. H6: MEDICAL DEVICE CODE:"2993" SHOULD BE ADDED. CLAIM#: (B)(4). Manufacturer Narrative: DATE OF EVENT: UNK. IMPLANT DATE: UNK. TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -10.50/1.0/079 | OS | Physician | 05/06/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14598561&pc=QCB | 2023826-2022-01454 | 2022-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 OF -11.0/1.5/104 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS OF DIFFERENT AXIS. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE: UNK. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: PRODUCT EVALUATION: "H3- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND HAPTIC BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS." CLAIM: (B)(4). | -11.0/1.5/104 | OS | Physician | 05/06/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14598562&pc=QCB | 2023826-2022-01453 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-06-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 -8.50/1.0/107 (SPHERE/CYLINDER/AXIS) OF IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). A LOW VAULT WAS REPORTED. THE LENS REMAINS IMPLANTED AND BEING MONITORED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. INVESTIGATION TYPE (B)(4): LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -8.50/1.0/107 | OS | Physician | 05/06/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14598924&pc=QCB | 2023826-2022-01734 | 2022-05-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -13.0/1.5/136 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES) B5-THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -13.0/1.5/136 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WAS OBSERVED. CAUSE REPORTED AS UNKNOWN. PATIENT IS HAPPY. D6B: (B)(6) 2022 . CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. DEVICE CODE (B)(4): OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). CLAIM #: (B)(4). | -13.0/1.5/136 | OD | Physician | 05/06/2022 | 2022 | 66 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14599233&pc=QCB | 2023826-2022-01735 | 2019-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -13.5/2.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2019. LENS ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2020 BUT THE PROBLEM WAS NOT RESOLVED. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -13.5/2.0/091 | OD | Physician | 05/06/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14605977&pc=QCB | 2023826-2022-01715 | 2021-11-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -8.5/1.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -8.5/1.0/094 | OD | Physician | 05/11/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14605993&pc=QCB | 2023826-2022-01718 | 2022-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5/1.0/072 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. CLAIM# (B)(4). | -9.5/1.0/072 | OS | Physician | 05/13/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14606020&pc=QCB | 2023826-2022-01721 | 2022-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.5/3.5/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. CLAIM# (B)(4). | -11.5/3.5/095 | OS | Physician | 05/13/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14606059&pc=QCB | 2023826-2022-01716 | 2022-03-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -7.0/1.5/125 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | -7.0/1.5/125 | OS | Physician | 05/11/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14620803&pc=QCB | 2023826-2022-01790 | 2022-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.00/+1.5/003 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -13.00/+1.5/003 | OS | Physician | 05/10/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14620977&pc=QCB | 2023826-2022-01789 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50/+2.5/087 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS EYE. THE REPORTER INDICATED THEY ARE HAVING UNDER CORRECTION PROBLEM AND IS NOT ABLE TO SEE THINGS CLEARLY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK, WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, DATE OF EVENT: UNK, IMPLANT DATE: UNK. EXPLANT DATE: UNK. INITIAL REPORTER NAME AND ADDRESS: UNK. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THE ISSUE HAS RESOLVED AND THERE IS NO PROBLEM NOW. (B)(4). | -10.50/+2.5/087 | NA | Other | 05/09/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14627811&pc=QCB | 2023826-2022-01826 | 2022-03-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -11.50/1.00/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULTING. THIS DID NOT RESOLVE THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. SEE MFR# 2023826-2022-01827 FOR REPLACEMENT LENS. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | -11.50/1.00/093 | OS | Physician | 06/06/2022 | 2022 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14627816&pc=QCB | 2023826-2022-01827 | 2022-03-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -12.50/1.50/080 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 AS A REPLACEMENT LENS TO A PREVIOUS EXCESSIVE VAULTING ISSUE. IT WAS REPORTED THAT THE PATIENT STILL HAD EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND IT WAS NOTED "A REPLACEMENT LENS WILL BE INSERTED IN THE FUTURE; NEXT SURGERY DATE IS UNDECIDED". THE PROBLEM WAS NOT RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -12.50/1.50/080 | OS | Physician | 06/06/2022 | 2022 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14627896&pc=QCB | 2023826-2022-01723 | 2022-01-28 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -7.5/1.0/97 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED, AND REPLACED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: B1: MALFUNCTION. B5: AL THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO_13.2; -7.5/1.0/97 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFTEYE (OS). THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS INTRAOPERATIVELY ON (B)(6) 2022. THERE WAS NO PATIENT INJURY. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. H1: MALFUNCTION. H6: IMPACT CODE: 2199. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -7.5/1.0/97 | OS | Physician | 05/16/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14627901&pc=QCB | 2023826-2022-01724 | 2022-02-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.5/1.5/89 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -10.5/1.5/89 | OD | Physician | 05/16/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14628717&pc=QCB | 2023826-2022-01803 | 2020-04-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5/2.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2020, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.5/2.0/090 | OS | Physician | 05/16/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14628828&pc=QCB | 2023826-2022-01804 | 2022-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF --7.5/0.5/171 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -7.5/0.5/171 | OS | Physician | 05/16/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14632836&pc=QCB | 2023826-2022-01638 | 2022-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -9.0/1.5/081 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. CLAIM# (B)(4). | -9.0/1.5/081 | OS | Physician | 05/16/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14632838&pc=QCB | 2023826-2022-01636 | 2022-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | N/A | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -08.00/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT:- UNK, ETHNICITY:- UNK, RACE- UNK, PMA/510K- THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -08.00/1.5/089 | OS | Physician | 05/16/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14632842&pc=QCB | 2023826-2022-01643 | 2022-03-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.0/1.5/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. CLAIM# (B)(4). | -9.0/1.5/101 | OS | Physician | 05/16/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14632846&pc=QCB | 2023826-2022-01635 | 2022-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.0/1.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2022 DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. CLAIM# (B)(4). | -9.0/1.5/094 | OS | Physician | 05/16/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14632847&pc=QCB | 2023826-2022-01649 | 2022-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.0/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK, ETHNICITY - UNK, RACE - UNK, (B)(4). CLAIM# (B)(4). | -13.0/1.5/091 | OD | Physician | 05/18/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14632849&pc=QCB | 2023826-2022-01659 | 2022-02-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.0/2.0/115 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4) - OFF-LABEL USE, UNDER 21 YRS OF AGE AT DATE OF IMPLANT. CLAIM# (B)(4). | -14.0/2.0/115 | OS | Physician | 05/18/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14632852&pc=QCB | 2023826-2022-01642 | 2022-04-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -11.0/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON 01/APR/2022 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK, ETHNICITY - UNK, RACE - UNK. CLAIM# (B)(4). | -11.0/1.0/089 | OS | Physician | 05/16/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14632857&pc=QCB | 2023826-2022-01648 | 2022-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.50/1.5/068 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.50/1.5/068 | OD | Physician | 05/18/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14632860&pc=QCB | 2023826-2022-01654 | 2022-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.0/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK, ETHNICITY - UNK. RACE - UNK. CLAIM# (B)(4). | -13.0/2.0/089 | OS | Physician | 05/19/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14632870&pc=QCB | 2023826-2022-01634 | 2021-12-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.00/1.0/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. CLAIM# (B)(4). | -8.00/1.0/074 | OS | Physician | 05/16/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14632886&pc=QCB | 2023826-2022-01655 | 2022-02-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -12.0/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. CLAIM# (B)(4). | -12.0/2.5/092 | OD | Physician | 05/18/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14632890&pc=QCB | 2023826-2022-01639 | 2022-03-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.50/1.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.50/1.5/087 | OD | Physician | 05/16/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14632894&pc=QCB | 2023826-2022-01647 | 2021-11-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -3.50/2.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. CLAIM# (B)(4). | -3.50/2.0/084 | OD | Physician | 05/18/2022 | 2022 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14632899&pc=QCB | 2023826-2022-01652 | 2021-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.5/1.5/169 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/1.5/169 | OD | Physician | 05/18/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14632927&pc=QCB | 2023826-2022-01640 | 2022-03-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.0/1.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -14.0/1.5/088 | OD | Physician | 05/16/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14632931&pc=QCB | 2023826-2022-01646 | 2022-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -11.00/2.00/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4). | -11.00/2.00/092 | OD | Physician | 05/18/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14636104&pc=QCB | 2023826-2022-01807 | 2022-04-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -8.5/3.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. ON (B)(6) 2022, A SHORTER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. (B)(4). CLAIM# (B)(4). | -8.5/3.5/092 | OD | Physician | 05/17/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14636206&pc=QCB | 2023826-2022-01810 | 2022-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.5/3.0/81 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. OFF LABEL USE- ACD <3.0. CLAIM# 724646. | -12.5/3.0/81 | OD | Physician | 05/19/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14636454&pc=QCB | 2023826-2022-01808 | 2022-04-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -7.5/3.5/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. ON (B)(6) 2022, A SHORTER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. OFF LABEL AGE<21. CLAIM# (B)(4). | -7.5/3.5/088 | OS | Physician | 05/17/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14636577&pc=QCB | 2023826-2022-01708 |  | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -4.5/+2.0/96 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK. CLAIM# (B)(4). | -4.5/+2.0/96 | OD | Physician | 05/10/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14639322&pc=QCB | 2023826-2022-01709 |  | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -4.0/+1.5/103 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | -4.0/+1.5/103 | OS | Physician | 05/10/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14639677&pc=QCB | 2023826-2022-01811 | 2022-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.5/1.0/081 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. (B)(4). | -10.5/1.0/081 | OS | Physician | 05/19/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14650320&pc=QCB | 2023826-2022-01737 | 2021-07-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -3.0/2.0/080 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM #: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). B5-THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -3.0/2.0/080 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. D6B- (B)(6) 2022. CLAIM# (B)(4). | -3.0/2.0/080 | OS | Physician | 05/10/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14650792&pc=QCB | 2023826-2022-01743 | 2022-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 OF -15.5/2.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS OF DIFFERENT AXIS DUE TO EXCESSIVE VAULT. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK. CLAIM# (B)(4). | -15.5/2.5/091 | OD | Physician | 05/10/2022 | 2022 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14650944&pc=QCB | 2023826-2022-01817 | 2022-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -8.5/2.5/166 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM #: (B)(4). | -8.5/2.5/166 | OD | Physician | 05/19/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14655313&pc=QCB | 2023826-2022-01822 | 2022-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -7.5/1.0/125 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY, FOR A SHORTER LENTH LENS DUE TO EXCESSIVE VAULT THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN.Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -7.5/1.0/125 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY, FOR A SHORTER LENTH LENS DUE TO EXCESSIVE VAULT THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: A4-A6:UNK INITIAL DUPLICATE. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | -7.5/1.0/125 | OS | Physician | 05/20/2022 | 2022 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14659132&pc=QCB | 2023826-2022-01747 | 2022-01-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 -13.0/1.0/115 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.0/1.0/115 | OD | Physician | 05/13/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14659625&pc=QCB | 2023826-2022-01784 | 2019-09-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6. -8.5/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2019 THE LENS WAS EXCHANGED DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE AND THE PROBLEM RESOLVED. CAUSE WAS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM #: (B)(4). | -8.5/1.0/085 | OS | Physician | 05/12/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14687835&pc=QCB | 2023826-2022-01854 | 2022-04-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -9.50/+1.50/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT HAD EXPERIENCE LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2022 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND THE LENS OPTIC TORN, THE HAPTIC TORN AND BROKEN. DIMENSIONAL INSPECTION COULD NOT BE PERFORMED DUE TO HAPTIC DAMAGE OF FOOT PLATES. CLAIM#: (B)(4). | -9.50/+1.50/90 | OD | Physician | 06/09/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14687959&pc=QCB | 2023826-2022-01857 | 2022-01-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -9.5/2.0/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORT AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -9.5/2.0/094 | OS | Physician | 06/11/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14687988&pc=QCB | 2023826-2022-01858 | 2022-04-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.0/1.0/24 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | -13.0/1.0/24 | OS | Physician | 06/11/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14688029&pc=QCB | 2023826-2022-01856 | 2022-01-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -9.5/1.5/085 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORT AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -9.5/1.5/085 | OD | Physician | 06/11/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14688801&pc=QCB | 2023826-2022-01845 |  | Injury | STAAR SURGICAL COMPANY | 2022-06-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.00/+3.5/055 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE SURGEON ROTATED (REPOSITIONED) THE LENS TO ITS ORIGINAL AXIS. AS FOR NOW EVERYTHING IS GOOD AND THE PATIENT DOESN'T HAVE ANY COMPLAINT. CLAIM # (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK. WEIGHT: UNK ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANT DATE: UNK. CLAIM #: (B)(4). | -13.00/+3.5/055 | OD | Physician | 05/17/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14691660&pc=QCB | 2023826-2022-01859 | 2022-05-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER THAT INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -7.0/1.5/074 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -7.0/1.5/074 | OD | Physician | 06/11/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14691739&pc=QCB | 2023826-2022-01875 | 2022-04-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -15.5/2.5/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON (B)(6) 2022 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -15.5/2.5/093 | OD | Physician | 06/13/2022 | 2022 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14697132&pc=QCB | 2023826-2022-01767 | 2022-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.5/2.5/092 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK, (B)(4). CLAIM# (B)(4). | -10.5/2.5/092 | OD | Physician | 05/25/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14697144&pc=QCB | 2023826-2022-01774 | 2022-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.5/1.0/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK, ETHNICITY- UNK, RACE- UNK. CLAIM# (B)(4). | -8.5/1.0/070 | OS | Physician | 05/27/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14697171&pc=QCB | 2023826-2022-01763 | 2022-05-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.5/2.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK. CLAIM# (B)(4). | -8.5/2.5/089 | OD | Physician | 05/23/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14697180&pc=QCB | 2023826-2022-01764 | 2022-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #: (B)(4). | -11.0/1.0/090 | OD | Physician | 05/23/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14697186&pc=QCB | 2023826-2022-01773 | 2022-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.0/0.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK, RACE- UNK, ETHNICITY- UNK. CLAIM# (B)(4). | -8.0/0.5/088 | OD | Physician | 05/27/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14697193&pc=QCB | 2023826-2022-01770 | 2022-05-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/2.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK. CLAIM# (B)(4). | -10.5/2.0/084 | OS | Physician | 05/26/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14697200&pc=QCB | 2023826-2022-01771 | 2022-05-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -07.00/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK. CLAIM# (B)(4). | -07.00/2.0/088 | OD | Physician | 05/27/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14697223&pc=QCB | 2023826-2022-01765 | 2022-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.0/1.0/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, RACE- UNK, ETHNICITY- UNK. CLAIM# (B)(4). | -10.0/1.0/073 | OS | Physician | 05/23/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14705910&pc=QCB | 2023826-2022-01847 | 2022-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.00/+2.5/108 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (LENS WAS IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. (B)(4). | -14.00/+2.5/108 | OD | Physician | 05/19/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14707091&pc=QCB | 2023826-2022-01751 | 2022-05-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON ATTEMPTED TO IMPLANT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS -12.5/+2.5/086 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS) BUT THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE ON (B)(6) 2022. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. THE LENS INTRAOPERATIVELY IMPLANTED; REMOVED AND REPLACED FOR AN IDENTICAL LENS. THIS RESOLVED THE PROBLEM. CAUSE IS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. . CLAIM#: (B)(4). | -12.5/+2.5/086 | OS | Physician | 05/17/2022 | 2022 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14707340&pc=QCB | 2023826-2022-01753 | 2022-04-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5/1.5/171 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE IS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.5/1.5/171 | OD | Physician | 05/19/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14716420&pc=QCB | 2023826-2022-01866 | 2022-05-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER STATED "DURING THE PROCESS OF LENS LOADING AND OPENING THE LENS BOTTLE, THE BOTTLE TIN WAS NOT FULLY OPENED, AND THE BOTTLE WAS BROKEN". THE 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -7.5/1.0/90 (SPHERE/CYLINDER/AXIS) WAS STILL IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. CAUSE IS UNKNOWN. Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, THE PROBABLE CAUSE OF THE FAILURE IS LIKELY DUE TO CHANGE IN THE MATERIAL OF THE LENS VIAL. AN ACTION PLAN, HEALTH HAZARD EVALUATION (HHE) ON 2ML VIAL BREAKAGE WAS INITIATED TO ASSESS AND IMPLEMENT CORRECTIONS AND PREVENTIVE ACTIONS FOR THE FAILURE IDENTIFIED IN THIS COMPLAINT. THE HHE FOR THE 2ML VIALS FOUND NO SEVERE THREAT TO PUBLIC HEALTH AND THE OBSERVED FREQUENCY OF THE REPORTED EVENT OF THE VIAL BREAKAGE IS OCCURRING WITHIN THE ACCEPTABLE RANGE. THE MANUFACTURER OF THE 2ML VIALS WAS ALSO ASKED TO SWITCH BACK TO THE PRIOR PACKAGING CONFIGURATION TO PREVENT FURTHER VIALS FROM BREAKING DURING USAGE OR SHIPPING. THUS, NO FURTHER ACTION IS REQUIRED AT THIS TIME. CLAIM# (B)(4). Manufacturer Narrative: H6: TYPE OF INVESTIGATION: 10: TESTING OF ACTUAL DEVICE. INVESTIGATION FINDINGS: THE VIAL WAS RETURNED BROKEN. NO LENS IN THE BOX. VISUAL INSPECTION FOUND THE VIAL BROKEN AT THE NECK. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, THE PROBABLE CAUSE OF THE FAILURE IS LIKELY DUE TO CHANGE IN THE PACKAGING CONFIGURATION OF THE LENS VIAL. AN ACTION PLAN, HEALTH HAZARD EVALUATION (HHE) ON 2ML VIAL BREAKAGE WAS INITIATED TO ASSESS AND IMPLEMENT CORRECTIONS AND PREVENTIVE ACTIONS FOR THE FAILURE IDENTIFIED IN THIS COMPLAINT. THE HHE FOR THE 2ML VIALS FOUND NO SEVERE THREAT TO PUBLIC HEALTH AND THE OBSERVED FREQUENCY OF THE REPORTED EVENT OF THE VIAL BREAKAGE IS OCCURRING WITHIN THE ACCEPTABLE RANGE. THE MANUFACTURER OF THE 2ML VIALS WAS ALSO ASKED TO SWITCH BACK TO THE PRIOR PACKAGING CONFIGURATION TO PREVENT FURTHER VIALS FROM BREAKING DURING USAGE OR SHIPPING. THUS, NO FURTHER ACTION IS REQUIRED AT THIS TIME. Manufacturer Narrative: CORRECTED DATA: H6; DEVICE HISTORY RECORD (DHR) REVIEW: DEVICE HISTORY RECORD REVIEW FOUND THAT BASED ON THE RESULTS OF THE INVESTIGATION, THE PROBABLE CAUSE OF THIS EVENT IS RELATED TO A CHANGE IN THE PACKAGING CONFIGURATION OF THE 2ML VIALS. HEALTH HAZARD EVALUATION (GP-1416) FOUND NO SEVERE THREAT TO PUBLIC HEALTH. AN ACTION PLAN, CAPA M22-04, HAS BEEN INITIATED TO ASSESS AND IMPLEMENT CORRECTIVE AND PREVENTATIVE ACTIONS FOR THE FAILURE IDENTIFIED IN THIS COMPLAINT. CORRECTIVE AND/OR PREVENTIVE ACTIONS ARE BEING ADDRESSED THROUGH CAPA M22-04. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | -7.5/1.0/90 | OS | Physician | 09/20/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14721037&pc=QCB | 2023826-2022-01823 | 2021-12-24 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 -10.5/2.0/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED AND REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | -10.5/2.0/086 | OS | Physician | 05/20/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14722039&pc=QCB | 2023826-2022-01936 | 2020-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -9.5/2.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.5/2.0/092 | OS | Physician | 06/15/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14722044&pc=QCB | 2023826-2022-01935 | 2022-03-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.0/1.0/084 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING AND ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THE PROBLEM RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.0/1.0/084 | OD | Physician | 06/16/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14722045&pc=QCB | 2023826-2022-01934 | 2022-04-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -12.0/2.0/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | -12.0/2.0/087 | OS | Physician | 06/16/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14722053&pc=QCB | 2023826-2022-01938 | 2022-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -6.0/2.5/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -6.0/2.5/90 | OD | Physician | 06/15/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14735210&pc=QCB | 2023826-2022-01867 | 2021-02-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -8.50/2.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 LENS REPOSITIONING WAS PERFORMED DUE TO VAULT AND LENS ROTATION. IT WAS REPORTED THAT AFTER LENS REPOSITIONING THE PROBLEM RESOLVED BUT THE LENS ROTATED AGAIN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: B5: ADDITIONAL INFORMATION RECEIVED: THE LENS WAS EXCHANGED AND PROBLEM RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: PRODUCT EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND THE OPTIC AND HAPTIC TORN ON THE LENS. DIMENSIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -8.50/2.0/093 | OD | Physician | 05/18/2022 | 2022 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14738638&pc=QCB | 2023826-2022-01868 | 2021-10-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -8.5/1.5/74 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). | -8.5/1.5/74 | OD | Physician | 05/19/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14738790&pc=QCB | 2023826-2022-01869 | 2021-10-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -8.0/1.0/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE EXCHANGE DID NOT RESOLVED THE PROBLEM. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: H6-INVESTIGATION CONCLUSION CODE 4310 ADDED TO INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-RACE: UNK. CLAIM#: (B)(4). | -8.0/1.0/91 | OS | Physician | 05/19/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14738969&pc=QCB | 2023826-2022-01870 | 2021-10-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -8.00/1.0/86 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022 AS A REPLACEMENT LENS TO ADDRESS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE PATIENT EXPERIENCED EXCESSIVE VAULT. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR SPHERICAL LENS OF SHORTER LENGTH AND DIFFERENT POWER. IT WAS REPORTED THAT THE EXCHANGE RESOLVED THE PROBLEM. THE REPORTER STATED "THE RESIDUAL ASTIGMATISM WAS TREATED BY BIOPTICS." SEE MFR # 2023826-2022-01869 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: WEIGHT - RACE: UNK. CLAIM# (B)(4). | -8.00/1.0/86 | OS | Physician | 05/19/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14739380&pc=QCB | 2023826-2022-01939 | 2022-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -11.5/1.0/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -11.5/1.0/92 | OS | Physician | 06/16/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14739387&pc=QCB | 2023826-2022-01940 | 2022-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -8.0/0.5/105 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -8.0/0.5/105 | OS | Physician | 06/16/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14739421&pc=QCB | 2023826-2022-01942 | 2021-10-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.0/1.5/110 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.0/1.5/110 | OS | Physician | 06/15/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14739422&pc=QCB | 2023826-2022-01941 | 2020-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -10.5/2.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.5/2.0/090 | OD | Physician | 06/15/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14739481&pc=QCB | 2023826-2022-01919 | 2022-01-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -7.50/5.0/92 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. LENS ROTATION WAS OBSERVED AND REFRACTIVE SURPRISE. THE LENS WAS REPOSITIONED ON (B)(6) 2021 AND THE PROBLEM WAS NOT RESOLVED. ON (B)(6) 2022 THE LENS WAS EXPLANTED. CAUSE WAS NOT REPORTED. Manufacturer Narrative: H6-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. H6- METHOD CODE 3331: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -7.50/5.0/92 | OD | Other Health Care Professional | 05/18/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14739585&pc=QCB | 2023826-2022-01920 | 2022-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -11.0/4.0/84 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME SIZE SPHERICAL MODEL LENS DUE TO LENS ROTATION. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: H6-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN/SPLIT AND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | -11.0/4.0/84 | OS | Physician | 05/18/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14745285&pc=QCB | 2023826-2022-01831 | 2022-05-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -8.5/1.5/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) /2022 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM#: (B)(4). | -8.5/1.5/098 | OD | Physician | 06/01/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14745286&pc=QCB | 2023826-2022-01829 | 2022-03-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: AGE - UNK. WEIGHT - UNK. ETHNICITY - UNK. (B)(4). | -11.0/2.0/089 | OS | Physician | 05/29/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14745312&pc=QCB | 2023826-2022-01835 | 2021-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.0/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM#: (B)(4). | -9.0/2.0/091 | OS | Physician | 06/02/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14745314&pc=QCB | 2023826-2022-01834 | 2022-01-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.0/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM#: (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.01/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS REMOVED ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. A SHORTER LENGTH LENS WAS IMPLANTED IN A SECOND SURGERY ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT REPORT AS UNKNOWN. H6 REPORTED AS 4629, CORRECTED TO 4627 AND 4629. CLAIM# (B)(4). | -11.0/1.0/085 | OS | Physician | 06/01/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14745365&pc=QCB | 2023826-2022-01836 | 2022-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5 12.1, -10.00/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. CLAIM# (B)(4). | -10.00/1.0/091 | OD | Physician | 06/03/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14760308&pc=QCB | 2023826-2022-01926 | 2022-05-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -4.0/+0.5/78 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(6). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2022 THE LENS WAS EXPLANTED; ON THIS SAME DATE A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. "PATIENT IS HAPPY". CLAIM# (B)(4). | -4.0/+0.5/78 | OS | Physician | 05/24/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14760343&pc=QCB | 2023826-2022-01966 |  | Injury | STAAR SURGICAL COMPANY | 2022-06-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -6.00/1.5/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION. ON (B)(6) 2022 THE LENS WAS EXPLANTED. A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THERE WAS NO PATIENT INJURY. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). | -6.00/1.5/088 | NA | Other Health Care Professional | 05/26/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14760348&pc=QCB | 2023826-2022-01927 | 2022-05-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -3.5/1.0/74 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - IT WAS LATER REPORTED THAT ON (B)(6) 2022 THE LENS WAS EXPLANTED; ON THIS SAME DATE A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. "PATIENT IS HAPPY". CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -3.5/1.0/74 | OD | Physician | 05/24/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14780631&pc=QCB | 2023826-2022-01955 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-06-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -09.50/+2.0/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD). THE LENS WILL BE REMOVED AND REPLACED WITH A SHORTER LENS. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. EXPLANT DATE: UNK. CLAIM #: (B)(4). | -09.50/+2.0/086 | OD | Other Health Care Professional | 05/25/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14780638&pc=QCB | 2023826-2022-01961 | 2022-05-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.00/+3.0/078 (SPHERE/CYLINDER/AXIS), WAS STUCK IN CARTRIDGE OR INJECTION SYSTEM. THERE WAS PATIENT CONTACT BUT NO INJURY. THE LENS WAS REPLACED WITH A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. AT POST-OP 1 DAY THE EYE WAS QUIET AND SETTLING WELL. THE LENS WAS THROWN AWAY. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM #: (B)(4). | -13.00/+3.0/078 | NA | Physician | 05/26/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14780691&pc=QCB | 2023826-2022-01958 | 2022-03-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/105 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2022 BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS EXPLANTED. THE LENS WAS REPLACED WITH A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE RESIDUAL ASTIGMATISM WAS CORRECTED BY LASER. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM #: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND ONE HAPTIC BROKEN. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- "THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/105 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO A LOW VAULT, LENS REPOSITIONING, AND LASER REFRACTIVE SURGERY. ON (B)(6) 2022 THE LENS WAS REPOSITIONED, BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL; SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. REPORTEDLY, "THE RESIDUAL ASTIGMATISM WAS CORRECTED WITH LASER." CLAIM# (B)(4). | -10.00/+1.5/105 | OD | Physician | 05/23/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14780726&pc=QCB | 2023826-2022-01957 | 2022-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -01.50/+2.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND REFRACTIVE (FUNCTIONAL) VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM #: (B)(4). | -01.50/+2.0/089 | OD | Physician | 05/23/2022 | 2022 | 42 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14781238&pc=QCB | 2023826-2022-01956 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-06-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -09.00/+2.0/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WILL BE REMOVED AND REPLACED WITH A SHORTER LENS. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANT DATE UNK. CLAIM # (B)(4). | -09.00/+2.0/082 | OS | Other Health Care Professional | 05/25/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14782626&pc=QCB | 2023826-2022-02017 | 2022-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.0/2.0/064 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -11.0/2.0/064 | OS | Physician | 06/16/2022 | 2022 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14782627&pc=QCB | 2023826-2022-02014 | 2022-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -6.5/2.0/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -6.5/2.0/094 | OS | Physician | 06/16/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14782636&pc=QCB | 2023826-2022-02018 | 2021-06-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -6.0/1.0/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2021 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM#: (B)(4). | -6.0/1.0/90 | OS | Physician | 06/16/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14782792&pc=QCB | 2023826-2022-02000 | 2022-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -14.0/2.00/92 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | -14.0/2.00/92 | OD | Physician | 05/23/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14782794&pc=QCB | 2023826-2022-02002 | 2022-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -9.5/2.5/79 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | -9.5/2.5/79 | OS | Physician | 05/23/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14794440&pc=QCB | 2023826-2022-02022 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -7.00/2.5/001 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). IT WAS REPORTED THAT THERE IS A PLANNED LENS EXPLANT. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). | -7.00/2.5/001 | OD | Physician | 05/27/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14794520&pc=QCB | 2023826-2022-02021 | 2022-05-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -9.5/2.0/074 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.5/2.0/074 | OS | Physician | 06/16/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14794648&pc=QCB | 2023826-2022-02020 | 2016-10-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -9.0/2.0/099 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2016. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2016 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.0/2.0/099 | OD | Physician | 06/16/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14794707&pc=QCB | 2023826-2022-01892 | 2021-12-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.0/1.5/172 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE IS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A VIAL. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | -10.0/1.5/172 | OD | Physician | 05/27/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14794765&pc=QCB | 2023826-2022-02005 | 2021-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -8.5/3.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | -8.5/3.0/090 | OD | Physician | 05/25/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14794783&pc=QCB | 2023826-2022-02006 | 2021-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.0/0.5/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6- 06/30/2024 SHOULD BE CORRECTED TO 07/18/2021. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H4- 09/09/2022 SHOULD BE CORRECTED TO 07/18/2021. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT-RACE: UNK. CLAIM#: (B)(4). | -10.0/0.5/094 | OS | Physician | 05/25/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14794875&pc=QCB | 2023826-2022-01893 | 2022-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM TMICL13.2 -11.0/2.5/033 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REPORTEDLY A SMUDGE WAS FOUND ON THE ICL. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5- "THE REPORTER INDICATED THAT A 13.2MM TMICL 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/2.5/033 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REPORTEDLY A SMUDGE WAS FOUND ON THE ICL. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR THE SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE STATUS IS REPORTED AS "PT STATES VISION IS PERFECT". THE REPORTER INDICATED THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY "SMUDGE NOTED ON ICL OS AT SLIT LAMP EXAM (B)(6) 22. NOTED AGAIN ON (B)(6) 22." H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER-RACE:UNK. DATE OF EVENT: UNK. CLAIM#: (B)(4). | -11.0/2.5/033 | OS | Physician | 05/27/2022 | 2022 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14805949&pc=QCB | 2023826-2022-02009 | 2022-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -13.5/2.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | -13.5/2.5/093 | OD | Physician | 05/26/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14805975&pc=QCB | 2023826-2022-02010 | 2022-05-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.5/2.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | -11.5/2.0/093 | OS | Physician | 05/27/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14805984&pc=QCB | 2023826-2022-02012 | 2022-05-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -5.0/0.5/75 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | -5.0/0.5/75 | OD | Physician | 05/27/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14806063&pc=QCB | 2023826-2022-02008 | 2021-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/084 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | -9.0/1.0/084 | OS | Physician | 05/26/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14806097&pc=QCB | 2023826-2022-02013 | 2022-04-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -13.0/2.0/096 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(4) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | -13.0/2.0/096 | OD | Physician | 05/27/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14807970&pc=QCB | 2023826-2022-01952 | 2022-05-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -16.5/3.0/101 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED, LENS REMAINS IMPLANTED. CAUSE OF EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT-RACE: UNK. CLAIM#: (B)(4). | -16.5/3.0/101 | OS | Physician | 05/25/2022 | 2022 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14829836&pc=QCB | 2023826-2022-01881 | 2022-02-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.0/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -8.0/2.5/092 | OS | Physician | 06/07/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14829843&pc=QCB | 2023826-2022-01883 | 2022-04-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5 12.1, -13.00/2.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). | -13.00/2.0/087 | OS | Physician | 06/10/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14829882&pc=QCB | 2023826-2022-01882 | 2022-04-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5 12.1, -12.50/2.0/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). | -12.50/2.0/074 | OD | Physician | 06/10/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14830720&pc=QCB | 2023826-2022-01967 | 2022-05-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -16.0/2.5/107 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH LENS ROTATION. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM #: (B)(4). | -16.0/2.5/107 | OD | Physician | 05/26/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14840986&pc=QCB | 2023826-2022-02048 | 2022-03-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -14.0/2.0/105 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: (B)(4). | -14.0/2.0/105 | OS | Physician | 05/29/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14840991&pc=QCB | 2023826-2022-02047 | 2022-03-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -6.51.0/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -6.51.0/088 | OD | Physician | 05/29/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14840993&pc=QCB | 2023826-2022-02046 | 2022-03-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -4.00/4.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -4.00/4.0/091 | OS | Physician | 05/30/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14841733&pc=QCB | 2023826-2022-01954 | 2021-02-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS EXPLANTED. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. POST-OP, THE UCVA WAS GOOD BUT THERE WAS GLARE. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. THE LENS WAS DISCARDED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- " THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER-11.5/+1.5/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED ROTATION NOT ASSOCIATED WITH A LOW VAULT. ON (B)(6) 2021 THE LENS WAS ROTATED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND REPLACED WITH A LENS OF THE SAME MODEL/ LENGTH AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS REPORTED AS "GOOD UCVA. C/O GLARE OD (TICL) VS. OS (SPHERICAL)". THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS THE DEVICE AND THE DEVICE FAILED TO PERFORM AS INTENDED. THE LENS WAS DISCARDED" CLAIM# (B)(4). | -11.50/+1.5/094 | OD | Other Health Care Professional | 06/01/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14841751&pc=QCB | 2023826-2022-01959 | 2022-04-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -09.00/+6.0/085 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2022. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5- ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR THE SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3: TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. RACE: UNK. ETHNICITY: UNK. EXPLANT DATE: UNK. CLAIM #: (B)(4). | -09.00/+6.0/085 | OS | Physician | 05/31/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14841802&pc=QCB | 2023826-2022-01962 | 2022-01-18 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.50/+3.5/091 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK ETHNICITY: UNK RACE: UNK (B)(4). CLAIM # (B)(4). | -12.50/+3.5/091 | OS | Physician | 05/29/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14846856&pc=QCB | 2023826-2022-02053 | 2022-03-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -6.0/5.0/94 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO:UNK. OFF LABEL USE- ACD <3.0. CLAIM# (B)(4). | -6.0/5.0/94 | OD | Physician | 05/29/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14846868&pc=QCB | 2023826-2022-02049 | 2022-03-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -14.5/2.0/066 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: (B)(4). | -14.5/2.0/066 | OD | Physician | 05/29/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14846886&pc=QCB | 2023826-2022-02050 | 2021-12-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -13.0/3.0/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.0/3.0/93 | OD | Physician | 05/29/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14846942&pc=QCB | 2023826-2022-02058 | 2022-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -8.5/1.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -8.5/1.0/092 | OS | Physician | 05/30/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14846944&pc=QCB | 2023826-2022-02051 | 2021-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.0/2.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.0/2.0/093 | OD | Physician | 05/29/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14846950&pc=QCB | 2023826-2022-02057 | 2022-03-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -4.5/0.5/054 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -4.5/0.5/054 | OD | Physician | 05/30/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14859851&pc=QCB | 2023826-2022-02062 | 2022-01-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -14.0/1.0/85 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022; THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: UNK. CLAIM#: (B)(4). | -14.0/1.0/85 | OD | Physician | 05/30/2022 | 2022 | 45 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14860021&pc=QCB | 2023826-2022-02063 | 2022-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -9.0/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: UNK OFF LABEL USE- ACD <3.0. CLAIM# (B)(4). | -9.0/2.0/089 | OS | Physician | 05/31/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14861997&pc=QCB | 2023826-2022-01897 | 2022-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -13.0/1.0/105 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS PATIENT RELATED-FACTOR. Manufacturer Narrative: PT INFO:UNK, IMPLANT DATE: UNK. CLAIM# (B)(4). | -13.0/1.0/105 | OD | Physician | 05/31/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862017&pc=QCB | 2023826-2022-02087 | 2022-03-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -10.50/+2.50/88 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2022 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. THEN ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: DEVICE EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4) - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.50/+2.50/88 | OD | Physician | 06/23/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862129&pc=QCB | 2023826-2022-02088 | 2021-02-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -15.00/+2.00/93 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE PATIENT EXPERIENCED LENS ROTATION. ON (B)(6) 2020 THE LENS WAS REPOSITIONED. BUT THIS DID NOT RESOLVE THE PROBLEM. THEN ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH/DIFFERENT MODEL LENS AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: B5 - LASER REFRACTIVE SURGERY WAS ALSO PERFORMED TO CORRECT THE PATIENTS RESIDUAL ASTIGMATISM, SHOULD HAVE ALSO BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.00/+2.00/93 | OD | Physician | 06/23/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862254&pc=QCB | 2023826-2022-02090 | 2022-05-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -8.50/+2.00/87 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH SAME MODEL LENS WITH NO PATIENT INJURY. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | -8.50/+2.00/87 | OS | Physician | 06/23/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862296&pc=QCB | 2023826-2022-02029 | 2022-04-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-06-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 -10.0/2.0/101 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS TORN DURING LOADING INTO THE INJECTOR. THERE WAS NO PATIENT CONTACT. ON (B)(6) 2022 A LENS OF DIFFERENT POWER WAS IMPLANTED INTO THE EYE. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | -10.0/2.0/101 | NA | Physician | 06/16/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862376&pc=QCB | 2023826-2022-01903 | 2022-04-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.00/2.5/073 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2017. LENS ROTATION NOT ASSOCIATED WITH A LOW VAULT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PT INFO: UNK. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.00/2.5/073 | OD | Physician | 06/09/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862524&pc=QCB | 2023826-2022-02024 | 2022-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -12.5/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON THE SAME DAY IN A SEPARATE SURGERY A LONGER LENGTH LENS OF DIFFERENT POWER WAS IMPLANTED. THIS RESOLVED THE PROBLEM. THE CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: H3 DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: UNK INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.5/2.0/089 | OS | Physician | 06/14/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862576&pc=QCB | 2023826-2022-02089 | 2022-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -13.00/+0.50/96 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON 19-NOV-2021. THE PATIENT EXPERIENCED LENS ROTATION. ON 15-MAY-2022 THE LENS WAS REPOSITIONED; BUT THIS DID NOT RESOLVE THE PROBLEM. ON 04-JUN-2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT MODEL AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B5 - BIOPTICS WAS ALSO PERFORMED TO CORRECT THE PATIENTS RESIDUAL ASTIGMATISM, SHOULD HAVE ALSO BEEN INCLUDED IN INITIAL MEDWATCH REPORT. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -13.00/+0.50/96 | OD | Physician | 06/23/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862622&pc=QCB | 2023826-2022-01901 | 2022-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 -9.50/2.0/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS OF DIFFERENT AXIS DUE TO EXCESSIVE VAULT. THE EXCHANGE LENS WAS IMPLANTED VERTICALLY. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PT INFO:UNK. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.50/2.0/087 | OS | Physician | 06/03/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862648&pc=QCB | 2023826-2022-01904 | 2022-04-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -14.00/2.5/084 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. LENS ROTATION NOT ASSOCIATED WITH A LOW VAULT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PT INFO: UNK. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: CLAIM# (B)(4). | -14.00/2.5/084 | OS | Physician | 06/09/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14862752&pc=QCB | 2023826-2022-02028 | 2022-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-06-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.0/2.5/082 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2022 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION. THIS RESOLVED THE PROBLEM. THE LENS REMAINS IMPLANTED. THE CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: UNK. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: ON (B)(6) 2022 THE LENS WAS REPLACED WITH A SIMILAR LENS DUE TO ROTATION NOT ASSOCIATED TO A LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT WAS A PATIENT-RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -10.0/2.5/082 | OS | Physician | 06/16/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14886286&pc=QCB | 2023826-2022-02096 | 2021-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -10.50/+2.00/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -10.50/+2.00/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. D6A- CORRECTED IMPLANT DATE FROM (B)(6) 2021 TO (B)(6) 2020. CLAIM# (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | -10.50/+2.00/91 | OD | Physician | 06/03/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14919321&pc=QCB | 2023826-2022-02064 | 2022-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -14.5/3.5/100 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LENS ROTATION WAS OBSERVED. ON (B)(6) 2021 THE LENS WAS REPOSITIONED AND THE PROBLEM WAS NOT RESOLVED. ON (B)(6) 2022 THE LENS WAS REMOVED AND A REPLACEMENT LENS OF SPHERICAL MODEL BUT SAME SIZE WAS USED. THE ASTIGMATISM DEGREE WAS CORRECTED BY LASER SURGERY. THE PROBLEM IS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. IMPACT (B)(4): LASER SURGERY. DEVICE CODE (B)(4): OFF-LABEL USE (ACD<3.00MM). CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: PRODUCT EVALUATION: LENS RETURNED DRY A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -14.5/3.5/100 | OD | Physician | 05/31/2022 | 2022 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14919819&pc=QCB | 2023826-2022-02079 | 2022-04-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+2.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (THE LENS WAS IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: TWO SIMILAR COMPLAINTS WERE REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.50/+2.0/089 | OS | Physician | 06/07/2022 | 2022 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14930541&pc=QCB | 2023826-2022-02112 | 2022-03-11 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -5.5/1.00/090 (SPHERE/CYLINDER/AXIS) WAS INTRAOPERATIVELY IMPLANTED, REMOVED, AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | -5.5/1.00/090 | OS | Physician | 06/06/2022 | 2022 | 41 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14930778&pc=QCB | 2023826-2022-02113 | 2022-03-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -12.00/1.00/092 (SPHERE/CYLINDER/AXIS) WAS INTRAOPERATIVELY IMPLANTED, REMOVED, AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, IN THE PATIENT'S LEFT EYE (OD) ON (B)(6) 2022. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. CLAIM#: (B)(4). | -12.00/1.00/092 | OS | Physician | 06/06/2022 | 2022 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14930905&pc=QCB | 2023826-2022-01972 | 2022-03-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.0/1.0/177 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS VERTICALLY DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -10.0/1.0/177 | OD | Physician | 06/11/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14930906&pc=QCB | 2023826-2022-01977 | 2022-04-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.5/1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -8.5/1.0/092 | OS | Physician | 06/11/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14930911&pc=QCB | 2023826-2022-01983 | 2022-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.0/1.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -9.0/1.5/088 | OS | Physician | 06/15/2022 | 2022 | 42 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14930919&pc=QCB | 2023826-2022-01990 | 2022-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.5/1.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. CLAIM# (B)(4). | -9.5/1.5/086 | OS | Physician | 06/16/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14930925&pc=QCB | 2023826-2022-01973 | 2022-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.5/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. CLAIM# (B)(4). | -11.5/1.0/087 | OS | Physician | 06/11/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14930928&pc=QCB | 2023826-2022-01995 | 2021-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -06.0/1.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. CLAIM# (B)(4). | -06.0/1.0/086 | OD | Physician | 06/16/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14930930&pc=QCB | 2023826-2022-02073 | 2021-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.00/1.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). | -10.00/1.5/094 | OS | Physician | 06/16/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14930934&pc=QCB | 2023826-2022-01985 | 2022-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.5/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -8.5/2.0/095 | OD | Physician | 06/15/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14930936&pc=QCB | 2023826-2022-01992 | 2021-01-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -11.5/2.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2021. FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT.THIS RESOLVED THE PROBLEM CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: UNDER 21 YRS OF AGE AT DATE OF IMPLANT. CLAIM# (B)(4). | -11.5/2.0/096 | OD | Physician | 06/16/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14930940&pc=QCB | 2023826-2022-01996 | 2021-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -05.0/1.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. ON (B)(6) 2021 A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -05.0/1.5/085 | OS | Physician | 06/16/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14930995&pc=QCB | 2023826-2022-01975 | 2022-03-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.0/2.0/171 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. CLAIM# (B)(4). | -9.0/2.0/171 | OS | Physician | 06/11/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14930998&pc=QCB | 2023826-2022-01981 | 2022-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.5/1.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). | -11.5/1.5/087 | OS | Physician | 06/15/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14931000&pc=QCB | 2023826-2022-01988 | 2022-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.0/1.5/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). | -12.0/1.5/107 | OD | Physician | 06/16/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14931009&pc=QCB | 2023826-2022-01999 | 2022-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -13.0/1.5/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.0/1.5/112 | OD | Physician | 06/16/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14931031&pc=QCB | 2023826-2022-01978 | 2021-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.5/2.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -11.5/2.0/087 | OS | Physician | 06/14/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14931036&pc=QCB | 2023826-2022-01971 | 2022-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.5/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -12.5/1.5/091 | OD | Physician | 06/13/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14931051&pc=QCB | 2023826-2022-01984 | 2022-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.0/1.5/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. CLAIM# (B)(4). | -8.0/1.5/109 | OD | Physician | 06/15/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14931060&pc=QCB | 2023826-2022-02074 | 2019-04-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/4.5/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2019. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. CLAIM# (B)(4). | -10.5/4.5/101 | OD | Physician | 06/16/2022 | 2022 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14931067&pc=QCB | 2023826-2022-01976 | 2022-03-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.5/0.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -7.5/0.5/085 | OD | Physician | 06/11/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14931071&pc=QCB | 2023826-2022-01974 | 2022-05-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.0/1.0/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. CLAIM# (B)(4). | -9.0/1.0/107 | OS | Physician | 06/11/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14931076&pc=QCB | 2023826-2022-01982 | 2022-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.5/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. CLAIM# (B)(4). | -13.5/1.0/091 | OD | Physician | 06/15/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14931077&pc=QCB | 2023826-2022-01987 | 2022-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.0/2.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. CLAIM# (B)(4). | -12.0/2.5/089 | OS | Physician | 06/16/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14931082&pc=QCB | 2023826-2022-01998 | 2022-01-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.0/1.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -10.0/1.0/083 | OS | Physician | 06/16/2022 | 2022 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14931084&pc=QCB | 2023826-2022-01989 | 2022-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.0/1.0/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. CLAIM# (B)(4). | -11.0/1.0/080 | OD | Physician | 06/16/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14931099&pc=QCB | 2023826-2022-01994 | 2021-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -6.5/2.5/076 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. (B)(4). | -6.5/2.5/076 | OD | Physician | 06/16/2022 | 2022 | 17 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14931110&pc=QCB | 2023826-2022-02072 | 2021-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/3.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. CLAIM# (B)(4). | -10.5/3.5/091 | OS | Physician | 06/16/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14932100&pc=QCB | 2023826-2022-02114 | 2022-05-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -11.0/2.0/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. IN A SUBSEQUENT SURGERY, LATER THAT DAY, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATON. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -11.0/2.0/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. CLAIM# (B)(4). | -11.0/2.0/086 | OS | Physician | 06/07/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14932161&pc=QCB | 2023826-2022-02115 | 2021-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5/2.0/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. (B)(4). | -9.5/2.0/086 | OD | Physician | 06/07/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14934906&pc=QCB | 2023826-2022-02103 | 2022-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.0/+1.0/123 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022 AND THE PATIENT EXPERIENCED LOW VAULT. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: ADDITIONAL INFORMATION: PROBLEM RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11/1.0/123 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULT WITH UNSPECIFIED ROTATION. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. PER ADDITIONAL INFORMATION, REPORTEDLY THE PATIENT IS HAPPY. PROBLEM RESOLVED. D9: RETURN DATE: 05-SEP-2023. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -11.0/+1.0/123 | OD | Physician | 09/23/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14939331&pc=QCB | 2023826-2022-01874 | 2021-11-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -8.5/1.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. IT WAS REPORTED THAT IN THE REPORTER OPINION " USER ERROR " WAS THE CAUSE OF THIS EVENT. HOWEVER, IT WAS REPORTED THAT THE DEVICE FAILED TO PERFORM AS INTENDED. FOR CAUSE DETAILS THE REPORTER STATED "VERY SMALL VAULT. SURGEON PREFERENCE CHANGE". FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) "NO CONCERNS" WAS REPORTED. Manufacturer Narrative: (B)(4). | -8.5/1.5/094 | OS | Physician | 06/06/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14940189&pc=QCB | 2023826-2022-02170 | 2022-03-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.50/+2.0/177 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE INITIAL AND REPLACEMENT LENSES WERE IMPLANTED VERTICALLY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN/BENT AND DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -16.50/+2.0/177 | OD | Physician | 06/11/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14940345&pc=QCB | 2023826-2022-02085 | 2022-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.00/+3.0/079 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. THE PATIENT IS HAPPY. Manufacturer Narrative: (B)(4). | -16.00/+3.0/079 | OS | Physician | 06/10/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14940370&pc=QCB | 2023826-2022-02075 |  | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.00/+1.0/095 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO THE LENS FLIPPED. THERE WAS NO PATIENT INJURY AND THE BACKUP LENS WAS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH: UNK, SEX: UNK, WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, DATE OF EVENT: UNK, IMPLANT DATE: UNK, EXPLANT DATE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | -09.00/+1.0/095 | NA | Other Health Care Professional | 06/07/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14940391&pc=QCB | 2023826-2022-02082 |  | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.50/+2.5/103 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE, THE LENS HAD A LOW VAULT AND HAD ROTATED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENS ON (B)(6) 2022. THE EXCHANGE WAS PERFORMED SUCCESSFULLY. CLAIM # (B)(4). | -13.50/+2.5/103 | OD | Physician | 06/09/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14940465&pc=QCB | 2023826-2022-02083 |  | Injury | STAAR SURGICAL COMPANY | 2022-07-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.50/+2.5/101 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE, THE LENS HAD A LOW VAULT AND HAD ROTATED. THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. RELEVANT TESTS/LABORATORY DATA: UNK, WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2022 FOR A LONGER LENS. THE LENS EXCHANGE WAS PERFORMED SUCCESSFULLY. CLAIM # (B)(4). | -12.50/+2.5/101 | OS | Physician | 06/09/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14949118&pc=QCB | 2023826-2022-02105 |  | Injury | STAAR SURGICAL COMPANY | 2022-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -7.50/3.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS MOVED. THE LENS WAS EXPLANTED AND A REPLACEMENT LENS OF LONGER LENGTH WAS LATER IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -7.50/3.5/090 | NA | Physician | 06/07/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14950205&pc=QCB | 2023826-2022-02068 | 2022-05-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 OF -10.0/2.0/080 (SPHERE/CYLINDER/AXIS) LENS WAS REPORTED TO TEAR/ BREAK DURING INJECTION/DELIVERY INTO THE RIGHT EYE (OD). ON (B)(6) 2022 THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. A REPLACEMENT LENS OF SAME MODEL AND SIZE WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED, CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | -10.0/2.0/080 | OD | Physician | 06/07/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14958674&pc=QCB | 2023826-2022-02069 | 2022-04-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -9.0/150/84 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME SIZE SPHERICAL MODEL LENS DUE TO LENS ROTATION. THE REPLACEMENT LENS RESOLVED THE PROBLEM. NO ADDITIONAL TREATMENT FOR RESIDUAL ASTIGMATISM. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H6-DEVICE EVALUATION: LENS RETURNED IN A MICRO-CENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -9.0/150/84 | OS | Physician | 06/09/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14959365&pc=QCB | 2023826-2022-02110 | 2021-12-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO_13.2, -13.5/2.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS BUT IT DID NOT RESOLVE THE PROBLEM. THE REPLACEMENT LENS WAS DAMAGED DURING THE OPERATION AND IT WAS REMOVED IN THE SAME OPERATION. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -13.5/2.5/080 | OS | Physician | 06/11/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14959579&pc=QCB | 2023826-2022-02111 | 2021-12-30 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO_12.6, -13.0/2.5/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021 AS A REPLACEMENT LENS TO ADDRESS EXCESSIVE VAULT BUT THE LENS WAS DAMGED DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY WITH NO PATIENT INJURY. ON (B)(6) 2022 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVE. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS UNKNOWN. ON (B)(6) 2022. THE VAULT VALUE WAS REPORTED TO 700. SEE MFR# 2023826-2022-02110 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.0/2.5/077 | OS | Physician | 06/11/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14959584&pc=QCB | 2023826-2022-02117 | 2022-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -13.5/3.0/92 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. IN A SUBSEQUENT SURGERY LATER THAT DAY, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: D9: 09/08/2021 SHOULD BE CORRECTED TO 06/20/2022. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H4: 06/20/2022 SHOULD BE CORRECTED TO 09/08/2021. CLAIM #(B)(4). Manufacturer Narrative: PT INFO: UNK. THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM# (B)(4). | -13.5/3.0/92 | OD | Physician | 06/10/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14959634&pc=QCB | 2023826-2022-02116 | 2022-03-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.5/2/68 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH, DIFFERENT MODEL LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE US. CLAIM#: (B)(4). | -12.5/2/68 | OS | Physician | 06/08/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14959788&pc=QCB | 2023826-2022-02171 | 2021-12-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.00/+2.5/078 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. A POSSIBLE CAUSE OF THE EVENT MAY HAVE BEEN DUE TO PATIENT RELATED FACTOR (CRYSTAL LENS RISE CLR). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM #: (B)(4). | -07.00/+2.5/078 | OS | Physician | 03/22/2022 | 2022 | 45 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14962883&pc=QCB | 2023826-2022-02178 | 2022-06-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -10.0/4.00/176 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE ON (B)(6) 2022. IT WAS NOTED THERE WAS CARTRIDGE FAILURE; AND STATED "EVEN IF YOU PULL THE LENS TO THE TIP OF THE CARTRIDGE; THE LENS WILL RETURN FROM THE NOZZLE. SO YOU CANNOT SET THE LENS. ALSO EVEN IF ONLY BSS IS USED IN THE CARTRIDGE; THE LENS WILL RETURN. THE LENS WAS IMPLANTED BY USING A CARTRIDGE FROM ANOTHER LOT. LENS REMAINS IMPLANTED. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "AFTER APPLYING 1ST BSS AND 2ND VISCOELASTIC SUBSTANCES TO THE CARTRIDGE IN THIS ORDER; THE LENS IS SET". Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: CARTIDGE DEVICE WAS RETURNED IN A ZIPLOC BAG. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE PRODUCT; RESIDUE IN THE NOSE CONE. H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | -10.0/4.00/176 | NA | Physician | 06/19/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14978206&pc=QCB | 2023826-2022-01913 | 2022-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -12.5/1.5/104 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS OF LONGER LENGTH DUE TO LOW VAULT WITH ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS OTHER, "SMALL LENS SIZE." Manufacturer Narrative: ADDITIONAL INFORMATION: (B)(4). | -12.5/1.5/104 | OD | Physician | 05/12/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14978541&pc=QCB | 2023826-2022-01914 | 2022-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/071 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS OF LONGER LENGTH DUE TO LOW VAULT WITH ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS OTHER, "SMALL LENS SIZE." Manufacturer Narrative: ADDITIONAL INFORMATION: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -12.0/1.5/071 | OS | Physician | 05/12/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14983229&pc=QCB | 2023826-2022-02207 | 2022-05-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.5/1.0/080 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE PATIENT'S EYE. INTRAOPERATIVELY THE LENS WAS REMOVED WITH NO PATIENT INJURY. ON (B)(6) 2022 A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED AT 54 DEGREES CLOCKWISE AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.5/1.0/080 | OS | Physician | 07/07/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14983334&pc=QCB | 2023826-2022-02205 | 2022-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.5/3.0/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON THE SAME DATE IN A SEPARATE SURGERY THE LENS WAS EXPLANTED. IT WAS NOTED "THE PATIENT IS UNDER CONSIDERATION; AND AT THE TIME OF THE SECOND OPERATION IS UNCERTAIN". THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -11.5/3.0/091 | OS | Physician | 07/07/2022 | 2022 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14985523&pc=QCB | 2023826-2022-02131 | 2022-01-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.0/3.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -15.0/3.0/091 | OD | Physician | 06/21/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14985524&pc=QCB | 2023826-2022-02132 | 2022-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -16.00/2.5/106 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -16.00/2.5/106 | OD | Physician | 06/21/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14985528&pc=QCB | 2023826-2022-02127 | 2022-01-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.0/2.0/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.0/2.0/105 | OD | Physician | 06/20/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14985530&pc=QCB | 2023826-2022-02126 | 2022-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/2.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # B)(4). Manufacturer Narrative: CORRECTION: B2 - REPORTED AS "HOSPITALIZATION (INITIAL OR PROLONGED)" - CORRECT TO "REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE". CLAIM# (B)(4). | -10.5/2.0/092 | OS | Physician | 06/20/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14985531&pc=QCB | 2023826-2022-02138 | 2022-06-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.0/2.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. CLAIM # (B)(4). | -11.0/2.5/088 | OD | Physician | 06/23/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14985539&pc=QCB | 2023826-2022-02139 | 2021-10-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.5/2.0/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. CLAIM # (B)(4). | -13.5/2.0/109 | OS | Physician | 06/23/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14985540&pc=QCB | 2023826-2022-02145 | 2022-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6, -9.5/1.5/152 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO LOW VAULT. A REPLACEMENT LENS WAS IMPLANTED ON 23/JUN/2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. (B)(4). CLAIM # (B)(4). | -9.5/1.5/152 | OS | Physician | 06/24/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14985544&pc=QCB | 2023826-2022-02151 | 2022-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.0/1.5/051 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5- " THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.0/1.5/051 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON THE SAME DAY, SEPARATE SURGERY THE LENS WAS REPLACED WITH A SAME MODEL/ LENGTH LENS, DIFFERENT AXIS. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN." CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: "THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/1.5/51 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 WHEN THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR THE SAME MODEL/LENGTH; DIFFERENT AXIS LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY "AFTER COMMUNICATION WITH SURGEON THE LENS WAS STUCK IN THE INCISION AREA THREE TIMES DURING IMPLANTATION SURGERY, AND THERE WERE SUSPECTED DAMAGE ON THE OPTIC ZONE OF ICL LENS DURING POST-OPERATION EXAMINATION SO THE SURGEON HAD DECIDED TO EXCHANGE THE LENS AFTER CAREFUL CONSIDERATION." CORRECTED DATA: HELATH EFFECT- IMPACT CODE: 4627 SHOULD BE REMOVED. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -9.0/1.5/051 | OD | Physician | 06/24/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=14985562&pc=QCB | 2023826-2022-02150 | 2022-04-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -6.0/3.0/006 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. CLAIM# (B)(4). | -6.0/3.0/006 | OD | Physician | 06/24/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15004305&pc=QCB | 2023826-2022-02223 | 2022-03-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 12.1MM VTICM5_12.1, -8.00/+2.5/004, IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. IT WAS REPORTED THAT THERE WAS PATIENT CONTACT WITH NO PATIENT INJURY. ON (B)(6) 2022 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -8.00/+2.5/004 | NA | Physician | 06/15/2022 | 2022 | 51 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15005302&pc=QCB | 2023826-2022-02228 | 2022-06-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 12.6MM VTICM5_12.6, -9.00/1.0/93, IMPLANTABLE COLLAMER LENS GOT STUCK IN THE PLUNGER WHILE IMPLANTING AND TORE. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY WITH NO PATIENT INJURY. ON (B)(6) 2022 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -9.00/1.0/93 | NA | Physician | 06/16/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15008799&pc=QCB | 2023826-2022-02202 | 2022-04-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -10.00/+1.0/078 (SPHERE/CYLINDER/AXIS) TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS PATIENT CONTACT BUT NO INJURY. THE LENS WAS REPLACED WITH A SAME LENGTH BUT DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.00/+1.0/078 | OS | Physician | 06/13/2022 | 2022 | 46 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15009494&pc=QCB | 2023826-2022-02286 |  | Injury | STAAR SURGICAL COMPANY | 2022-07-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -13.0/+1.0/107 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. IT WAS REPORTED THAT THE WRONG LENS WAS IMPLANTED DUE TO THE CUSTOMER ORDERED THE WRONG LENS REGARDING THE SEQ AND LENS. THE PATIENT EXPERIENCED NO ADVERSE EVENTS. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT DIFFERENT POWER LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS USER ERROR. Manufacturer Narrative: CLAIM# (B)(4). | -13.0/+1.0/107 | OS | Physician | 07/13/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15010209&pc=QCB | 2023826-2022-02234 | 2022-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/110 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT THE PROBLEM DID NOT RESOLVE. ON (B)(6) 2022 THE LENS WAS EXPLANTED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -9.0/1.0/110 | OD | Physician | 06/17/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15010330&pc=QCB | 2023826-2022-02235 | 2022-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/077 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. ON (B)(6) 2022 THE LENS WAS EXPLANTED. CAUSE OF EVENT WAS UNKNOWN. THE REPORTER STATED " AFTER THE LENS OF THE PATIENTS LEFT EYE ROTATED, HE WAS VERY DISSATISFIED AND UNWILLING TO ACCEPT THE REPLACEMENT OF THE LENS, SO HE VOLUNTARILY REMOVED THE LENS." Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.0/1.0/077 | OS | Physician | 06/17/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15020740&pc=QCB | 2023826-2022-02287 | 2022-02-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6; -9.50/1.00/77 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS; THAT WAS INTENDED FOR THE PATIENT'S RIGHT EYE (OD); HAD TORE DURING LOADING INTO THE CARTRIDGE. THIS OCCURRED ON (B)(6) 2022. ON THE SAME DATE A REPLACEMENT LENS OF THE SAME MODEL/LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.50/1.00/77 | OD | Physician | 07/06/2022 | 2022 | 45 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15021588&pc=QCB | 2023826-2022-02035 | 2022-10-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 OF -10.5/4.5/013 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. REPORTEDLY THE LENS WAS TOO BIG. HIS RESOLVED THE PROBLEM. THE LENS REMAINS IMPLANTED. AN EXPLANT HAD BEEN PLANNED. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 OF -10.5/4.5/013 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 AND REPLACED ON THE SAME DAY IN A SEPARATE SURGERY WITH A SHORTER LENGTH LENS OF A DIFFERENT MODEL DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE IS DUE TO DEVICE - LENS WAS TOO LARGE. H1: SERIOUS INJURY. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN DRY IN A VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC. THE LENS RETURNED WAS NOT THE SIZE/MODEL STATED IN THE CLAIM DUE TO DISCREPANT INFORMATION SHARED BY THE REPORTER. THE DIMENSIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. HOWEVER, AFTER FURTHER CLARIFICATION ABOUT THE RETURNED LENS THE DIMENSIONAL INSPECTION WOULD PASS. CLAIM# (B)(4). Manufacturer Narrative: AGE OR DATE OF BIRTH, SEX, WEIGHT, ETHNICITY, RACE:UNK. DATE OF EVENT: UNK. IMPLANT DATE: UNK. CLAIM# (B)(4). | -10.5/4.5/013 | OD | Physician | 06/16/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15021590&pc=QCB | 2023826-2022-02036 | 2022-10-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 OF -13.00/4.5/001 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S EYE. REPORTEDLY THE LENS WAS TOO BIG. HIS RESOLVED THE PROBLEM. THE LENS REMAINS IMPLANTED. AN EXPLANT HAD BEEN PLANNED. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 OF -13.00/4.5/001 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 AND REPLACED ON THE SAME DAY IN A SEPARATE SURGERY WITH A SHORTER LENGTH LENS OF A DIFFERENT MODEL DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE IS DUE TO DEVICE - LENS WAS TOO LARGE. CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. THE LENS RETURNED WAS NOT THE SIZE/MODEL STATED IN THE CLAIM DUE TO DISCREPANT INFORMATION SHARED BY THE REPORTER. THE DIMENSIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. HOWEVER, AFTER FURTHER CLARIFICATION ABOUT THE RETURNED LENS THE DIMENSIONAL INSPECTION WOULD PASS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -13.00/4.5/001 | OS | Physician | 06/16/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15022647&pc=QCB | 2023826-2022-02268 |  | Injury | STAAR SURGICAL COMPANY | 2022-07-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.1MM TMICL 12.1 IMPLANTABLE COLLAMER LENS OF -12.5/2.0/085 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS ROTATION WAS OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE REPORTED AS DEVICE. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: B5 - LENS WAS EXPLANTED DUE TO LENS ROTATION WITH INSUFFICIENT VAULT. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -12.5/2.0/085 | OD | Other Health Care Professional | 06/15/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15029508&pc=QCB | 2023826-2022-02292 | 2019-09-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -5.5/0.5/081 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. ON (B)(6) 2019, THE LENS WAS EXCHANGED FOR A LONGER LENGTH, SPHERICAL LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS PATIENT RELATED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. CLAIM# (B)(4). | -5.5/0.5/081 | OD | Physician | 06/16/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15029514&pc=QCB | 2023826-2022-02291 | 2022-11-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -8.0/1.5/163 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. CLAIM# (B)(4). | -8.0/1.5/163 | OD | Physician | 06/16/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15029515&pc=QCB | 2023826-2022-02293 | 2021-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -5.5/1.1/117 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: CORRECTED TO: "THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -15.5/1.0/117 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN." CLAIM #(B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK. (B)(4). CLAIM# (B)(4). | -15.5/1.0/117 | OD | Physician | 06/16/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15032617&pc=QCB | 2023826-2022-02180 | 2022-04-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -15.5/2.5/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5- "THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -15.5/2.5/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS IMPLANTED VERTICALLY AND THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN." CLAIM# (B)(4). | -15.5/2.5/091 | OD | Physician | 06/20/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15032775&pc=QCB | 2023826-2022-02181 | 2021-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -7.5/1.5/172 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -7.5/1.5/172 | OD | Physician | 06/20/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15032889&pc=QCB | 2023826-2022-01441 | 2022-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -14.0/5.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PMA/510K: THIS PRODUCT IS NOT MARKETED IN THE U.S. CLAIM# (B)(4). | -14.0/5.0/090 | OS | Physician | 04/12/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15033416&pc=QCB | 2023826-2022-02297 | 2022-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.5/2.0/098 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -11.5/2.0/098 | OS | Physician | 07/11/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15033419&pc=QCB | 2023826-2022-02016 | 2021-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -17.5/1.0/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT HAD EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -17.5/1.0/89 | OD | Physician | 06/16/2022 | 2022 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15042130&pc=QCB | 2023826-2022-02163 | 2021-12-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.5/1.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. ON (B)(6) 2022 A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -12.5/1.5/088 | OS | Physician | 06/29/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15042168&pc=QCB | 2023826-2022-02154 | 2022-05-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.5/2.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.5/2.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: (B)(4). | -11.5/2.0/086 | OD | Physician | 06/27/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15042170&pc=QCB | 2023826-2022-02158 | 2022-05-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.0/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.0/1.5/091 | OS | Physician | 06/27/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15042171&pc=QCB | 2023826-2022-02166 | 2022-03-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -10.00/2.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). | -10.00/2.0/087 | OD | Physician | 06/30/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15042183&pc=QCB | 2023826-2022-02153 | 2022-05-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.5/2.0/177 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - "REPORTEDLY, THE VERTICAL POSITION OF 13.2 WAS PLACED FIRST, AND THE HORIZONTAL POSITION OF 12.6 WAS ADJUSTED AT THE HEIGHT OF THE ARCH." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -11.5/2.0/177 | OD | Physician | 06/27/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15042188&pc=QCB | 2023826-2022-02156 | 2022-06-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.5/1.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.5/1.5/094 | OD | Physician | 06/27/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15042191&pc=QCB | 2023826-2022-02160 | 2022-06-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7, -10.0/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.0/2.0/091 | OD | Physician | 06/28/2022 | 2022 | 16 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15042193&pc=QCB | 2023826-2022-02168 | 2022-05-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6, -14.00/3.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE. CLAIM# (B)(4). | -14.00/3.5/092 | OD | Physician | 06/30/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15042233&pc=QCB | 2023826-2022-02140 | 2022-06-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -6.50/1.5/179 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -6.50/1.5/179 | OD | Physician | 06/23/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15042291&pc=QCB | 2023826-2022-02155 | 2022-05-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.5/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -8.5/1.0/089 | OS | Physician | 06/27/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15042347&pc=QCB | 2023826-2022-02164 | 2022-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/2.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.0/2.0/092 | OD | Physician | 06/29/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15046795&pc=QCB | 2023826-2022-02212 | 2022-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+4.0/087 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: (B)(4). | -12.0/+4.0/087 | OS | Physician | 07/06/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15046937&pc=QCB | 2023826-2022-02214 | 2022-05-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/087 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: (B)(4). | -10.0/+1.5/087 | OS | Physician | 07/06/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15047022&pc=QCB | 2023826-2022-02215 | 2022-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+1.0/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -7.5/+1.0/094 | OS | Physician | 07/05/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15047034&pc=QCB | 2023826-2022-02213 | 2022-05-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+3.0/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: (B)(4). | -12.0/+3.0/094 | OD | Physician | 07/06/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15047197&pc=QCB | 2023826-2022-02211 | 2021-12-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.0/087 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: (B)(4). | -7.0/+1.0/087 | OS | Physician | 07/06/2022 | 2022 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15047778&pc=QCB | 2023826-2022-02218 | 2022-05-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.50/+2.00/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: (B)(4). | -12.50/+2.00/091 | OD | Physician | 07/07/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15047987&pc=QCB | 2023826-2022-02219 | 2022-05-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.50/+1.0/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: (B)(4). | -11.50/+1.0/091 | OS | Physician | 07/07/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15054822&pc=QCB | 2023826-2022-02183 | 2022-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.0/1.5/098 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. IN A SUBSEQUENT SURGERY LATER THAT DAY, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. OFF LABEL USE: ACD <3.0. CLAIM#: (B)(4). | -11.0/1.5/098 | OD | Physician | 06/20/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15061057&pc=QCB | 2023826-2022-02330 | 2022-07-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -15.5/2.0/67 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS HAD TORN DURING INJECTION/DELIVERY INTO THE PATIENT'S EYE. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A BACK UP LENS OF LONGER LENGTH AND THE PROBLEM WAS RESOLVED. THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | -15.5/2.0/67 | OD | Physician | 07/12/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15061076&pc=QCB | 2023826-2022-02327 | 2022-06-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -10.5/0.5/169 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME MODEL BUT DIFFERENT LENGTH LENS; WITH A SMALL CYLINDER CHANGE AND NOTED "THE CYLINDER WAS ADJUSTED DURING THE SECOND SURGERY WHICH DID NOT AFFECT THE PATIENT'S VISUAL ACUITY AND VISUAL QUALITY". THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.5/0.5/169 | OS | Physician | 06/23/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15061535&pc=QCB | 2023826-2022-02185 | 2022-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.5/1.5/085 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: A2- (B)(6) 2024 SHOULD BE CORRECTED TO (B)(6) 1991. CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNK. CLAIM# (B)(4). | -9.5/1.5/085 | OS | Physician | 06/20/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15065196&pc=QCB | 2023826-2022-02186 | 2021-11-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -14.5/1.5/103 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. REPORTEDLY, "1. THE NEW LENS WAS IMPLANTED VERTICALLY. 2. THE REASONS FOR DIFFERENCE OF REFRACTIVE POWER BETWEEN TWO LENS ARE AS FOLLOWS: 1. SEQ WERE ALMOST THE SAME; 2. AT FIRST, THE PATIENT COULDN'T WAIT FOR THE SURGERY WHILE A TICL WITH CYLINDER OF 1.50D WAS IN STOCK. THE CYLINDER POWER OF THE NEW LENS WAS THE SAME AS MANIFEST REFRACTIVE RESULT." Manufacturer Narrative: WEIGHT, RACE, ETHNICITY: UNK. CLAIM#: (B)(4). | -14.5/1.5/103 | OS | Physician | 06/21/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15067932&pc=QCB | 2023826-2022-02277 | 2022-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF -17.50/3.0/098 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXPLANTED. A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED THE NEXT DAY AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - CORRECTED TO: "ON (B)(6) 2022 THE LENS WAS EXCHANGED." IN INITIAL MDR . D6B - EXPLANT DATE (B)(6) 2022 SHOULD BE CORRECTED TO INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: B5 - CORRECTED TO: "THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -17.50/3.0/098 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022." IN THE INITIAL MDR. D4 - CORRECTED TO: "MODEL # VTICMO12.6" IN THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | -17.50/3.0/098 | OD | Physician | 06/16/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15069023&pc=QCB | 2023826-2022-02298 | 2022-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+5.0/098 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -13.0/+5.0/098 | OS | Physician | 07/10/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15069026&pc=QCB | 2023826-2022-02302 | 2021-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+2.0/105 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2021 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -13.0/+2.0/105 | OS | Physician | 07/11/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15070101&pc=QCB | 2023826-2022-02278 |  | Injury | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -15.0/2.0/088 (SPHERE/CYLINDER/AXIS) LENS TORE ON INSERTION BEING IMPLANTED INTO THE PATIENT'S EYE. THE LENS WAS REMOVED AND THE BACK UP LENS WAS USED INSTEAD WHICH RESULTED IN A SUCCESSFUL SURGERY. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -15.0/2.0/088 | NA | Other Health Care Professional | 06/16/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15070308&pc=QCB | 2023826-2022-02313 | 2021-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -05.00/+1.5/180 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS EXPLANTED. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -05.00/+1.5/180 | OD | Physician | 06/20/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15070353&pc=QCB | 2023826-2022-02314 | 2022-04-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -08.50/+1.0/171 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN, BENT, AND DEFORMED TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -08.50/+1.0/171 | OS | Physician | 06/21/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15070445&pc=QCB | 2023826-2022-02323 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.5/1.5/072 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 AS A REPLACEMENT LENS TO CORRECT LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM, THE REPORTER STATED "THE VAULT IS TOO LOW". PER UPDATED INFORMATION THE LENS REMAINS IMPLANTED, THE REPORTER STATED " PATIENT IS CURRENTLY DOING WELL. THE SURGEON DECIDED TO CONTINUE FOLLOW-UP." IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR # 2023826-2022-02321 FOR INITIAL LENS IMPLANT. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -12.5/1.5/072 | OS | Physician | 06/20/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15070716&pc=QCB | 2023826-2022-02321 | 2022-03-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -12.5/1.5/92 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. IT WAS REPORTED THAT THIS EXCHANGE DID NOT RESOLVE THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -12.5/1.5/92 | OS | Physician | 06/20/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15070982&pc=QCB | 2023826-2022-02324 | 2022-04-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -12.0/4.0/165 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON THE SAME DATE THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS THAT WAS IMPLANTED VERTICALLY AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -12.0/4.0/165 | OD | Physician | 06/21/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15071069&pc=QCB | 2023826-2022-02193 | 2021-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -17.0/2.5/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -17.0/2.5/91 | OD | Physician | 12/02/2021 | 2021 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15071070&pc=QCB | 2023826-2022-02194 | 2021-05-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF - 13.5/2.5/92 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED, AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION:UNK. CLAIM# (B)(4). Manufacturer Narrative: B2:MALFUNCTION. B5:AL THE REPORTER INDICATED THAT A 13.2MM VTICMO_13.2; -13.5/2.5/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED WITH A SHORTER LENGTH LENS INTRAOPERATIVELY ON (B)(6) 2021 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. NO PATIENT INJURY REPORTED. CAUSE REPORTED AS UNKNOWN. H1:MALFUNCTION. H6:IMPACT CODE 2199. CLAIM # (B)(4). | - 13.5/2.5/92 | OD | Physician | 09/06/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15071071&pc=QCB | 2023826-2022-02331 | 2021-05-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF - 15.0/4.5/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | - 15.0/4.5/93 | OS | Physician | 06/21/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15071072&pc=QCB | 2023826-2022-02191 | 2022-02-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5/2.0/14 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.5/2.0/14 | OD | Physician | 06/21/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15071073&pc=QCB | 2023826-2022-02332 | 2022-01-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF - 8.0/1.5/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | - 8.0/1.5/086 | OD | Physician | 06/21/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15071074&pc=QCB | 2023826-2022-02190 | 2021-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.5/1.0/110 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.5/1.0/110 | OD | Physician | 06/21/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15071075&pc=QCB | 2023826-2022-02333 | 2022-01-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF - 8.5/2.0/94 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | - 8.5/2.0/94 | OS | Physician | 06/21/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15071082&pc=QCB | 2023826-2022-02192 | 2021-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -15.0/1.5/175 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS A PATIENT RELATED FACTOR. REPORTEDLY THE STATUS OF THE EYE'S VISION IS 0.5. Manufacturer Narrative: (B)(4). | -15.0/1.5/175 | OS | Physician | 06/21/2022 | 2022 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15071172&pc=QCB | 2023826-2022-02334 | 2022-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/96 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.0/1.0/96 | OD | Physician | 06/21/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15077705&pc=QCB | 2023826-2022-02317 | 2022-05-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -03.50/+1.5/171 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. POST-OP THERE WAS MILD WOUND EDEMA, WHICH HAD CLEARED UP. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -03.50/+1.5/171 | OS | Physician | 06/22/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15077941&pc=QCB | 2023826-2022-02339 | 2022-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.0/2.0/073 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - RACE: UNK. CLAIM# (B)(4). | -10.0/2.0/073 | OD | Physician | 06/24/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15077944&pc=QCB | 2023826-2022-02337 | 2022-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.0/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - RACE: UNK. CLAIM# (B)(4). | -9.0/2.0/089 | OD | Physician | 06/23/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15077945&pc=QCB | 2023826-2022-02336 | 2022-05-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -7.5/2.0/085 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - RACE: UNK. CLAIM# (B)(4). | -7.5/2.0/085 | OD | Physician | 06/23/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15077966&pc=QCB | 2023826-2022-02338 | 2022-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.0/2.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO:UNK. CLAIM# (B)(4). | -9.0/2.0/092 | OD | Physician | 06/23/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15084692&pc=QCB | 2023826-2022-02318 | 2022-06-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -10.00/+5.5/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2022 BUT THE PROBLEM WAS NOT RESOLVED. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE SURGEON REPOSITIONED (ROTATED) THE LENS AGAIN AND PATIENT WAS SEEING 6/6 ON DAY 1. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.00/+5.5/087 | OS | Physician | 06/23/2022 | 2022 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15084724&pc=QCB | 2023826-2022-02319 | 2022-06-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.00/+4.5/062, TORE DURING LOADING. THE LENS TORE WHEN THE SURGEON CONNECTED THE CARTRIDGE TO THE INJECTOR. THE LENS WAS NOT IMPLANTED BECAUSE IT WAS NOTED UNDER THE MICROSCOPE THE LENS WAS TORN. THERE WAS NO PATIENT CONTACT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "PATIENT VERY SATISFIED- NO SYMPTOMS". CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4) . | -11.00/+4.5/062 | NA | Physician | 06/23/2022 | 2022 | 46 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15089498&pc=QCB | 2023826-2022-02390 | 2022-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.5/1.0/003(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. REPORTEDLY, "VERTICALLY POSITION IMPLANTATION." Manufacturer Narrative: WEIGHT: UNKNOWN, ETHINICITY:UNKNOWN, RACE:UNKNOWN. CLAIM#(B)(4). | -12.5/1.0/003 | OS | Physician | 06/24/2022 | 2022 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15089629&pc=QCB | 2023826-2022-02388 | 2022-01-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -7.5/1.0/68(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -7.5/1.0/68 | OS | Physician | 06/24/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15089634&pc=QCB | 2023826-2022-02389 | 2022-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -10.0/2.5/092(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6)2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS OTHER. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHINICITY: UNKNOWN, RACE:UNKNOWN. OFF LABEL USE- ACD<3.0 CLAIM# (B)(4). | -10.0/2.5/092 | OD | Physician | 06/24/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15090110&pc=QCB | 2023826-2022-02394 | 2022-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -16.0/3.5/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6)2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHINICITY: UNKNOWN, RACE:UNKNOWN. CLAIM# (B)(4). | -16.0/3.5/091 | OD | Physician | 06/24/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15099144&pc=QCB | 2023826-2022-02431 | 2022-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -10.0/1.0/180 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH/MODEL LENS BUT IMPLANTED HORIZONTALLY; IT WAS NOTED" THE CRYSTAL CHANGES FROM VERTICAL TO HORIZONAL". THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -10.0/1.0/180 | OD | Physician | 06/27/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15099198&pc=QCB | 2023826-2022-02432 | 2022-03-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -17.0/2.0/107 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH/MODEL LENS BUT IMPLANTED VERTICALLY. IT WAS NOTED "THE CRYSTAL CHANGES FROM HORIZONTAL TO VERTICAL". THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -17.0/2.0/107 | OD | Physician | 06/27/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15099763&pc=QCB | 2023826-2022-02400 | 2022-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.5/3.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION: UNK. CLAIM# 725117. | -11.5/3.0/091 | OS | Physician | 06/27/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15100030&pc=QCB | 2023826-2022-02434 | 2022-07-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.5/1.0/074 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED. IT WAS NOTED "AFTER LENS REMOVAL; PATIENT CONSIDERS NOT IMPLANTING". THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | -9.5/1.0/074 | OS | Physician | 07/22/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15106415&pc=QCB | 2023826-2022-02243 | 2022-01-22 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6, -4.00/1.50/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. (B)(4). | -4.00/1.50/080 | OD | Physician | 07/06/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15106478&pc=QCB | 2023826-2022-02260 | 2022-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED VERTICALLY A 12.6MM VTICMO12.6, -06.0/6.0/166 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -06.0/6.0/166 | OD | Physician | 07/08/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15106624&pc=QCB | 2023826-2022-02251 | 2022-03-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.5/1.5/115 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(4) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(4) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -10.5/1.5/115 | OD | Physician | 07/07/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15106650&pc=QCB | 2023826-2022-02239 | 2022-05-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.0/3.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.0/3.0/095 | OD | Physician | 07/06/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15106657&pc=QCB | 2023826-2022-02237 | 2022-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.0/3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.0/3.0/093 | OS | Physician | 07/04/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15106672&pc=QCB | 2023826-2022-02248 | 2022-02-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -6.00/2.0/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -6.00/2.0/100 | OS | Physician | 07/06/2022 | 2022 | 50 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15106682&pc=QCB | 2023826-2022-02252 | 2022-03-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -09.0/1.5/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -09.0/1.5/079 | OS | Physician | 07/07/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15106688&pc=QCB | 2023826-2022-02256 | 2022-03-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED VERTICALLY A 12.6MM VTICMO12.6, -13.0/1.5/050 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. (B)(4). | -13.0/1.5/050 | OD | Physician | 07/08/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15106697&pc=QCB | 2023826-2022-02263 | 2022-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/1.5/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. (B)(4). | -11.0/1.5/110 | OS | Physician | 07/08/2022 | 2022 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15106705&pc=QCB | 2023826-2022-02259 | 2022-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -15.00/1.0/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. CLAIM # (B)(4). | -15.00/1.0/102 | OD | Physician | 07/08/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15106710&pc=QCB | 2023826-2022-02258 | 2022-03-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED VERTICALLY A 13.2MM VTICMO13.2, -15.5/3.0/173 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -15.5/3.0/173 | OD | Physician | 07/08/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15106811&pc=QCB | 2023826-2022-02253 | 2022-03-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.0/2.0/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | -12.0/2.0/103 | OS | Physician | 07/08/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15106816&pc=QCB | 2023826-2022-02255 | 2022-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -06.5/1.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -06.5/1.0/086 | OD | Physician | 07/08/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15108035&pc=QCB | 2023826-2022-02407 | 2022-03-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -8.0/1.5/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.0/1.5/094 | OS | Physician | 06/29/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15108058&pc=QCB | 2023826-2022-02406 | 2022-01-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.5/1.5/078 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -9.5/1.5/078 | OD | Physician | 06/29/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15108123&pc=QCB | 2023826-2022-02402 | 2022-03-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -4.5/1.0/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. REPORTEDLY, "THE SECOND SLICE IS VERTICAL." Manufacturer Narrative: (B)(4). | -4.5/1.0/91 | OD | Physician | 06/28/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15108258&pc=QCB | 2023826-2022-02420 | 2022-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -09.50/+1.5/006 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD). THE LENS WAS REPORTED AS LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS WITH DIOPTER -9.5/+1.5/06 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) IN (B)(6) 2019. THE PATIENT EXPERIENCED A LOW VAULT. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR SAME MODEL/LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. REPORTEDLY "THE EXCHANGE DID NOT RESULT ANY SIGNIFICANT IMPROVEMENT. THE PATIENT IS CLOSELY BEING MONITORED. FURTHER EXPLANTATIONS ARE NOT PLANNED." THE REPORTER INDICATED THE CAUSE OF THE EVENT IS UNKNOWN." CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT CLAIM #: (B)(4). | -09.50/+1.5/006 | OD | Physician | 07/01/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15108263&pc=QCB | 2023826-2022-02423 | 2022-04-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -06.50/+1.5/180 (SPHERE/CYLINDER/AXIS), TORE DURING LOADING AND THERE WAS NO PATIENT CONTACT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT INFO: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) Manufacturer Narrative: CORRECTED DATA: B5- " THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/1.5/180 (SPHERE/CYLINDER/AXIS) WHEN THE LENS TORE/BROKE BEFORE/DURING LOADING. THE LENS WAS NOT IMPLANTED. REPORTEDLY "THE LENS TORE DURING LOADING" AND THE DAMAGE WAS NOTED "WHEN LOADING THE LENS INTO". THE REPORTER INDICATED THE CAUSE OF THE EVENT IS UNKNOWN. ON (B)(6) 2022 A REPLACEMENT LENS OF THE SAME MODEL/LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. REPORTEDLY THE STATUS OF THE EYE IS " GOOD". " CLAIM# (B)(4). | -06.50/+1.5/180 | NA | Physician | 06/30/2022 | 2022 | 45 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15108319&pc=QCB | 2023826-2022-02418 |  | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -09.50/+5.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2022. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: "THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.50/+5.00/93 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. REPORTEDLY "DR. (B)(6) TRIED TO REPOSITION THE LENS AT THE END OF ON (B)(6). THE RESULTS ARE SIMILAR TO THE EXPECTED REFRACTION PROVIDED BY THE STAAR MEDICAL AFFAIR TEAM. +0.75-1.5@25 MEASURED BY ARK; UCVA : 0.6+. THE UCVA WAS SLIGHTLY IMPROVED FROM 0.5 TO 0.6+." THE LENS REMAINS IMPLANTED." CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -09.50/+5.0/093 | OS | Physician | 06/28/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15108324&pc=QCB | 2023826-2022-02416 | 2022-03-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -16.50/+2.0/097 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: B5: REPORTEDLY, PROBLEM WAS NOT RESOLVED. ICL WAS EXPLANTED (B)(6) 2023 DUE TO ROTATION AND IOP SPIKES. CLAIM# (B)(4). | -16.50/+2.0/097 | OD | Other Health Care Professional | 06/28/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15108409&pc=QCB | 2023826-2022-02428 | 2022-06-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). | -11.50/+2.0/091 | OD | Physician | 06/29/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15108454&pc=QCB | 2023826-2022-02417 | 2022-03-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS, -16.50/+1.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5- "THE LENS WAS REPLACED WITH A SAME MODEL/ LENGTH LENS AND THE PROBLEM WAS RESOLVED" SHOULD BE CORRECTED TO "THE LENS WAS REPLACED WITH A DIFFERENT MODEL/ SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED." CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.50/1.5/092 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A DIFFERENT MODEL/ SAME LENGTH LENS AND THE PROBLEM WAS NOT RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. " CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | -16.50/+1.5/092 | OS | Other Health Care Professional | 06/28/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15108472&pc=QCB | 2023826-2022-02419 | 2022-06-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -02.50/+1.0/179 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON HAD CHOSEN THE WRONG LENS POWER AND THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4). Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2022 FOR ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. CLAIM#: (B)(4). | -02.50/+1.0/179 | OS | Physician | 06/27/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15108534&pc=QCB | 2023826-2022-02422 | 2022-05-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.00/+3.5/090 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: B5: THE PROBLEM WAS RESOLVED WITH THE REMOVAL OF THE INITIAL LENS. THE SURGEON DOES NOT PLAN TO IMPLANT A REPLACEMENT LENS. THE PATIENTS POST-OP CONDITION IS NORMAL. CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC STRETCHED OUT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). | -12.00/+3.5/090 | OD | Physician | 06/28/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15108826&pc=QCB | 2023826-2022-02456 | 2022-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.5/2.5/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -13.5/2.5/088 | OD | Physician | 07/25/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15108892&pc=QCB | 2023826-2022-02489 | 2022-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -9.5/1.5/082 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (ACD < 3.0MM) CLAIM#: (B)(4). | -9.5/1.5/082 | OD | Physician | 07/25/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15108928&pc=QCB | 2023826-2022-02487 | 2022-07-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.5/1.0/174 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | -10.5/1.0/174 | OD | Physician | 07/25/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15108957&pc=QCB | 2023826-2022-02488 | 2022-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.5/4.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | -13.5/4.5/089 | OD | Physician | 07/25/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15108959&pc=QCB | 2023826-2022-02486 | 2021-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -15.0/4.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -15.0/4.0/090 | OD | Physician | 07/25/2022 | 2022 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15109075&pc=QCB | 2023826-2022-02440 | 2022-06-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 WHILE OPENING THE LENS VIAL FOR A 12.1MM VTICMO12.1, -9.50/1.5/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, THE LENS VIAL WAS BROKEN DURING THE UNCAPPING PROCESS. THE LENS WAS IMPLANTED. NO PATIENT OR SURGEON INJURY HAVE BEEN REPORTED. THE PROBLEM WAS REPORTED AS RESOLVED. Manufacturer Narrative: RETURNED PRODUCT EVALUATION FOUND THE LENS VIAL BROKEN AT THE NECK. NO LENS WAS RETURNED. LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. BASED ON THE RESULTS OF THE INVESTIGATION, THE PROBABLE CAUSE OF THE FAILURE IS LIKELY DUE TO CHANGE IN THE PACKAGING CONFIGURATION OF THE LENS VIALS. AN ACTION PLAN, HEALTH HAZARD EVALUATION (GP-1416), ON 2ML VIAL BREAKAGE WAS INITIATED. THE HHE FOUND NO SEVERE THREAT TO PUBLIC HEALTH AND THE OBSERVED FREQUENCY OF THE REPORTED EVENT OF THE VIAL BREAKAGE IS OCCURRING WITHIN THE EXPECTED RANGE WITH THE RISK ALSO BEING IN AN ACCEPTABLE RANGE. NIPRO HAS AGREED TO SEND FUTURE VIAL SHIPMENTS TO STAAR USING THE PRIOR PACKAGING CONFIGURATION TO PREVENT FURTHER VIALS FROM BRAKING DURING USAGE OR SHIPPING. THUS, NO FURTHER ACTION IS REQUIRED AT THIS TIME. (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: TYPE OF INVESTIGATION CODE; 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, THE PROBABLE CAUSE OF THIS EVENT IS RELATED TO A CHANGE IN THE PACKAGING CONFIGURATION OF THE 2ML VIALS. THE HEALTH HAZARD EVALUATION (GP-1416) FOUND NO SEVERE THREAT TO PUBLIC HEALTH. AN ACTION PLAN, CAPA M22-04, HAS BEEN INITIATED TO ASSESS AND IMPLEMENT CORRECTIVE AND PREVENTATIVE ACTIONS FOR THE FAILURE IDENTIFIED IN THIS COMPLAINT. CLAIM# (B)(4). | -9.50/1.5/111 | NA | Physician | 06/27/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15109129&pc=QCB | 2023826-2022-02437 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-07-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Use of Device Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 12.1MM, VTICMO12.1, -12.5/3.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. THERE WAS NO PATIENT INJURY. ON (B)(6) 2022 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT ID: (B)(6). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4) | -12.5/3.5/087 | NA | Physician | 06/27/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15109175&pc=QCB | 2023826-2022-02439 | 2022-05-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 13.2MM, VTICMO13.2, -18.00/1.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS GOT STUCK IN THE CARTRIDGE OR INJECTION SYSTEM AND TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY. ON (B)(6) 2022, A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: CORRECTED DATA: B5 - IT WAS ADDITIONALLY REPORTED THAT ANTERIOR CHAMBER IRRIGATION/EVACUATION OF VISCO FLUIDS WAS PERFORMED. - THIS SHOULD HAVE BEEN INCLUDED. CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.00/1.0/094 | NA | Physician | 06/27/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15117065&pc=QCB | 2023826-2022-02375 | 2022-02-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+2.5/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -8.0/+2.5/091 | OD | Physician | 07/13/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15117131&pc=QCB | 2023826-2022-02376 | 2022-02-22 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/112 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) CLAIM# (B)(4). | -9.5/+1.0/112 | OS | Physician | 07/13/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15117148&pc=QCB | 2023826-2022-02374 | 2022-06-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+2.5/096 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.0/+2.5/096 | OD | Physician | 07/12/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15117182&pc=QCB | 2023826-2022-02379 | 2022-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.0/+2.5/078 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -17.0/+2.5/078 | OD | Physician | 07/15/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15117722&pc=QCB | 2023826-2022-02421 | 2022-07-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -09.50/+0.5/011 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE LENS WAS REPORTED AS LOW VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "THE LENS REMAINS IMPLANTED AND THE PROBLEM WAS RESOLVED. REPORTEDLY "THE PATIENT IS CLOSELY MONITORED. AN EXCHANGE MAY BE PLANNED IN 2023." CLAIM# (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | -09.50/+0.5/011 | OS | Physician | 07/01/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15117749&pc=QCB | 2023826-2022-02424 | 2022-04-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -05.50/+2.0/152 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE PROBLEM WAS RESOLVED WITH THE IMPLANT OF THE REPLACEMENT LENS. CLAIM # (B)(4). Manufacturer Narrative: (B)(4).CLAIM # (B)(4). | -05.50/+2.0/152 | OD | Physician | 07/01/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15118196&pc=QCB | 2023826-2022-02410 | 2022-03-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -2.5/0.5/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -2.5/0.5/089 | OS | Physician | 07/01/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15118197&pc=QCB | 2023826-2022-02409 | 2021-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.0/1.5/106 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4).CLAIM# (B)(4). | -10.0/1.5/106 | OS | Physician | 07/01/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15118198&pc=QCB | 2023826-2022-02413 | 2022-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.0/2.5/114 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -12.0/2.5/114 | OD | Physician | 07/01/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15118199&pc=QCB | 2023826-2022-02412 | 2022-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -18.0/1.5/074 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -18.0/1.5/074 | OS | Physician | 07/01/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15118268&pc=QCB | 2023826-2022-02284 | 2022-04-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -11.50/1.5/080 (SPHERE/CYLINDER/AXIS) TORE/BROKE DURING LOADING THE LENS INTO THE INJECTOR. THE LENS WAS NOT IMPLANTED. A REPLACEMENT LENS OF SPHERICAL MODEL WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5 - -11.50/1.5/080 SHOULD BE CORRECTED TO -11.50/1.5/079. CLAIM#:(B)(4). Manufacturer Narrative: (B)(4). | -11.50/1.5/080 | NA | Physician | 06/16/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15118307&pc=QCB | 2023826-2022-02342 | 2022-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -12.0/3.5/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. LENS ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2022 BUT THE PROBLEM WAS NOT RESOLVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -12.0/3.5/095 | OD | Physician | 06/20/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15118446&pc=QCB | 2023826-2022-02343 | 2022-06-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -10.0/4.0/105 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD). ON (B)(6) 2022 THE LENS WAS IMPLANTED; REMOVED AND REPLACED INTRA-OPERATIVELY FOR THE BACKUP LENS DUE TO THE TRAILING HAPTIC OF THE ICL BROKE. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS DEVICE. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | -10.0/4.0/105 | OD | Physician | 06/20/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15118494&pc=QCB | 2023826-2022-02344 | 2021-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MMVTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -15.0/2.0/56 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE(OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS REPOSITIONED BUT DID NOT RESOLVE. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS ICL SIZE. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | -15.0/2.0/56 | OS | Physician | 06/21/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15126961&pc=QCB | 2023826-2022-02385 | 2022-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+2.5/083 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.5/+2.5/083 | OD | Physician | 07/17/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15126963&pc=QCB | 2023826-2022-02386 | 2022-03-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.5/150 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE LENS WAS IMPLANTED VERTICALLY. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+1.5/150 | OD | Physician | 07/17/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15126964&pc=QCB | 2023826-2022-02414 | 2022-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+2.0/102 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE LENS WAS IMPLANTED VERTICALLY. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.5/+2.0/102 | OS | Physician | 07/17/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15126965&pc=QCB | 2023826-2022-02493 | 2022-02-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+2.5/095 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -13.5/+2.5/095 | OS | Physician | 07/19/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15126970&pc=QCB | 2023826-2022-02492 | 2022-02-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+2.5/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -13.5/+2.5/095 | OD | Physician | 07/19/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15126973&pc=QCB | 2023826-2022-02499 | 2022-05-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+2.0/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -12.5/+2.0/093 | OS | Physician | 07/22/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15129671&pc=QCB | 2023826-2022-02447 | 2022-06-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF -10.0/+1.0/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED UPSIDE DOWN INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 . THE LENS WAS IMPLANTED, REMOVED AND REPLACED INTRAOPERATIVELY WITH NO PATIENT INJURY. PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.0/+1.0/095 | OD | Physician | 06/30/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15131485&pc=QCB | 2023826-2022-02426 | 2022-01-27 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.50/+1.0/092 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS WAS INSERTED UPSIDE DOWN IN THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS TORN DURING REMOVAL PROCESS. THE LENS WAS REPLACED WITH A DIFFERENT MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -16.50/+1.0/092 | OS | Physician | 07/01/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15131519&pc=QCB | 2023826-2022-02429 | 2022-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -07.00/+1.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). Manufacturer Narrative: A2: DATE OF BIRTH: (B)(6) 2004 CLAIM # (B)(4). | -07.00/+1.0/093 | OD | Physician | 07/01/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15132750&pc=QCB | 2023826-2022-02452 | 2022-04-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 12.1MM, VTICM5_12.1, -12.5/1.0/134 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT WITH NO PATIENT INJURY. ON (B)(6) 2022 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: B5- "THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/1.0/134 (SPHERE/CYLINDER/AXIS) INTO THE RIGHT EYE (OD) WHEN THE LENS TORE/BROKE BEFORE/DURING LOADING ON (B)(6) 2022. THE LENS WAS NOT IMPLANTED. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. ON (B)(6) 2022 A REPLACEMENT LENS OF THE SAME MODEL/LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.5/1.0/134 | OD | Physician | 07/01/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15134788&pc=QCB | 2023826-2022-02515 | 2022-07-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1; -6.50/+2.5/098 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS; HAD BECAME CRIMPED AND WOULD NOT OPEN. IT WAS ALSO REPORTED THAT THE HAPTICS CAME OFF; THE LENS WAS NOT IMPLANTED. A BACK UP LENS WAS USED. Manufacturer Narrative: TYPE OF INVESTIGATION: (B)(4) - NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -6.50/+2.5/098 | NA | Other Health Care Professional | 07/26/2022 | 2022 | 27
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15134841&pc=QCB | 2023826-2022-02506 | 2022-06-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-07-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6; -13.5/2.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING WITH ROTATION. ON (B)(6) 2022 THE LENS WAS EXPLANTED; LATER ON THIS SAME DATE A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -13.5/2.5/091 | OD | Physician | 07/28/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15136762&pc=QCB | 2023826-2022-02346 | 2022-05-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-07-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -11.5/1.0/112 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). ON (B)(6) 2022 THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY DUE TO LENS ROTATION. A REPLACEMENT LENS OF LONGER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. REPORTER STATES CAUSE AS REPEATED LENS ROTATION SEEN PEROPERATIVELY." Manufacturer Narrative: ADDITIONAL INFORMATION: H3: PRODUCT EVALUATION: LENS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: UNK. (B)(4). CLAIM # (B)(4). | -11.5/1.0/112 | OS | Physician | 06/22/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15140023&pc=QCB | 2023826-2022-02357 | 2022-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7, -09.0/5.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -09.0/5.0/084 | OS | Physician | 07/11/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15140042&pc=QCB | 2023826-2022-02347 | 2022-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.0/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -09.0/1.0/087 | OD | Physician | 07/11/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15140048&pc=QCB | 2023826-2022-02356 | 2022-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -9.50/4.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: INFORMATION CORRECTION: A2 - REPORTED AS (B)(6) 2022 - CORRECT TO (B)(6) 2002. CLAIM# (B)(4). | -9.50/4.5/090 | OD | Physician | 07/11/2022 | 2022 | 2 MO | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15140053&pc=QCB | 2023826-2022-02362 | 2021-06-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.5/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -8.5/1.0/091 | OS | Physician | 07/13/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15140056&pc=QCB | 2023826-2022-02368 | 2022-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -10.5/1.5/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -10.5/1.5/111 | OS | Physician | 07/15/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15140070&pc=QCB | 2023826-2022-02351 | 2022-05-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.5/1.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -13.5/1.5/105 | OD | Physician | 07/11/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15140077&pc=QCB | 2023826-2022-02363 | 2021-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.0/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO EXCESSIVE VAULT. ON (B)(6) 2021 A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.0/2.0/091 | OS | Physician | 07/13/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15140082&pc=QCB | 2023826-2022-02353 | 2021-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.5/2.0/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.5/2.0/075 | OD | Physician | 07/11/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15140091&pc=QCB | 2023826-2022-02365 | 2022-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2, -15.00/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6). 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT:UNKNOWN ,ETHINICITY: UNKNOWN, RACE:UNKNOWN. CLAIM# (B)(4). | -15.00/1.5/090 | OD | Physician | 07/15/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15140098&pc=QCB | 2023826-2022-02367 | 2022-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.5/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT:UNKNOWN, ETHNICITY:UNKNOWN, RACE:UNKNOWN. CLAIM # (B)(4). | -10.5/1.5/090 | OD | Physician | 07/15/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15140120&pc=QCB | 2023826-2022-02366 | 2022-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2, -15.50/1.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). | -15.50/1.5/086 | OS | Physician | 07/15/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15142671&pc=QCB | 2023826-2022-02519 | 2022-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -2.0/1.5/63 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE REPORTED AS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). B5- THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -2.0/1.5/63 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. PATIENT IS HAPPY. D6B- (B)(6) 2022. CLAIM# (B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM #: (B)(4). | -2.0/1.5/63 | OD | Physician | 07/02/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15149888&pc=QCB | 2023826-2022-02612 | 2022-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -10.5/2.0/077 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LENS ROTATION WHICH WAS NOTED 8 MONTHS AFTER IMPLANTATION. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. IT WAS NOTED THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR AND INDICATED "PATIENT HAD LOOSE SUSPENSORY LIGAMENT". Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | -10.5/2.0/077 | OS | Physician | 07/20/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15149958&pc=QCB | 2023826-2022-02611 | 2022-01-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Patient-Device Incompatibility; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -8.0/2.0/114 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED LENS ROTATION WHICH WAS NOTED 8 MONTHS AFTER IMPLANTATION. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. IT WAS NOTED THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR AND INDICATED "PATIENT HAD LOOSE SUSPENSORY LIGAMENT". Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -8.0/2.0/114 | OD | Physician | 07/20/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15153281&pc=QCB | 2023826-2022-02272 | 2021-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -15.50/1.5/35 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (ODO ON (B)(6) 2021. EXCESSIVE VAULT WAS OBSERVED. ON (B)(6) 2021 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -15.50/1.5/35 | OD | Physician | 06/15/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15153723&pc=QCB | 2023826-2022-02514 | 2022-07-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50/+2.0/110 (SPHERE/CYLINDER/AXIS), TORE DURING LOADING AND THERE WAS NO PATIENT CONTACT. A REPLACEMENT LENS WAS NOT IMPLANTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN WITH DEBRIS THROUGHOUT THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -07.50/+2.0/110 | NA | Physician | 07/08/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15153817&pc=QCB | 2023826-2022-02508 | 2022-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -06.50/+4.0/038 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. THE PATIENT IS UNHAPPY AS THE RESIDUAL IS +0.75/-2.25/180 AND THE PATIENT IMPROVING TO 20/20. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDTIONAL DATA: B5- "THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -06.50/+4.0/038 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON 11/JUN/2022. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. REPORTEDLY "THE PATIENT IS UNHAPPY AS THE RESIDUAL IS +0.75/-2.25/180 AND THE PATIENT IMPROVING TO 20/20." THE LENS WAS EXPLANTED. ON 19/SEPT./2022 A REPLACEMENT LENS OF THE SAME MODEL/LENGTH WAS IMPLANTED. REPORTEDLY "POST OP DETAILS ON 20/09/2022- VAULT-1CT- ACCEPTANCE -0.75/-1.070 & 6/6." CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT/DEFORMED AND OPTIC DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -06.50/+4.0/038 | OS | Physician | 07/04/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15159616&pc=QCB | 2023826-2022-02527 | 2022-06-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF -11.5/1.5/154 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED; REMOVED AND REPLACED INTRA-OPERATIVELY ON (B)(6)2022 FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. PATIENT IS RECOVERING WELL. CAUSE IS REPORTED AS A PATIENT RELATED FACTOR. DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: CLAIM#: (B)(4). | -11.5/1.5/154 | NA | Other Health Care Professional | 06/24/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15160387&pc=QCB | 2023826-2022-02615 | 2022-03-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -7.00/+2.5/85 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED LOW VAULTING WITH LENS ROTATION. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -7.00/+2.5/85 | OD | Physician | 08/02/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15160708&pc=QCB | 2023826-2022-02529 | 2022-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6 VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -12.5/2.0/011 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR THE SAME MODEL AND SIZE LENS DUE TO LOW VAULT. THE REPLACEMENT LENS WAS IMPLANTED VERTICALLY AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN AND BENT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT INFORMATION:UNK. (B)(4). CLAIM# (B)(4). | -12.5/2.0/011 | OD | Physician | 06/24/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15167212&pc=QCB | 2023826-2022-02588 | 2022-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.5/+2.5/100 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. IT WAS REPORTED THAT THE PATIENT WAS DISSATISFIED (FELLOW EYE HAD ANGLE CLOSURE) AND THE LENS WAS EXPLANTED ON (B)(6) 2022. IT WAS REPORTED THAT THE PROBLEM RESOLVED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID WITH RESIDUE/DEBRIS ON PRODUCT IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS.CLEAR RESIDUE THROUGHOUT THE LENS WAS OBSERVED. CLAIM #(B)(4). Manufacturer Narrative: B1: ADVERSE EVENT. CLAIM#: (B)(4). | -11.5/+2.5/100 | OD | Physician | 07/05/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15171917&pc=QCB | 2023826-2022-02594 | 2022-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -16.0/1.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON 13-JUN-2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO LOW VAULT WITH ROTATION AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -16.0/1.5/089 | OD | Physician | 07/08/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15172075&pc=QCB | 2023826-2022-02595 | 2022-06-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -6.0/2.5/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC DEFORMED TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -6.0/2.5/088 | OD | Physician | 07/08/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15176574&pc=QCB | 2023826-2022-02574 | 2022-03-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -11.5/1.5/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED VERTICALLY INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LENS ROTATION WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H3: PRODUCT EVALUATION: LENS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND THE HAPTIC TORN AND BROKEN. DIMENSIONAL INSPECTION WAS UNABLE TO BE PERFORMED DUE TO SIGNIFICANT LENS DAMAGE. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -11.5/1.5/086 | OD | Physician | 07/06/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15178228&pc=QCB | 2023826-2022-02619 | 2022-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MMVTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -11/3.5/19 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE LEFT EYE (OS) OF A PATIENT WITH PRE-EXISTING KERATOCONUS ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REPOSITIONED AND EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK. CLAIM# (B)(4). | -11/3.5/19 | OS | Physician | 07/06/2022 | 2022 | 44 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15178247&pc=QCB | 2023826-2022-02603 | 2022-07-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00/2.5/132 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE INCORRECT LENS WAS IMPLANTED. THE LENS WAS REPOSITIONED ON (B)(6) 2022 AND THE PROBLEM WAS RESOLVED. POST-OP THE STATUS OF THE EYE WAS WNL, IOL AT 79 DEGREES AS PLANNED POST ROTATION. MR POST ROTATION (D1): -0.5 -0.75 @ 10 = 20/20-2. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -18.00/2.5/132 | OS | Physician | 07/12/2022 | 2022 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15178694&pc=QCB | 2023826-2022-02596 | 2022-07-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 THE DOCTOR FOUND A 12.6MM, VTICMO12.6, -9.0/1.5/030 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS BROKEN AFTER OPENING THE VIAL. THERE WAS NO PATIENT CONTACT. A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 25-OCT-2022. H3: DEVICE EVALUATION: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE PRODUCT. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN AND BENT. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | -9.0/1.5/030 | NA | Physician | 07/08/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15179790&pc=QCB | 2023826-2022-02579 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM 13.2 IMPLANTABLE COLLAMER LENS OF -6.50/1.0/084 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD). THE REPORTER STATES THE LENS IS PLANNED TO BE REMOVED FOR AN UNKNOWN REASON. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -6.50/1.0/084 | OD | Physician | 07/06/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15179828&pc=QCB | 2023826-2022-02620 | 2022-04-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -12.0/3.5/92 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPOSITIONED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCATED TO LOW VAULT. THE PROBLEM WAS NOT RESOLVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SPHERIC MODEL, SAME LENGTH LENS, AND THE PROBLEM WAS RESOLVED. REPORTEDLY, "THE RESIDUAL ASTIGMATISM WAS CORRECTED BY LASER. Manufacturer Narrative: H3: DEVICE EVALUATON: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INPSECTION FOUND FIBERS ON A NO VISIBLE DAMAGED LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT:UNKNOWN, ETHINICITY: UNKNOWN, SEX:UNKNOWN. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -12.0/3.5/92 | OD | Physician | 07/07/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15179858&pc=QCB | 2023826-2022-02580 | 2022-07-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -16.0/2.5/094 (SPHERE/CYLINDER/AXIS) LENS TEAR/BREAK DURING LOADING; AFTER OPENING VIAL. THE LENS WAS NOT IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THE CAUSE IS REPORTED AS DEVICE; HYPOTHETICAL FORCEPS ISSUE. Manufacturer Narrative: B5 - BACKUP LENS WAS IMPLANTED ON (B)(6) 2022 AND PROBLEM WAS RESOLVED. STATUS OF EYE "PATIENT IS VERY SATISFIED". (B)(4). Manufacturer Narrative: INVESTIGATION TYPE: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -16.0/2.5/094 | OS | Physician | 07/07/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15180054&pc=QCB | 2023826-2022-02621 | 2022-05-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1 MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -16.5/2.5/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPOSITIONED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THE PROBLEM WAS NOT RESOLVED. LATER THAT DAY, THE LENS WAS EXCHANGED FOR A SPHERIC MODEL, SAME LENGTH LENS, AND THE PROBLEM WAS RESOLVED. REPORTEDLY, "THE RESIDUAL ASTIGMATISM WAS CORRECTED BY LASER. Manufacturer Narrative: CORRECTED DATA: B3: DATE OF EVENT (B)(6) 2022 SHOULD BE CORRECTED TO SAY (B)(6) 2022. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGED TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -16.5/2.5/93 | OS | Physician | 10/31/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15180314&pc=QCB | 2023826-2022-02623 | 2022-07-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030017 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022, A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -13.5/3.5/87 (SPHERE/CYLINDER/AXIS) LENS TORE/BREAK INJECTION/DELIVERY INTO THE RIGHT EYE (OD). THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. THE LENS WAS INTRAOPERATIVELY REMOVED AND REPLACED WITH A REPLACEMENT LENS. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: PATIENT IDENTIFICATION: UNKNOWN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -13.5/3.5/87 | OD | Physician | 07/08/2022 | 2022 | 52 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15180447&pc=QCB | 2023826-2022-02626 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-08-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -4.50/1.0/006 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B1-ADVERSE EVENT OR PRODUCT PROBLEM: "ADVERSE EVENT" SHOULD BE REMOVED AND CORRECTED TO STATE "PRODUCT PROBLEM". B2-OUTCOMES ATTRIBUTED TO ADVERSE EVENT: "REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE" SHOULD BE OMITTED FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -4.50/1.0/006 (SPHERE/CYLINDER/AXIS) WAS INTRAOPERATIVELY IMPLANTED, REMOVED, AND REPLACED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). | -4.50/1.0/006 | OD | Physician | 10/24/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15182866&pc=QCB | 2023826-2022-02460 | 2022-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -14.0/1.5/022 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | -14.0/1.5/022 | OD | Physician | 07/18/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15182867&pc=QCB | 2023826-2022-02459 | 2022-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -07.0/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -07.0/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION IN IRIDOCORNEAL ANGLES AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. H6 - HEALTH EFFECT CLINICAL COIDE REPORTED AS 4582 - CORRECT TO 4581. H6 - HEALTH EFFECT IMPACT CODE 4581: SIGNIFICANT REDUCTION OF IRIDO-CONEAL ANGLES. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | -07.0/1.5/092 | OS | Physician | 07/18/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15182871&pc=QCB | 2023826-2022-02461 | 2022-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -08.0/1.0/150 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -08.0/1.0/150 | OS | Physician | 07/18/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15182875&pc=QCB | 2023826-2022-02468 | 2022-04-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -06.5/1.0/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -06.5/1.0/072 | OS | Physician | 07/18/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15182878&pc=QCB | 2023826-2022-02458 | 2022-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -05.0/1.5/123 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -05.0/1.5/123 | OS | Physician | 07/18/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15182879&pc=QCB | 2023826-2022-02473 | 2022-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -09.5/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -09.5/1.0/089 | OS | Physician | 07/21/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15182885&pc=QCB | 2023826-2022-02470 | 2022-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.5/3.0/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED VERTICALLY WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.5/3.0/074 | OS | Physician | 07/18/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15182887&pc=QCB | 2023826-2022-02463 | 2022-06-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -09.5/0.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON 05/JUN/2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -09.5/0.5/085 | OD | Physician | 07/18/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15182888&pc=QCB | 2023826-2022-02474 | 2022-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -17.5/1.5/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH SPHERICAL LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -17.5/1.5/110 | OD | Physician | 07/21/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15182890&pc=QCB | 2023826-2022-02476 | 2022-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -12.5/2.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -12.5/2.5/089 | OD | Physician | 07/21/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15182896&pc=QCB | 2023826-2022-02481 | 2022-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.5/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED VERTICALLY WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT UNK. ETHNICITY UNK. RACE UNK. (B)(4). CLAIM # (B)(4). | -10.5/1.5/092 | OS | Physician | 07/21/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15182898&pc=QCB | 2023826-2022-02484 | 2022-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.0/1.5/061 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -09.0/1.5/061 | OS | Physician | 07/22/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15182901&pc=QCB | 2023826-2022-02480 | 2022-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -08.00 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - REPORTED AS 12.6MM VICMO12.6, -08.00 DIOPTER LENS - CORRECT TO 12.1MM VTICMO12.1, -9.5/1.5/112 LENS. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. DEVICE PROBLEM CODE: (B)(4). CLAIM # (B)(4). | -8 | OD | Physician | 07/21/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15182903&pc=QCB | 2023826-2022-02483 | 2022-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -09.5/4.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. CLAIM# (B)(4). | -09.5/4.5/090 | OD | Physician | 07/22/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15182909&pc=QCB | 2023826-2022-02482 | 2022-07-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -11.0/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | -11.0/2.5/093 | OS | Physician | 07/21/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15191541&pc=QCB | 2023826-2022-02675 | 2022-07-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -16.5/4.0/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS IMPLANTED UPSIDE DOWN. INTRAOPERATIVELY THE LENS WAS REMOVED WITH NO PATIENT INJURY. A REPLACEMENT LENS WAS IMPLANTED AT A LATER DATE AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -16.5/4.0/092 | OD | Physician | 08/03/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15191857&pc=QCB | 2023826-2022-02693 | 2022-03-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6 MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -12.5/2.0/001 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS THAT WAS IMPLANTED HORIZONTALLY DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -12.5/2.0/001 | OD | Physician | 07/10/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15191886&pc=QCB | 2023826-2022-02692 | 2022-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -12.5/5.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS THAT WAS IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). | -12.5/5.0/091 | OD | Physician | 07/10/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15192184&pc=QCB | 2023826-2022-02707 | 2022-07-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -9.5/1.5/20 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND IT WAS NOTED THERE IS NO PLANS FOR A REPLACEMENT LENS AT A LATER DATE. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). | -9.5/1.5/20 | OD | Physician | 08/08/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15192543&pc=QCB | 2023826-2022-02708 | 2022-04-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -10/1.5/077 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS USER ERROR. Manufacturer Narrative: (B)(4). | -10/1.5/077 | OS | Physician | 08/03/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15195880&pc=QCB | 2023826-2022-02606 | 2022-06-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+2.5/172 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (LENS IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.00/+2.5/172 | NA | Physician | 07/15/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15195909&pc=QCB | 2023826-2022-02609 | 2021-01-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50/+2.5/010 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED LATER THAT SAME DAY DUE TO EXCESSIVE VAULTING AND THE PROBLEM WAS RESOLVED. THE LENS WAS NOT REPLACED WITH ANOTHER LENS DUE TO THE PATIENT WAS PREGNANT. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: PRODUCT EVALUATION: LENS RETURNED DRY A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY:UNK. RACE: UNK. (B)(4). CLAIM # (B)(4). | -07.50/+2.5/010 | OD | Physician | 07/12/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15195917&pc=QCB | 2023826-2022-02610 | 2022-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+1.0/170 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT/DEFORMED. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT:UNKNOWN, ETHINICITY: UNKNOWN, RACE: UNKNOWN. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.00/+1.0/170 | OS | Physician | 07/12/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15196107&pc=QCB | 2023826-2022-02698 | 2022-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -12.5/+1.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: B3: DATE OF EVENT - 29-JUN-2022 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -12.5/+1.5/093 | OD | Physician | 07/11/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15196182&pc=QCB | 2023826-2022-02700 | 2022-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -7.0/+1.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS LATER EXPLANTED AND REPLACED FOR SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -7.0/+1.5/092 | OS | Physician | 07/11/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15199945&pc=QCB | 2023826-2022-02676 | 2022-04-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -18.0/3.0/091 (SPHERE/CYLINDER/AXIS) TORE/ BROKE DURING INJECTION/ DELIVERY INTO THE PATIENT'S RIGHT EYE (OD). THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY. ON A LATER DATE A REPLACEMENT LENS OF THE SAME MODEL AND SIZE WAS IMPLANTED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). | -18.0/3.0/091 | OD | Physician | 07/11/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15204581&pc=QCB | 2023826-2022-02743 | 2022-07-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -9.0/4.5/084 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED. IT WAS INDICATED THERE IS NO PLANS FOR A REPLACEMENT LENS. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -9.0/4.5/084 | OS | Physician | 08/09/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15204699&pc=QCB | 2023826-2022-02744 | 2022-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -14.50/2.50/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE PATIENT EXPERIENCED LOW VAULTING WITH LENS ROTATION WHICH WAS ASSESSED ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS NOTED AS "SMALLER LENS CAUSED ROTATION". Manufacturer Narrative: CLAIM#: (B)(4). | -14.50/2.50/90 | OD | Physician | 08/09/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15208919&pc=QCB | 2023826-2022-02641 | 2022-07-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -07.00/+3.0/097 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING AND THE PROBLEM WAS RESOLVED. THE PATIENT DID NOT WANT ANOTHER ICL IMPLANTED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM#(B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS AND RESIDUE THROUGHOUT THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -07.00/+3.0/097 | NA | Physician | 07/22/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15210704&pc=QCB | 2023826-2022-02625 | 2022-07-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -18.0/2.0/110 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS INTRAOPERATIVELY REMOVED DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. REPORTEDLY, "MAINTAIN PREOPERATIVE STATE." Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -18.0/2.0/110 | OD | Physician | 07/11/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15217202&pc=QCB | 2023826-2022-02748 | 2022-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -11.0/3.0/071 (SPHERE/CYLINDER/AXIS) WAS INTRAOPERATIVELY IMPLANTED AND REMOVED ON (B)(6) 2022, DUE TO A LENS TEAR NEAR THE HASHMARK. THE REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. NO PATIENT INJURY WAS REPORTED. REPORTEDLY, "I COULD NOT VISUALLY SEE IF THE TEAR WAS ALREADY IN THE LENS, IF THE TEAR WAS FORMED BY THE INSERTER, OR IF THE TEAR WAS CAUSED BY THE OCUCOAT CANNULA." Manufacturer Narrative: CORRECTED DATA: G2: "FOREIGN" SHOULD BE REMOVED. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). | -11.0/3.0/071 | NA | Physician | 07/13/2022 | 2022 | 29
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15217845&pc=QCB | 2023826-2022-02749 | 2022-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -13.0/2.0/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AND REMOVED ON (B)(6) 2022 , DUE TO THE LENS FLIPPED UPSIDE DOWN . THE REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. NO PATIENT INJURY WAS REPORTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: G2- "FOREIGN" SHOULD BE REMOVED. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -13.0/2.0/087 | NA | Physician | 07/13/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15218330&pc=QCB | 2023826-2022-02705 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: REPORTEDLY, A 12.6MM TMICL12.6 -16.00/1.5/068 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL, LENGTH AND POWER LENS ON AN UNKNOWN DATE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#:(B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN VIAL. VISUAL INSPECTION FOUND A HAPTIC AND OPTIC TORN. CLAIM #(B)(4). Manufacturer Narrative: B4: CORRECTED TO 05-JUL-2022 IN INITIAL MDR. B5: CORRECTED TO: REPORTEDLY, A 12.6MM TMICL12.6 -16.00/1.5/068 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL, LENGTH AND POWER LENS ON AN UNKNOWN DATE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. G3: CORRECTED TO 05-JUL-2022 IN INITIAL MDR CLAIM#: (B)(4). | -16.00/1.5/068 | OS | Other Health Care Professional | 07/05/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15218482&pc=QCB | 2023826-2022-02706 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: REPORTEDLY, A 12.6MM TMICL12.6 -16.00/1.0/106 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL, LENGTH AND POWER LENS ON AN UNKNOWN DATE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -16.00/1.0/106 | OD | Other Health Care Professional | 07/14/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15218853&pc=QCB | 2023826-2022-02679 | 2022-01-27 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 -14.5/1.5/109 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED WITHIN THE SAME SURGERY AND REPLACED WITH A SHORTER LENGTH LENS OF SIMILAR DIOPTER DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. NO PATIENT INJURY. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINTS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE CODE: 1494, OFF-LABEL USE, AGE<21. (B)(4). | -14.5/1.5/109 | OD | Other Health Care Professional | 07/13/2022 | 2022 | 17 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15220735&pc=QCB | 2023826-2022-02759 | 2022-05-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 13.2MM, VTICMO13.2, -7.0/1.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE LEFT EYE (OS). THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY, WITH NO PATIENT INJURY. A REPLACEMENT LENS WAS IMPLANTED AT A LATER DATE AND THE PROBLEM RESOLVED. Manufacturer Narrative: (B)(4). | -7.0/1.5/087 | OS | Physician | 07/15/2022 | 2022 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15225743&pc=QCB | 2023826-2022-02549 | 2022-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -09.0/1.0/065 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. CLAIM# (B)(4). | -09.0/1.0/065 | OS | Physician | 07/25/2022 | 2022 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15225745&pc=QCB | 2023826-2022-02554 | 2022-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/1.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. (B)(4). CLAIM# (B)(4). | -11.0/1.5/083 | OS | Physician | 07/25/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15225752&pc=QCB | 2023826-2022-02557 | 2022-06-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.5/3.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -12.5/3.0/092 | OD | Physician | 07/25/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15225758&pc=QCB | 2023826-2022-02562 | 2022-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/2.5/111 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.5/2.5/111 | OD | Physician | 07/25/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15225759&pc=QCB | 2023826-2022-02568 | 2022-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.5/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. CLAIM# (B)(4). | -09.5/2.5/091 | OD | Physician | 07/28/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15225761&pc=QCB | 2023826-2022-02553 | 2022-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/1.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. SEE MFR # 2022-02552 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: (B)(4). | -10.5/1.5/095 | OD | Physician | 07/25/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15225762&pc=QCB | 2023826-2022-02551 | 2022-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.0/3.0/076 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.0/3.0/076 | OD | Physician | 07/25/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15225764&pc=QCB | 2023826-2022-02558 | 2022-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.5/2.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.5/2.0/092 | OD | Physician | 07/25/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15225766&pc=QCB | 2023826-2022-02555 | 2022-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -07.00/1.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -07.00/1.0/088 | OD | Physician | 07/25/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15225767&pc=QCB | 2023826-2022-02570 | 2022-05-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -08.0/2.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. CLAIM # (B)(4). | -08.0/2.5/086 | OD | Physician | 07/28/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15225773&pc=QCB | 2023826-2022-02567 | 2022-01-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.0/1.5/176 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.0/1.5/176 | OD | Physician | 07/28/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15225777&pc=QCB | 2023826-2022-02566 | 2022-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.0/1.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. CLAIM# (B)(4). | -09.0/1.5/083 | OD | Physician | 07/28/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15225784&pc=QCB | 2023826-2022-02552 | 2022-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.5/1.5/110 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS DID NOT RESOLVE THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, THE ARCH HEIGHT HAS NO OBVIOUS CHANGE WHEN THE 13.2 CRYSTAL IS REPLACED. SEE MFR # 2022-02553 FOR CLAIM THAT RESOLVED PROBLEM. Manufacturer Narrative: (B)(4). | -10.5/1.5/110 | OD | Physician | 07/25/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15225785&pc=QCB | 2023826-2022-02556 | 2022-07-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/1.5/061 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -11.0/1.5/061 | OD | Physician | 07/25/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15225788&pc=QCB | 2023826-2022-02559 | 2022-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.5/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.5/2.0/088 | OS | Physician | 07/25/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15225789&pc=QCB | 2023826-2022-02561 | 2022-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.0/1.5/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. DEVICE PROBLEM CODE: (B)(4). CLAIM # (B)(4). | -09.0/1.5/071 | OS | Physician | 07/25/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15225799&pc=QCB | 2023826-2022-02569 | 2022-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.0/2.5/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.0/2.5/072 | OS | Physician | 07/28/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15225803&pc=QCB | 2023826-2022-02571 | 2022-05-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -08.0/2.5/068 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -08.0/2.5/068 | OS | Physician | 07/28/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15230247&pc=QCB | 2023826-2022-02687 | 2022-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -14.0/1.5/174 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR THE SAME MODEL AND SIZE LENS DUE TO LOW VAULT. THE REPLACEMENT LENS WAS IMPLANTED VERTICALLY AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: H3-DEVICE EVALUATION: SHOULD BE CORRECTED TO STATE, "THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS". CLAIM# (B)(4). | -14.0/1.5/174 | OD | Physician | 07/15/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15230360&pc=QCB | 2023826-2022-02688 | 2022-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -13.0/1.5/168 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR THE SAME MODEL AND SIZE LENS DUE TO LOW VAULT. THE REPLACEMENT LENS WAS IMPLANTED VERTICALLY AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -13.0/1.5/168 | OS | Physician | 07/15/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15233654&pc=QCB | 2023826-2022-02732 | 2022-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+2.5/112 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (LENS IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: (B)(4). | -09.00/+2.5/112 | OS | Physician | 07/25/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15233674&pc=QCB | 2023826-2022-02728 | 2022-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.00/+1.0/140 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT REPLACED WITH ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: PRODUCT EVALUATION: LENS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -16.00/+1.0/140 | OD | Physician | 07/25/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15233693&pc=QCB | 2023826-2022-02731 | 2022-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -05.50/+1.5/073 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2022 BUT THE PROBLEM WAS NOT RESOLVED SO THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH A LONGER LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -05.50/+1.5/073 | OS | Physician | 07/25/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15233776&pc=QCB | 2023826-2022-02736 | 2022-07-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50/+1.0/109 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). | -06.50/+1.0/109 | OD | Physician | 07/26/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15233832&pc=QCB | 2023826-2022-02733 | 2022-07-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+0.5/111 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY AND RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.50/+0.5/111 | OS | Physician | 07/26/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15233866&pc=QCB | 2023826-2022-02739 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AND REMOVED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -05.50/+1.0/101 (SPHERE/CYLINDER/AXIS) FROM THE PATIENTS EYE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS IMPLANTED IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -05.50/+1.0/101 | OD | Physician | 07/26/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15233869&pc=QCB | 2023826-2022-02737 | 2022-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS RIGHT EYE (OD). THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.00/+2.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING AND THE PROBLEM WAS RESOLVED. THE PATIENT IS IN GOOD CONDTION. THE CAUSE OF THE EVENT WAS UNKNOWN. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: D2: COMMON DEVICE NAME: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). | -10.00/+2.0/093 | OD | Physician | 07/26/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237379&pc=QCB | 2023826-2022-02808 | 2022-06-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 12.6MM, VTICMO12.6, -8.5/6.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE LEFT EYE (OS). THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY, WITH NO PATIENT INJURY. A REPLACEMENT LENS WAS IMPLANTED AT A LATER DATE AND THE PROBLEM RESOLVED. Manufacturer Narrative: (B)(4). | -8.5/6.0/082 | OS | Physician | 07/18/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237529&pc=QCB | 2023826-2022-02690 | 2022-06-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.5/0.5/075 (SPHERE/CYLINDER/AXIS) TORE/ BROKE DURING INJECTION/ DELIVERY INTO THE PATIENT'S LEFT EYE (OS). ON (B)(6) 2022 THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY. LATER ON 14-JUN-2022 A REPLACEMENT LENS OF THE SAME MODEL AND SIZE WAS IMPLANTED. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.5/0.5/075 | OS | Physician | 07/17/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237530&pc=QCB | 2023826-2022-02691 | 2022-06-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -6.5/1.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR THE SAME MODEL AND SIZE LENS DUE TO LOW VAULT. THE REPLACEMENT LENS WAS IMPLANTED VERTICALLY AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: PRODUCT EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND THE HAPTIC BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -6.5/1.0/094 | OD | Physician | 07/17/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237535&pc=QCB | 2023826-2022-02809 | 2022-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS -3.5/0.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS LATER EXPLANTED AND REPLACED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM RESOLVED. CAUSE WAS REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS -3.5/0.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS LATER EXPLANTED AND REPLACED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION AND THE PROBLEM RESOLVED. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(6). CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS -3.5/0.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS LATER EXPLANTED AND REPLACED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION AND THE PROBLEM RESOLVED. CAUSE WAS REPORTED AS UNKNOWN. H6: MEDICAL DEVICE PROBLEM CODE: 2923. CLAIM# (B)(4). | -3.5/0.5/093 | OD | Physician | 07/18/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237942&pc=QCB | 2023826-2022-02532 | 2022-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+0.5/169 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -8.0/+0.5/169 | OS | Physician | 07/22/2022 | 2022 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237943&pc=QCB | 2023826-2022-02534 | 2022-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.0/+0.5/102 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -5.0/+0.5/102 | OD | Physician | 07/22/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237945&pc=QCB | 2023826-2022-02539 | 2022-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.5/085 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.5/+1.5/085 | OD | Physician | 07/25/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237946&pc=QCB | 2023826-2022-02538 | 2022-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.5/095 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.5/+1.5/095 | OS | Physician | 07/25/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237947&pc=QCB | 2023826-2022-02537 | 2022-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.0/+2.0/091 | OD | Physician | 07/25/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237948&pc=QCB | 2023826-2022-02535 | 2022-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+0.5/169 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.5/4.0/095 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LENS OF THE SAME MODEL/SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. REPORTEDLY "THERE WAS NO ORIGINAL POWER AT THAT TIME SO AN EXISTING LENS ADJACENT TO THE SPHERO-CYLINDRICAL AXIS WAS SELECTED FOR IMPLANTATION." CAUSE OF THE EVENT REPORTED AS UNKNOWN." CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -8.0/+0.5/169 | OS | Physician | 07/22/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237950&pc=QCB | 2023826-2022-02536 | 2022-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+2.5/088 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.5/+2.5/088 | OS | Physician | 07/25/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237952&pc=QCB | 2023826-2022-02540 | 2022-06-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+4.0/100 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIED (UDI)#- (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). | -11.5/+4.0/100 | OS | Physician | 10/20/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237959&pc=QCB | 2023826-2022-02541 | 2022-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+4.0/103 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+4.0/103 | OS | Physician | 07/26/2022 | 2022 | 47 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237960&pc=QCB | 2023826-2022-02546 | 2022-05-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.00/+1.00/078 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -6.00/+1.00/078 | OS | Physician | 07/26/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237962&pc=QCB | 2023826-2022-02709 | 2022-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.0/002 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE LENS WAS IMPLANTED VERTICALLY. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.0/+2.0/002 | OD | Physician | 07/27/2022 | 2022 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237963&pc=QCB | 2023826-2022-02710 | 2022-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+2.0/178 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/+2.0/178 | OS | Physician | 07/27/2022 | 2022 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237965&pc=QCB | 2023826-2022-02545 | 2022-05-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.0/068 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.0/+1.0/068 | OD | Physician | 05/21/2022 | 2022 | 42 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237968&pc=QCB | 2023826-2022-02714 | 2022-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+1.5/075 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -14.0/+1.5/075 | OD | Physician | 07/28/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237970&pc=QCB | 2023826-2022-02712 | 2022-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+1.5/097 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.5/+1.5/097 | OD | Physician | 07/28/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237981&pc=QCB | 2023826-2022-02717 | 2022-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+2.0/070 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D4: UNIQUE IDENTIFIER: (B)(4). D9: DEVICE AVAILABLE FOR EVALUATION: YES. RETURNED TO MANUFACTURER 11/07/2022. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+2.0/070 | OD | Physician | 08/01/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237982&pc=QCB | 2023826-2022-02715 | 2022-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.0/070 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -8.5/+1.0/070 | OD | Physician | 07/29/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237983&pc=QCB | 2023826-2022-02720 | 2022-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+2.5/083 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: D4 - UNIQUE IDENTIFIER (UDI)# : (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #: (B)(4). | -13.5/+2.5/083 | OS | Physician | 08/02/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237984&pc=QCB | 2023826-2022-02723 | 2022-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.0/+3.5/098 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: (B)(4). | -16.0/+3.5/098 | OS | Physician | 08/02/2022 | 2022 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237985&pc=QCB | 2023826-2022-02722 | 2022-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.50/+2.5/076 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -12.50/+2.5/076 | OD | Physician | 08/02/2022 | 2022 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237986&pc=QCB | 2023826-2022-02721 | 2022-06-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.0/008 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -7.0/+1.0/008 | OS | Physician | 08/02/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237987&pc=QCB | 2023826-2022-02716 | 2022-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.0/080 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -11.5/+1.0/080 | OD | Physician | 07/29/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15237988&pc=QCB | 2023826-2022-02718 | 2022-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+1.5/110 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.0/+1.5/110 | OD | Physician | 08/01/2022 | 2022 | 45 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15245895&pc=QCB | 2023826-2022-02784 | 2022-06-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50/+4.5/078 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENT'S LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: D4 - UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). | -09.50/+4.5/078 | OS | Physician | 07/29/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15245906&pc=QCB | 2023826-2022-02740 | 2022-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+2.5/106 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -10.50/+2.5/106 | OD | Physician | 07/27/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15245932&pc=QCB | 2023826-2022-02741 | 2022-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -13.00/+5.0/093 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS TORE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -13.00/+5.0/093 | OD | Physician | 07/27/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15245960&pc=QCB | 2023826-2022-02729 | 2022-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.50/+0.5/066 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS NOT REPLACED WITH ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: PRODUCT EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -16.50/+0.5/066 | OS | Physician | 07/25/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15253742&pc=QCB | 2023826-2022-02830 | 2022-05-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -9.0/+2.0/74 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. IT WAS REPORTED THAT THERE WAS NO PATIENT INJURY AND THAT THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | -9.0/+2.0/74 | OS | Physician | 07/19/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15254458&pc=QCB | 2023826-2022-02819 | 2022-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+3.0/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: D4 - UNIQUE IDENTIFIFER (UDI)# : (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #: (B)(4). Manufacturer Narrative: B3:CORRECTED TO DATE OF EVENT (B)(6) 2022 IN INITIAL MDR. CLAIM # (B)(4). | -12.0/+3.0/092 | OD | Physician | 06/15/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15254493&pc=QCB | 2023826-2022-02823 | 2022-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+3.50/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+3.50/094 | OS | Physician | 08/04/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15254516&pc=QCB | 2023826-2022-02811 | 2022-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.50/+2.0/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -11.50/+2.0/090 | OD | Physician | 08/02/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15254569&pc=QCB | 2023826-2022-02820 | 2022-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.5/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: D4 - UNQUE IDENTIFIER (UDI) # (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+2.5/093 | OS | Physician | 08/03/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15255766&pc=QCB | 2023826-2022-02832 | 2022-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -13.5/1.5/107 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW WITH ROTATION. IT WAS REPORTED THAT THE EXCHANGE LENS DID NOT RESOLVE THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: PRODUCT EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -13.5/1.5/107 | OD | Physician | 07/20/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15255857&pc=QCB | 2023826-2022-02754 | 2022-02-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030017 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022, A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -14.5/2.0/91 (SPHERE/CYLINDER/AXIS) LENS TORE AND WAS DAMAGED DURING LOADING. THERE WAS NO PATIENT CONTACT. A REPLACEMENT LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE AND THE PROBLEM WAS RESOLVED. REPORTEDLY, STATUS OF THE EYE IS, "GOOD." Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. WORK ORDER SEARCH: CLAIM# (B)(4). IS REGARDING AN EXCESSIVE VAULT EXCHANGE WHICH IS UNRELATED TO THIS COMPLAINT. NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -14.5/2.0/91 | OD | Physician | 07/20/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15256458&pc=QCB | 2023826-2022-02833 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-08-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -13.5/1.5/106 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 AS A REPLACEMENT LENS TO ADDRESS LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. THE REPORTER STATED THAT PATIENTS VISION UCVA IS 20/20 AND THAT THE DOCTOR DECIDED TO CONTINUE TO MONITOR. SEE MFR # 2023826-2022-02832 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). (B)(4). | -13.5/1.5/106 | OD | Physician | 07/20/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15256527&pc=QCB | 2023826-2022-02831 | 2022-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.5/2.0/66 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO LOW VAULT WITH ROTATION AND LATER REPLACED WITH A SPHERICAL LENS OF LONGER LENGTH. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED TO BE USER ERROR AND OTHER. Manufacturer Narrative: CLAIM#: (B)(4). | -12.5/2.0/66 | OS | Physician | 07/20/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15257548&pc=QCB | 2023826-2022-02834 | 2022-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -7.00/+5.5/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PROBLEM RESOLVED. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -7.00/+5.5/089 | OD | Physician | 07/22/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15257831&pc=QCB | 2023826-2022-02835 | 2022-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -7.50/+6.0/077 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. IT WAS REPORTED THAT REFRACTIVE TREATMENT WAS PERFORMED, LENS REMAINS IMPLANTED.THE REPORTER STATED THAT ORIENTATION OF TORIC ICL IS STABLE. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5: PER UPDATED INFORMATION, THE PROBLEM IS REPORTED AS RESOLVED. REPORTER STATES THAT "THE CLINIC PLANS A LASER TOUCH-UP." CLAIM#: (B)(4). Manufacturer Narrative: B5: THE LENS WAS EXCHANGED ON (B)(6) 2023 WITH SAME LENGTH BUT DIFFERENT DIOPTER LENS AND PROBLEM WAS RESOLVED. REPORTEDLY, "THE CLINIC AND THE PATIENT DECIDED TO PROCEED A LENS EXCHANGE INSTEAD THE LASER TOUCH-UP." CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -7.50/+6.0/077 | OS | Physician | 07/22/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15262399&pc=QCB | 2023826-2022-02746 | 2022-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -5.5/1.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS NOT RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: OFF LABEL USE- <21. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS FOUND. CLAIM# (B)(4). Manufacturer Narrative: B5: PROBLEM WAS RESOLVED. CLAIM# (B)(4). | -5.5/1.0/091 | OS | Physician | 07/22/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15262434&pc=QCB | 2023826-2022-02755 | 2022-02-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030017 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022, A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -14.0/1.50/81 (SPHERE/CYLINDER/AXIS) LENS TORE AND WAS DAMAGED DURING LOADING. THERE WAS NO PATIENT CONTACT. A REPLACEMENT LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE AND THE PROBLEM WAS RESOLVED. REPORTEDLY, STATUS OF THE EYE IS, "GOOD." Manufacturer Narrative: CORRECTED DATA: B1- "ADVERSE EVENT OR PRODUCT PROBLEM: ADVERSE EVENT" SHOULD BE CORRECTED TO "PRODUCT PROBLEM". CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -14.0/1.50/81 | OS | Physician | 10/18/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15265887&pc=QCB | 2023826-2022-02837 | 2022-05-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT WHILE LOADING A 13.2MM VTICM5_13.2; -10.5/3.0/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE INJECTOR; THE LENS HAD TORE. THIS OCCURRED ON (B)(6) 2022. A BACK UP LENS WAS IMPLANTED ON THIS SAME DATE AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.5/3.0/91 | NA | Physician | 08/18/2022 | 2022 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15269538&pc=QCB | 2023826-2022-02840 | 2022-06-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.5/+2.0/112 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -17.5/+2.0/112 | OS | Physician | 08/09/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15269595&pc=QCB | 2023826-2022-02841 | 2022-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.0/085 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI) #(B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). | -10.0/+1.0/085 | OD | Physician | 08/09/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15270427&pc=QCB | 2023826-2022-02209 | 2022-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+3.0/002 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED ON 25-JUN-2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -14.0/+3.0/002 | OD | Physician | 07/06/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15272004&pc=QCB | 2023826-2022-02853 | 2022-06-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 13.2MM, VTICM513.2 -13.5/+3.00/100 (SPHERE/CYLINDER/AXIS) IMPLATABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE RIGHT (OD) EYE. THERE WAS NO PATIENT CONTACT OR INJURY. CAUSE OF EVENT IS UNKNOWN. IT WAS REPORTED THAT THERE IS NO PLAN TO IMPLANT ANOTHER LENS. Manufacturer Narrative: EXPIRATION DATE: 12-31-2024. (B)(4). | -13.5/+3.00/100 | OD | Physician | 07/26/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15272054&pc=QCB | 2023826-2022-02848 | 2022-05-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -8.5/2.0/012 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(6). (B)(4). EXPIRATION DATE: 09-30-2024. (B)(4). Manufacturer Narrative: D4: EXPIRATION DATE: 30-SEP-2024 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). | -8.5/2.0/012 | OS | Physician | 07/25/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15272072&pc=QCB | 2023826-2022-02851 | 2022-06-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.5/2.0/074 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: ZOUMINGXUAN. (B)(4). EXPIRATION DATE: 12-31-2024. CLAIM#: (B)(4). | -11.5/2.0/074 | OS | Physician | 07/25/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15272116&pc=QCB | 2023826-2022-02849 | 2022-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.0/3.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: MEDICAL DEVICE PROBLEM CODE: (B)(4). EXPIRATION DATE: 02-28-2025 CLAIM#: (B)(4). | -10.0/3.0/090 | OS | Physician | 07/25/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15272128&pc=QCB | 2023826-2022-02844 | 2022-02-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.0/1.0/4 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). EXPIRATION DATE: (B)(6). 2024. CLAIM#: (B)(4). | -15.0/1.0/4 | OS | Physician | 07/25/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15272140&pc=QCB | 2023826-2022-02850 | 2022-06-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.5/2.5/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: EXPIRATION DATE: 11-30-2024. (B)(4). | -9.5/2.5/093 | OS | Physician | 07/25/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15272186&pc=QCB | 2023826-2022-02852 | 2022-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.5/1.5/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). EXPIRATION DATE: 02-28-2025. CLAIM#: (B)(4). | -11.5/1.5/095 | OS | Physician | 07/25/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15276382&pc=QCB | 2023826-2022-02825 | 2022-05-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.5/1.5/130 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED WITH LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: D4: EXPIRATION DATE- 12/31/2024. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: PRODUCT EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: D4: MODEL- SHOULD BE CORRECTED TO VTICMO_12.6. CLAIM#(B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -12.5/1.5/130 | OS | Physician | 10/27/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15277193&pc=QCB | 2023826-2022-02803 | 2022-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -8.5/2.5/169 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR THE SAME MODEL AND SIZE LENS DUE TO LOW VAULT. THE INITIAL LENS WAS PLACED VERTICALLY AND THE REPLACEMENT LENS WAS PLACED HORIZONTALLY. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -8.5/2.5/169 | OD | Physician | 07/25/2022 | 2022 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15277266&pc=QCB | 2023826-2022-02804 | 2022-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -12.0/1.0/73 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR THE SAME MODEL AND SIZE LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE LENS WAS IMPLANTED VERTICALLY. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -12.0/1.0/73 | OD | Physician | 07/25/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15280271&pc=QCB | 2023826-2022-02876 | 2022-02-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.5/1.5/072 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H6: THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. CLAIM# (B)(4). Manufacturer Narrative: D4: EXPIRATION DATE: 30-SEP-2024 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -12.5/1.5/072 | OS | Physician | 07/25/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15280272&pc=QCB | 2023826-2022-02829 | 2022-06-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.5/2.0/086(SPHERE/CYLINDER/AXIS) TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED WITH THE REPLACEMENT LENS AND THE PROBLEM WAS RESOLVED. NO PATIENT INJURY WAS REPORTED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/2.0/086 | OS | Physician | 07/25/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15282213&pc=QCB | 2023826-2022-02646 | 2022-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.5/3.0/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. CLAIM # (B)(4). | -10.5/3.0/078 | OS | Physician | 08/02/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15282216&pc=QCB | 2023826-2022-02645 | 2022-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/3.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. CLAIM# (B)(4). | -11.0/3.5/094 | OD | Physician | 08/01/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15282220&pc=QCB | 2023826-2022-02648 | 2022-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -130/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT UNKNOWN. ETHNICITY UNKNOWN. RACE UNKNOWN. CLAIM# (B)(4). | -130/2.5/091 | OS | Physician | 08/02/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15282222&pc=QCB | 2023826-2022-02654 | 2022-06-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.0/1.0/165 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/1.0/165 SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED A LOW VAULT WITHOUT ROTATION. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR THE SAME MODEL/SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS REPORTED AS "OK". THE CAUSE OF THE EVENT IS UNKNOWN." CLAIM #(B)(4). | -7.0/1.0/165 | OD | Physician | 08/02/2022 | 2022 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15282223&pc=QCB | 2023826-2022-02649 | 2022-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -12.5/3.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -12.5/3.5/093 | OD | Physician | 08/02/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15282225&pc=QCB | 2023826-2022-02650 | 2022-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -12.5/4.0/167 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED VERTICALLY WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -12.5/4.0/167 | OS | Physician | 08/02/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15282227&pc=QCB | 2023826-2022-02656 | 2022-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.5/4.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. (B)(4). | -11.5/4.0/086 | OS | Physician | 08/03/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15282228&pc=QCB | 2023826-2022-02657 | 2022-03-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.5/1.0/106 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS DID NOT RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. REPORTEDLY, PATIENT'S INTRAOCULAR ANATOMY IS ABNORMAL AND THE VAUL REMAINS LOW AFTER REPLACING THE LENS. SEE MFR # 2022-02658 FOR CLAIM TO RESOLVE PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -15.5/1.0/106 | OD | Physician | 08/03/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15282231&pc=QCB | 2023826-2022-02658 | 2022-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.5/1.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. MFR # (B)(4) FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | -15.5/1.0/096 | OD | Physician | 08/03/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15282234&pc=QCB | 2023826-2022-02653 | 2022-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.5/3.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, CLAIM# (B)(4). | -14.5/3.0/090 | OD | Physician | 08/03/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15282238&pc=QCB | 2023826-2022-02659 | 2022-04-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.0/1.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6). THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6). THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -14.0/1.5/088 | OS | Physician | 08/03/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15282239&pc=QCB | 2023826-2022-02660 | 2022-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.0/1.0/139 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. (B)(4). | -10.0/1.0/139 | OS | Physician | 08/03/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15282242&pc=QCB | 2023826-2022-02664 | 2022-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -09.5/1.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. (B)(4). Manufacturer Narrative: D4: EXPIRATION DATE (B)(6) 2025, SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). | -09.5/1.5/088 | OD | Physician | 08/03/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15282244&pc=QCB | 2023826-2022-02665 | 2022-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.5/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -11.5/2.5/092 | OS | Physician | 08/03/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15282247&pc=QCB | 2023826-2022-02670 | 2022-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -12.5/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. (B)(4). | -12.5/2.0/091 | OD | Physician | 08/03/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15282248&pc=QCB | 2023826-2022-02673 | 2022-06-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7, -12.0/1.5/118 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -12.0/1.5/118 | OS | Physician | 08/04/2022 | 2022 | 45 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15282249&pc=QCB | 2023826-2022-02661 | 2022-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -18.0/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. CLAIM# (B)(4). | -18.0/1.5/089 | OS | Physician | 08/03/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15282250&pc=QCB | 2023826-2022-02669 | 2022-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.5/3.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D4: EXPIRATION DATE: (B)(6) 2024. D9: DEVICE AVAILABLE FOR EVALUATION?: YES. RETURNED (B)(6) 2022. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: A2: AGE AT TIME OF EVENT, DATE OF BIRTH- SHOULD BE CORRECTED TO STATE: (B)(6) 1991. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. (B)(4). CLAIM# (B)(4). | -10.5/3.0/092 | OS | Physician | 08/03/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15282251&pc=QCB | 2023826-2022-02671 | 2022-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -14.5/3.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK, (B)(4). CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -14.5/3.5/087 | OS | Physician | 08/03/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15282255&pc=QCB | 2023826-2022-02672 | 2022-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -18.0/4.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D4: EXPIRATION DATE: 31-MAR-2025. D9: DEVICE AVAILABLE FOR EVALUATION?: YES. RETURNED: 21-OCT-2022. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. (B)(4). | -18.0/4.0/094 | OS | Physician | 08/04/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15283194&pc=QCB | 2023826-2022-02655 | 2022-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.0/1.5/109 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. CLAIM# (B)(4). | -8.0/1.5/109 | OS | Physician | 08/03/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15285522&pc=QCB | 2023826-2022-02881 | 2022-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM513.7 IMPLANTABLE COLLAMER LENS OF -6.5/2.5/11 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED INTO THE RIGHT EYE (OD) OF A PATIENT WITH PRE-EXISTING AMBLIOPIA ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REPOSITIONED DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. IT WAS THEN EXCHANGED FOR A SHORTER LENGTH LENS. THE REPORTER NOTED THAT SECONDARY INTERVENTION WAS NOT NECESSARY. Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. OFF-LABEL - PRE-EXISTING AMBLIOPIA. (B)(4). | -6.5/2.5/11 | OD | Physician | 07/26/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15285553&pc=QCB | 2023826-2022-02880 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -9.0/2/77 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. STATUS OF THE EYE IS REPORTED AS "RESOLVED." CAUSE OF THE EVENT IS USER ERROR. REPORTEDLY, "LENS TOO LONG." Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK. CLAIM# (B)(4). | -9.0/2/77 | OD | Physician | 07/26/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15285646&pc=QCB | 2023826-2022-02674 | 2022-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Improper or Incorrect Procedure or Method | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -08.0/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS USER ERROR. DEVICE DID NOT FAIL TO PERFORM AS INTENDED. REPORTEDLY, LENS TOO SHORT. Manufacturer Narrative: ADDITIONAL INFORMATION: D4:EXPIRATION DATE: 31-AUG-2024. D9: DEVICE AVAILABLE FOR EVALUATION?: YES. RETURNED :17-AUG-2022. H6: MEDICAL DEVICE PROBLEM CODE: 1494- OFF LABEL USE; ACD<3.00MM. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. CLAIM# (B)(4). | -08.0/1.5/092 | OS | Physician | 08/03/2022 | 2022 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15285781&pc=QCB | 2023826-2022-02522 | 2022-05-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF -5.50/5.50/113 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE REPORTED AS DEVICE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). | -5.50/5.50/113 | OD | Physician | 07/01/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15285963&pc=QCB | 2023826-2022-02523 | 2022-05-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF -5.50/6.0/019 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 . EXCESSIVE VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. CAUSE REPORTED AS DEVICE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). | -5.50/6.0/019 | OS | Physician | 07/01/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15294004&pc=QCB | 2023826-2022-02805 |  | Injury | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -8.5/1.0/084 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS USER ERROR; DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: H6 - MEDICAL DEVICE PROBLEM CODE: 1494 (OFF-LABEL USE: ACD<3.0MM) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -8.5/1.0/084 | OD | Physician | 07/26/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15296606&pc=QCB | 2023826-2022-02788 | 2022-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -07.00/+1.5/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED LATER THAT SAME DAY DUE TO EXCESSIVE VAULTING AND THE PROBLEM WAS RESOLVED. THE LENS WAS NOT REPLACED WITH ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: EXPIRATION DATE: 31-MAY-2025. NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -07.00/+1.5/087 | OD | Physician | 08/01/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15296615&pc=QCB | 2023826-2022-02792 | 2023-03-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED AN ICL IMPLANTABLE COLLAMER LENS, INTO THE PATIENTS LEFT EYE (OS). THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: B5: THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.00/+4.0/173 (SPHERE/CYLINDER/AXIS), ON (B)(6) 2022. H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CORRECTED DATA: D2: COMMON DEVICE NAME: PHAKIC TORIC INTRAOCULAR LENS, PRODUCT CODE: QCB. CLAIM # (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (B)(4) SHOULD BE ADDED TO THE INITAL MDR. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE AND RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE ON THE LENS WITH FIBERS ON THE LENS SURFACE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. B3 - DATE OF EVENT: 19-MAR-2022 SHOULD BE ADDED TO THE INITIAL MDR. B5 - "ON (B)(6) 2023 THE LENS WAS EXPLANTED DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE (B)(6) 2023 SHOULD BE ADDED TO INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4627 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | -07.00/+4.0/173 | OS | Physician | 08/02/2022 | 2022 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15296629&pc=QCB | 2023826-2022-02789 | 2022-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -08.00/+1.5/091 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. THE LENS WAS NOT REPLACED WITH ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, EXPIRATION DATE: 31-MAY-2025. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5- "THE REPORTER INDICATED THE SURGEON INSERTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -08.00/+1.5/091 (SPHERE/CYLINDER/AXIS) ON (B)(6) 2022. ON THE SAME DAY SEPARATE SURGERY THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. THE LENS WAS NOT REPLACED WITH ANOTHER ICL LENS. THE CAUSE OF THE EVENT WAS UNKNOWN." CLAIM# (B)(4). | -08.00/+1.5/091 | NA | Physician | 08/01/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15298359&pc=QCB | 2023826-2022-02907 | 2022-06-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 12.6MM, VTICM512.6 -13.5/+4.00/91 (SPHERE/CYLINDER/AXIS) IMPLATABLE COLLAMER LENS TORE/BROKE DURING LOADING. THERE WAS NO PATIENT CONTACT. A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. EXPIRATION DATE: 03-31-2025. CLAIM# (B)(4). | -13.5/+4.00/91 | NA | Physician | 07/26/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15298871&pc=QCB | 2023826-2022-02918 | 2022-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -6.0/0.5/176 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON THE SAME DATE AS IMPLANTATION; IN A SEPARATE SURGERY THE LENS WAS EXPLANTED. ON (B)(6) 2022 A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN AND BENT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: EXPIRATION DATE: 09-30-2024. (B)(4). | -6.0/0.5/176 | OD | Physician | 08/19/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15299000&pc=QCB | 2023826-2022-02921 | 2022-07-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -10.0/4.0/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS TORN. THE REPORTER STATED " EITHER DEFECT OR DURING LOAD ". NO PATIENT CONTACT. Manufacturer Narrative: EXPIRATION DATE: 12/31/2023. CLAIM#: (B)(4). Manufacturer Narrative: B5- PER UPDATED INFORMATION THE LENS WAS NOTED TO BE DAMAGED AFTER LOADING AND VIEWING IN CARTRIDGE UNDER MICROSCOPE. A BACKUP LENS WAS SUCCESSFULLY IMPLANTED. CLAIM # (B)(4). | -10.0/4.0/095 | NA | Physician | 07/26/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15299598&pc=QCB | 2023826-2022-02924 | 2022-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM013.2 IMPLANTABLE COLLAMER LENS OF -14.50/+4.50/127 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. A REPLACEMENT LENS OF LONGER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE WAS REPORTED AS UNKNOWN. IT WAS REPORTED THAT AT PATIENTS LAST VISIT IOP WAS 28MMHG AND THAT THIS IS DUE TO MEDICATION AND IS NOT RELATED TO THE ICL PROCEDURE. Manufacturer Narrative: EXPIRATION DATE: 09/30/2022. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED WITH MOISTURE IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). | -14.50/+4.50/127 | OD | Physician | 07/28/2022 | 2022 | 58 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15299684&pc=QCB | 2023826-2022-02929 | 2022-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -14.0/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: EXPIRATION DATE: 12/31/2024. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -14.0/2.5/092 | OS | Physician | 07/29/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15303042&pc=QCB | 2023826-2022-02926 | 2022-07-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 13.2MM, VTICM5_13.2, -8.00/2.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED AND REPLACED INTRAOPERATIVELY DUE TO THE DOCTOR NOTICING A MISSING HAPTIC AFTER IMPLANTING. IT WAS REPORTED THAT THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: EXPIRATION DATE: 12/31/2023. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED A VIAL WITH LIQUID. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM# (B)(4). | -8.00/2.0/090 | NA | Physician | 07/29/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15307046&pc=QCB | 2023826-2022-02885 | 2022-07-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 OF -7.50/1.0/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS FOUND TO BE TORN/BROKE DURING INJECTION DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THE LENS WAS NOT IMPLANTED.NO PATIENT CONTACT; THE LENS DID NOT TOUCH THE EYE. NO PATIENT INJURY. A REPLACEMENT LENS OF THE SAME MODEL/SIZE AND POWER WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: (B)(4). | -7.50/1.0/92 | OD | Physician | 07/28/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15307163&pc=QCB | 2023826-2022-02886 | 2021-12-31 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -8.50/1.50/65 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME SIZE SPHERICAL MODEL LENS DUE TO LENS ROTATION. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: B5 - THE LENS WAS REPOSITIONED ON 03 JAN 2022 WHICH DID NOT RESOLVE THE ISSUE. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -8.50/1.50/65 | OS | Physician | 07/28/2022 | 2022 | 41 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15307871&pc=QCB | 2023826-2022-02855 | 2022-03-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -08.50/+1.0/083 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: EXPIRATION DATE: 31-AUG-2024. (B)(4). CLAIM # (B)(4). | -08.50/+1.0/083 | OS | Physician | 08/08/2022 | 2022 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15307872&pc=QCB | 2023826-2022-02854 | 2022-03-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -08.50/+3.5/088 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPORTED AS LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2022 FOR ANOTHER SAME MODEL/LENGTH LENS. THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPIRATION DATE: 31-JUL-2023. EXPLANT DATE: NA. (B)(4). CLAIM # (B)(4). | -08.50/+3.5/088 | OD | Physician | 08/08/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15316814&pc=QCB | 2023826-2022-02957 | 2022-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO_13.2; -6/+1/114 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCE EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXPLANTED. LATER ON THIS SAME DATE A REPLACEMENT LENS OF SHORTER LENGTH/DIFFERENT MODEL WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED "THE RELATION BETWEEN THE W-W AND THE SULCUS WAS DIFFERENT THEN USUAL". Manufacturer Narrative: EXPIRATION DATE: 05-31-2025. CLAIM#: (B)(4). | -6/+1/114 | OD | Physician | 08/01/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15316845&pc=QCB | 2023826-2022-02958 | 2022-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -6/+0.5/78 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCE EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXPLANTED. LATER ON THIS SAME DATE A REPLACEMENT LENS OF SHORTER LENGTH/DIFFERENT MODEL WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED "THE RELATION BETWEEN THE W-W AND THE SULCUS WAS DIFFERENT THEN USUAL". Manufacturer Narrative: D4 - EXPIRATION DATE : 03-31-2025. CLAIM#: (B)(4). Manufacturer Narrative: EXPIRATION DATE: 05-31-2025. CLAIM#: (B)(4). | -6/+0.5/78 | OS | Physician | 08/01/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15316846&pc=QCB | 2023826-2022-02959 | 2022-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE PATIENT EXPERIENCED HIGH VAULTING AND NOTED "WRONG LENS CALCULATION IN OCOS". LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -7.00/+1.0/035 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022, THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS USER ERROR; WITH NO DEVICE RELATED ISSUES AND NOTED "USER SELECTED WRONG LENS POWER; OCOS WORKED WELL". CLAIM#: (B)(4). | -7.00/+1.0/035 | OD | Physician | 08/02/2022 | 2022 | 46 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15316966&pc=QCB | 2023826-2022-02947 | 2022-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5/2.0/013 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS OF THE SAME LENGTH AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: H3: PRODUCT EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: EXPIRATION DATE: 03/31/2025. CLAIM# (B)(4). | -11.5/2.0/013 | OD | Physician | 08/01/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15321169&pc=QCB | 2023826-2022-02858 | 2022-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -09.50/+2.5/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2019 THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: B5- "THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -09.50/+2.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SPHERICAL MODEL/SAME LENGTH LENS AND LRI WAS PERFORMED. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS UNKNOWN." H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPIRATION DATE: 31-OCT-2021. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.50/+2.5/091 | OD | Physician | 08/09/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15321178&pc=QCB | 2023826-2022-02857 | 2022-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -13.50/+5.0/088 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2022 THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: EXPIRATION DATE: 31-MAY-2025. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -13.50/+5.0/088 | OS | Physician | 08/09/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15321205&pc=QCB | 2023826-2022-02856 | 2022-06-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.00/+3.0/085 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2021 THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS REPLACED WITH A DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5- "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -3.0/+3.0/85 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED ROTATION NOT ASSOCIATED WITH A LOW VAULT. ON 17/JUNE/2021 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LENS OF THE SPHERICAL MODEL/ SAME LENGTH AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS UNKNOWN." DEVICE EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS IN A RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN AND BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK; ETHNICITY: UNK; RACE: UNK. EXPIRATION DATE: 31-OCT-2023. (B)(4). THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CLAIM # (B)(4). | -13.00/+3.0/085 | OD | Physician | 08/09/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15325100&pc=QCB | 2023826-2022-02974 | 2022-08-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7; -10.0/2.0/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS HAD TORE DURING LOADING. INTRAOPERATIVELY A BACK UP LENS OF THE SAME MODEL/LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THIS OCCURRED ON (B)(6) 2022. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. IT WAS UNKNOWN WHEN THE DAMAGE WAS NOTED. Manufacturer Narrative: EXPIRATION DATE: 10/31/2023. LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.0/2.0/093 | NA | Other Health Care Professional | 08/25/2022 | 2022 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15325284&pc=QCB | 2023826-2022-02962 | 2022-01-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -15.5/0.5/67 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. STATUS OF THE EYE IS REPORTED AS, "ALL FINE" AND "PATIENT IS HAPPY." Manufacturer Narrative: CLAIM#: (B)(4). | -15.5/0.5/67 | OD | Physician | 08/01/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15325286&pc=QCB | 2023826-2022-02961 | 2022-01-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -15.0/1.0/53 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. STATUS OF THE EYE IS REPORTED AS, "ALL FINE" AND "PATIENT IS HAPPY." Manufacturer Narrative: CLAIM#: (B)(4). | -15.0/1.0/53 | OS | Physician | 08/01/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15325378&pc=QCB | 2023826-2022-02964 | 2022-04-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-08-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.0/1.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED WITH LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(6). CLAIM# (B)(4). | -12.0/1.0/091 | OS | Physician | 08/02/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15334063&pc=QCB | 2023826-2022-02865 | 2022-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.5/1.0/133 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. EXPIRATION DATE IS 03/31/2025. CLAIM# (B)(4). | -12.5/1.0/133 | OD | Physician | 08/15/2022 | 2022 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15334064&pc=QCB | 2023826-2022-02863 | 2021-12-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.0/5.5/160 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED VERTICALLY WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: EXPIRATION DATE IS 08/31/2024. (B)(4). | -14.0/5.5/160 | OS | Physician | 08/14/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15334068&pc=QCB | 2023826-2022-02869 | 2022-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.5/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: EXPIRATION DATE IS 02/28/2025. (B)(4). | -13.5/2.0/090 | OS | Physician | 08/15/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15334069&pc=QCB | 2023826-2022-02866 | 2022-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.0/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. EXPIRATION DATE IS 03/31/2025. CLAIM# (B)(4). | -11.0/2.0/090 | OS | Physician | 08/15/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15334075&pc=QCB | 2023826-2022-02870 | 2022-02-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.50/2.0/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. EXPIRATION DATE IS 12/31/2024. (B)(4). | -09.50/2.0/071 | OD | Physician | 08/16/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15334076&pc=QCB | 2023826-2022-02867 | 2022-05-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -08.5/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. EXPIRATION DATE IS 05/31/2024. (B)(4). | -08.5/1.5/092 | OS | Physician | 08/15/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15334079&pc=QCB | 2023826-2022-02872 | 2022-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.0/3.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. EXPIRATION DATE IS 05/31/2024. CLAIM# (B)(4). | -12.0/3.0/083 | OD | Physician | 08/16/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15334080&pc=QCB | 2023826-2022-02871 | 2022-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.0/1.0/106 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: EXPIRATION DATE IS 03/31/2025. (B)(4). | -10.0/1.0/106 | OS | Physician | 08/16/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15337162&pc=QCB | 2023826-2022-02894 | 2022-06-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -15.50/2.0/026 (SPHERE/CYLINDER/AXIS) LENS TEAR/BREAK DURING LOADING (B)(6) 2022. THE LENS WAS NOT IMPLANTED INTO THE PATIENTS LEFT EYE (OS). THE BACKUP LENS LOADED AND IMPLANTED WITH NO ISSUE; PATIENT AND SURGEON SATISFIED WITH OUTCOME. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. " NOTED TEAR IN LENS FOOTPLATE ON LOADING." Manufacturer Narrative: EXPIRATION DATE: 31-MAR-2025. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -15.50/2.0/026 | OS | Physician | 08/02/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15342024&pc=QCB | 2023826-2022-02902 | 2022-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A TORIC 13.2MM SIZE IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN THE PATIENT'S RIGHT EYE. REPORTEDLY, "PT HAS BEEN NOTED TO HAVE A SHALLOW ACD SINCE SHORTLY AFTER SURGERY; I SAW HER TODAY AND NOTED VERY HIGH VAULT WITH SHALLOW AC." THE LENS WAS LATER EXCHANGED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS WITH A -17.0/2.5/087 (SPHERE/CYLINDER/AXIS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED DUE TO EXCESSIVE VAULTING. IN A SUBSEQUENT SURGERY LATER THAT DAY, A SMALLER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "THIS IS A CASE WHEREBY WTW MEASUREMENTS BASED ON 3 MEASUREMENT DEVICES WERE CONSISTENT PREOP. IT IS ASSUMED THAT THE ACTUAL SULCUS DIAMETER WAS SMALLER THAN PREDICTED BY WTW." STATUS OF THE EYE IS REPORTED AS, "DOING WELL. VAULT NORMALIZED. NOTE THAT THIS EYE IS AMBLYOPIC WITH 20/50 BCVA PRE-ICL." CLAIM # (B)(4). | -17.0/2.5/087 | OD | Physician | 07/28/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15342343&pc=QCB | 2023826-2022-02895 |  | Injury | STAAR SURGICAL COMPANY | 2022-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF -5.50/2.0/127 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED DUE TO SIZE. THE LENS WAS REPLACED ON AN UNKNOWN DATE WITH ANOTHER STAAR LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: DEVICE: (B)(6) 2024. NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | -5.50/2.0/127 | OD | Physician | 08/03/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15342351&pc=QCB | 2023826-2022-02896 |  | Injury | STAAR SURGICAL COMPANY | 2022-09-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF -7.50/2.0/048 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED DUE TO SIZE. THE LENS WAS REPLACED ON AN UNKNOWN DATE WITH ANOTHER STAAR LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED.Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF -7.50/2.0/048 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED DUE TO SIZE. THE LENS WAS REPLACED ON AN UNKNOWN DATE WITH ANOTHER STAAR LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(6) 2024. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF -7.50/2.0/048 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REMOVED DUE TO SIZE. THE LENS WAS REPLACED ON AN UNKNOWN DATE WITH ANOTHER STAAR LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. D6B: EXPLANT DATE-( B)(6) 2022. CLAIM# (B)(4). | -7.50/2.0/048 | OS | Physician | 08/03/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15345331&pc=QCB | 2023826-2022-03009 | 2022-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: REPORTEDLY, A 12.6MM VTICM5_12.6 -7.50/1.0/093 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS EYE ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 AND REPLACED WITH A SHORTER LENGTH LENS ON AN UNKNOWN DATE. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: EXPIRATION DATE: 03/31/2024. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). | -7.50/1.0/093 | NA | Other Health Care Professional | 08/03/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15345333&pc=QCB | 2023826-2022-02955 | 2022-05-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 12.6MM VTICMO12.6, -17.5/2.0/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS GOT STUCK IN CARTRIDGE DURING INJECTION AND TORE. THERE WAS PATIENT CONTACT WITH NO PATIENT INJURY. IN THE REPORTERS OPINION THE CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: EXPIRATION DATE: 01/31/2025. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -17.5/2.0/074 | NA | Physician | 08/04/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15351340&pc=QCB | 2023826-2022-03012 | 2022-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -4.50/5.5/076 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION AND DECENTRATION. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: (EXPIRATION DATE): 08/31/2023. CLAIM#: (B)(4). Manufacturer Narrative: B5- PER UPDATED INFORMATION THE PROBLEM RESOLVED, AND THE PATIENT IS FINE. H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A BROKEN/BENT HAPTIC. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | -4.50/5.5/076 | OD | Physician | 08/04/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15352430&pc=QCB | 2023826-2022-03021 | 2022-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS). LENS TEAR/ BREAK WAS REPORTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -14.0/1.5/123 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE; THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS WITH SIMILAR POWER AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. D4- (B)(6) 2024. H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -14.0/1.5/123 | OS | Physician | 08/03/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15352469&pc=QCB | 2023826-2022-03016 | 2022-07-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -14.50/3.0/088 (SPHERE/CYLINDER/AXIS) LENS TEAR/BREAK DURING LOADING ON (B)(6) 2022. THE LENS WAS NOT IMPLANTED INTO THE PATIENTS LEFT EYE (OS). NO PATIENT CONTACT REPORTED; LENS DID NOT TOUCH THE EYE. A REPLACEMENT LENS OF LONGER LENGTH WITH SIMILAR POWER WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE REPORTED AS UNKNOWN. "THE LENS WAS TEAR DURING LOADING." Manufacturer Narrative: EXPIRATION DATE: 28-EB-2025. (B)(4). | -14.50/3.0/088 | OS | Physician | 08/05/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15366720&pc=QCB | 2023826-2022-03041 | 2022-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7; -15.5/2.0/75 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS; BUT IMPLANTED VERTICALLY. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: EXPIRATION DATE: 05-31-2022. (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: 21-OCT-2022. H3-DEVICE EVALUATION: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN DEVICE SPECIFICATIONS. CLAIM# (B)(4). | -15.5/2.0/75 | OS | Physician | 08/31/2022 | 2022 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15366795&pc=QCB | 2023826-2022-03024 | 2022-08-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VICMO13.2 IMPLANTABLE COLLAMER LENS OF -14.0/3.5/88 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED UPSIDE DOWN INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED, AND REPLACED FOR A SAME SIZE, DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS REPORTED AS USER ERROR WITH NO PATIENT INJURY. REPORTEDLY, "LENS BROKE WHEN BEING EXPLANTED." STATUS OF THE EYE IS REPORTED AS "OK." Manufacturer Narrative: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -14.0/3.5/88 | OS | Physician | 08/04/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15369850&pc=QCB | 2023826-2022-03033 | 2022-06-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 13.2MM, VTICM5_13.2, -9.50/1.00/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE BEFORE LOADING. IT WAS REPORTED THAT THE DAMAGE WAS NOTED AFTER OPENING THE VIAL. THERE WAS NO PATIENT CONTACT. ON (B)(6) 2022 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS THE DEVICE; HOWEVER, IT WAS REPORTED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -9.50/1.00/107 | NA | Physician | 08/08/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15374929&pc=QCB | 2023826-2022-03027 |  | Injury | STAAR SURGICAL COMPANY | 2022-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A TORIC IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN THE PATIENT'S EYE. LOW VAULT WAS REPORTED AND A LENS EXCHANGED IS PLANNED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A VTICM5_12.1; -7.00/2.00/054 (SPHERE/DIAMETER/AXIS) DIOPTER IMPLANTABLE COLLAMER LENS WAS IMPLANTED IN THE PATIENT'S EYE ON (B)(6) 2021. LOW VAULT WAS REPORTED WITHOUT ROTATION. REPORTEDLY THE PATIENT DECIDED ON GOING BACK TO WEARING GLASSES. ON (B)(6) 2023 THE LENS WAS EXPLANTED. PROBLEM RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -7.00/2.00/054 | NA | Physician | 08/10/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15375203&pc=QCB | 2023826-2022-03035 | 2022-08-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: =THE REPORTER INDICATED THAT ON (B)(6) 2022 A 13.2MM, VTICMO13.2, -9.50/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING LOADING, THERE WAS NO PATIENT CONTACT. ON (B)(6) 2022 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS UNKNOWN. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -9.50/1.0/091 | NA | Physician | 08/09/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15375501&pc=QCB | 2023826-2022-03017 | 2022-04-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -7.5/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS USER ERROR; DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: MANUFACTURE DATE: 25-JAN-2025. CLAIM# (B)(4). | -7.5/2.0/089 | OD | Physician | 08/08/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15375508&pc=QCB | 2023826-2022-03018 | 2022-04-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -8.5/3.0/102 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS USER ERROR; DEVICE DID NOT FAIL TO PERFORM AS INTENDED. Manufacturer Narrative: EXPIRATION DATE: 28-FEB-2025. CLAIM# (B)(4). | -8.5/3.0/102 | OS | Physician | 08/08/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15381582&pc=QCB | 2023826-2022-03036 | 2022-07-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(4) 2022 A 13.2MM, VTICMO13.2, -10.50/3.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE BEFORE/DURING LOADING, THERE WAS NO PATIENT CONTACT. ON 30-JUL-2022 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.50/3.0/091 | NA | Physician | 08/09/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15386460&pc=QCB | 2023826-2022-03031 |  | Injury | STAAR SURGICAL COMPANY | 2022-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 WITH A, -11.50/2.0/071 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE. IT IS REPORTED, "LENS TO BIG THEY WILL REMOVE AND IMPLANT A SMALLER SIZE." Manufacturer Narrative: PATIENT INFORMATION UNK. WORK ORDER SEARCH FOUND NO SIMLIAR COMPLAINTS. CLAIM# (B)(4). Manufacturer Narrative: B5: LENS WAS EXCHANGED ON (B)(6) 2022 WITH A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. REPORTEDLY, "LOOKS GREAT." CAUSE WAS REPORTED AS PATIENT-RELATED FACTOR AND DEVICE FAILED TO PERFORM AS INTENDED. PATIENT DID NOT CONFORM TO SIZING NOMOGRAM. CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A ZIPLOC. VISUAL INSPECTION FOUND RESIDUE ON A HAPTIC TORN LENS. LENS MODEL WAS NOT A VTICM5. CLAIM# (B)(4). | -11.50/2.0/071 | NA | Physician | 08/12/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15386687&pc=QCB | 2023826-2022-03032 |  | Injury | STAAR SURGICAL COMPANY | 2022-09-08 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 WITH A, -10.50/2.0/094 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE. IT IS REPORTED, "LENS TOO BIG WILL REMOVE TO SMALLER." Manufacturer Narrative: B5 - LENS WAS EXCHANGED ON (B)(6) 2022 WITH A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. REPORTEDLY, "LOOKS GREAT." CAUSE WAS REPORTED AS PATIENT-RELATED FACTOR AND DEVICE FAILED TO PERFORM AS INTENDED. PATIENT DID NOT CONFORM TO SIZING NOMOGRAM. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A ZIPLOC BAG. VISUAL INSPECTION FOUND RESIDUE ON A NO VISIBLE DAMAGED LENS. CLAIM# (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNKNOWN. WORK ORDER SEARCH FOUND NO SIMLIAR COMPLAINTS. CLAIM# (B)(4). | -10.50/2.0/094 | NA | Physician | 08/12/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15392156&pc=QCB | 2023826-2022-03039 | 2022-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICM5_12.1, -11.5/3.0/101 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED AFTER ICL REMOVAL (DUE TO EXCESSIVE VAULT) OF THE FELLOW EYE. IT WAS REPORTED THAT THE PROBLEM RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND A BROKEN HAPTIC. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM #(B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.5/3.0/101 | OD | Physician | 08/12/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15395970&pc=QCB | 2023826-2022-03057 | 2022-06-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -16.0/2.0/083 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE. INTRAOPERATIVELY THE LENS WAS REMOVED WITH NO PATIENT INJURY. ON (B)(6) 2022 A REPLACEMENT LENS OF THE SAME LENGTH/MODEL AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -16.0/2.0/083 | OS | Physician | 09/05/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15395984&pc=QCB | 2023826-2022-03058 | 2022-03-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-09 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.0/4.5/010 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT IMPLANTED HORIZONTALLY. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9:20-DEC-2022. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT AND RESIDUE ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -10.0/4.5/010 | OD | Physician | 09/05/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15398212&pc=QCB | 2023826-2022-03062 | 2022-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -12.50/2.0/075 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON 01-AUG-2022. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF 12.50/2.0/075 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LOW VAULT WAS OBSERVED. ON (B)(4) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. PATIENT IS HAPPY. CLAIM#: (B)(4). Manufacturer Narrative: UNK. LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). | -12.50/2.0/075 | OS | Physician | 08/12/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15398213&pc=QCB | 2023826-2022-03061 | 2022-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -15.0/2.0/118 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -15.0/2.0/118 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LOW VAULT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. PATIENT IS HAPPY. CLAIM# (B)(4). Manufacturer Narrative: UNK. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). | -15.0/2.0/118 | OD | Physician | 08/12/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15407313&pc=QCB | 2023826-2022-03065 | 2022-07-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.0/1.0/112 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT IMPLANTED HORIZONTALLY; INSTEAD OF VERTICALLY LIKE THE INITIAL LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN, BENT, AND DEFORMED TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | -9.0/1.0/112 | OD | Physician | 09/05/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15407353&pc=QCB | 2023826-2022-03068 | 2022-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -8.0/2.0/172 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXPLANTED. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9:20-DEC-2022. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT AND RESIDUE ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | -8.0/2.0/172 | OS | Physician | 09/08/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15414370&pc=QCB | 2023826-2022-03088 |  | Injury | STAAR SURGICAL COMPANY | 2022-09-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -13.5/3.5/93 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4). | -13.5/3.5/93 | OD | Physician | 09/07/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15414371&pc=QCB | 2023826-2022-03107 | 2022-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -13.5/+1.5/85 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING WITH ROTATION. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH/DIFFERENT MODEL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM# (B)(4). | -13.5/+1.5/85 | OS | Physician | 09/07/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15414372&pc=QCB | 2023826-2022-03106 | 2022-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -13.5/+1.5/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING WITH ROTATION. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH/DIFFERENT MODEL LENS AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4). | -13.5/+1.5/92 | OD | Physician | 09/07/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15418191&pc=QCB | 2023826-2022-02979 | 2022-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.5/2.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: EXPIRATION DATE IS 02/28/2025. (B)(4). | -14.5/2.5/095 | OD | Physician | 08/22/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15418192&pc=QCB | 2023826-2022-02980 | 2021-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.0/2.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(4) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -15.0/2.5/094 | OD | Physician | 08/22/2022 | 2022 | 17 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15418193&pc=QCB | 2023826-2022-02978 | 2022-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -11.0/2.50/120 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. REPORTEDLY, EYE QUIET, LENS IN POSITION, AND PATIENT, SURGEON HAPPY. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. CLAIM# (B)(4). | -11.0/2.50/120 | OD | Physician | 08/22/2022 | 2022 | 47 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15418195&pc=QCB | 2023826-2022-02983 | 2022-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -08.5/0.5/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. DEVICE PROBLEM CODE: (B)(4). CLAIM # (B)(4). | -08.5/0.5/107 | OD | Physician | 08/22/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15418197&pc=QCB | 2023826-2022-02986 | 2022-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -09.5/1.5/014 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -09.5/1.5/014 | OS | Physician | 08/21/2022 | 2022 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15418198&pc=QCB | 2023826-2022-02984 | 2022-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -09.5/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -09.5/2.5/093 | OD | Physician | 08/22/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15418200&pc=QCB | 2023826-2022-02987 | 2022-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/2.0/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTION DATA; B5 - REPORTED AS THE SPHERE OF THE LENS "081" - CORRECT TO "84." CLAIM# (B)(4). | -11.0/2.0/084 | OS | Physician | 08/22/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15418201&pc=QCB | 2023826-2022-02991 | 2022-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.5/2.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -09.5/2.5/090 | OD | Physician | 08/24/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15418204&pc=QCB | 2023826-2022-02993 | 2022-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.5/1.5/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | -09.5/1.5/075 | OS | Physician | 08/24/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15418205&pc=QCB | 2023826-2022-02982 | 2022-02-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -07.0/1.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -07.0/1.0/094 | OD | Physician | 08/22/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15418206&pc=QCB | 2023826-2022-02981 | 2022-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.5/3.0/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -16.5/3.0/079 | OS | Physician | 08/22/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15418207&pc=QCB | 2023826-2022-02994 | 2022-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.0/1.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.0/1.5/088 | OD | Physician | 08/24/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15418208&pc=QCB | 2023826-2022-02998 | 2022-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -08.5/1.0/106 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. DEVICE PROBLEM CODE: (B)(4). CLAIM # (B)(4). | -08.5/1.0/106 | OD | Physician | 08/24/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15418210&pc=QCB | 2023826-2022-02997 | 2022-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.0/2.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.0/2.0/086 | OD | Physician | 08/24/2022 | 2022 | 17 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15418213&pc=QCB | 2023826-2022-02990 | 2022-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -07.0/1.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -07.0/1.5/086 | OD | Physician | 08/24/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15418214&pc=QCB | 2023826-2022-03000 | 2022-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.5/2.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY:UNK. RACE:UNK. (B)(4). | -14.5/2.5/087 | OS | Physician | 08/25/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15418216&pc=QCB | 2023826-2022-03004 | 2022-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.5/3.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS ON (B)(6) 2022 DUE TO LOW VAULT. THIS RESOLVE THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. SEE MFR #2022-03003 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. (B)(4). CLAIM# (B)(4). | -15.5/3.5/088 | OS | Physician | 08/26/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15418219&pc=QCB | 2023826-2022-02996 | 2022-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.0/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. DEVICE PROBLEM CODE: (B)(4). CLAIM # (B)(4). | -11.0/2.0/089 | OS | Physician | 08/24/2022 | 2022 | 17 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15418222&pc=QCB | 2023826-2022-03003 | 2022-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -15.5/3.5/079 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULT AND REPLACED WITH A LONGER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2022. THIS DID NOT RESOLVE THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. SEE MFR #2022-03004 FOR CLAIM THAT RESOLVED PROBLEM. Manufacturer Narrative: DEVICE EVALUATION: H3- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDU/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. (B)(4). | -15.5/3.5/079 | OS | Physician | 08/26/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15418223&pc=QCB | 2023826-2022-03002 | 2022-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -15.0/3.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. (B)(4). CLAIM# (B)(4). | -15.0/3.0/089 | OD | Physician | 08/26/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15423887&pc=QCB | 2023826-2022-03077 | 2022-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -8.5/1.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LOW VAULT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -8.5/1.0/093 | OD | Physician | 08/22/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15428378&pc=QCB | 2023826-2022-03150 | 2022-08-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 THE DOCTOR FOUND A 13.2MM VTICMO13.2, -15.0/+2.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS, WITHOUT AXIS ALIGNMENT MARKS AFTER IMPLANTING THE LENS INTO THE PATIENTS RIGHT EYE (OD). LENS REMAINS IMPLANTED. PATIENT CONDITION IS NORMAL, IT WAS REPORTED THAT THE POSTOPERATIVE FOLLOW-UP HAD NORMAL VISUAL ACUITY AND NO REFRACTIVE ABNORMALITY. CAUSE OF EVENT WAS REPORTED TO BE THE DEVICE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#(B)(4). Manufacturer Narrative: H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). | -15.0/+2.5/094 | OD | Physician | 08/16/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15428628&pc=QCB | 2023826-2022-03093 | 2022-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+2.0/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -13.0/+2.0/090 | OS | Physician | 08/22/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15428629&pc=QCB | 2023826-2022-02936 | 2022-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.0/178 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -7.0/+1.0/178 | OS | Physician | 08/17/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15428632&pc=QCB | 2023826-2022-03095 | 2022-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+1.5/075 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). | -13.0/+1.5/075 | OD | Physician | 08/22/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15428686&pc=QCB | 2023826-2022-02937 | 2022-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/+2.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -14.5/+2.0/091 | OD | Physician | 08/17/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15428694&pc=QCB | 2023826-2022-02941 | 2022-06-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.0/+1.5/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE CODE (B)(4). CLAIM# (B)(4). | -16.0/+1.5/090 | OS | Physician | 08/18/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15428701&pc=QCB | 2023826-2022-03091 | 2022-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+1.0/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: D4 - UNIQUE IDENTIFIER (UDI)# : (B)(4), SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: DEVICE CODE (B)(4). CLAIM# (B)(4). | -7.5/+1.0/093 | OS | Physician | 08/22/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15428705&pc=QCB | 2023826-2022-02940 | 2022-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.0/+0.5/000 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE LENS WAS LATER REPLACED WITH A SHORTER LENGTH LENS ON (B)(6) 2022 AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D4: UNIQUE IDENTIFIER: (B)(4). D9: DEVICE AVAILABLE FOR EVALUATION? YES. RETURNED: 21-OCT-2022. CLAIM#: (B)(4). | -5.0/+0.5/000 | OS | Physician | 08/19/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15428707&pc=QCB | 2023826-2022-03092 | 2022-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/+1.5/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D4: UNIQUE IDENTIFIER: (B)(4). CLAIM#: (B)(4). | -6.0/+1.5/090 | OD | Physician | 08/22/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15428721&pc=QCB | 2023826-2022-02935 | 2022-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+1.5/015 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: DEVICE CODE (B)(4). CLAIM# (B)(4). | -7.5/+1.5/015 | OD | Physician | 08/17/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15428725&pc=QCB | 2023826-2022-03089 | 2022-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.5/+1.0/088 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -4.5/+1.0/088 | OD | Physician | 08/18/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15433328&pc=QCB | 2023826-2022-03153 | 2022-08-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -11.5/1.5/129 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. WHILE IMPLANTING THE LENS; THE LENS HAD BECAME TORN DUE TO A CARTRIDGE ISSUE. IT WAS NOTED "LENS WAS CAUGHT BETWEEN THE FTP AND THE INNER WALL OF THE CARTRIDGE AND DAMAGED". THERE WAS NO PATIENT INJURY. THE LENS REMAINS IMPLANTED. THE PROBLEM IS NOT RESOLVED. THE REPORTER INDICATED THE PATIENT WILL BE MONITORED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND ADDITIONALLY IT WAS NOTED "LENS EJECTS SIDEWAYS; PART OF THE LENS GETS CAUGHT IN THE PLUNGER AND THE LENS DOES NOT COME OUT SMOOTHLY". Manufacturer Narrative: FIVE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#(B)(4). | -11.5/1.5/129 | OS | Physician | 08/23/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15433348&pc=QCB | 2023826-2022-03155 | 2022-08-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8.50/3.00/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. WHILE IMPLANTING THE LENS; THE LENS HAD BECAME TORN DUE TO A CARTRIDGE ISSUE. IT WAS NOTED "LENS WAS CAUGHT BETWEEN THE FTP AND THE INNER WALL OF THE CARTRIDGE AND DAMAGE; THE LENS FLOWING BACKWARDS OR BEHAVING ERRATICALLY". ADDITIONALLY THE LENS WAS INVERTED WHEN INJECTED. THERE WAS NO PATIENT INJURY. THE LENS REMAINS IMPLANTED. THE PROBLEM IS NOT RESOLVED. THE REPORTER INDICATED THE PATIENT WILL BE MONITORED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND ADDITIONALLY IT WAS NOTED "LENS EJECTS SIDEWAYS; PART OF THE LENS GETS CAUGHT IN THE PLUNGER AND THE LENS DOES NOT COME OUT SMOOTHLY". Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: 4110 - DEVICE WORK ORDER SEARCH: FIVE SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -8.50/3.00/95 | OD | Physician | 08/23/2022 | 2022 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15433388&pc=QCB | 2023826-2022-03156 | 2022-08-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-15 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON ATTEMPTED TO IMPLANT A 12.6MM VTICM5_12.6; -7.5/+3.0/79 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE WITH NO PATIENT CONTACT. INTRAOPERATIVELY A REPLACEMENT LENS OF THE SAME LENGTH/MODEL WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM# (B)(4). | -7.5/+3.0/79 | OS | Physician | 09/07/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15440840&pc=QCB | 2023826-2022-03147 | 2022-07-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 12.6MM, VTICMO12.6, -12.5/2.5/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE RIGHT EYE (OD). THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY, WITH NO PATIENT INJURY. A REPLACEMENT LENS WAS IMPLANTED AT A LATER DATE AND THE PROBLEM RESOLVED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT ON (B)(6) 2022 A 12.6MM, VTICMO12.6, -12.5/2.5/100 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS GOT STUCK IN THE CARTRIDGE OR INJECTION SYSTEM. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY, WITH NO PATIENT INJURY. ON (B)(6) 2022 A REPLACEMENT LENS WAS IMPLANTED, AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. CLAIM # (B)(4). | -12.5/2.5/100 | OD | Physician | 08/17/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15441551&pc=QCB | 2023826-2022-03080 | 2022-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -7.5/1.0/86 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LOW VAULT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. STATUS OF EYE: " MODERATE POST SURGERY EDEMA, SURGEON AWARE, PATIENT FOLLOW UP IN 1 WEEK." Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -7.5/1.0/86 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LOW VAULT WITH ROTATION WAS OBSERVED. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CLAIM #(B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -7.5/1.0/86 | OS | Physician | 08/24/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15442277&pc=QCB | 2023826-2022-03152 | 2022-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -11.0/+3.5/94 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS OF SAME LENGTH. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: CORRECTED DATA: H6- MEDICAL DEVICE PROBLEM CODE: 1494- OFF LABEL USE (AGE<21 AT TIME OF IMPLANT). CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT IN A MICRO-CENTRIFUGE VIAL. VISUAL INSPECTION FOUND A DEFORMED OPTIC AND HAPTIC. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -11.0/+3.5/94 | OD | Physician | 08/19/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15442545&pc=QCB | 2023826-2022-03179 | 2022-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -14.0/3.0/95 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING WITH ROTATION. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6 - "MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT)" SHOULD BE REMOVED FROM THE INITIAL MDR FOR CORRECTION. | -14.0/3.0/95 | OD | Physician | 10/10/2022 | 2022 | 47 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15442550&pc=QCB | 2023826-2022-03178 | 2022-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-16 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -11.50/+1.00/122 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS HAD TORE DURING INJECTION INTO THE EYE; THE LENS WAS STUCK IN THE CARTRIDGE OR INJECTION SYSTEM. THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. INTRAOPERATIVELY THE LENS WAS REMOVED AND AN INCISION ENLARGEMENT WAS ALSO PERFORMED. THE CAUSE OF THE EVENT WAS REPORTED AS USER ERROR. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | -11.50/+1.00/122 | OD | Physician | 09/05/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15449607&pc=QCB | 2023826-2022-03099 | 2022-07-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+1.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -7.5/+1.0/093 | OD | Physician | 08/22/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15449613&pc=QCB | 2023826-2022-03104 | 2022-08-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+0.5/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). | -11.0/+0.5/092 | OD | Physician | 08/24/2022 | 2022 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15449614&pc=QCB | 2023826-2022-03105 | 2022-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.5/087 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). | -10.5/+1.5/087 | OS | Physician | 08/24/2022 | 2022 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15449617&pc=QCB | 2023826-2022-03160 | 2022-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+3.0/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). | -11.5/+3.0/090 | OD | Physician | 08/25/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15449648&pc=QCB | 2023826-2022-03098 | 2022-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.0/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4), SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -7.0/+1.0/090 | OS | Physician | 08/22/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15449682&pc=QCB | 2023826-2022-03161 | 2022-06-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.5/088 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON 27-JUN-2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - 00841542125826 SHOULD BE ADDED TO THE INITIAL MDR. | -9.5/+1.5/088 | OD | Physician | 08/25/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15449696&pc=QCB | 2023826-2022-03100 | 2022-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+3.0/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -11.5/+3.0/092 | OD | Physician | 08/22/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15449701&pc=QCB | 2023826-2022-03097 | 2022-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0/+2.0/156 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -15.0/+2.0/156 | OD | Physician | 08/22/2022 | 2022 | 43 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15449725&pc=QCB | 2023826-2022-03101 | 2022-05-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+2.0/086 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -8.0/+2.0/086 | OS | Physician | 08/22/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15449735&pc=QCB | 2023826-2022-03158 | 2022-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+2.0/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -13.0/+2.0/090 | OD | Physician | 08/25/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15452272&pc=QCB | 2023826-2022-03197 | 2022-06-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -10.0/3.5/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2022 THE LENS WAS REPOSITIONED WHICH WAS NOTED AS RESOLVING THE PROBLEM; BUT THE LENS WAS STILL EXCHANGED WITH A SAME LENGTH BUT DIFFERENT MODEL LENS (SPHERICAL) AND THE PROBLEM WAS RESOLVED. IT WAS INDICATED "THE PATIENT REQUESTED SURGERY EVEN AFTER FULL COMMUNICATION BEFORE SURGERY. THE LENS ROTATION OCCURRED AFTER OPERATION; THE VICMO LENS WAS REPLACED AFTER REPOSITIONING. THE EYES REACHED 20/16 AND 20/12 AFTER THE OPERATION; THE PATIENT WAS QUITE SATISFIED". THE CAUSE OF THE EVENT WAS REPORTED AS THE PATIENT RELATED FACTOR; AND NOTED "SENSORY LIGAMENT WAS SLACK". Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -10.0/3.5/087 | OS | Physician | 09/06/2022 | 2022 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15452273&pc=QCB | 2023826-2022-03198 | 2022-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -12.0/2.5/111 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2022 THE LENS WAS REPOSITIONED WHICH WAS NOTED AS RESOLVING THE PROBLEM; BUT THE LENS WAS STILL EXCHANGED WITH A SAME LENGTH BUT DIFFERENT MODEL LENS (SPHERICAL) AND THE PROBLEM WAS RESOLVED. IT WAS INDICATED "THE PATIENT REQUESTED SURGERY EVEN AFTER FULL COMMUNICATION BEFORE SURGERY. THE LENS ROTATION OCCURRED AFTER OPERATION; THE VICMO LENS WAS REPLACED AFTER REPOSITIONING. THE EYES REACHED 20/16 AND 20/12 AFTER THE OPERATION; THE PATIENT WAS QUITE SATISFIED". THE CAUSE OF THE EVENT WAS REPORTED AS THE PATIENT RELATED FACTOR; AND NOTED "SENSORY LIGAMENT WAS SLACK". Manufacturer Narrative: CLAIM#:(B)(4). Manufacturer Narrative: D9: 07-NOV-2022. H3-DEVICE EVALUATION: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -12.0/2.5/111 | OD | Physician | 09/06/2022 | 2022 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15458791&pc=QCB | 2023826-2022-03115 | 2022-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.5/0.5/113 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT. A SHORTER LENGTH LENS WAS IMPLANTED IN A SECOND SURGERY ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -14.5/0.5/113 | OD | Physician | 08/29/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15458792&pc=QCB | 2023826-2022-03110 | 2022-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.00/3.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -16.00/3.0/095 | OD | Physician | 08/29/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15458795&pc=QCB | 2023826-2022-03119 | 2022-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.0/6.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.0/6.0/093 | OD | Physician | 08/30/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15458800&pc=QCB | 2023826-2022-03127 | 2022-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.0/1.5/067 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHINICITY: UNKNOWN, RACE: UNKNOWN. CLAIM# (B)(4). | -12.0/1.5/067 | OD | Physician | 09/01/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15458807&pc=QCB | 2023826-2022-03121 | 2022-03-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -10.5/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/1.5/089 | OD | Physician | 08/31/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15458811&pc=QCB | 2023826-2022-03124 | 2022-07-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6, -6.50/1.0/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A LONGER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S RIGHT EYE (OD) DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -6.50/1.0/099 | OD | Physician | 09/01/2022 | 2022 | 42 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15458821&pc=QCB | 2023826-2022-03111 | 2022-06-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -09.0/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -09.0/1.5/093 | OS | Physician | 08/29/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15458825&pc=QCB | 2023826-2022-03112 | 2022-06-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.5/1.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -11.5/1.5/087 | OD | Physician | 08/28/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15458827&pc=QCB | 2023826-2022-03117 | 2022-07-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -5.0/2.5/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE PROBLEM CODE: (B)(4) - OFF-LABEL USE, ACD<3.0MM. CLAIM# (B)(4). | -5.0/2.5/077 | OD | Physician | 08/30/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15458829&pc=QCB | 2023826-2022-03118 | 2022-07-14 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -6.5/2.0/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN.. Manufacturer Narrative: CLAIM# (B)(4). | -6.5/2.0/103 | OS | Physician | 08/30/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15458832&pc=QCB | 2023826-2022-03123 | 2022-08-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6, -12.0/1.0/035 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A LONGER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S RIGHT EYE (OD) DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -12.0/1.0/035 | OD | Physician | 09/01/2022 | 2022 | 43 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15458837&pc=QCB | 2023826-2022-03128 | 2022-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.0/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -12.0/1.5/089 | OS | Physician | 09/01/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15458864&pc=QCB | 2023826-2022-03114 | 2022-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/3.0/161 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN. CLAIM# (B)(4). | -11.0/3.0/161 | OD | Physician | 08/29/2022 | 2022 | 44 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15458865&pc=QCB | 2023826-2022-03130 | 2022-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/1.5/179 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY:UNKNOWN RACE: UNKNOWN. CLAIM# (B)(4). | -10.5/1.5/179 | OS | Physician | 09/01/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15459771&pc=QCB | 2023826-2022-03237 | 2022-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1; -12.50/1.5/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS "SMALL LENS". Manufacturer Narrative: (B)(4). | -12.50/1.5/093 | OD | Physician | 09/01/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15462027&pc=QCB | 2023826-2022-03137 | 2022-08-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5/2.5/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS DUE TO EXCESSIVE VAULT. REPORTEDLY, LENS WAS IMPLANTED VERTICALLY. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY. VISUAL INSPECTION FOUND THE LENS HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. OFF LABEL USE- ACD <3.0. CLAIM# (B)(4). | -9.5/2.5/095 | OD | Physician | 08/22/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15465149&pc=QCB | 2023826-2022-03168 | 2022-07-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.5/111 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D4: UNIQUE IDENTIFIER: (B)(4). D9: DEVICE AVAILABLE FOR EVALUATION?: YES. RETURNED 25-OCT-2022. CLAIM#(B)(4). Manufacturer Narrative: (B)(4). | -10.00/+1.5/111 | OD | Physician | 08/22/2022 | 2022 | 17 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15465181&pc=QCB | 2023826-2022-03171 | 2022-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -11.50/+1.0/055 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2022 BUT THE PROBLEM WAS NOT RESOLVED AND THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT: UNK. PATIENT ETHNICITY: UNK. PATIENT RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: PRODUCT EVALUATION: LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL OBSERVATION FOUND THE LENS HAPTIC BROKEN, BENT AND DEFORMED. DIMENSIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -11.50/+1.0/055 | OS | Physician | 08/24/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15467986&pc=QCB | 2023826-2022-03139 | 2022-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -12.50/2.0/107 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED AND A LARGER LENGTH LENS WAS IMPLANTED DUE TO TOP ROTATION. Manufacturer Narrative: B5: REPORTEDLY, "PATIENT IS DOING FINE." CLAIM #(B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS. (B)(4). | -12.50/2.0/107 | NA | Physician | 08/24/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15468396&pc=QCB | 2023826-2022-03253 | 2022-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-22 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -11.5/1.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS THAT WAS IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -11.5/1.0/092 | OD | Physician | 08/23/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15469690&pc=QCB | 2023826-2022-03186 | 2022-08-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS WITH DIOPTER -11.0/+1.5/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. IN A SEPARATE SURGERY THAT SAME DAY THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - D4: UNIQUE IDENTIFIER (UDI) # (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | -11.0/+1.5/094 | OD | Physician | 08/31/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15469697&pc=QCB | 2023826-2022-03187 | 2022-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.5/089 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). | -9.5/+1.5/089 | OD | Physician | 09/01/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15469913&pc=QCB | 2023826-2022-03189 | 2022-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+1.5/073 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -13.5/+1.5/073 | OS | Physician | 09/01/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15470250&pc=QCB | 2023826-2022-03183 | 2022-08-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.5/074 INTO THE PATIENT'S LEFT EYE (OS) ON 17-AUG-2022. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+1.5/074 | OS | Physician | 08/29/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15470330&pc=QCB | 2023826-2022-03188 | 2022-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+2.0/088 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# (B)(4), SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4), Manufacturer Narrative: CLAIM# (B)(4). | -13.5/+2.0/088 | OD | Physician | 09/01/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15474139&pc=QCB | 2023826-2022-03245 | 2022-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-23 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.0/3.5/008 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED AND IN A SUBSEQUENT SURGERY LATER THAT DAY, A SAME SIZE LENS WAS IMPLANTED DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK. OFF LABEL USE- ACD <3.0. CLAIM# (B)(4). | -11.0/3.5/008 | OS | Physician | 08/26/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15477465&pc=QCB | 2023826-2022-03262 | 2022-08-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -8.0/1.0/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS IMPLANTED, REMOVED AND REPLACED WITH A SHORTER LENGTH LENS INRAOPERATIVELY DUE TO LOW VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -8.0/1.0/088 | OS | Physician | 08/26/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15477978&pc=QCB | 2023826-2022-03259 | 2022-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -6.0/0.5/165 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS OF DIFFERENT POWER THAT WAS IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT/DEFORMED TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -6.0/0.5/165 | OD | Physician | 08/25/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15478058&pc=QCB | 2023826-2022-03172 | 2022-08-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON REPORTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.50/+3.0/093 (SPHERE/CYLINDER/AXIS) WAS MISSING THE LENS POSITIONING MARKINGS. THE LENS WAS IMPLANTED ON (B)(6) 2022 AND AT ONE DAY POST-OP, SHOWED GOOD VISUAL ACUITY (THE PATIENT SEES 1.0). THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION: D4: UNIQUE IDENTIFIER NUMBER: (B)(4). CLAIM #(B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM #: (B)(4). | -07.50/+3.0/093 | NA | Physician | 08/24/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15478238&pc=QCB | 2023826-2022-03246 | 2022-06-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM513.7 IMPLANTABLE COLLAMER LENS OF -12.5/2.0/174 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED. REPORTEDLY, "THE PATIENT EXPERIENCED HYPER VAULT. I USED THE ICL LENGTH SUGGESTED BY THE SOFTWARE, THE OPTIC CALCULATION WAS INCORRECT" Manufacturer Narrative: (B)(4). | -12.5/2.0/174 | OS | Physician | 08/26/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15478885&pc=QCB | 2023826-2022-03258 | 2022-08-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.5/+1.5/073 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. PATIENT EXPERIENCED EXCESSIVE VAULT, LENS REMAINS IMPLANTED. PER UPDATED INFORMATION INTERVENTION IS NOT PLANNED, PATIENT WILL BE MONITORED CLOSELY. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.5/+1.5/073 | OD | Physician | 08/25/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15479500&pc=QCB | 2023826-2022-03255 | 2022-08-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 12.6MM, VTICMO12.6, -5.5/0.5/114 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE RIGHT EYE (OD). THE LENS WAS IMPLANTED, REMOVED AND REPLACED INTRAOPERATIVELY WITH NO PATIENT INJURY AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT THE REPLACEMENT LENS WAS IMPLANTED IN VERTICAL POSITION. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). | -5.5/0.5/114 | OD | Physician | 08/24/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15480923&pc=QCB | 2023826-2022-03283 | 2022-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: REPORTEDLY, "UNSTERILE." Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: B5: THE REPORTER STATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS WITH A, -7.0/2.0/114 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "PATIENT IS HAPPY." CLAIM# (B)(4). Manufacturer Narrative: INITIAL MEDWATCH SUBMISSION: B1 SHOULD HAVE BEEN 'PRODUCT PROBLEM (MALFUNCTION)' B2 IS NOT APPLICABLE FIRST SUPPLEMENTAL MEDWATCH SUBMISSION: B1: ADVERSE EVENT B2: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE CLAIM# (B)(4). | -7.0/2.0/114 | OD | Physician | 08/26/2022 | 2022 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15491969&pc=QCB | 2023826-2022-03317 | 2022-09-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT WHILE THE SURGEON ATTEMPTED TO IMPLANT A 12.6MM VTICMO12.6; -9.0/1.5/123 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) THE LENS TORE DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT WITH NO PATIENT INJURY. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A BACK UP LENS OF THE SAME LENGTH/MODEL. THIS OCCURRED ON (B)(6) 2022. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS USER ERROR WITH DEVICE CAUSALITY; IT WAS NOTED "WHEN TRYING TO FIX PLACEMENT BEGAN TO BRING IT OUT TO REINSERT IT AND IT TORE; THEN PLACED BACK UP LENS". Manufacturer Narrative: CORRECTED DATA: B5- " THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/1.5/123 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 WHEN THE LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE EYE. REPORTEDLY "THE ICL WAS IMPLANTED UPSIDE DOWN AND DURING THE PROCESS OF REMOVAL THE ICL THEN TORE." THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. THE LENS WAS IMPLANTED; REMOVED AND REPLACED INTRAOPERATIVELY WITH THE SAME MODEL/ LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS REPORTED AS "MILD CORNEAL EDEMA BUT RESOLVING PLANT TO SEE HIM AGAIN TOMORROW". THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS USER ERROR AND THE DEVICE FAILED TO PERFORM AS INTENDED. REPORTEDLY "WHEN TRYING TO FIX PLACEMENT BEGAN TO BRING IT OUT TO REINSERT IT AND IT TORE THEN PLACED BACK UP LENS". CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- "THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS" SHOULD BE CORRECTED TO "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS" CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -9.0/1.5/123 | OD | Other Health Care Professional | 09/15/2022 | 2022 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15495515&pc=QCB | 2023826-2022-02631 | 2022-07-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -11.0-2.5/089 (SPHERE/CYLINDER/AXIS) TORE/ BROKE DURING INJECTION/ DELIVERY INTO THE PATIENT'S RIGHT EYE (OD). THE LENS WAS IMPLANTED; REMOVED AND REPLACED INTRA-OPERATIVELY ON (B)(6) 2022. A REPLACEMENT LENS OF THE SAME MODEL AND SIZE LENS AND THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: H6-DEVICE CODE: 2993 IN PREVIOUS MDR IS CORRECTED TO 3189. CLAIM# (B)(4). | -11.0-2.5/089 | OD | Physician | 07/08/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15495983&pc=QCB | 2023826-2022-02797 | 2022-06-16 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -13.0/1.5/080 (SPHERE/CYLINDER/AXIS) TORE/ BROKE DURING INJECTION/ DELIVERY INTO THE PATIENT'S RIGHT EYE (OD). ON (B)(6) 2022 THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY. LATER ON (B)(6) 2022 A REPLACEMENT LENS OF THE SAME MODEL AND SIZE WAS IMPLANTED. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.0/1.5/080 | OD | Physician | 07/20/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496571&pc=QCB | 2023826-2022-03200 | 2022-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.0/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -11.0/1.5/092 | OS | Physician | 09/05/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496577&pc=QCB | 2023826-2022-03207 | 2022-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.5/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.5/1.0/089 | OD | Physician | 09/05/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496591&pc=QCB | 2023826-2022-03216 | 2022-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. CLAIM# (B)(4). | -10.5/1.5/093 | OD | Physician | 09/05/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496602&pc=QCB | 2023826-2022-03221 | 2022-06-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.5/1.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.5/1.5/082 | OS | Physician | 09/05/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496610&pc=QCB | 2023826-2022-03227 | 2022-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.0/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -18.0/1.0/091 | OD | Physician | 09/06/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496618&pc=QCB | 2023826-2022-03204 | 2022-02-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.5/1.5/071 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). | -11.5/1.5/071 | OS | Physician | 09/05/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496621&pc=QCB | 2023826-2022-03201 | 2022-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.5/1.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -09.5/1.5/086 | OS | Physician | 09/05/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496638&pc=QCB | 2023826-2022-03205 | 2022-07-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.5/1.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFTHT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. (B)(4). | -11.5/1.0/088 | OS | Physician | 09/05/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496652&pc=QCB | 2023826-2022-03208 | 2022-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.0/1.5/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4), UNDER 21 YRS OF AGE AT DATE OF IMPLANT. CLAIM # (B)(4). | -13.0/1.5/072 | OD | Physician | 09/05/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496654&pc=QCB | 2023826-2022-03210 | 2022-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.5/1.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.5/1.5/082 | OD | Physician | 09/05/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496672&pc=QCB | 2023826-2022-03199 | 2022-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.5/2.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. CLAIM# (B)(4). | -15.5/2.5/090 | OD | Physician | 09/05/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496688&pc=QCB | 2023826-2022-03223 | 2022-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.0/1.5/013 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -09.0/1.5/013 | OS | Physician | 09/05/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496689&pc=QCB | 2023826-2022-03229 | 2022-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.5/1.0/147 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -10.5/1.0/147 | OD | Physician | 09/06/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496692&pc=QCB | 2023826-2022-03217 | 2022-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED IN AN OBIQUE POSITION A 13.2MM VTICMO13.2, -10.0/1.0/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -10.0/1.0/105 | OD | Physician | 09/05/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496698&pc=QCB | 2023826-2022-03225 | 2022-03-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -08.5/1.5/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -08.5/1.5/073 | OD | Physician | 09/06/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496699&pc=QCB | 2023826-2022-03211 | 2022-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/3.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/3.0/088 | OS | Physician | 09/05/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496704&pc=QCB | 2023826-2022-03235 | 2022-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.5/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: UNK CLAIM# 726181. | -16.5/2.0/093 | OD | Physician | 09/06/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496715&pc=QCB | 2023826-2022-03212 | 2022-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.5/3.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.5/3.0/093 | OD | Physician | 09/05/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496716&pc=QCB | 2023826-2022-03219 | 2022-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.5/3.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -11.5/3.5/095 | OD | Physician | 09/05/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496718&pc=QCB | 2023826-2022-03222 | 2022-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.5/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. CLAIM# (B)(4). | -10.5/2.0/089 | OS | Physician | 09/05/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496744&pc=QCB | 2023826-2022-03214 | 2022-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.5/3.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. CLAIM# (B)(4). | -14.5/3.5/093 | OD | Physician | 09/05/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496746&pc=QCB | 2023826-2022-03220 | 2022-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -09.0/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. CLAIM# (B)(4). | -09.0/2.0/088 | OD | Physician | 09/05/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496748&pc=QCB | 2023826-2022-03226 | 2022-04-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.5/3.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED OBLIQUELY WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. CLAIM# (B)(4). | -11.5/3.5/092 | OS | Physician | 09/06/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496778&pc=QCB | 2023826-2022-03206 | 2022-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -06.5/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED VERTICALLY WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -06.5/1.0/087 | OD | Physician | 09/05/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496785&pc=QCB | 2023826-2022-03209 | 2022-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.0/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -18.0/2.0/095 | OD | Physician | 09/05/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496796&pc=QCB | 2023826-2022-03213 | 2022-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.5/2.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). | -12.5/2.5/090 | OD | Physician | 09/05/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496813&pc=QCB | 2023826-2022-03224 | 2022-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.5/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. (B)(4). CLAIM# (B)(4). | -10.5/1.0/091 | OD | Physician | 09/05/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15496815&pc=QCB | 2023826-2022-03230 | 2022-08-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.0/1.5/1020 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -09.0/1.5/1020 | OD | Physician | 09/06/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15498852&pc=QCB | 2023826-2022-03338 | 2022-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS -10.0/1.0/137 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM RESOLVED. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -10.0/1.0/137 | OS | Physician | 08/29/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15498908&pc=QCB | 2023826-2022-03325 | 2022-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.00/+3.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2022 BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK; ETHNICITY: UNK; RACE: UNK. (B)(4). | -12.00/+3.0/094 | OS | Physician | 08/31/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15499148&pc=QCB | 2023826-2022-03372 | 2022-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -7.0/1.0/115 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME MODEL/LENGTH LENS BUT IMPLANTED VERTICALLY. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND FIBRE ON GLASS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -7.0/1.0/115 | OS | Physician | 09/26/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15499480&pc=QCB | 2023826-2022-03341 | 2022-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS -8.0 DIOPTER WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO LOW VAULT. ON (B)(6) 2022 A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- "-8.0" SHOULD BE CORRECTED TO "-8.00/2.5/144 (SPHERE/CYLINDER/AXIS)" . CLAIM# (B)(4). | -8.00/2.5/144 | OS | Physician | 08/29/2022 | 2022 | 42 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15499651&pc=QCB | 2023826-2022-03354 | 2022-02-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -8.0/3.0/85 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). | -8.0/3.0/85 | OD | Physician | 08/29/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15499653&pc=QCB | 2023826-2022-03355 | 2022-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -13.0/3.0/003 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). | -13.0/3.0/003 | OS | Physician | 08/29/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15506107&pc=QCB | 2023826-2022-03252 | 2022-08-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.0/2.5/086 (SPHERE/CYLINDER/AXIS) WAS INTRAOPERATIVELY IMPLANTED AND REMOVED DUE TO A TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. NO PATIENT INJURY WAS REPORTED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: WEIGHT-RACE:UNK. OFF LABEL - ACD<3.0 <21. CLAIM# (B)(4). | -12.0/2.5/086 | OS | Physician | 08/24/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15506302&pc=QCB | 2023826-2022-03328 | 2022-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.00/+2.5/109 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS EYE ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO LENS ROTATION. THE LENS WAS REPLACED WITH A DIFFERENT LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD). THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID, IN A SPECIMEN CUP. VISUAL INSPECTION FOUND A HAPTIC TORN. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.00/+2.5/109 | OD | Other Health Care Professional | 09/12/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15509002&pc=QCB | 2023826-2022-03348 | 2022-08-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-09-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -10.0/2.0/101 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS THAT WAS IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: 16-FEB-2023. H3-DEVICE EVALUATION: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED DRY A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -10.0/2.0/101 | OS | Physician | 09/01/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15522870&pc=QCB | 2023826-2022-03367 | 2022-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -12.5/1.5/169 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - RACE: UNK. (B)(4). CLAIM# (B)(4). | -12.5/1.5/169 | OS | Physician | 09/04/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15524123&pc=QCB | 2023826-2022-03368 | 2022-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -13.0/1.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: B5: " THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.00/1.0/90 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED A LOW VAULT WITH ROTATION. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR THE SAME MODEL/LONGER LENGTH LENS; IMPLANTED VERTICALLY AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT IS UNKNOWN." CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -13.0/1.0/090 | OS | Physician | 09/05/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529034&pc=QCB | 2023826-2022-03195 | 2022-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/110 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE NEXT DAY THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. THE LENS WAS LATER REPLACED WITH A SHORTER LENGTH LENS ON (B)(6) 2022 AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -10.0/+1.5/110 | OS | Physician | 09/02/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529038&pc=QCB | 2023826-2022-03266 | 2022-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0/+2.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# (B)(4), SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | -15.0/+2.0/093 | OD | Physician | 09/05/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529040&pc=QCB | 2023826-2022-03190 | 2022-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.5/107 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). | -10.5/+2.5/107 | OD | Physician | 09/01/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529049&pc=QCB | 2023826-2022-03193 | 2021-12-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+2.0/100 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: (B)(4). | -12.5/+2.0/100 | OD | Physician | 09/01/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529051&pc=QCB | 2023826-2022-03271 | 2022-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+0.5/133 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# (B)(4), SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | -7.5/+0.5/133 | OD | Physician | 09/05/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529055&pc=QCB | 2023826-2022-03269 | 2022-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.0/+2.0/073 INTO THE PATIENT'S LEFT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# (B)(4), SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | -16.0/+2.0/073 | OS | Physician | 09/05/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529061&pc=QCB | 2023826-2022-03276 | 2022-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+3.0/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -14.0/+3.0/090 | OD | Physician | 09/04/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529093&pc=QCB | 2023826-2022-03387 | 2022-06-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.5/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS WHICH WAS IMPLANTED VERTICALLY DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.5/+1.5/092 | OS | Physician | 09/05/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529101&pc=QCB | 2023826-2022-03278 | 2022-08-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+1.0/78 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -8.0/+1.0/78 | OS | Physician | 09/05/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529110&pc=QCB | 2023826-2022-03192 | 2022-06-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.0/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# (B)(4), SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). | -10.5/+2.0/089 | OS | Physician | 09/01/2022 | 2022 | 17 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529118&pc=QCB | 2023826-2022-03196 | 2022-07-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.0/055 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# (B)(4), SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). | -10.5/+1.0/055 | OD | Physician | 09/02/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529139&pc=QCB | 2023826-2022-03267 | 2022-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+2.5/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# (B)(4), SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | -14.0/+2.5/091 | OS | Physician | 09/05/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529152&pc=QCB | 2023826-2022-03264 | 2022-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+2.0/111 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# (B)(4), SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | -14.0/+2.0/111 | OD | Physician | 09/05/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529180&pc=QCB | 2023826-2022-03384 | 2022-07-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.5/088 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. IN A SEPARATE SURGERY THAT DAY A SHORTER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -11.5/+1.5/088 | OS | Physician | 09/05/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529185&pc=QCB | 2023826-2022-03273 | 2022-06-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+3.0/079 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# (B)(4), SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | -9.5/+3.0/079 | OS | Physician | 09/05/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529189&pc=QCB | 2023826-2022-03385 | 2022-07-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.0/105 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. IN A SEPARATE SURGERY THAT DAY A SHORTER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -11.5/+1.0/105 | OD | Physician | 09/05/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529192&pc=QCB | 2023826-2022-03194 | 2022-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.50/+2.0/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). | -8.50/+2.0/090 | OD | Physician | 09/01/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529215&pc=QCB | 2023826-2022-03380 | 2022-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.5/087 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). | -9.0/+1.5/087 | OD | Physician | 09/04/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529218&pc=QCB | 2023826-2022-03382 | 2022-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+3.0/086 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). | -10.0/+3.0/086 | OS | Physician | 09/04/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529220&pc=QCB | 2023826-2022-03389 | 2022-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+1.0/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -14.0/+1.0/095 | OD | Physician | 09/05/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529249&pc=QCB | 2023826-2022-03277 | 2022-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/+0.5/107 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -5.5/+0.5/107 | OD | Physician | 09/04/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529310&pc=QCB | 2023826-2022-03191 | 2022-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.0/074 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# (B)(4), SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). | -11.0/+2.0/074 | OS | Physician | 09/01/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529317&pc=QCB | 2023826-2022-03265 | 2022-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.5/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -7.0/+1.5/093 | OD | Physician | 09/05/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529340&pc=QCB | 2023826-2022-03268 | 2022-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+0.5/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# (B)(4), SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | -7.5/+0.5/092 | OD | Physician | 09/05/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529385&pc=QCB | 2023826-2022-03381 | 2022-08-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/+2.0/073 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). | -14.5/+2.0/073 | OD | Physician | 09/05/2022 | 2022 | 47 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529392&pc=QCB | 2023826-2022-03386 | 2022-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.5/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.5/+1.5/091 | OD | Physician | 09/05/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529516&pc=QCB | 2023826-2022-03293 | 2022-06-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -14.5/3.0/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -14.5/3.0/082 | OS | Physician | 09/15/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529534&pc=QCB | 2023826-2022-03299 | 2022-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.5/2.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.5/2.5/090 | OS | Physician | 09/15/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529559&pc=QCB | 2023826-2022-03298 | 2022-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.5/2.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.5/2.5/092 | OS | Physician | 09/15/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15529567&pc=QCB | 2023826-2022-03302 | 2022-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.5/2.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. (B)(4). CLAIM# (B)(4). | -16.5/2.5/086 | OD | Physician | 09/15/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15530976&pc=QCB | 2023826-2022-03294 | 2022-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.0/3.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.0/3.5/093 | OD | Physician | 09/15/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15531045&pc=QCB | 2023826-2022-03288 | 2022-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.5/2.0/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. CLAIM# (B)(4). | -9.5/2.0/072 | OS | Physician | 09/14/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15531786&pc=QCB | 2023826-2022-03297 | 2022-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.0/1.0/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. (B)(4). CLAIM# (B)(4). | -10.0/1.0/108 | OS | Physician | 09/15/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15532310&pc=QCB | 2023826-2022-03303 | 2022-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.5/0.5/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.5/0.5/070 | OD | Physician | 09/15/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15534202&pc=QCB | 2023826-2022-03429 | 2022-08-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF A -11.5/2.0/093 (SPHERE/CYLINDER/AXIS) TORE BEFORE/DURING LOADING ON (B)(6) 2022. ON (B)(6) 2022, A REPLACEMENT LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) AND THE PROBLEM WAS RESOLVED. THE CAUSE IS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -11.5/2.0/093 | OD | Physician | 09/05/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15534366&pc=QCB | 2023826-2022-03459 |  | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -8.50/3.0/080 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE PATIENT HAD EXCESSIVE VAULTING AND IT WAS INDICATED THAT THE SURGEON EXPLANTED THE LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). | -8.50/3.0/080 | OD | Physician | 09/08/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15534459&pc=QCB | 2023826-2022-03430 | 2021-01-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -1.5/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED WITH LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: B5: CORRECTED TO - "THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -1.5/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED WITH LOW VAULT AND REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN." CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC BENT/ DEFORMED TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -1.5/2.0/089 | OS | Physician | 09/05/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15536366&pc=QCB | 2023826-2022-03433 | 2022-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/074 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS A PATIENT REALTED FACTOR. Manufacturer Narrative: (B)(4). | -12.0/1.5/074 | OD | Physician | 09/06/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15539804&pc=QCB | 2023826-2022-03439 | 2022-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -8.0/2.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED WITH LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "13.2MM/T1009553 FOR VERTICAL PLACEMENT." Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY, RACE:UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -8.0/2.0/093 | OD | Physician | 09/07/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15539809&pc=QCB | 2023826-2022-03440 | 2022-08-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.0/2.0/082 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS A PATIENT RELATED FACTOR. STATUS OF THE EYE IS REPORTED AS, "GOOD." Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK. CLAIM# (B)(4). | -11.0/2.0/082 | OD | Physician | 09/07/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15540391&pc=QCB | 2023826-2022-03441 | 2022-08-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030017 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022, A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -13.0/1.0/84 (SPHERE/CYLINDER/AXIS) LENS TORE AND WAS DAMAGED DURING LOADING. LENS WAS NOT IMPLANTED. REPORTEDLY, STATUS OF THE EYE IS, "GOOD," AND "THE SURGERY DATE FOR THE REPLACEMENT LENS HAS NOT YET BEEN DETERMINED." Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: B1: ADVERSE EVENT SHOULD BE OMITTED AND REPLACED WITH PRODUCT PROBLEM FOR CORRECTION. CLAIM #(B)(4). | -13.0/1.0/84 | NA | Physician | 09/07/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15540392&pc=QCB | 2023826-2022-03442 | 2022-08-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030017 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022, A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -12.5/1.0/100 (SPHERE/CYLINDER/AXIS) LENS TORE AND WAS DAMAGED DURING LOADING. LENS WAS NOT IMPLANTED. REPORTEDLY, STATUS OF THE EYE IS, "GOOD," AND "THE SURGERY DATE FOR THE REPLACEMENT LENS HAS NOT YET BEEN DETERMINED." Manufacturer Narrative: CORRECTED DATA: B1: ADVERSE EVENT SHOULD BE OMITTED AND REPLACED WITH PRODUCT PROBLEM FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, & RACE :UNK. CLAIM# (B)(4) IS REGARDING AN EXCESSIVE VAULT EXCHANGE WHICH IS UNRELATED TO THIS COMPLAINT. NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -12.5/1.0/100 | NA | Physician | 09/07/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15540744&pc=QCB | 2023826-2022-03461 | 2022-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS -13.5/3.0/096 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO LOW VAULT. A REPLACEMENT LENS OF LONGER LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS -13.5/3.0/096 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6)2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. PROBLEM RESOLVED. CAUSE WAS REPORTED AS UNKNOWN. CLAIM # (B)(4). | -13.5/3.0/096 | OD | Physician | 09/05/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15540746&pc=QCB | 2023826-2022-03401 | 2022-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -8.0/1.0/164 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM RESOLVED. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5: "THE REPORTER INDICATED THAT THE SURGEON VERTICALLY IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.00/1.0/164 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXPLANTED DUE TO LOW VAULT WITHOUT ROTATION. ON THE SAME DAY SEPARATE SURGERY THE LENS WAS REPLACED WITH A LENS OF THE SAME MODEL/ LONGER LENGTH LENS IMPLANTED HORIZONTALLY AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN." CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -8.0/1.0/164 | OS | Physician | 09/05/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15540829&pc=QCB | 2023826-2022-03443 | 2022-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.5/3.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A REPLACEMENT LENS DUE TO A LENS TILT. REPORTEDLY, "THE ICL LENS TILT." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -9.5/3.0/093 | OD | Physician | 09/07/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15540838&pc=QCB | 2023826-2022-03462 |  | Injury | STAAR SURGICAL COMPANY | 2022-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICMO13.2, -15.50/3.5/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED AND REPLACED WITH BACK UP LENS BECAUSE IT FLIPPED DURING INSERTION. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: CORRECTED DATA: B1: ADVERSE EVENT SHOULD REPLACE PRODUCT PROBLEM FOR CORRECTION. H6: HEALTH EFFECT- IMPACT CODE: 2199 SHOULD BE OMMITTED AND REPLACED WITH 4627 FOR CORRECTION. CLAIM# (B)(4). | -15.50/3.5/092 | NA | Physician | 09/05/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15540965&pc=QCB | 2023826-2022-03448 | 2022-03-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -13.5/2.5/94 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS THE DEVICE. REPORTEDLY, STATUS OF THE EYE IS "VISION WELL," AND "ICL STABLE." Manufacturer Narrative: (B)(4). | -13.5/2.5/94 | OD | Physician | 09/07/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15541001&pc=QCB | 2023826-2022-03392 | 2022-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+2.5/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. IN A SEPARATE SURGERY THAT DAY A SHORTER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D4: UNIQUE DEVICE IDENTIFIER: (B)(4). CLAIM: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+2.5/092 | OS | Physician | 09/05/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15541002&pc=QCB | 2023826-2022-03394 | 2022-06-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/102 INTO THE PATIENT'S LEFT EYE (OS) ON 03-JUN-2022. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D4: UNIQUE IDENTIFIER (UDI) NUMBER: (B)(4). CLAIM (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: 10/21/2022. H3: PRODUCT EVALUATION: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -10.0/+1.5/102 | OS | Physician | 09/05/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15541006&pc=QCB | 2023826-2022-03390 | 2022-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+2.0/088 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -13.0/+2.0/088 | OD | Physician | 09/05/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15546506&pc=QCB | 2023826-2022-03510 | 2022-03-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/068 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.0/+1.5/068 | OD | Physician | 09/04/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15546929&pc=QCB | 2023826-2022-03313 | 2022-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -8.5/1.0/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -8.5/1.0/094 | OD | Physician | 09/05/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15547044&pc=QCB | 2023826-2022-03314 | 2022-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. . ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: B2 SECTION IN INITIAL MDR (LIFE-THREATNING) IS NOT APPLICALE. IT WAS CHECKED IN ERROR. CLAIM# (B)(4). Manufacturer Narrative: D4: CORRECTED TO: SERIAL NUMBER: (B)(6) IN INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -9.0/1.0/91 | OS | Physician | 09/05/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15547064&pc=QCB | 2023826-2022-03449 | 2022-08-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -13.0/2.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO A PATIENT ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022. A SHORTER LENGTH LENS WAS THEN IMPLANTED. Manufacturer Narrative: B5: A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -13.0/2.5/094 (SPHERE/CYLINDER/AXIS) WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION AND BLURRED VISION. CAUSE OF THE DEVICE IS OTHER. THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. REPORTEDLY, "LENS WAS TOO SMALL AND ROTATED." PROBLEM WAS RESOLVED. STATUS OF THE EYE IS REPORTED AS, "LOOKS GOOD!" CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -13.0/2.5/093 | NA | Physician | 09/09/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15547398&pc=QCB | 2023826-2022-03315 | 2022-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF -13.5/5.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS BUT WAS IMPLANTED VERTICALLY. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: A2: CORRECTED TO: AGE AT TIME OF EVENT, DATE OF BIRTH: ON (B)(6) 2002 IN THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE AND RESIDUE/DEBRIS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -13.5/5.0/092 | OS | Physician | 09/05/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15547403&pc=QCB | 2023826-2022-03464 | 2022-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.0/170 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D4: UNIQUE IDENTIFIER (UDI) NUMBER: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -8.5/+1.0/170 | OD | Physician | 09/06/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15547482&pc=QCB | 2023826-2022-03466 | 2022-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+1.0/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -7.5/+1.0/093 | OS | Physician | 09/06/2022 | 2022 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15547533&pc=QCB | 2023826-2022-03465 | 2022-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+1.0/085 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -7.5/+1.0/085 | OD | Physician | 09/06/2022 | 2022 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15547735&pc=QCB | 2023826-2022-03468 | 2022-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.5/111 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -9.0/+1.5/111 | OD | Physician | 09/06/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15547890&pc=QCB | 2023826-2022-03454 | 2022-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/81 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. EXCESSIVE VAULTING WAS OBSERVED. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/81 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULTING WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. NO PATIENT INJURY REPORTED. CLAIM#: (B)(4). | -9.0/1.0/81 | OS | Physician | 09/09/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15547892&pc=QCB | 2023826-2022-03453 | 2022-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/74(SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULTING WAS OBSERVED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: AL THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/1.0/074 INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REPLACED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT NOT REPORTED. NO PATIENT INJURY REPORTED. CLAIM #(B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -9.0/1.0/74 | OD | Physician | 09/09/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15552142&pc=QCB | 2023826-2022-03515 | 2022-03-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Optical Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -12.5/2.5/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT WITH GLARE/HALOES. THE PROBLEM RESOLVED. CAUSE REPORTED AS DEVICE AND PATIENT RELATED FACTOR. Manufacturer Narrative: CLAIM# (B)(4). | -12.5/2.5/91 | OD | Physician | 09/06/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15552464&pc=QCB | 2023826-2022-03473 | 2022-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.0/070 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). | -7.0/+1.0/070 | OD | Physician | 09/07/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15552490&pc=QCB | 2023826-2022-03472 | 2022-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+1.0/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.0/+1.0/089 | OS | Physician | 09/07/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15552525&pc=QCB | 2023826-2022-03471 | 2022-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.5/+1.0/091 | OD | Physician | 09/07/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15552537&pc=QCB | 2023826-2022-03469 | 2022-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.0/088 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE LENS WAS IMPLANTED VERTICALLY. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -11.0/+2.0/088 | OD | Physician | 09/07/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15555239&pc=QCB | 2023826-2022-03312 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -8.0/1.0/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS AN EXCHANGE LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021 TO CORRECT REFRACTIVE SURPRISE. THE PROBLEM WAS NOT RESOLVED. "THE PATIENT IS 45 YEARS OLD, PRESBYOPIA -1.50D, PROFESSIONAL ACCOUNTANT, USUALLY SEE MORE CLOSE, THE MAIN COMPLAINT OF NEAR VISION IS STILL POOR VISION TO SOLVE THE NEAR VISION, SO 2 OPERATIONS." THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR# 2023826-2022-03516 FOR INITIAL LENS. Manufacturer Narrative: ADDITIONAL INFORMATION: PREVIOUS MDR SUBMITTED IS NOT APPLICABLE. UPON FURTHER REVIEW IT HAS BEEN DETERMINED THAT THE CLAIM DOES NOT MEET DEFINITION OF A COMPLAINT. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -8.0/1.0/095 | OD | Physician | 09/06/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15555753&pc=QCB | 2023826-2022-03477 | 2022-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+1.5/088 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -8.0/+1.5/088 | OS | Physician | 09/07/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15555892&pc=QCB | 2023826-2022-03476 | 2022-05-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.5/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.0/+1.5/092 | OD | Physician | 09/07/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15555954&pc=QCB | 2023826-2022-03479 | 2022-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+3.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). | -14.0/+3.0/093 | OD | Physician | 09/07/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15556087&pc=QCB | 2023826-2022-03512 | 2022-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-06 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -12.00/1.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LOW VAULT WITH ROTATION, IT WAS REPORTED THAT THE REPOSITIONING RESOLVED THE PROBLEM BUT AFTER A FEW MONTHS IT ROTATED AGAIN. THE LENS WAS REPOSITIONED FOR A SECOND TIME ON (B)(6) 2022 AND A PLAN EXCHANGED IS PENDING. IT WAS REPORTED THAT THE PATIENT HAS 6/12 VISION, LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -12.00/1.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2020. ON (B)(6) 2021 THE LENS WAS REPOSITIONED DUE TO LOW VAULT WITH ROTATION, IT WAS REPORTED THAT THE REPOSITIONING RESOLVED THE PROBLEM BUT AFTER A FEW MONTHS IT ROTATED AGAIN. THE LENS WAS REPOSITIONED FOR A SECOND TIME ON (B)(6) 2022. THE LENS WAS EXCHANGED FOR A LONGER LENGTH ON (B)(6) 2022 AND THE PROBLEM WAS RESOLVED CLAIM#: (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -12.00/1.0/092 | OS | Physician | 09/06/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15558242&pc=QCB | 2023826-2022-03470 | 2022-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.0/087 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE LENS WAS IMPLANTED VERTICALLY. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.0/+1.0/087 | OS | Physician | 09/07/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15559455&pc=QCB | 2023826-2022-03518 | 2022-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -3.50/3.0/160 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REFRACTIVE SURPRISE WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2022 AND THE PROBLEM WAS NOT RESOLVED. "ICL ROTATED FROM 11° TO 25° IN (B)(6) 2022, WHICH REDUCED THE REFRACTIVE ERROR FROM (+1.0/-1.0/110°) TO (+0.5/-0.5/65°), AND THE UNCORRECTED VISION INCREASED FROM 0.7- TO 0.9-." THE LENS REMAINS IMPLANTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -3.50/3.0/160 | OS | Other Health Care Professional | 09/07/2022 | 2022 | 42 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15559518&pc=QCB | 2023826-2022-03524 | 2022-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+2.0/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI) #(B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). | -14.0/+2.0/091 | OS | Physician | 09/08/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15562218&pc=QCB | 2023826-2022-03456 | 2022-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -7.5/1.0/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -7.5/1.0/087 | OS | Physician | 09/14/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15562532&pc=QCB | 2023826-2022-03457 | 2022-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.0/1.0/0142 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON(B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: PATIENT WEIGHT,ETHNICITY AND RACE UNK. CLAIM# (B)(4). | -12.0/1.0/0142 | OS | Physician | 07/20/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15562768&pc=QCB | 2023826-2022-03528 | 2022-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.5/087 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -9.5/+1.5/087 | OS | Physician | 09/08/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15565244&pc=QCB | 2023826-2022-03333 | 2022-08-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON REMOVED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS, -10.00/+3.0/100 (SPHERE/CYLINDER/AXIS), FROM THE VIAL AND UPON INSPECTION NOTED THE LENS WAS DAMAGED. THE LENS WAS NOT USED AND THERE WAS NO PATIENT CONTACT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE, IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON NOTED A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.00/3.0/100 (SPHERE/CYLINDER/AXIS) WAS "DAMAGED" ON 03/AUG/2022. THE DAMAGE OCCURRED DURING LOADING AND WAS NOTED WHILE REMOVING FROM THE VIAL. THERE WAS NO PATIENT CONTACT. REPORTEDLY "THE DOCTOR FOUND THE LENS DAMAGED WHEN SHE TOOK IT BY FOAM TIP FROM THE VIAL. THEY NORMALLY SEE THE LENS THROUGH MICROSCOPE BEFORE LOADING. SO, THE DOCTOR FOUND THE LENS WAS "DAMAGED" AT THE BEGINNING." ON THE SAME DAY A LENS OF THE SAME MODEL/LENGTH WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "GOOD VISION AFTER USING THE BACK UP LENS (SAME DIOPTER AND CYLINDER)". THE CAUSE OF THE EVENT WAS THE DEVICE. CAUSE DETAILS: "FOUND THAT LENS TEAR/BREAK WHILE REMOVING FROM THE VIAL, DID NOT IMPLANT THIS LENS". CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM #: (B)(4). | -10.00/+3.0/100 | OD | Physician | 09/14/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15565516&pc=QCB | 2023826-2022-03335 | 2022-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+2.5/096 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | -10.50/+2.5/096 | OS | Physician | 09/15/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15565664&pc=QCB | 2023826-2022-03334 | 2022-08-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.50/+1.0/087 (SPHERE/CYLINDER/AXIS), FROM THE VIAL AND UPON INSPECTION NOTED THE LENS WAS DAMAGED. THE LENS WAS NOT USED AND THERE WAS NO PATIENT CONTACT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: ADDITIONAL DATA: H6 - WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. B5: THE REPORTER INDICATED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.50/+1.0/087 (SPHERE/CYLINDER/AXIS) MAY HAVE BEEN DAMAGED DURING LOADING. THERE WAS NO PATIENT CONTACT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS DUE TO THE DEVICE. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION WAS UNABLE TO CONFIRM IF LENS WAS DAMAGED DUE TO LENS BEING STUCK IN CARTRIDGE. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | -08.50/+1.0/087 | NA | Physician | 09/14/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15566587&pc=QCB | 2023826-2022-03336 | 2022-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+1.5/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: WEIGHT: UNK; ETHNICITY: UNK; RACE: UNK. (B)(4). | -18.00/+1.5/092 | OD | Physician | 09/16/2022 | 2022 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15567500&pc=QCB | 2023826-2022-03569 | 2022-09-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6)2022 A 13.7MM, VTICMO13.7, -13.0/3.0/85 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS OVERTURN DURING OPERATION OF THE LEFT EYE (OS). IT WAS REPORTED THAT THE LENS BROKE DURING REMOVAL. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY WITH NO PATIENT INJURY. A REPLACEMENT LENS WAS IMPLANTED BUT IT DID NOT RESOLVE THE PROBLEM. CAUSE OF EVENT WAS REPORTED TO BE USER ERROR. Manufacturer Narrative: CLAIM #(B)(4). | -13.0/3.0/85 | OS | Physician | 09/09/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15567509&pc=QCB | 2023826-2022-03570 | 2022-09-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 VTICMO13.2, -13.0/3.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS OVERTURN DURING OPERATION OF THE LEFT EYE (OS). IT WAS REPORTED THAT THE LENS BROKE DURING REMOVAL. THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY WITH NO PATIENT INJURY. CAUSE OF EVENT WAS REPORTED TO BE USER ERROR. SEE MFR # 2023826-2022-03569 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: CLAIM # (B)(4). | -13.0/3.0/083 | OS | Physician | 09/09/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15567581&pc=QCB | 2023826-2022-03572 | 2022-09-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -6.50/2.00/113 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. IT WAS REPORTED THAT THE LENS WAS IMPLANTED, REMOVED AND REPLACED WITH A SHORTER LENGTH LENS INTRAOPERATIVELY DUE TO EXCESSIVE VAULT. PROBLEM RESOLVED. FOR CAUSE DETAILS THE REPORTER STATED "SIZE 12.6 WAS TOO LARGE, SWAPPED TO SIZE 12.1." Manufacturer Narrative: H6 - DEVICE CODE 1494: OFF-LABEL USE (ACD <3.0MM). CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -6.50/2.00/113 | OS | Physician | 09/07/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15567591&pc=QCB | 2023826-2022-03566 | 2022-08-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL12.6 -8.00/1.0/120 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS DID NOT SIT WELL AND WAS REPLACED WITH A SAME MODEL, POWER AND LENGTH LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -8.00/1.0/120 | NA | Physician | 09/07/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15567699&pc=QCB | 2023826-2022-03534 | 2021-10-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/085 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -10.0/+1.5/085 | OS | Physician | 09/09/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15567781&pc=QCB | 2023826-2022-03571 | 2021-10-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS -7.00/+1.50/75 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021 THE LENS WAS EXPLANTED DUE TO LOW VAULT WITH ROTATION. THE LENS WAS LATER REPLACED WITH A SAME LENGTH LENS AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.00/+1.50/75 | OS | Physician | 09/09/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15567788&pc=QCB | 2023826-2022-03533 | 2022-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.5/160 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). | -9.0/+1.5/160 | OS | Physician | 09/09/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15567814&pc=QCB | 2023826-2022-03567 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 -13.50/2.5/098 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS MISSING CHUNK FROM PERIPHERY AND WAS NOT IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN VIAL. VISUAL INSPECTION FOUND ONE OF THE HAPTICS TORN. THE TORN PIECE WAS NOT RETURNED. H6- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -13.50/2.5/098 | NA | Physician | 09/08/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15567877&pc=QCB | 2023826-2022-03584 | 2022-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICH12.1 IMPLANTABLE COLLAMER LENS -0.50/+4.50/27 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A DIFFERENT AXIS LENS DUE TO AXIS ENCODING ERROR AND THE PROBLEM RESOLVED. USER ERROR WAS REPORTED TO BE THE CAUSE OF THE EVENT. Manufacturer Narrative: B5: CORRECTED TO: THE REPORTER INDICATED THAT A 12.1MM VTICH12.1 IMPLANTABLE COLLAMER LENS +0.50/+4.50/030 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A DIFFERENT AXIS LENS DUE TO AXIS ENCODING ERROR AND THE PROBLEM RESOLVED. USER ERROR WAS REPORTED TO BE THE CAUSE OF THE EVENT. D6A: CORRECTED TO ON (B)(6) 2022. D6B: CORRECTED TO ON (B)(6) 2022. H6 - MEDICAL DEVICE PROBLEM CODE 1494 (OFF-LABEL USE: ACD<3.0MM). CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -0.50/+4.50/27 | OD | Physician | 09/07/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15568080&pc=QCB | 2023826-2022-03563 | 2022-08-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-07 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6)2022 A 12.6MM, VTICM5_12.6, -9.50/4.0/86 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS FOUND BROKEN BEFORE LOADING, THERE WAS NO PATIENT CONTACT. A BACKUP LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT WAS THE DEVICE. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -9.50/4.0/86 | NA | Physician | 09/07/2022 | 2022 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15569954&pc=QCB | 2023826-2022-03423 | 2022-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -09.0/1.5/069 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -09.0/1.5/069 | OD | Physician | 09/22/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15570021&pc=QCB | 2023826-2022-03422 | 2022-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -09.0/1.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -09.0/1.0/095 | OS | Physician | 09/22/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15570107&pc=QCB | 2023826-2022-03406 | 2022-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -09.0/0.5/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION - UNK. CLAIM# (B)(4). | -09.0/0.5/077 | OD | Physician | 09/19/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15570274&pc=QCB | 2023826-2022-03404 | 2022-08-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1, -5.5/1.0/102 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED WITH A LONGER LENGTH LENS INTRAOPERATIVELY ON 20/AUG/2022 FROM PATIENT'S RIGHT EYE (OD) DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT,ETHNICITY,RACE:UNK (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: LENS WAS RETURNED IN MICRO-CENTRIFUGE VIAL, WITH MOISTURE ON THE PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS.DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -5.5/1.0/102 | OD | Physician | 09/19/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15571117&pc=QCB | 2023826-2022-03411 | 2022-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -07.0/5.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). UNDER 21 YRS OF AGE AT DATE OF IMPLANT. (B)(4). | -07.0/5.5/092 | OS | Physician | 09/19/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15571130&pc=QCB | 2023826-2022-03405 | 2022-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -8.0/1.0/017 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -8.0/1.0/017 | OS | Physician | 09/19/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15571227&pc=QCB | 2023826-2022-03424 | 2022-05-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -08.5/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -08.5/2.0/089 | OS | Physician | 09/23/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15571673&pc=QCB | 2023826-2022-03407 | 2022-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -05.5/1.0/127 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK; RACE: UNK. (B)(4). | -05.5/1.0/127 | OD | Physician | 09/19/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15573124&pc=QCB | 2023826-2022-03425 | 2022-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.5/1.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -09.5/1.0/094 | OD | Physician | 09/23/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15574073&pc=QCB | 2023826-2022-03416 | 2022-09-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -09.0/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTION: H1 - REPORTED AS SERIOUS INJURY - CORRECT TO MALFUNCTION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -09.0/1.5/090 | OD | Physician | 09/21/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15574091&pc=QCB | 2023826-2022-03419 | 2022-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-10 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.0/3.0/137 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.0/3.0/137 | OS | Physician | 09/22/2022 | 2022 | 17 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15574564&pc=QCB | 2023826-2022-03415 | 2022-09-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -7.0/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT,ETHNICITY AND RACE: UNK. (B)(4). Manufacturer Narrative: CORRECTION: H1 - REPORTED AS SERIOUS INJURY - CORRECT TO MALFUNCTION. CLAIM# (B)(4). | -7.0/1.0/090 | OS | Physician | 09/21/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15574888&pc=QCB | 2023826-2022-03413 | 2022-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.5/4.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.5/4.0/089 | OD | Physician | 09/20/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15576219&pc=QCB | 2023826-2022-03412 | 2022-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.00/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, (B)(4). | -11.00/2.0/090 | OS | Physician | 09/20/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15577121&pc=QCB | 2023826-2022-03426 | 2022-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -09.5/4.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -09.5/4.0/083 | OS | Physician | 09/23/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15577145&pc=QCB | 2023826-2022-03421 | 2022-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -07.0/1.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -07.0/1.0/092 | OS | Physician | 09/22/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15577183&pc=QCB | 2023826-2022-03420 | 2022-09-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -18.0/5.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -18.0/5.5/082 | OS | Physician | 09/22/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15577360&pc=QCB | 2023826-2022-03599 | 2022-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+2.5/083 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -12.0/+2.5/083 | OD | Physician | 09/13/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15577391&pc=QCB | 2023826-2022-03618 | 2022-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -12.5/+6.0/103 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THE PROBLEM RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4).: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -12.5/+6.0/103 | OD | Physician | 10/06/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15581843&pc=QCB | 2023826-2022-03603 | 2022-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+0.5/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). | -7.0/+0.5/091 | OD | Physician | 09/15/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15581875&pc=QCB | 2023826-2022-03605 | 2022-07-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.5/077 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE (B)(4). CLAIM# (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). | -10.0/+2.5/077 | OS | Physician | 09/15/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15581876&pc=QCB | 2023826-2022-03602 | 2022-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+1.5/088 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). | -7.5/+1.5/088 | OD | Physician | 09/15/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15581889&pc=QCB | 2023826-2022-03604 | 2022-07-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+3.0/084 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -9.5/+3.0/084 | OD | Physician | 09/15/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15584540&pc=QCB | 2023826-2022-03621 | 2022-09-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -14.0/+5.0/084 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXPLANTED. ON THIS SAME DATE A SHORTER LENGTH LENS WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: MEDICAL DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -14.0/+5.0/084 | OS | Physician | 10/06/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15584544&pc=QCB | 2023826-2022-03622 | 2022-06-22 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-11 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -10.5/2.0/084 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE. INTRAOPERATIVELY THE LENS WAS REMOVED WITH NO PATIENT INJURY. ON 17-JUL-2022 A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -10.5/2.0/084 | OS | Physician | 10/09/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15588171&pc=QCB | 2023826-2022-03575 |  | Injury | STAAR SURGICAL COMPANY | 2022-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -9.50/1.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. LOW VAULT WITH LENS ROTATION WAS OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE REPORTED AS USER ERROR. DEVICE DID NOT FAIL TO PERFORM AS INTENDED. ICL LOW DIAMETER ICL SELECTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5- "ON 22/NOV THE LENS WAS EXCHANGED WITH THE SAME MODEL/LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS REPORTED AS "DOING WELL". CLAIM# (B)(4). | -9.50/1.0/090 | OS | Other Health Care Professional | 09/23/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15588821&pc=QCB | 2023826-2022-03591 | 2022-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -3.00/+0.50/66, IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022 AND THE PATIENT EXPERIENCED LOW VAULT, LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5- PER UPDATED INFORMATION ON 19-SEP-2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | -3.00/+0.50/66 | OD | Physician | 09/13/2022 | 2022 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15591414&pc=QCB | 2023826-2022-03638 | 2022-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -14.0/2.5/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO THE SURGEON REPORTS THE PATIENT FEAR COMPLICATIONS AND INSISTED ON THE REMOVAL OF THE ICL. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -14.0/2.5/095 | OD | Physician | 10/09/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15591462&pc=QCB | 2023826-2022-03639 | 2022-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -16.0/3.0/079 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO THE SURGEON REPORTS THE PATIENT FEAR COMPLICATIONS AND INSISTED ON THE REMOVAL OF THE ICL. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -16.0/3.0/079 | OS | Physician | 10/09/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15591775&pc=QCB | 2023826-2022-03492 | 2022-06-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.50/+0.5/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6)2022 . THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS EXPLANTED ON (B)(6) 2022. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | -12.50/+0.5/091 | OD | Physician | 09/21/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15591785&pc=QCB | 2023826-2022-03486 | 2022-08-31 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -11.00/+3.0/070 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE PATEINT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: D4 - UNIQUE IDENTIFIER (UDI) # : (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: PATIENT INFORMATION:UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # 726328. | -11.00/+3.0/070 | OS | Physician | 09/19/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15591852&pc=QCB | 2023826-2022-03494 | 2022-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -12.50/+2.0/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM # (B)(4). | -12.50/+2.0/091 | OS | Physician | 09/20/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15591879&pc=QCB | 2023826-2022-03488 | 2022-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -06.00/+1.5/003 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). | -06.00/+1.5/003 | OD | Physician | 09/20/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15591943&pc=QCB | 2023826-2022-03487 | 2022-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-12 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.00+4.0/041 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2022 BUT THE PROBLEM WAS NOT RESOLVED SO THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING WELL AND HAPPY AFTER EXCHANGE. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -08.00+4.0/041 | OS | Physician | 09/19/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15594494&pc=QCB | 2023826-2022-03595 | 2022-09-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM512.6, -11.0/4.0/015 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULT WITH ROTATION, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: AL THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -11.0/4.0/015 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULT WITH ROTATION. REPORTEDLY THE LENS WAS EXPLANTED AT AN UNKNOWN DATE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. D9: 31-DEC-2022. H3: DEVICE EVALUATION: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. H6: 4627 DEVICE EXPLANTED. 10. 4310. CLAIM#: (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.0/4.0/015 | OD | Physician | 09/13/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15601445&pc=QCB | 2023826-2022-03641 | 2022-03-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN -16.5/1.0/066 (SPHERE/CYLINDER/AXIS)IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2021. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: H6 - INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -16.5/1.0/066 | OS | Other Health Care Professional | 09/15/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15601489&pc=QCB | 2023826-2022-03640 | 2022-03-04 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-13 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN -16.5/1.5/180 (SPHERE/CYLINDER/AXIS)IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 -16.5/1.5/115 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -16.5/1.5/180 | OD | Other Health Care Professional | 09/15/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15608122&pc=QCB | 2023826-2022-03610 | 2022-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+2.0/084 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -9.5/+2.0/084 | OD | Physician | 09/16/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15608127&pc=QCB | 2023826-2022-03612 | 2022-05-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.5/102 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+1.5/102 | OD | Physician | 09/16/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15608153&pc=QCB | 2023826-2022-03609 | 2022-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+1.5/095 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). | -12.5/+1.5/095 | OS | Physician | 09/16/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15608169&pc=QCB | 2023826-2022-03606 | 2022-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.0/+2.0/082 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). | -16.0/+2.0/082 | OS | Physician | 09/16/2022 | 2022 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15608175&pc=QCB | 2023826-2022-03611 | 2022-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+3.5/087 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE CODE (B)(4). CLAIM# (B)(4). | -12.5/+3.5/087 | OD | Physician | 09/16/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15609999&pc=QCB | 2023826-2022-03615 | 2022-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.5/+5.0/008 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). | -17.5/+5.0/008 | OD | Physician | 09/16/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15610008&pc=QCB | 2023826-2022-03616 | 2022-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.0/110 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). | -8.5/+1.0/110 | OD | Physician | 09/16/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15610067&pc=QCB | 2023826-2022-03614 | 2022-08-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/075 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4). SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -10.0/+1.5/075 | OD | Physician | 09/16/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15610084&pc=QCB | 2023826-2022-03613 | 2022-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/+1.0/172 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -6.0/+1.0/172 | OD | Physician | 09/16/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15611544&pc=QCB | 2023826-2022-03645 | 2022-10-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-14 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -7.5/1.0/151 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -7.5/1.0/151 | OS | Physician | 10/11/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15620756&pc=QCB | 2023826-2022-03498 | 2022-05-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-17 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.00/+3.0/082 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO THE LENS TORE DURING DELIVERY INTO THE EYE. THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK; ETHNICITY: UNK; AND WEIGHT: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -11.00/+3.0/082 | OS | Physician | 09/19/2022 | 2022 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15620863&pc=QCB | 2023826-2022-03655 | 2022-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.5/+1.0/071 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -17.5/+1.0/071 | OD | Physician | 09/20/2022 | 2022 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15620866&pc=QCB | 2023826-2022-03650 | 2022-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.5/089 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+1.5/089 | OD | Physician | 09/19/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15620868&pc=QCB | 2023826-2022-03652 | 2022-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE (B)(4). CLAIM# (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). | -10.0/+1.5/090 | OS | Physician | 09/20/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15620877&pc=QCB | 2023826-2022-03654 | 2022-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.5/+3.0/160 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -17.5/+3.0/160 | OS | Physician | 09/20/2022 | 2022 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15620892&pc=QCB | 2023826-2022-03653 | 2022-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -9.5/+1.0/091 | OD | Physician | 09/20/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15620909&pc=QCB | 2023826-2022-03499 | 2022-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+1.5/007 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED HORIZONTALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -09.00/+1.5/007 | OD | Physician | 09/19/2022 | 2022 | 43 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15621434&pc=QCB | 2023826-2022-03671 |  | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -14.00/3.0/139 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. IT WAS INDICATED THE LENS WAS EXPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS WAS RETURNED IN A SPECIMEN CUP WITH MOISTURE. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN. (B)(4). | -14.00/3.0/139 | NA | Other Health Care Professional | 09/20/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15624387&pc=QCB | 2023826-2022-03536 | 2022-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.0/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT IN A SECOND SURGERY ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -12.0/1.0/089 | OS | Physician | 09/26/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15624388&pc=QCB | 2023826-2022-03535 | 2022-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.5/0.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -09.5/0.5/105 | OD | Physician | 09/26/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15624392&pc=QCB | 2023826-2022-03546 | 2022-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.0/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -12.0/2.0/093 | OD | Physician | 09/27/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15624395&pc=QCB | 2023826-2022-03541 | 2022-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.0/1.0/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. CLAIM# (B)(4). | -10.0/1.0/097 | OD | Physician | 09/27/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15624396&pc=QCB | 2023826-2022-03543 | 2022-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -09.0/2.0/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -09.0/2.0/112 | OD | Physician | 09/27/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15624399&pc=QCB | 2023826-2022-03539 | 2022-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -10.5/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.5/1.5/090 | OS | Physician | 09/26/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15624402&pc=QCB | 2023826-2022-03542 | 2022-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.0/1.0/003 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -11.0/1.0/003 | OD | Physician | 09/27/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15624403&pc=QCB | 2023826-2022-03552 | 2022-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.0/1.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. CLAIM# (B)(4). | -10.0/1.5/087 | OS | Physician | 09/27/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15624404&pc=QCB | 2023826-2022-03538 | 2022-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -12.0/1.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -12.0/1.0/093 | OD | Physician | 09/26/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15624406&pc=QCB | 2023826-2022-03545 | 2022-02-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.0/1.5/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -10.0/1.5/084 | OS | Physician | 09/27/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15624409&pc=QCB | 2023826-2022-03549 | 2022-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.5/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS VERTICALLY DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. (B)(4). CLAIM# (B)(4). | -10.5/1.5/089 | OD | Physician | 09/27/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15624411&pc=QCB | 2023826-2022-03558 | 2022-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -08.5/1.5/097 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -08.5/1.5/097 | OS | Physician | 09/27/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15624412&pc=QCB | 2023826-2022-03548 | 2022-06-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -07.5/3.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. (B)(4). CLAIM# (B)(4). | -07.5/3.0/085 | OD | Physician | 09/27/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15624415&pc=QCB | 2023826-2022-03560 | 2022-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -12.5/2.0/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. CLAIM# (B)(4). | -12.5/2.0/105 | OD | Physician | 09/30/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15624417&pc=QCB | 2023826-2022-03555 | 2022-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -09.5/3.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -09.5/3.5/090 | OD | Physician | 09/27/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15624420&pc=QCB | 2023826-2022-03551 | 2022-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.5/2.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. CLAIM# (B)(4). | -12.5/2.0/086 | OD | Physician | 09/27/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15624425&pc=QCB | 2023826-2022-03561 | 2022-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -08.0/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. (B)(4). CLAIM# (B)(4). | -08.0/1.0/089 | OD | Physician | 09/30/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15624427&pc=QCB | 2023826-2022-03556 | 2022-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 MM VTICMO12.6, -09.0/1.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -09.0/1.5/094 | OS | Physician | 09/27/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15624429&pc=QCB | 2023826-2022-03554 | 2022-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.5/1.0/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.5/1.0/072 | OS | Physician | 09/27/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15624436&pc=QCB | 2023826-2022-03562 | 2022-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -08.5/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. (B)(4). | -08.5/1.5/090 | OS | Physician | 09/30/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15624597&pc=QCB | 2023826-2022-03544 | 2022-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VICMO12.6, -09.5/2.5/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTION: B5 - INITIAL LENS REPORTED AS VICMO12.6 - CORRECT TO VTICMO12.6. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -09.5/2.5/073 | OS | Physician | 09/27/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15624614&pc=QCB | 2023826-2022-03559 | 2022-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -10.5/1.5/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTION: B3 - REPORTED AS 08-AUG-2022 - CORRECT TO 22-AUG-2022. B4 - REPORTED AS 27-SEP-2022 - CORRECT TO 30-SEP-2022. CLAIM# (B)(4). | -10.5/1.5/074 | OS | Physician | 09/27/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15629229&pc=QCB | 2023826-2022-03495 | 2022-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.50/+1.0/089 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO MEASUREMENT ERROR IN MR (PER THE SURGEON, THERE MIGHT BE AN ERROR IN MR BY THE MACHINE OR THE PRACTITIONER). THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -06.50/+1.0/089 | OD | Physician | 09/23/2022 | 2022 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15629366&pc=QCB | 2023826-2022-03504 | 2022-09-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.50/+1.0/069 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -17.50/+1.0/069 | NA | Physician | 09/20/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15629378&pc=QCB | 2023826-2022-03497 | 2022-04-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -09.00/+2.5/103 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL AND REFRACTION (FUNCTIONAL) VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -09.00/+2.5/103 | OD | Physician | 09/23/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15629684&pc=QCB | 2023826-2022-03657 | 2022-08-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.5/109 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.5/+1.5/109 | OS | Physician | 09/20/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15629685&pc=QCB | 2023826-2022-03659 | 2022-09-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-18 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+3.0/087 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT AND STATED "BIGGER SIZE LENS". Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -13.5/+3.0/087 | OD | Physician | 09/20/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15635994&pc=QCB | 2023826-2022-03706 |  | Injury | STAAR SURGICAL COMPANY | 2022-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Patient Device Interaction Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50/+6.0/133 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS REPLACED WITH A LONGER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL, WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS AND RESIDUE ON THE FOOTPLATE AND HAPTIC. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+6.0/133 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED A LOW VAULT WITH ROTATION. ON (B)(6) 2022 THE LENS EXCHANGED THE LENS WITH THE SAME MODEL; SIMILAR POWER AND LONGER LENGTH. THE CAUSE OF THE EVENT WAS A PATIENT RELATED EVENT AND THE DEVICE FAILED TO PERFORM AS INTENDED. REPORTEDLY "TORIC ICL ROTATED DUE TO UNDERSIZING". CLAIM#: (B)(4). Manufacturer Narrative: PT INFO: UNK. EVENT DATE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.50/+6.0/133 | OD | Physician | 09/21/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15636058&pc=QCB | 2023826-2022-03709 | 2022-06-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -5.5/1.0/90 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE PATIENT EXPERIENCED LOW VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND IT WAS NOTED "THE SURGERY SCHEDULE FOR THE REPLACEMENT LENS HAS NOT YET BEEN DETERMINED". THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC TORN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -5.5/1.0/90 | OD | Physician | 10/13/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15636061&pc=QCB | 2023826-2022-03710 | 2022-06-29 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -5.0/1.0/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND IT WAS NOTED "THE SURGERY SCHEDULE FOR THE REPLACEMENT LENS HAS NOT YET BEEN DETERMINED". THE PROBLEM IS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -5.0/1.0/89 | OS | Physician | 10/13/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15636257&pc=QCB | 2023826-2022-03503 | 2022-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00/+4.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO INACCURATE BIOLOGICAL MEASUREMENT DATA. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | -14.00/+4.0/091 | OS | Physician | 09/20/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15636267&pc=QCB | 2023826-2022-03496 | 2022-06-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -06.50/+1.0/087 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO MEASUREMENT ERROR IN MR (PER THE SURGEON, THERE MIGHT BE AN ERROR IN MR BY THE MACHINE OR THE PRACTITIONER). THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT: UNK, ETNINICITY: UNK, RACE: UNK. WORK ORDER SEARCH: ANOTHER SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). | -06.50/+1.0/087 | OS | Physician | 09/23/2022 | 2022 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15636449&pc=QCB | 2023826-2022-03660 | 2022-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.0/008 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON 13-SEP-2022 FOR A LENS TWO SIZES LONGER IN LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). | -11.0/+2.0/008 | OD | Physician | 09/21/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15636451&pc=QCB | 2023826-2022-03666 | 2022-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.0/082 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - #(B)(4). SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). | -11.0/+2.0/082 | OD | Physician | 09/21/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15636453&pc=QCB | 2023826-2022-03661 | 2022-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-19 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.0/180 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LENS TWO SIZES LONGER IN LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+2.0/180 | OS | Physician | 09/21/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15639495&pc=QCB | 2023826-2022-03667 | 2022-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+1.0/078 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -7.5/+1.0/078 | OD | Physician | 09/21/2022 | 2022 | 44 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15640338&pc=QCB | 2023826-2022-03668 | 2022-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.0/+2.0/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -16.0/+2.0/093 | OS | Physician | 09/23/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15640340&pc=QCB | 2023826-2022-03663 | 2022-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -10.0/+2.0/091 | OD | Physician | 09/22/2022 | 2022 | 17 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15640345&pc=QCB | 2023826-2022-03670 | 2022-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+4.0/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -13.5/+4.0/093 | OS | Physician | 09/23/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15640346&pc=QCB | 2023826-2022-03669 | 2022-06-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-20 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022 . WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: CLAIM#(B)(4). | -7.0/+1.0/093 | OD | Physician | 09/23/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15651827&pc=QCB | 2023826-2022-03629 | 2022-08-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -4.0/5.5/075 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE LEFT EYE (OS) OF A PATIENT WITH PRE-EXISTING CORENAL OR REFRACTIVE SURGERY, A HISTORY OF KERATITIS, AND PREVIOUS DALK, ON (B)(6) 2022. LOW VAULT WITHOUT ROTATION WAS OBSERVED. REPORTEDLY, "LENS TO BE EXCHANGED." Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. PATIENT PRE-EXISTING CONDITIONS: PREVIOUS CORNEAL OR REFRACTIVE SURGERY, HISTORY OF KERATITIS, PREVIOUS DALK. CLAIM# (B)(4). | -4.0/5.5/075 | OS | Physician | 09/21/2022 | 2022 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15651830&pc=QCB | 2023826-2022-03701 | 2022-09-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -13.00/+1.00/122 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM# (B)(4). | -13.00/+1.00/122 | OD | Physician | 09/23/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15651946&pc=QCB | 2023826-2022-03631 | 2022-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.0/1.5/077 (SPHERE/CYLINDER/AXIS), WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED CAUSE IS REPORTED AS OTHER. Manufacturer Narrative: (B)(4). | -10.0/1.5/077 | OD | Physician | 09/21/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15652233&pc=QCB | 2023826-2022-03509 | 2022-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+1.5/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2022 BUT THE PROBLEM WAS NOT RESOLVED SO THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM # (B)(4). Manufacturer Narrative: B5 - "THE REPLACEMENT LENS WAS IMPLANTED VERTICALLY." SHOULD BE ADDED TO THE INITIAL MDR. D4 - UNIQUE IDENTIFIER (UDI)# (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -09.00/+1.5/074 | OS | Physician | 09/22/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15653306&pc=QCB | 2023826-2022-03508 | 2022-06-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -09.00/+1.0/104 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM # (B)(4). | -09.00/+1.0/104 | OD | Physician | 09/22/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15653423&pc=QCB | 2023826-2022-03764 |  | Injury | STAAR SURGICAL COMPANY | 2022-10-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11/+1.00/62 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). | -11/+1.00/62 | OS | Physician | 09/23/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15654701&pc=QCB | 2023826-2022-03783 | 2022-07-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-21 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -11.5/1.5/99 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE PATIENT EXPERIENCED LOW VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: DATE RETURNED: 03-NOV-2022. H3: DEVICE EVALUATION: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -11.5/1.5/99 | OS | Physician | 10/13/2022 | 2022 | 46 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15659546&pc=QCB | 2023826-2022-03715 | 2022-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+4.0/179 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -11.5/+4.0/179 | OD | Physician | 09/26/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15659554&pc=QCB | 2023826-2022-03718 | 2022-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+2.5/176 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). | -7.5/+2.5/176 | OD | Physician | 09/27/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15659564&pc=QCB | 2023826-2022-03712 | 2022-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.5/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LENS TWO SIZES SHORTER IN LENGTH DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -10.0/+2.5/091 | OD | Physician | 09/26/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15659576&pc=QCB | 2023826-2022-03716 | 2022-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.5/+2.5/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). | -15.5/+2.5/095 | OD | Physician | 09/27/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15659593&pc=QCB | 2023826-2022-03714 | 2022-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-24 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+0.5/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). | -6.5/+0.5/094 | OD | Physician | 09/26/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15664689&pc=QCB | 2023826-2022-03785 | 2022-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+3.0/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 . THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -14.0/+3.0/094 | OS | Physician | 09/27/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15664691&pc=QCB | 2023826-2022-03786 | 2022-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.0/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -7.0/+1.0/093 | OS | Physician | 09/27/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15665926&pc=QCB | 2023826-2022-03676 | 2022-05-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6, -10.5/1.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. CLAIM # (B)(4). | -10.5/1.5/089 | OD | Physician | 10/04/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15665940&pc=QCB | 2023826-2022-03683 | 2022-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.5/2.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED IN A SECOND SURGERY ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -11.5/2.0/085 | OS | Physician | 10/07/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15665941&pc=QCB | 2023826-2022-03679 | 2022-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.0/1.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.0/1.0/088 | OD | Physician | 10/07/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15665943&pc=QCB | 2023826-2022-03687 | 2022-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -09.5/2.5/077 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED DUE TO LOW VAULT ON 14/JUL/2022. IN A SECOND SURGERY ON (B)(6) 2022, A LONGER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -09.5/2.5/077 | OS | Physician | 10/07/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15665944&pc=QCB | 2023826-2022-03680 | 2022-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2, -11.0/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED IN A SECOND SURGERY ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -11.0/2.0/088 | OS | Physician | 10/07/2022 | 2022 | 17 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15665945&pc=QCB | 2023826-2022-03689 | 2022-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -09.5/1.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED VERTICALLY WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. (B)(4). CLAIM# (B)(4). | -09.5/1.0/094 | OD | Physician | 10/07/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15665946&pc=QCB | 2023826-2022-03682 | 2022-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -13.0/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.0/1.5/091 | OS | Physician | 10/07/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15665947&pc=QCB | 2023826-2022-03681 | 2022-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -12.0/1.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. (B)(4). | -12.0/1.5/094 | OD | Physician | 10/07/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15665949&pc=QCB | 2023826-2022-03692 | 2022-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -08.5/2.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -08.5/2.5/095 | OS | Physician | 10/07/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15665953&pc=QCB | 2023826-2022-03684 | 2022-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.0/2.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. (B)(4). CLAIM# (B)(4). | -10.0/2.5/085 | OD | Physician | 10/07/2022 | 2022 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15665955&pc=QCB | 2023826-2022-03686 | 2022-10-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.0/1.5/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. (B)(4). CLAIM# (B)(4). | -12.0/1.5/084 | OD | Physician | 10/07/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15665956&pc=QCB | 2023826-2022-03688 | 2022-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.5/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.5/1.5/092 | OD | Physician | 10/07/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15665958&pc=QCB | 2023826-2022-03693 | 2022-08-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.5/2.0/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -14.5/2.0/112 | OS | Physician | 10/07/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15665959&pc=QCB | 2023826-2022-03691 | 2022-07-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -08.5/2.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. CLAIM# (B)(4). | -08.5/2.0/085 | OD | Physician | 10/07/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15665960&pc=QCB | 2023826-2022-03690 | 2022-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.0/2.0/015 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -12.0/2.0/015 | OD | Physician | 10/07/2022 | 2022 | 45 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15665962&pc=QCB | 2023826-2022-03694 | 2022-08-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.0/1.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK. (B)(4). CLAIM# (B)(4). | -14.0/1.5/094 | OD | Physician | 10/07/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15667543&pc=QCB | 2023826-2022-03784 | 2022-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-25 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -9.5/2.5/175 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE PATIENT EXPERIENCED LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND IMPLANTED HORIZONTAL; OPPOSITE OF THE INITIAL LENS VERTICAL IMPLANT POSITION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: MEDICAL DEVICE PROBLEM CODE (B)(4). CLAIM#: (B)(4). | -9.5/2.5/175 | OS | Physician | 10/19/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15676234&pc=QCB | 2023826-2022-03840 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: IT WAS REPORTED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -13.50/1.5/107 (SPHERE/CYLINDER/AXIS) TORE DURING LOADING. THERE WAS NO PATIENT CONTACT. A BACKUP LENS WAS IMPLANTED SUCCESSFULLY. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.50/1.5/107 | NA | Physician | 09/26/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15677441&pc=QCB | 2023826-2022-03773 | 2021-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.00/+2.5/087 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN. CLAIM # (B)(4). | -07.00/+2.5/087 | OD | Physician | 09/27/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15677454&pc=QCB | 2023826-2022-03772 | 2022-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.50/+3.0/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: B5 - THE LENS WAS EXCHANGED ON (B)(6) 2022 DUE TO REFRACTIVE SURPRISE WITH SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE LENS OPTIC SPLIT AND THE HAPTIC TORN AND BROKEN. OVERALL LENGTH (DIAMETER) VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. UNABLE TO PERFORM WIDTH INSPECTION DUE TO DAMAGED STATE OF LENS. CLAIM#: (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.50/+3.0/092 | OS | Physician | 09/26/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15677472&pc=QCB | 2023826-2022-03771 | 2021-12-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -13.50/+3.5/091 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2020. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | -13.50/+3.5/091 | OD | Physician | 09/26/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15677508&pc=QCB | 2023826-2022-03769 | 2022-04-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.00/+2.5/092 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY, IN A MICROCENTRIFUGE VIAL, WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BROKEN. DIMENSIONAL VERIFICATION WAS PERFORMED AND THE LENS WAS FOUND TO BE IN SPECIFICATION. CLAIM#: (B)(4). | -11.00/+2.5/092 | OS | Physician | 09/26/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15677509&pc=QCB | 2023826-2022-03770 | 2021-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+3.5/001 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.50/+3.5/001 | OS | Physician | 09/26/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15677514&pc=QCB | 2023826-2022-03768 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50/+1.5/013 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6)2022. THIS LENS WAS THE REPLACEMENT LENS FOR MFR REPORT # 2023826-2022-02420 BUT THE EXCHANGE DID NOT RESULT ANY SIGNIFICANT IMPROVEMENT (LOW VAULT). THE PATIENT IS BEING CLOSELY MONITORED. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. EXPLANT DATE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). | -09.50/+1.5/013 | OD | Physician | 09/27/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15677895&pc=QCB | 2023826-2022-03842 | 2022-05-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VICMO13.2; -11.0/2.0/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE. INTRAOPERATIVELY THE LENS WAS REMOVED. THERE WAS NO PATIENT INJURY. ON (B)(6) 2022 A REPLACEMENT LENS OF THE SAME MODEL/LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -11.0/2.0/094 | OS | Physician | 10/20/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15677896&pc=QCB | 2023826-2022-03841 | 2022-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -13.5/1.5/072 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING. LATER ON THIS SAME DATE THE LENS WAS EXPLANTED. IT WAS ADDITIONAL NOTED THERE WILL BE A SECOND LENS IMPLANTATION BUT THERE IS NO CONFIRMED DATE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -13.5/1.5/072 | OS | Physician | 10/20/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15677905&pc=QCB | 2023826-2022-03789 | 2022-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/068 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). | -9.5/+1.0/068 | OD | Physician | 09/28/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15677908&pc=QCB | 2023826-2022-03793 | 2022-06-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-26 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+1.0/172 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LENS TWO SIZES SHORTER IN LENGTH DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -6.5/+1.0/172 | OS | Physician | 09/28/2022 | 2022 | 42 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15678925&pc=QCB | 2023826-2022-03844 | 2022-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -9.5/1.5/73 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6)2022 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.5/1.5/73 | OS | Physician | 09/27/2022 | 2022 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15679001&pc=QCB | 2023826-2022-03847 | 2022-09-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6)2022 A 13.7MM, VTICM5_13.7, -8.50/1.5/084 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING LOADING. THERE WAS NO PATIENT CONTACT. ON (B)(6)2022 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -8.50/1.5/084 | NA | Physician | 09/28/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15683667&pc=QCB | 2023826-2022-03778 | 2022-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+5.5/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: (B)(4): UNK . ETHNICITY: UNK. RACE: UNK. CLAIM # (B)(4). | -11.50/+5.5/095 | OS | Physician | 09/27/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15683843&pc=QCB | 2023826-2022-03779 | 2022-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2022 BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED . THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC DEFORMED TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -10.50/+1.0/090 | OS | Physician | 09/28/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15685019&pc=QCB | 2023826-2022-03796 | 2022-06-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/106 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -10.0/+1.5/106 | OD | Physician | 09/29/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15685120&pc=QCB | 2023826-2022-03801 | 2022-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+2.0/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -14.0/+2.0/090 | OS | Physician | 09/30/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15685133&pc=QCB | 2023826-2022-03798 | 2022-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+2.5/176 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -8.0/+2.5/176 | OD | Physician | 09/29/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15685170&pc=QCB | 2023826-2022-03797 | 2022-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+3.0/163 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -9.5/+3.0/163 | OS | Physician | 09/29/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15685196&pc=QCB | 2023826-2022-03795 | 2022-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+3.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -7.0/+3.0/093 | OD | Physician | 09/29/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15685213&pc=QCB | 2023826-2022-03800 | 2022-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.5/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). | -9.0/+1.5/089 | OS | Physician | 09/30/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15685252&pc=QCB | 2023826-2022-03799 | 2022-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+2.0/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). | -9.0/+2.0/094 | OD | Physician | 09/30/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15686866&pc=QCB | 2023826-2022-03577 | 2022-09-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF -13.0/3.0/058 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. LENS ROTATION NOT ASSOCIATED TO LOW VAULT WAS OBSERVED. LENS ROTATED 90 DEGREES. ON (B)(6) 2022 THE LENS WAS REMOVED AND ON SAME DAY DIFFERENT SURGERY A SPHERICAL LENS WAS IMPLANTED OF LONGER LENGTH AND THE PROBLEM RESOLVED. PATIENT WILL REQUIRE LASIK SURGERY TO ADDRESS THE ASTIGMATISM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN STUCK IN TAPE DRY. VISUAL INSPECTION FOUND HAPTIC TORN AND TAPE RESIDUE THROUGHOUT THE LENS. CLAIM#: (B)(4). | -13.0/3.0/058 | OD | Physician | 09/27/2022 | 2022 | 43 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15686873&pc=QCB | 2023826-2022-03578 | 2022-09-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF -11.50/4.0/111 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. LENS ROTATION NOT ASSOCIATED TO LOW VAULT WAS OBSERVED. LENS ROTATED 90 DEGREES. ON (B)(6) 2022 THE LENS WAS REMOVED AND ON SAME DAY DIFFERENT SURGERY A SPHERICAL LENS WAS IMPLANTED OF LONGER LENGTH AND THE PROBLEM RESOLVED. PATIENT WILL REQUIRE LASIK SURGERY TO ADDRESS THE ASTIGMATISM. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED STUCK IN TAPE DRY. VISUAL INSPECTION FOUND HAPTIC TORN AND TAPE RESIDUE THROUGHOUT THE LENS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -11.50/4.0/111 | OS | Physician | 09/27/2022 | 2022 | 43 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15686937&pc=QCB | 2023826-2022-03579 | 2022-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -10.0/4.5/009 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LOW VAULT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL AND LENGTH LENS WITH DIFFERENT POWER. THE REPORTER STATES THE LENS WAS PLACED VERTICALLY INTO THE EYE FOR THE FIRST TIME AND THE REPLACEMENT WAS CHANGED TO HORIZONTAL POSITION. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -10.0/4.5/009 | OD | Physician | 09/27/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15686955&pc=QCB | 2023826-2022-03580 | 2022-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -10.0/2.5/002 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LOW VAULT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL AND LENGTH LENS WITH DIFFERENT POWER. THE REPORTER STATES THE LENS WAS PLACED VERTICALLY INTO THE EYE FOR THE FIRST TIME AND THE REPLACEMENT WAS CHANGED TO HORIZONTAL PLACEMENT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC/ HAPTIC DEFORMED . DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -10.0/2.5/002 | OS | Physician | 09/27/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15687049&pc=QCB | 2023826-2022-03582 |  | Injury | STAAR SURGICAL COMPANY | 2022-10-27 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF -10.0/1.0/153 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LOW VAULT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND THE PROBLEM WAS RESOLVED. PATIENT IS HAPPY. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. THE RETURNED LENS IS NOT THE SIZE STATED IN THE CLAIM. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -10.0/1.0/153 | OD | Physician | 09/28/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15690549&pc=QCB | 2023826-2022-03850 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-10-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -8.50/2.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022 AS A REPLACEMENT LENS TO ADDRESS ASTIGMATISM (HIGH CYLINDER POWER). IT WAS REPORTED THAT THE PROBLEM RESOLVED BUT FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) " LOW VAULT" WAS REPORTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5-PER UPDATED INFORMATION THE REPORTER STATED THAT AFTER POST-OPERATIVE APPOINTMENT THE VAULT WAS DOCUMENTED AS THE OD 500UM AND THE SURGEON DETERMINED THAT MEDICAL INTERVENTION IS NOT NECESSARY AT THIS TIME. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: 1494- OFF-LABEL USE SHOULD BE ADDED FOR ACD<3.00MM. ADDITIONAL INFORMATION: CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -8.50/2.5/090 | OD | Physician | 09/29/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15690890&pc=QCB | 2023826-2022-03853 | 2022-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO_12.6 IMPLANTABLE COLLAMER LENS -17.0/2.5/109 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2022 THE LENS WAS EXPLANED DUE TO EXCESSIVE VAULT. THE EXPLANT RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. THE REPORTER STATED THE SURGEON WONT BE IMPLANTING ANOTHER LENS. Manufacturer Narrative: H3 - LENS WAS RETURNED IN A MICRO-CENTRIFUGE VIAL AND DRY WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUDN THE LENS HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -17.0/2.5/109 | OS | Physician | 09/30/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15690996&pc=QCB | 2023826-2022-03854 | 2022-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO_12.6 IMPLANTABLE COLLAMER LENS -18.0/3.0/075 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANED DUE TO EXCESSIVE VAULT. THE EXPLANT RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. THE REPORTER STATED THE SURGEON WONT BE IMPLANTING ANOTHER LENS. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -18.0/3.0/075 | OD | Physician | 09/30/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15691052&pc=QCB | 2023826-2022-03851 | 2022-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS -7.5/1.0/70 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) IN JULY OF 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. THE REPORTER STATED " WE SPECULATE THAT THE REASON FOR THE HIGH VAULT WAS DUE TO INCORRECT LENS SIZE SELECTION." Manufacturer Narrative: IMPLANT DATE: JULY 2022. CLAIM# (B)(4). | -7.5/1.0/70 | OS | Physician | 09/29/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15693830&pc=QCB | 2023826-2022-03780 | 2021-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-28 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -10.50/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | -10.50/+2.0/090 | OD | Physician | 09/29/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15694614&pc=QCB | 2023826-2022-03782 | 2022-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.50/+1.5/066 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.50/+1.5/066 | OS | Physician | 09/29/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15696145&pc=QCB | 2023826-2022-03781 | 2022-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-29 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.00/+1.5/106 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -14.00/+1.5/106 | OD | Physician | 09/29/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15697662&pc=QCB | 2023826-2022-03890 |  | Injury | STAAR SURGICAL COMPANY | 2022-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A SURGEON IMPLANTED A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -13.5/2.0/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED UPSIDE DOWN, AND TORE WHILE EXPLANTING. THE BACKUP LENS WAS THEN IMPLANTED. Manufacturer Narrative: PATIENT IDENTIFIER, AGE OR DATE OF BIRTH, SEX, WEIGHT, ETHNICITY, RACE: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: G2: REPORT SOURCE: FOREIGN SHOULD BE OMMITTED FOR CORRECTION. H3: DEVICE EVALUATED BY MANUFACTURER- NO SHOULD BE SELECTED FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H1: TYPE OF REPORABLE EVENT-MALFUNCTION SHOULD BE OMMITTED AND SERIOUS INJURY SHOULD BE SELECTED FOR CORRECTION. H6: HEALTH EFFECT: IMPACT CODE: 4627 SHOULD BE ADDED FOR DEVICE EXPLANTATION. CLAIM# (B)(4). | -13.5/2.0/088 | NA | Physician | 10/05/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15697666&pc=QCB | 2023826-2022-03889 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -13.5/1.5/087 (SPHERE/CYLINDER/AXIS) WAS TORN WHEN THE SURGEON DELIEVERED IT INTO THE PATIENT'S LEFT EYE (OS), WITH THE LIOLI INJECTOR. THE LENS WAS REMOVED. Manufacturer Narrative: PATIENT IDENTIFIER, AGE OR DATE OF BIRTH, SEX, WEIGHT, ETHNICITY, RACE: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS. CLAIM# (B)(4). | -13.5/1.5/087 | OS | Physician | 10/03/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15697669&pc=QCB | 2023826-2022-03888 | 2022-09-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -13.0/1.0/091 (SPHERE/CYLINDER/AXIS) WAS TORN WHEN THE SURGEON DELIEVERED IT INTO THE PATIENT'S RIGHT EYE (OD) WITH THE LIOLI INJECTOR. THE LENS WAS REMOVED. Manufacturer Narrative: PATIENT IDENTIFIER, AGE OR DATE OF BIRTH, SEX, WEIGHT, ETHNICITY, RACE:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS. CLAIM #(B)(4) 1D COMPLAINT NOT SIMILAR TO THIS COMPLAINT. CLAIM# (B)(4). Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -13.0/1.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED UPSIDE DOWN. PATIENT CONTACT BUT NO PATIENT INJURY. LENS WAS IMPLANTED, REMOVED AND REPLACED INTRAOPERATIVELY ON (B)(6) 2022 AND PROBLEM WAS RESOLVED. CAUSE OF EVENT WAS REPORTED AS "USER ERROR". THE DEVICE DID NOT FAIL TO PERFORM. "SURGEON IMPLANT ERROR" WAS REPORTED. CLAIM#: (B)(4). | -13.0/1.0/091 | OD | Physician | 12/29/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15697671&pc=QCB | 2023826-2022-03893 | 2022-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-10-30 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -6.0/3.0/11 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2022, THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT, SIGNIFICANT REDUCTION OF ENDOTHELIAL CELL COUNT OR SIGNIFICANT ENDOTHELIAL CELL LOSS, AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. CAUSE IS REPORTED AS THE DEVICE. REPORTEDLY, "SIZING PROBLEM, OUTLIER TO THE ONLINE CALCULATOR." STATUS OF THE EYE IS REPORTED AS, "VISION RETURNED TO PRE-ICL LEVEL, EYE QUIET, 2 CLOCK HOUR IRIS ATROPY FROM THE EXPLANT SURGERY, CORNEAL ENDOTHELIAL CELL DROPPED FROM 3044 (PREOP) TO 2499." THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: D4: VTICMO_13.2 SHOULD BE CORRECTED TO VTICM5_13.2 . ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -6.0/3.0/11 | OS | Physician | 10/06/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15712296&pc=QCB | 2023826-2022-03911 | 2022-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.5/+1.5/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -4.5/+1.5/093 | OD | Physician | 10/03/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15712297&pc=QCB | 2023826-2022-03912 | 2020-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.0/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+2.0/094 | OS | Physician | 10/03/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15712748&pc=QCB | 2023826-2022-03900 | 2022-08-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -8.0/2.5/150 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME SIZE LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN.STATUS OF THE EYE IS REPORTED AS, "OK." Manufacturer Narrative: CORRECTED DATA: H6- INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. DEVICE EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC/HAPTIC DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -8.0/2.5/150 | OD | Physician | 10/08/2022 | 2022 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15712751&pc=QCB | 2023826-2022-03899 | 2022-05-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-01 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.0/1.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED, AND REPLACED FOR A SMALLER LENGTH LENS DUE TO A LENS TEAR DURING LOADING. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: B1 - CORRECTED TO: PRODUCT PROBLEM (E.G., DEFECTS/MALFUNCTIONS) IN THE INITIAL MDR. B2 - REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE SHOULD BE REMOVED FROM INITIAL MDR. H1 - CORRECTED TO: MALFUNCTION IN THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: D4: CORRECTED TO- MODEL # VTICMO_13.2 IN THE INITIAL MDR. CLAIM # (B)(4) Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE THROUGHOUT THE LENS IN A VIAL . VISUAL INSPECTION FOUND THE HAPTIC TORN TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -9.0/1.5/092 | OD | Physician | 10/07/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15713725&pc=QCB | 2023826-2022-03877 | 2022-10-06 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/126 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO THE BACK HAPTIC OF THE LENS WAS HUNG UP IN THE INJECTOR AND WAS DAMAGED. THERE WAS NO PATIENT INJURY NOTED. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS STABLE. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.00/+1.0/126 | NA | Other Health Care Professional | 10/06/2022 | 2022 | 48 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15713869&pc=QCB | 2023826-2022-03882 | 2022-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -04.50/+1.5/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: D4- UNIQUE IDENTIFIER (UDI) #- (B)(4). CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -04.50/+1.5/091 | OS | Physician | 10/03/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15713881&pc=QCB | 2023826-2022-03881 | 2022-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.50/+2.5/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK; ETHNICITY: UNK; RACE: UNK. (B)(4). Manufacturer Narrative: D4- UNIQUE IDENTIFIER (UDI)#- (B)(4). H6- INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. DEVICE EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -13.50/+2.5/091 | OD | Physician | 10/03/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15713906&pc=QCB | 2023826-2022-03879 | 2022-08-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.00/+3.0/091 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE LENS IS STABLE AND PATIENT IS HAPPY. THE CAUSE WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5-"THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.00/3.0/091 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED A LOW VAULT WITH ROTATION. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH THE DIFFERENT MODEL/LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. REPORTEDLY "LENS IS STABLE AFTER EXCHANGING WITH THE BIGGER SIZE LENS. PATIENT IS HAPPY". THE REPORTER INDICATED THE CAUSE OF THE EVENT IS UNKNOWN." CLAIM# (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | -13.00/+3.0/091 | OD | Physician | 10/03/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15713988&pc=QCB | 2023826-2022-03920 | 2022-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 -7.5/1.5/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS' LEFT EYE (OS) ON (B)(6)2022 . THE PATIENT EXPERIENCED EXCESSIVE VAULT, LENS REMAINS IMPLANTED. IT WAS REPORTED THAT THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. FOR CAUSE DETAILS THE REPORTER STATES " ICL CALCULATION SOFTWARE AND/OR WTW FROM MY DEVICE (ALCON BIOGRAPH) OVER SIZING?". IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5-PER UPDATED INFORMATION ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS. THE REPORTER STATES "THE EXCHANGE WAS UNEVENTFUL". CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | -7.5/1.5/094 | OS | Physician | 10/03/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15714110&pc=QCB | 2023826-2022-03724 | 2022-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.5/2.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY, AND RACE- UNK. (B)(4). | -13.5/2.5/094 | OS | Physician | 10/08/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15714117&pc=QCB | 2023826-2022-03727 | 2022-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.00/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -15.00/2.0/088 | OD | Physician | 10/08/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15714118&pc=QCB | 2023826-2022-03730 | 2022-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.5/1.5/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. CLAIM# (B)(4). | -13.5/1.5/075 | OS | Physician | 10/08/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15714121&pc=QCB | 2023826-2022-03734 | 2022-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.00/0.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -10.00/0.5/083 | OD | Physician | 10/08/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15714122&pc=QCB | 2023826-2022-03733 | 2022-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.0/2.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK, ETHNICITY - UNK, RACE - UNK. (B)(4). | -10.0/2.5/089 | OS | Physician | 10/08/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15714128&pc=QCB | 2023826-2022-03738 | 2022-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.0/3.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. CLAIM# (B)(4). | -14.0/3.0/091 | OS | Physician | 10/08/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15714129&pc=QCB | 2023826-2022-03737 | 2022-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.5/2.0/142 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.5/2.0/142 | OS | Physician | 10/08/2022 | 2022 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15714135&pc=QCB | 2023826-2022-03745 | 2022-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -07.0/2.0/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. CLAIM# (B)(4). | -07.0/2.0/072 | OS | Physician | 10/11/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15714142&pc=QCB | 2023826-2022-03751 | 2022-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -08.50/5.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. CLAIM # (B)(4). | -08.50/5.0/087 | OS | Physician | 10/11/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15714146&pc=QCB | 2023826-2022-03747 | 2022-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -04.5/1.0/143 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. (B)(4). | -04.5/1.0/143 | OD | Physician | 10/11/2022 | 2022 | 46 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15714150&pc=QCB | 2023826-2022-03752 | 2022-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -09.0/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. CLAIM# (B)(4). | -09.0/2.0/091 | OD | Physician | 10/12/2022 | 2022 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15714153&pc=QCB | 2023826-2022-03749 | 2022-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -10.50/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -10.50/2.5/093 | OS | Physician | 10/11/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15714156&pc=QCB | 2023826-2022-03750 | 2022-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6, -7.0/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. CLAIM # (B)(4). | -7.0/1.5/093 | OD | Physician | 10/11/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15714162&pc=QCB | 2023826-2022-03758 | 2022-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.5/3.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.5/3.0/092 | OS | Physician | 10/12/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15714165&pc=QCB | 2023826-2022-03760 | 2021-02-23 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -6.5/1.0/055 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -6.5/1.0/055 | OS | Physician | 05/24/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15714170&pc=QCB | 2023826-2022-03757 | 2022-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. CLAIM# (B)(4). | -10.5/2.5/091 | OD | Physician | 10/12/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15714173&pc=QCB | 2023826-2022-03763 | 2022-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/2.0/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK, ETHNICITY- UNK, RACE - UNK. (B)(4). | -10.5/2.0/110 | OD | Physician | 10/14/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15714177&pc=QCB | 2023826-2022-03761 | 2022-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.5/1.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. CLAIM# (B)(4). | -13.5/1.0/095 | OS | Physician | 10/14/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15714226&pc=QCB | 2023826-2022-03741 | 2022-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.0/4.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.0/4.0/089 | OS | Physician | 10/11/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15714239&pc=QCB | 2023826-2022-03746 | 2022-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -06.5/1.5/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -06.5/1.5/078 | OS | Physician | 10/11/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15714251&pc=QCB | 2023826-2022-03762 | 2022-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.5/0.5/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. (B)(4). | -12.5/0.5/112 | OD | Physician | 10/14/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15714371&pc=QCB | 2023826-2022-03902 | 2022-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -13.5/3.0/091(SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO LOW VAULT. THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXCHANGED AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: D4: MODEL NUMBER: SHOULD BE CORRECTED TO VTICMO_12.6. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN AND SPLIT AND THE HAPTIC TORN. DIMENSIONAL INSPECTION WAS UNABLE TO BE PERFORMED DUE TO SIGNIFICANT LENS DAMAGE. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -13.5/3.0/091 | OS | Physician | 10/08/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15714375&pc=QCB | 2023826-2022-03901 | 2022-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.0/2.0/004 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. IN A SUBSEQUENT SURGERY LATER THAT DAY, THE LENS WAS REMOVED AND A SAME SIZE REPLACEMENT LENS WAS IMPLANTED VERTICALLY DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "FIRST VERTICAL POSITION." THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -9.0/2.0/004 | OD | Physician | 10/08/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15714381&pc=QCB | 2023826-2022-03905 | 2022-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -8.5/3.0/015 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS INTRAOPERATIVELY REMOVED AND A SAME SIZE REPLACEMENT LENS WAS IMPLANTED VERTICALLY DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "THE LENS EXTRACTED WAS IMPLANTED VERTICALLY." THE STATUS OF THE EYE REPORTED AS, "OK." THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6- INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. DEVICE EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -8.5/3.0/015 | OS | Physician | 10/07/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15714832&pc=QCB | 2023826-2022-03891 | 2022-07-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -13.5/1.0/067 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. IN A SUBSEQUENT SURGERY LATER THAT DAY, THE LENS WAS REMOVED DUE TO LENS DISLOCATION/SUBLUXATION. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS A PATIENT RELATED FACTOR. REPORTEDLY, "THIS PATIENT HAD SOME PROBLEMS WITH THE SUSPENSORY LIGAMENT, WHICH CAUSED THE IMPLANTED LENS TO DEVIATE." Manufacturer Narrative: (B)(4). | -13.5/1.0/067 | OD | Physician | 10/11/2022 | 2022 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15716524&pc=QCB | 2023826-2022-03862 |  | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -12.50/2.5/110 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM#: (B)(4). | -12.50/2.5/110 | OD | Physician | 10/03/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15716563&pc=QCB | 2023826-2022-03887 | 2022-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -04.00/+1.0/008 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO THE WRONG LENS WAS CALCULATED. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS DUE TO POORLY DETERMINED INITIAL REFRACTION. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: 12/15/2022. H3: PRODUCT EVALUATION: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM (B)(4). Manufacturer Narrative: CORRECTED DATA: D4: UNIQUE IDENTIFIER NUMBER- (B)(4). CLAIM# (B)(4). | -04.00/+1.0/008 | OD | Physician | 10/07/2022 | 2022 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15716679&pc=QCB | 2023826-2022-03861 |  | Injury | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -11.0/2.5/110 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM#: (B)(4). | -11.0/2.5/110 | OS | Physician | 10/03/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15716999&pc=QCB | 2023826-2022-03973 | 2022-09-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -10.5/3.5/001 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. INTRAOPERATIVELY THE LENS WAS REMOVED. IT WAS REPORTED THE ICL COULD NOT BE POSITIONED IN THE CENTER AND NOTED "THE PATIENT HAD CATARACT SURGERY; THE TICL CAN'T BE CENTERED". THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND THE DEVICE FAILED TO PERFORM AS INTENDED. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). | -10.5/3.5/001 | OD | Physician | 11/01/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15719162&pc=QCB | 2023826-2022-03863 | 2022-09-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-02 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -13.0/2.0/087 (SPHERE/CYLINDER/AXIS) LENS TEAR/BREAK DURING LOADING. THE LENS WAS NOT IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). THE SURGERY WAS POSTPONED. THE CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM #(B)(4). | -13.0/2.0/087 | OS | Physician | 10/03/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15723057&pc=QCB | 2023826-2022-03977 | 2022-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-03 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -12.5/2.0/070 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2022 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS BUT SPHERICAL LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -12.5/2.0/070 | OS | Physician | 11/02/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15725342&pc=QCB | 2023826-2022-03918 | 2022-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.0/108 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.5/+1.0/108 | OD | Physician | 10/07/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15727784&pc=QCB | 2023826-2022-03919 | 2022-08-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+1.5/084 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -12.0/+1.5/084 | OS | Physician | 10/07/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15729092&pc=QCB | 2023826-2022-03916 | 2022-08-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.0/075 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022 . WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM #(B)(4). | -10.5/+2.0/075 | OD | Physician | 10/07/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15733854&pc=QCB | 2023826-2022-03924 | 2022-08-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -13.0/3.5/065 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND LATER REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM RESOLVED. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.0/3.5/065 | OD | Physician | 10/05/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15737073&pc=QCB | 2023826-2022-03927 | 2022-09-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 13.2MM, VTICM5_13.2, -13.00/1.5/080 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING LOADING. THERE WAS NO PATIENT CONTACT. ON (B)(6) 2022 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | -13.00/1.5/080 | NA | Physician | 10/07/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15738259&pc=QCB | 2023826-2022-03926 | 2022-06-17 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 13.2MM, VTICM5_13.2, -13.00/1.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE DURING LOADING. THERE WAS NO PATIENT CONTACT. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | -13.00/1.5/083 | NA | Physician | 10/07/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15740042&pc=QCB | 2023826-2022-03864 | 2022-09-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement; No Apparent Adverse Event; Ejection Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -12.50/3.5/089 (SPHERE/CYLINDER/AXIS) LENS TORE/BROKE DURING INJECTION INTO THE EYE ON (B)(6)2022. THE LENS WAS NOT IMPLANTED INTO THE PATIENT BUT IT DID HAVE PATIENT CONTACT WITH NO PATIENT INJURY. REPORTEDLY THE DOCTOR STOPPED AND RESTARTED INJECTING DUE TO THE INJECTOR RESISTANCE AND THAT PROBABLY DID NOT HELP. ON THE SAME DAY SAME SURGERY THE BACKUP LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE. " DOING GREAT, SHOWING SLIGHTLY MORE CYL THAN EXPECTED BUT IT'S STILL EARLY." CAUSE REPORTED AS DEVICE AND USER ERROR. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT A 13.2MM VTICM5_ 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.50/3.5/089 (SPHERE/CYLINDER/AXIS) GOT STUCK IN THE CARTRIDGE/INJECTION SYSTEM; THEN THE LENS WAS IMPLANTED UPSIDE DOWN ON (B)(6)2022. REPORTEDLY THE SURGEON "DID NOT LIKE HOW THE PRIMARY LENS LOOKED WHEN TRYING TO IMPLANT IT". THERE WAS PATIENT CONTACT BUT NO PATIENT INJURY. THE LENS WAS IMPLANTED; REMOVED AND REPLACED INTRAOPERATIVELY WITH A BACK UP LENS OF THE SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. REPORTEDLY THE LENS TORE/BROKE WHEN THE DOCTOR REMOVED THE LENS. THE STATUS OF THE EYE WAS REPORTED AS "DOING GREAT; SHOWING SIGNIFICANTLY MORE CYL THAN EXPECTED BUT IT'S STILL EARLY". THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS THE DEVICE/USER ERROR AND THE DEVICE FAILED TO PERFORM AS INTENDED. REPORTEDLY "I STOPPED AND RESTARTED INJECTING DUE TO INJECTOR RESISTANCE AND THAT PROBABLY DIDN'T HELP". CLAIM# (B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#(B)(4). | -12.50/3.5/089 | NA | Physician | 10/05/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15740285&pc=QCB | 2023826-2022-03865 | 2022-08-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-04 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 OF -11.0/1.0/069 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE/BROKE DURING LOADING ON (B)(6)2022 . THE LENS WAS NOT IMPLANTED INTO THE PATIENTS LEFT EYE (OS). A REPLACEMENT LENS OF THE SAME MODEL/SIZE AND SIMILAR POWER WAS IMPLANTED ON (B)(6)2022 AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#(B)(4). | -11.0/1.0/069 | OS | Physician | 10/05/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15741000&pc=QCB | 2023826-2022-03868 | 2021-11-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -9.5/2.5/085 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPOSITIONED AND THAT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIMILAR POWER DUE TO LENS ROTATION. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -9.5/2.5/085 | OS | Physician | 10/07/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15741001&pc=QCB | 2023826-2022-03867 | 2022-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -10/.5/3.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPOSITIONED TWICE AND THAT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -10/.5/3.0/089 | OS | Physician | 10/07/2022 | 2022 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15741002&pc=QCB | 2023826-2022-03869 | 2022-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -13.0/1.5/109 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL AND LENGTH LENS WITH DIFFERENT POWER. THE REPORTER STATES THE REPLACEMENT LENS WAS PLACED VERTICALLY AND THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -13.0/1.5/109 | OD | Physician | 10/07/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15741079&pc=QCB | 2023826-2022-03871 | 2022-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -13.5/3.0/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED AND THAT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE REPLACEMENT LENS IS HORIZONTALLY IMPLANTED AND TURNED CLOCKWISE 30 °. Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BROKEN/BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -13.5/3.0/095 | OD | Physician | 10/07/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15741082&pc=QCB | 2023826-2022-03872 | 2022-08-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-05 04:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -11.0/3.0/092 (SPHERE/CYLINDER/AXIS) TORE/ BROKE DURING INJECTION/ DELIVERY INTO THE PATIENT'S RIGHT EYE (OD). THE LENS WAS IMPLANTED; REMOVED ON (B)(6)2022 . A REPLACEMENT LENS WAS LATER IMPLANTED OF THE SAME MODEL/SIZE AND POWER. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM #(B)(4). | -11.0/3.0/092 | OD | Physician | 10/07/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15741999&pc=QCB | 2023826-2022-04026 | 2022-09-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+3.5/042 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -9.5/+3.5/042 | OD | Physician | 10/08/2022 | 2022 | 53 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15742000&pc=QCB | 2023826-2022-04025 | 2022-09-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/170 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.0/+1.5/170 | OS | Physician | 10/08/2022 | 2022 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15742003&pc=QCB | 2023826-2022-04029 | 2022-10-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/+1.5/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -14.5/+1.5/091 | OS | Physician | 10/08/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15742006&pc=QCB | 2023826-2022-04027 | 2022-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+3.0/084 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE (B)(4). CLAIM# (B)(4). | -6.5/+3.0/084 | OS | Physician | 10/08/2022 | 2022 | 53 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15742009&pc=QCB | 2023826-2022-04032 | 2022-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.5/+3.0/073 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LENS TWO SIZES SHORTER IN LENGTH DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -15.5/+3.0/073 | OS | Physician | 10/08/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15742014&pc=QCB | 2023826-2022-04033 | 2022-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+3.0/095 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.5/+3.0/095 | OS | Physician | 10/08/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15742016&pc=QCB | 2023826-2022-04034 | 2022-07-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.5/084 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+2.5/084 | OD | Physician | 10/08/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15748465&pc=QCB | 2023826-2022-04049 | 2022-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AN IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). THE PATIENT EXPERIENCED LOW VAULTING. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5 - THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.5/+6.0/111 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING. LENS REMAINS IMPLANTED. H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.5/+6.0/111 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING WITHOUT ROTATION. ON (B)(6) 2022 THE LENS WAS REPLACED WITH A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. CLAIM# (B)(6). | -10.5/+6.0/111 | OD | Physician | 10/10/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15749535&pc=QCB | 2023826-2022-03979 | 2022-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2 MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -10.50/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D4: UNIQUE DEVICE NUMBER: (B)(4). | -10.50/+2.0/090 | OS | Physician | 10/08/2022 | 2022 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15749539&pc=QCB | 2023826-2022-03982 | 2022-08-15 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINED IMPLANTED. THE LENS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2022-03981. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: B5: POST-OP, THE PATIENTS VISUAL ACUITY IS 20/20. THE VAULT VALUE IS 211 AND THE LENS IS TRANSPARENT. ALTHOUGH THE VAULT IS A LITTLE BIT LOW, IT IS RELATIVELY STABLE. THE DOCTOR IS STILL ON OBSERVING. CLAIM # (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFER (UDI)# - (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.50/+2.0/095 | OD | Physician | 10/08/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15749650&pc=QCB | 2023826-2022-04036 | 2022-02-27 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5/2.0/111 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. (B)(6) JUN-2022 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS OF LONGER LENGTH AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.5/2.0/111 | OS | Physician | 10/09/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15749762&pc=QCB | 2023826-2022-03981 | 2022-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/107 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS BUT THE VAULT WAS STILL LOW. SEE MFR. REPORT # 2023826-2022-03982 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.50/+2.0/107 | OD | Physician | 10/08/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15749925&pc=QCB | 2023826-2022-04054 | 2022-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.0/067 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.0/+2.0/067 | OD | Physician | 10/09/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15749930&pc=QCB | 2023826-2022-04050 | 2022-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+3.0/095 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.5/+3.0/095 | OS | Physician | 10/09/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15749931&pc=QCB | 2023826-2022-04063 | 2022-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LENS TWO SIZES LONGER IN LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.5/+1.0/094 | OD | Physician | 10/09/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15749932&pc=QCB | 2023826-2022-04058 | 2022-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -18.0/+1.5/167 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS WITH OPPOSITE TORIC MERIDIAN DUE TO LOW VAULT. THE LENS WAS IMPLANTED VERTICALLY AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -18.0/+1.5/167 | OD | Physician | 10/09/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15749935&pc=QCB | 2023826-2022-04059 | 2022-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/073 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+1.0/073 | OS | Physician | 10/09/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15749941&pc=QCB | 2023826-2022-04064 | 2022-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+3.5/174 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.0/+3.5/174 | OD | Physician | 10/09/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15749942&pc=QCB | 2023826-2022-04062 | 2022-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+1.5/107 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -8.0/+1.5/107 | OS | Physician | 10/09/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15750043&pc=QCB | 2023826-2022-04056 | 2022-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.5/089 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/+1.5/089 | OD | Physician | 10/09/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15750044&pc=QCB | 2023826-2022-04052 | 2022-06-30 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS WITH DIOPTER -13.0/+2.0/174 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022 . WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM #(B)(4). | -13.0/+2.0/174 | OD | Physician | 10/09/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15750045&pc=QCB | 2023826-2022-04053 | 2022-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+2.5/086 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -12.5/+2.5/086 | OS | Physician | 10/09/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15751548&pc=QCB | 2023826-2022-04037 | 2022-03-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -13.5/2.0/081 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL AND LENGTH LENS. IT WAS REPORTED THAT THE REPLACEMENT LENS WAS IMPLANTED VERTICALLY AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9- 11-28-2022. H3-RETURNED PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -13.5/2.0/081 | OD | Physician | 10/09/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15756583&pc=QCB | 2023826-2022-03987 | 2022-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -07.00/+2.0/060 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). | -07.00/+2.0/060 | OS | Physician | 10/09/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15759231&pc=QCB | 2023826-2022-03814 | 2022-09-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.5/1.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -7.5/1.5/093 | OD | Physician | 10/20/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15759234&pc=QCB | 2023826-2022-03818 | 2022-09-20 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6, -8.0/1.0/115 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -8.0/1.0/115 | OS | Physician | 10/20/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15759237&pc=QCB | 2023826-2022-03810 | 2021-01-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -7.0/1.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2021. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -7.0/1.0/094 | OD | Physician | 10/18/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15759238&pc=QCB | 2023826-2022-03808 | 2022-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -06.5/2.5/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -06.5/2.5/093 | OS | Physician | 10/18/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15759244&pc=QCB | 2023826-2022-03805 | 2022-10-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.0/3.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.0/3.5/095 | OD | Physician | 10/17/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15759245&pc=QCB | 2023826-2022-03807 | 2022-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -06.5/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY, AND RACE- UNK. CLAIM# (B)(4). | -06.5/2.0/093 | OD | Physician | 10/18/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15759247&pc=QCB | 2023826-2022-03819 | 2022-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.5/1.5/176 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -10.5/1.5/176 | OS | Physician | 10/20/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15759248&pc=QCB | 2023826-2022-03817 | 2022-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.5/1.0/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(6). CLAIM# (B)(4). | -11.5/1.0/074 | OD | Physician | 10/20/2022 | 2022 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15759249&pc=QCB | 2023826-2022-03822 | 2022-01-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.5/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTION DATA: D4 - REPORTED INCORRECTLY CORRECTION IN/SEE D4. H4 - REPORTED AS (B)(6) 2022 CORRECT TO (B)(6) 2021. CLAIM# (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE- UNK. CLAIM# (B)(4). | -10.5/1.0/089 | OD | Physician | 10/20/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15759251&pc=QCB | 2023826-2022-03812 | 2022-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -09.5/1.5/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTION DATA: B4: REPORTED AS 18-OCT-2022 - CORRECT TO 21-OCT-2022. G3: REPORTED AS 18-OCT-2022 - CORRECT TO 21-OCT-2022. CLAIM#: (B)(4). | -09.5/1.5/070 | OD | Physician | 10/18/2022 | 2022 | 16 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15759252&pc=QCB | 2023826-2022-03811 | 2022-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.0/1.0/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: CORRECTION DATA: B4 - REPORTED AS (B)(6) 2022 - CORRECT TO (B)(6) 2022. G3 - REPORTED AS (B)(6) 2022 - CORRECT TO (B)(6) 2022. CLAIM# (B)(4). | -10.0/1.0/107 | OD | Physician | 10/18/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15759253&pc=QCB | 2023826-2022-03823 | 2022-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON VERTICALLY IMPLANTED A 13.7MM VTICMO13.7, -11.0/1.5/007 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED HORIZONTALLY WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/1.5/007 | OD | Physician | 10/20/2022 | 2022 | 41 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15759255&pc=QCB | 2023826-2022-03820 | 2022-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.5/1.0/146 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/1.0/146 | OS | Physician | 10/20/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15759260&pc=QCB | 2023826-2022-03821 | 2022-06-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.0/1.5/025 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.0/1.5/025 | OD | Physician | 10/20/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15759271&pc=QCB | 2023826-2022-03826 | 2022-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.0/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.0/1.5/091 | OS | Physician | 10/21/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15761749&pc=QCB | 2023826-2022-03989 | 2022-10-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | . | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.00/+1.5/100 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS HAD EXCESSIVE VAULTING AND REMAINED IMPLANTED. CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PATIENT WEIGHT: UNK. PATIENT ETHNICITY: UNK. PATIENT RACE: UNK. EXPLANT DATE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). Manufacturer Narrative: B5: THE LENS WAS REMOVED AND REPLACED ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_ 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -07.00/+1.5/100 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. ON THE SAME DAY SEPARATE SURGERY THE LENS WAS REPLACED WITH THE SAME MODEL/SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS REPORTED AS "THE LENS WAS REPLACED SUCCESSFULLY NOVEMBER (B)(6) 2022" AND "VAULT AFTER ONE DAY OF SURGERY 262 MICRAS". THE REPORTER INDICATED THE CAUSE OF THE EVENT IS UNKNOWN AND THE DEVICE FAILED TO PERFORM AS INTENDED. REPORTEDLY "SIZING OUTLIER/ UNUSUAL ANATOMY". CLAIM# (B)(4). | -07.00/+1.5/100 | OD | Physician | 10/11/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15762626&pc=QCB | 2023826-2022-03985 | 2022-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+3.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND AN OPTIC DEFORMED AND HAPTIC BENT TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -09.50/+3.0/094 | OD | Physician | 10/10/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15762711&pc=QCB | 2023826-2022-03986 | 2022-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.50/+3.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -09.50/+3.0/088 | OS | Physician | 10/10/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15762726&pc=QCB | 2023826-2022-03990 | 2022-06-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -13.00/+1.0/081 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT: UNK. PATIENT ETHNICITY: UNK. PATIENT RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -13.00/+1.0/081 | OS | Physician | 10/12/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15763563&pc=QCB | 2023826-2022-04066 | 2022-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+2.0/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -9.0/+2.0/094 | OS | Physician | 10/10/2022 | 2022 | 38 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15764224&pc=QCB | 2023826-2022-04069 | 2022-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.50/+1.0/082 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.50/+1.0/082 | OS | Physician | 10/10/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15764329&pc=QCB | 2023826-2022-04068 | 2022-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.50/+1.0/086 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.50/+1.0/086 | OD | Physician | 10/10/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15764678&pc=QCB | 2023826-2022-04044 | 2022-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -10.5/2.0/020 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS LATER EXCHANGED WITH A SAME LENGTH LENS THAT WAS IMPLANTED HORIZONTALLY DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -10.5/2.0/020 | OD | Physician | 10/12/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15764714&pc=QCB | 2023826-2022-04077 | 2022-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -7.50/1.0/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -7.50/1.0/89 | OS | Physician | 11/02/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15764923&pc=QCB | 2023826-2022-03995 | 2022-04-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -13.0/2.5/084 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED AND THAT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL/SIZE LENS WITH SIMILAR POWER AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: H4: DEVICE MANUFACTURING DATE: 26-DEC-2021 SHOULD BE CORRECTED TO 28-DEC-2021. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN/BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | -13.0/2.5/084 | OS | Physician | 10/11/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15764926&pc=QCB | 2023826-2022-03997 | 2022-03-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -8.5/1.0/147 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. REPORTEDLY THE PROBLEM WAS RESOLVED. NO PLANS TO IMPLANT ANOTHER LENS. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE CODE (B)(4) OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -8.5/1.0/147 | OS | Physician | 10/11/2022 | 2022 | 57 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15764995&pc=QCB | 2023826-2022-03996 | 2022-03-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -11.0/1.0/034 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6)2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT. REPORTEDLY THE PROBLEM WAS RESOLVED. NO PLANS TO IMPLANT ANOTHER LENS. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -11.0/1.0/034 | OD | Physician | 10/11/2022 | 2022 | 57 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15765831&pc=QCB | 2023826-2022-04071 | 2022-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+2.0/087 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -12.0/+2.0/087 | OD | Physician | 10/11/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15765833&pc=QCB | 2023826-2022-04072 | 2022-06-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -18.0/+3.5/108 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -18.0/+3.5/108 | OS | Physician | 10/11/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15765839&pc=QCB | 2023826-2022-04074 | 2022-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTCMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+0.5/110 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -8.0/+0.5/110 | OD | Physician | 10/11/2022 | 2022 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15773786&pc=QCB | 2023826-2022-04041 | 2022-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -10.0/2.0/091 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6)2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. IT WAS REPORTED THAT THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE/ FIBERS ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.0/2.0/091 | OD | Physician | 10/11/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15777037&pc=QCB | 2023826-2022-04104 | 2022-09-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+1.5/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -7.5/+1.5/093 | OD | Physician | 10/12/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15777287&pc=QCB | 2023826-2022-04105 | 2022-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+2.0/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+2.0/090 | OD | Physician | 10/12/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15777309&pc=QCB | 2023826-2022-04101 | 2022-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.0/+1.5/092 | OD | Physician | 10/12/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15779177&pc=QCB | 2023826-2022-04042 | 2022-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -10.0/2.5/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. DURING SURGERY THE DOCTOR NOTED THAT THE LENS HAD NO AXIS ALIGNMENT MARKS. ON (B)(6) 2022 THE LENS WAS EXCHANGED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). TYPE OF INVESTIGATION - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER (S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | -10.0/2.5/087 | OD | Physician | 10/12/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15779962&pc=QCB | 2023826-2022-04047 | 2022-08-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS -11.00/1.5/085 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 IT WAS REPORTED THAT LOW VAULT WAS OBSERVED AND THE LENS WAS REMOVED AND REPLACED WITH A LONGER LENGTH LENS INTRAOPERATIVELY DUE TO LOW VAULT. FOR STATUS OF THE EYE (SIGNS/SYMPTOMS) THE REPORTER STATED " THE LENS THAT WAS REPLACED DID NOT STABILIZE THE VAULTING AND WAS LATER REMOVED FROMT HE EYE". Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND OPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). | -11.00/1.5/085 | OD | Physician | 10/13/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15780126&pc=QCB | 2023826-2022-04048 | 2022-08-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -11.00/1.5/72 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 AS A REPLACEMENT LENS TO ADDRESS LOW VAULT BUT IT DID NOT STABILIZE THE VAULT. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND THE PROBLEM RESOLVED. "NEXT OPERATION DATE UNDECIDED" WAS REPORTED. SEE MFR # 2023826-2022-04047 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(6). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -11.00/1.5/72 | OD | Physician | 10/13/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15782581&pc=QCB | 2023826-2022-03953 | 2022-09-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF A -15.0/2.0/85 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REPLACED WITH A SAME SIZE LENS DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT, BUT NO PATIENT INJURY. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PATIENT ETHNICITY,WEIGHT, RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -15.0/2.0/85 | OS | Physician | 10/12/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15782602&pc=QCB | 2023826-2022-04113 | 2022-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.00/2.5/114 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). ON (B)(6) 2022 THE LENS WAS EXPLANTED AND ON THIS SAME DATE REPLACED WITH A LONGER LENGTH LENS AND DIFFERENT LENS POWER. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: B5: LENS WAS EXCHANGED ON (B)(6) 2022 WITH A LONGER LENGTH LENS DUE TO LOW VAULT AND PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS REPORTED AS "USER ERROR"; THE DEVICE DID NOT FAIL TO PERFORM. LENS SIZE WRONG, A LONGER SIZED LENS RESOLVED THE ISSUE. CLAIM#: (B)(4). | -10.00/2.5/114 | OD | Physician | 10/16/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15782639&pc=QCB | 2023826-2022-04114 | 2022-08-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -9.00/3.5/072 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). ON (B)(6)2022 THE LENS WAS EXPLANTED AND ON THIS SAME DATE REPLACED WITH A LONGER LENGTH LENS AND DIFFERENT LENS POWER. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: B5: LENS WAS EXCHANGED ON (B)(6) 2022 WITH A LONGER LENGTH LENS DUE TO LOW VAULT AND PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS NOT REPORTED. CLAIM#: (B)(4). | -9.00/3.5/072 | OS | Physician | 10/16/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15784677&pc=QCB | 2023826-2022-04081 | 2022-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.00/+1.0/091 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS BUT THE PROBLEM WAS NO RESOLVED. SEE MFR. REPORT # 2023826-2022-04082 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT,ETHNICITY AND RACE: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -13.00/+1.0/091 | OS | Physician | 10/15/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15784708&pc=QCB | 2023826-2022-04086 | 2022-09-13 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON WAS ATTEMPTING TO INSERT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.50/+1.0/091 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) BUT THE LENS BECAME STUCK IN THE CARTRIDGE AND TORE. THERE WAS NO PATIENT CONTACT OR INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: (B)(4). | -15.50/+1.0/091 | OD | Physician | 10/17/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15784720&pc=QCB | 2023826-2022-04088 | 2022-07-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50/+2.5/072 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO DOWNSIZING. THE LENS WAS REPLACED WITH A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PATIENT AGE,SEX, WEIGHT, ETHNICITY, RACE: UNK. DATE OF EVENT: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. B5: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.50/+2.5/072 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -09.50/+2.5/072 | OD | Physician | 10/16/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15791331&pc=QCB | 2023826-2022-04090 | 2022-07-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.00/+1.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO DOWNSIZING. THE LENS WAS REPLACED WITH A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL DATA: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AN ANTERIOR CHAMBER DEPTH (ACD) LESS THAN 3.0MM. CORRECTED DATA: B1: PRODUCT PROBLEM. B5: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.00/+1.5/094 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. H1: MALFUNCTION. CLAIM # (B)(4). Manufacturer Narrative: AGE: UNK. SEX: UNK. WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DATE OF EVENT: UNK. CLAIM # (B)(4). | -08.00/+1.5/094 | OD | Physician | 10/16/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15792404&pc=QCB | 2023826-2022-03308 | 2022-08-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICH 13.2 IMPLANTABLE COLLAMER LENS OF -8.5/2.5/087 (SPHERE/CYLINDER/AXIS) TORE/ BROKE DURING INJECTION/ DELIVERY INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS IMPLANTED; REMOVED INTRA-OPERATIVELY ON (B)(6) 2022. A REPLACEMENT LENS WAS NOT IMPLANTED. Manufacturer Narrative: B5 - CORRECTED TO: 'THE REPORTER INDICATED THAT A 13.2MM VTICH 13.2 IMPLANTABLE COLLAMER LENS OF 8.5/2.5/087 (SPHERE/CYLINDER/AXIS) TORE/ BROKE DURING INJECTION/ DELIVERY INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS IMPLANTED; REMOVED INTRA-OPERATIVELY ON (B)(6) 2022. A REPLACEMENT LENS WAS NOT IMPLANTED." IN THE INITIAL MDR. H6 - MEDICAL DEVICE PROBLEM CODE: 1494 (OFF-LABEL USE: ACD<3.0MM) SHOULD BE ADDED TO THE INTIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: B5 - A BACKUP LENS WAS IMPLANTED ON (B)(6) 2022 BUT PROBLEM WAS NOT RESOLVED. EXCESSIVE VAULTING WAS REPORTED. CLAIM# (B)(4). Manufacturer Narrative: INVESTIGATION TYPE : LENS WORK ORDER SEARCH. NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -8.5/2.5/087 | OD | Physician | 09/01/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15793019&pc=QCB | 2023826-2022-04082 | 2022-07-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -13.00/+1.0/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS REPORTED AS HAVING LOW VAULTING AND REMAINED IMPLANTED. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2022-04081. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | -13.00/+1.0/094 | OS | Physician | 10/15/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15793440&pc=QCB | 2023826-2022-04119 | 2022-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.0/+1.5/096 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -17.0/+1.5/096 | OD | Physician | 10/15/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15793449&pc=QCB | 2023826-2022-04116 | 2022-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/074 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+1.0/074 | OS | Physician | 10/15/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15793654&pc=QCB | 2023826-2022-04089 | 2022-07-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50/+3.0/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO DOWNSIZING AND SLIGHT CHANGE OF LENS RX. THE LENS WAS REPLACED WITH A SHORTER LENS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: ADDITIONAL DATA: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. CORRECTED DATA: B1: PRODUCT PROBLEM. B5: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -10.50/+3.0/086 (SPHERE/CYLINDER/AXIS), DURING THE SAME SURGERY DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. H1: MALFUNCTION. CLAIM # (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -10.50/+3.0/086 | OS | Physician | 10/16/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15794319&pc=QCB | 2023826-2022-04115 | 2022-10-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -7.00/+1.5/89 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - "REPORTEDLY, "ALL IS FINE! THE SURGEON PLANS A CLOSED MONITORING. PATIENT IS HAPPY. THE PROBLEM IS RESOLVED." " SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | -7.00/+1.5/89 | OD | Physician | 10/19/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15794532&pc=QCB | 2023826-2022-04117 | 2022-07-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+1.0/079 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -7.5/+1.0/079 | OS | Physician | 10/15/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15795262&pc=QCB | 2023826-2022-04003 | 2022-09-08 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Appropriate Term/Code Not Available | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -5.00/1.5/079 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON AN UNKNOWN DATE. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS FOR AN UNKNOWN REASON. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: B5-THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -5.0/1.5/079 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS IMPLANTED; REMOVED AND EXCHANGED INTRA-OPERATIVELY FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS USER ERROR. "A WRONG SIZED LENS WAS ORDERED." H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM). CLAIM# (B)(4). | -5.00/1.5/079 | OD | Physician | 10/16/2022 | 2022 | 1 MO | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15798625&pc=QCB | 2023826-2022-04106 | 2022-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+2.0/77 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+2.0/77 | OS | Physician | 10/12/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15801812&pc=QCB | 2023826-2022-03956 | 2022-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.5/4.0/101 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). ON (B)(6) 2022, THE LENS WAS REMOVED AND REPLACED FOR A LARGER LENGTH LENS DUE TO RECOMMENDATION. Manufacturer Narrative: ADDITIONAL INFORMATON: B5: LENS WAS IMPLANTED ON (B)(6) 2022 AND EXCHANGED ON (B)(6) 2022. CAUSE OF THE EVENT WAS REPORTED AS "USER ERROR," AND "LENS CHOSEN WAS TOO SHORT. A LONGER SIZED LENS RESOLVED THE ISSUE." PROBLEM WAS RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATED BY MFR: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4). | -11.5/4.0/101 | OD | Physician | 10/16/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15801842&pc=QCB | 2023826-2022-04091 | 2022-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -04.50/+1.0/085 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# (B)(4), SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4), Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | -04.50/+1.0/085 | OS | Physician | 10/18/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15801880&pc=QCB | 2023826-2022-03955 | 2022-06-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -11.0/3.5/112 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). ON (B)(6) 2022, THE LENS WAS REMOVED AND REPLACED FOR A DIFFERENT MODEL, LARGER LENGTH LENS DUE TO RECOMMENDATION. Manufacturer Narrative: DEVICE EVALUATED BY MFR: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: LENS WAS IMPLANTED ON (B)(6) 2022 AND EXCHANGED ON (B)(6) 2022. REPORTEDLY, "CONTINUE TO MONITOR." PROBLEM WAS NOT RESOLVED. CLAIM#: (B)(4). | -11.0/3.5/112 | OS | Physician | 10/16/2022 | 2022 | 28
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15801972&pc=QCB | 2023826-2022-04129 | 2022-11-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED 12.6MM VTICM5_12.6; -12.50/3.0/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS TORE DURING INJECTION/DELIVERY INTO THE EYE. INTRAOPERATIVELY THE LENS WAS REMOVED. IT WAS ADDITIONALLY NOTED "LENS GOT STUCK IN PLUNGER". THERE WAS NO PATIENT INJURY. ON(B)(6)2022 A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | -12.50/3.0/92 | OD | Physician | 11/10/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15801981&pc=QCB | 2023826-2022-04093 | 2022-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -04.50/+1.0/075 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: D4 - UNIQUE IDENTIFIER (UDI) #: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | -04.50/+1.0/075 | OS | Physician | 10/18/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15802126&pc=QCB | 2023826-2022-03958 | 2022-05-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.0/3.5/076 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REPOSITIONED. THE PROBLEM WAS NOT RESOLVED. ON 20-SEP-2022, THE LENS WAS EXCHANGED FOR A SPHERICAL MODEL LENS DUE TO LENS ROTATION NOT ASSOCIATED WITH LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "FIVE MONTHS AFTER THE FIRST IMPLANTATION OF TICL LENS, THE PATIENT FOUND LENS ROTATION IN THE LEFT EYE AFTER SURGERY, AND DECIDED TO PERFORM TICL ALIGNMENT SURGERY AFTER THE SURGEON'S INTERVIEW. THE LENS ROTATED AGAIN 3 MONTHS AFTER THE OPERATION. AFTER CONSULTATION BETWEEN SURGEONS AND EXPERT PROFESSORS, IT WAS DECIDED TO REPLACE ICL CRYSTALS, AND THE ASTIGMATISM WAS CORRECTED BY LASER SURGERY. Manufacturer Narrative: ADDITIONAL INFORMATION: D9-RETURN DATE: (B)(6)2023 H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | -9.0/3.5/076 | OS | Physician | 10/17/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15804276&pc=QCB | 2023826-2022-03933 | 2022-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.0/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.0/2.0/089 | OS | Physician | 10/24/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15804279&pc=QCB | 2023826-2022-03935 | 2022-08-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -11.50/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. CLAIM # (B)(4). | -11.50/1.0/091 | OD | Physician | 10/24/2022 | 2022 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15804281&pc=QCB | 2023826-2022-03936 | 2022-08-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -11.0/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE: UNK. CLAIM# (B)(4). | -11.0/1.0/089 | OS | Physician | 10/24/2022 | 2022 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15804282&pc=QCB | 2023826-2022-03938 | 2022-08-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2, -11.0/3.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S RIGHT EYE (OD) DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. CLAIM# (B)(4). | -11.0/3.0/092 | OD | Physician | 10/26/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15804285&pc=QCB | 2023826-2022-03942 | 2022-03-01 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.5/1.0/07289 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. CLAIM# (B)(4). | -10.5/1.0/07289 | OD | Physician | 10/26/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15804286&pc=QCB | 2023826-2022-03948 | 2022-09-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.5/1.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/1.5/087 | OS | Physician | 10/27/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15804287&pc=QCB | 2023826-2022-03947 | 2022-08-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED VERTICALLY A 13.2MM VTICMO13.2, -14.5/3.0/167 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTION DATA. D6B - REPORTED AS (B)(6) 2022 - CORRECT TO (B)(6) 2022 CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -14.5/3.0/167 | OD | Physician | 10/27/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15804289&pc=QCB | 2023826-2022-03945 | 2022-09-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -17.5/2.5/004 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNKNOWN. (B)(4). CLAIM# (B)(4). | -17.5/2.5/004 | OD | Physician | 10/27/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15804290&pc=QCB | 2023826-2022-03950 | 2022-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/1.5/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.0/1.5/107 | OS | Physician | 10/27/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15804295&pc=QCB | 2023826-2022-03951 | 2022-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -5.0/0.5/173 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -5.0/0.5/173 | OD | Physician | 10/27/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15807347&pc=QCB | 2023826-2022-04006 | 2022-09-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -16.0/4.0/090 (SPHERE/CYLINDER/AXIS) TORE/ BROKE DURING INJECTION/ DELIVERY INTO THE PATIENT'S RIGHT EYE (OD). THE LENS WAS IMPLANTED; REMOVED ON (B)(6) 2022. A REPLACEMENT LENS WAS LATER IMPLANTED OF THE SAME MODEL/SIZE AND POWER. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -16.0/4.0/090 | OD | Other Health Care Professional | 10/17/2022 | 2022 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15807853&pc=QCB | 2023826-2022-04098 | 2022-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+1.0/075 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: D4: UNIQUE IDENTIFIER NUMBER: (B)(4), SHOULD BE ADDED FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | -12.00/+1.0/075 | OD | Physician | 10/21/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15808950&pc=QCB | 2023826-2022-04099 | 2022-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.50/+1.5/110 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | -08.50/+1.5/110 | OD | Physician | 10/21/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15809320&pc=QCB | 2023826-2022-04097 | 2022-03-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+1.5/094 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS REPLACED WITH A LONGER LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: D4: UNIQUE IDENTIFIER NUMBER (B)(4). ADDITIONAL INFORMATION: D9: DEVICE AVAILABLE FOR EVALUATION: YES: DATE RETURNED- (B)(6) 2022. CLAIM # (B)(4). | -11.50/+1.5/094 | OD | Physician | 10/21/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15809894&pc=QCB | 2023826-2022-03962 | 2022-05-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -12.5/3.5/102 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2022 AND (B)(6) 2018, THE LENS WAS REPOSITIONED DUE TO LOW VAULT WITH ROTATION. ROTATION NOT RESOLVED. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SPHERICAL LENS. THE PROBLEM WAS NOT RESOLVED. CAUSE IS REPORTED AS UNKNOWN. REPORTEDLY, "THE PATIENT IMPLANTED THE CRYSTAL ON (B)(6) 2018, ROTATED FOUR YEARS LATER, AND STILL ROTATED AFTER TWO ADJUSTMENTS ON (B)(6) 2022 AND (B)(6) 2022, SO THE ICL OF THE SAME MODEL WAS REPLACED, AND THE ARCH HEIGHT WAS STILL HIGH LOW, ICL REPLACEMENT PERFORMED ON 2022.10.10." Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D6A: ON (B)(6) 2018. CLAIM#: (B)(4). | -12.5/3.5/102 | OD | Physician | 10/18/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15812711&pc=QCB | 2023826-2022-04095 | 2022-07-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.00/+0.5/072 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE SECOND PIECE OF ICL HAD NO EFFECT ON THE PATIENT'S VISION. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN. CLAIM # (B)(4). | -14.00/+0.5/072 | OD | Physician | 10/20/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15812940&pc=QCB | 2023826-2022-03946 | 2022-08-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.00/4.5/169 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE LENS WAS REPLACED VERTICALLY WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT:UNKNOWN, ETHNICITY: UNKNOWN, RACE: UNKNOWN. CLAIM# (B)(4). | -16.00/4.5/169 | OS | Physician | 10/27/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15813182&pc=QCB | 2023826-2022-04181 | 2022-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/+3.5/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LENS THREE SIZES LONGER IN LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -14.5/+3.5/092 | OS | Physician | 10/19/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15813378&pc=QCB | 2023826-2022-04174 | 2022-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -5.50/+1.0/080 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING. ON 14-OCT-2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM#: (B)(4). | -5.50/+1.0/080 | OS | Physician | 10/21/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15819895&pc=QCB | 2023826-2022-04183 | 2022-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT 13.2MM VTICMO13.2 -16.5/5.0/070 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS THAT WAS IMPLANTED VERTICALLY AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. PRODUCT RETURNED DAMAGED. VISUAL INSPECTION FOUND THE HAPTIC BROKEN/BENT. DIMENSIONAL INSPECTION FOUND THE OVERALL LENGTH WITHIN SPECIFICATIONS BUT WIDTH COULD NOT BE PERFORMED DUE TO DAMAGED STATE OF THE LENS. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -16.5/5.0/070 | OS | Physician | 10/18/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15821020&pc=QCB | 2023826-2022-03963 | 2022-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF -12.5/3.5/102 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2022 AND (B)(6) 2022, THE LENS WAS REPOSITIONED DUE TO LOW VAULT WITH ROTATION. ROTATION NOT RESOLVED. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SPHERICAL LENS. THE PROBLEM WAS NOT RESOLVED. CAUSE IS REPORTED AS UNKNOWN. REPORTEDLY, "THE PATIENT IMPLANTED THE CRYSTAL ON (B)(6) 2018, ROTATED FOUR YEARS LATER, AND STILL ROTATED AFTER TWO ADJUSTMENTS ON (B)(6) 2022 AND (B)(6) 2022, SO THE ICL OF THE SAME MODEL WAS REPLACED, AND THE ARCH HEIGHT WAS STILL HIGH LOW, ICL REPLACEMENT PERFORMED ON (B)(6) 2022." Manufacturer Narrative: CORRECTED DATA: B5: "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VICMO 12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 AS A REPLACEMENT LENS TO RESOLVE LOW VAULT WITH ROTATION. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL/LONGER LENGTH LENS DUE TO "ARCH HEIGHT WAS STILL HIGH LOW" AND LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN." H6- HEALTH EFFECT- IMPACT CODE: 4628 AND MEDICAL DEVICE PROBLEM CODE: 2923 SHOULD BE REMOVED. CLAIM#: (B)(4). Manufacturer Narrative: OFF LABEL USE- ACD <3.0. CLAIM# (B)(4). | -12.5/3.5/102 | OD | Physician | 10/18/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15822810&pc=QCB | 2023826-2022-04185 | 2022-05-24 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 12.6MM, VTICMO12.6 -9.0/0.5/075 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE LEFT EYE (OS). THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY, WITH NO PATIENT INJURY. ON (B)(6) 2022 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.0/0.5/075 | OS | Physician | 10/19/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15822858&pc=QCB | 2023826-2022-04187 | 2022-04-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 12.6MM, VTICMO12.6 -10.5/1.5/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE LEFT EYE (OS). THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY, WITH NO PATIENT INJURY. ON (B)(6) 2022 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT IS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.5/1.5/078 | OS | Physician | 10/20/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15823030&pc=QCB | 2023826-2022-04194 | 2022-08-09 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -15.0/2.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022.THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY WITH NO PATIENT INJURY. ON (B)(6) 2022 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -15.0/2.0/087 | OD | Physician | 10/21/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15823100&pc=QCB | 2023826-2022-04186 | 2022-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -8.5/1.5/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT THE REPLACEMENT LENS WAS PLACED IN A VERTICAL POSITION. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.5/1.5/090 | OD | Physician | 10/19/2022 | 2022 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15823158&pc=QCB | 2023826-2022-04192 | 2022-06-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -13.0/2.0/135 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM RESOLVED. IT WAS REPORTED THAT THE REPLACEMENT LENS WAS IMPLANTED HORIZONTALLY . CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.0/2.0/135 | OS | Physician | 10/21/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15823363&pc=QCB | 2023826-2022-04189 | 2022-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATE THAT A 12.1MM IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) AND THE PATIENT EXPERIENCED LOW VAULT, LENS REMAINS IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. B5- PER UPDATED INFORMATION A 12.1MM VTICMO12.1 -9.0/1.5/089 IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6)2018. ON (B)(6)2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO OBSERVATION OF LOW VAULT. THIS EXCHANGE RESOLVED THE PROBLEM. PER REPORTER THE CAUSE OF THIS EVENT IS DUE TO PATIENT RELATED FACTORS. THE REPORTER ALSO STATED THAT IT IS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. CLAIM # (B)(4). Manufacturer Narrative: (EXPIRATION DATE): UNK, NO SERIAL NUMBER REPORTED. (DEVICE MANUFACTURING DATE): UNK, NO SERIAL NUMBER REPORTED. CLAIM# (B)(4). | -9.0/1.5/089 | OD | Physician | 10/20/2022 | 2022 | 43 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15826415&pc=QCB | 2023826-2022-04200 | 2022-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.0/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.0/+1.0/093 | OS | Physician | 10/20/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15826661&pc=QCB | 2023826-2022-04100 | 2022-08-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -08.00/+1.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: D4: UNIQUE IDENTIFIER NUMBER: (B)(4). D9: DEVICE AVAILABLE FOR EVALUATION: YES-DATE RETURNED: 20-DEC-2022. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). | -08.00/+1.0/088 | OS | Physician | 10/21/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15826669&pc=QCB | 2023826-2022-04096 | 2022-09-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT DURING THE LOADING OF A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -11.00/+2.0/094 (SPHERE/CYLINDER/AXIS), THE LENS WAS DAMAGED. THERE WAS NO PATIENT CONTACT. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS DUE TO BROKEN FORCEPS. Manufacturer Narrative: PATIENT WEIGHT: UNK. PATIENT ETHNICITY: UNK. PATIENT RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -11.00/+2.0/094 | NA | Physician | 10/20/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15826925&pc=QCB | 2023826-2022-04197 | 2022-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+2.0/091 | OD | Physician | 10/20/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15827225&pc=QCB | 2023826-2022-04198 | 2022-05-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+1.0/108 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -6.5/+1.0/108 | OS | Physician | 10/20/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15827241&pc=QCB | 2023826-2022-04201 | 2022-09-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/075 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -9.5/+1.0/075 | OD | Physician | 10/20/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15829267&pc=QCB | 2023826-2022-04191 | 2022-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6 -10.0/1.5/95 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE WAS REPORTED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS REPORTD AS UNKNOWN, REPORTER STATED LENS OR CARTRIDGE OR INJECTOR MAY BE DEFECTIVE. FOR CAUSE DETAILS THE REPORTER STATED " AFTER INSERTION OF THE LENS INTO THE EYE, ONE OF THE FOUR HAPTICS WAS DAMAGED". ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 -10.00/1.5/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON 15-AUG-2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE AND THE PROBLEM RESOLVED. NO INJURY TO THE PATIENT WAS REPORTED. CAUSE OF EVENT WAS UNKNOWN. DEVICE EVALUATION: LENS WAS RETURNED IN LIQUID, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTICS BROKEN. MEDICAL DEVICE PROBLEM CODE: PREVIOUS CODE NOT APPLICABLE, THE DEVICE DID NOT DISLODGE OR DISLOCATE (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.0/1.5/95 | OD | Physician | 10/21/2022 | 2022 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15829785&pc=QCB | 2023826-2022-04208 | 2022-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.0/133 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+2.0/133 | OS | Physician | 10/21/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15829798&pc=QCB | 2023826-2022-04203 | 2022-07-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.0/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE: (B)(4). CLAIM# (B)(4). | -10.0/+2.0/090 | OD | Physician | 10/20/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15829800&pc=QCB | 2023826-2022-04205 | 2022-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+1.5/080 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). (ACD < 3.0MM). CLAIM#: (B)(4). | -12.0/+1.5/080 | OS | Physician | 10/20/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15829802&pc=QCB | 2023826-2022-04206 | 2022-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/+1.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -5.5/+1.0/091 | OD | Physician | 10/21/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15829803&pc=QCB | 2023826-2022-04202 | 2022-06-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.5/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -9.5/+1.5/090 | OS | Physician | 10/20/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15830454&pc=QCB | 2023826-2022-04199 | 2022-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.0/149 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LENS TWO SIZES LONGER IN LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+1.0/149 | OS | Physician | 10/20/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15830515&pc=QCB | 2023826-2022-03968 | 2022-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.5/1.0/131 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED WITH LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "AFTER THE PATIENT'S RIGHT EYE WAS IMPLANTED WITH TICL CRYSTAL FOR THE FIRST TIME, THE CRYSTAL ROTATION WAS RECHECKED BY 90 DEGREES ONE DAY AFTER THE OPERATION. AFTER ONE MONTH OF OBSERVATION, THE RECHECKED CRYSTAL POSITION WAS NOT SIGNIFICANTLY DIFFERENT FROM THE LAST RECHECK. AFTER THE SURGEON'S INTERVIEW, HE DECIDED TO CUSTOMIZE A SPECIAL AXIAL CRYSTAL FOR TICL REPLACEMENT." Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN/BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNKNOWN, ETHNICITY:UNKNOWN, RACE:UNKNOWN. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -11.5/1.0/131 | OD | Physician | 10/21/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15830521&pc=QCB | 2023826-2022-03972 | 2022-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Patient-Device Incompatibility | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -15.5/1.5/170 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) AS A REPLACEMENT LENS ON (B)(6) 2022. THE PROBLEM WAS NOT RESOLVED FOR THE EXCHANGE. REPORTEDLY, "VISUAL GHOSTING," AND "TAKE OUT THE CRYSTAL." Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B1: ADVERSE EVENT SHOULD BE CORRECTED TO PRODUCT PROBLEM. B2: OUTCOMES ATTRIBUTED TO ADVERSE EVENT- SHOULD BE OMITTED FOR CORRECTION. D9: DEVICE AVAILABLE FOR EVALUATION: YES; RETURN DATE: 20-DEC-2022. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC TORN AND BENT WITH FOREIGN MATERIAL ON THE LENS, SPECIFICALLY DEBRIS THROUGHOUT THE LENS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -15.5/1.5/170 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) AS A REPLACEMENT LENS ON (B)(6) 2022. THE LENS WAS EXPLANTED AT AN UNKNOWN DATE. THE PROBLEM WAS NOT RESOLVED FOR THE EXCHANGE. REPORTEDLY, "VISUAL GHOSTING," AND "TAKE OUT THE CRYSTAL." CLAIM# (B)(4). | -15.5/1.5/170 | OD | Physician | 10/21/2022 | 2022 | 42 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15830623&pc=QCB | 2023826-2022-04210 | 2022-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.5/+1.5/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 . THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -15.5/+1.5/089 | OS | Physician | 10/21/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15830649&pc=QCB | 2023826-2022-04209 | 2022-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -10.0/+1.5/093 | OS | Physician | 10/21/2022 | 2022 | 16 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15830676&pc=QCB | 2023826-2022-04204 | 2022-09-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+2.0/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. IN A SEPARATE SURGERY THAT DAY THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. IN A THIRD SURGERY THAT SAME DAY A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE CODE (B)(4): OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM#: (B)(4). | -12.5/+2.0/089 | OS | Physician | 10/20/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15830765&pc=QCB | 2023826-2022-04015 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -11.50/1.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON AN UNKNOWN DATE. THE DOCTOR IS UNDECIDED IF THE LENS WILL BE EXCHANGED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.50/1.5/093 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD). REPORTEDLY "DR LIM WANTS TO KNOW IF SHE SHOULD CONTINUE TO MONITOR OR EXCHANGE FOR A LONGER LENS." CLAIM#: (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -11.50/1.5/090 | OD | Physician | 10/19/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15830820&pc=QCB | 2023826-2022-04014 | 2022-10-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -12.00/2.0/124 (SPHERE/CYLINDER AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022 . LOW VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE REPORTED AS UNKNOWN. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: INVESTIGATION TYPE : LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -12.00/2.0/124 | OS | Physician | 10/19/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15831960&pc=QCB | 2023826-2022-04017 | 2022-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO 13.7 IMPLANTABLE COLLAMER LENS OF -12.00/1.0/82 (SPHERE/CYLINDER AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS ROTATION WAS OBSERVED. LENS WAS REPOSITIONED AND THE PROBLEM WAS NOT RESOLVED. ON (B)(6) 2022 THE THE LENS WAS EXCHANGED FOR A SPHERICAL LENS OF THE SAME SIZE AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. LENS RETURNED WITH ALL FOOTPLATES BROKEN/DAMAGED. VISUAL INSPECTION FOUND HAPTIC BROKEN. UNABLE TO PERFORM DIMENSIONAL INSPECTION DUE TO HAPTIC DAMAGE. CLAIM# (B)(4). | -12.00/1.0/82 | OS | Physician | 10/20/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15831963&pc=QCB | 2023826-2022-04019 | 2022-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -9.5/2.0/168 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. LOW VAULT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND ON THE SAME DAY A REPLACEMENT LENS OF LONGER LENGTH LENS WAS IMPLANTED. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. STATUS OF EYE "THE IMPLANT IS A VERTICAL TABLET ONE SIZE LARGER." Manufacturer Narrative: (B)(4). | -9.5/2.0/168 | OS | Physician | 10/21/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15832029&pc=QCB | 2023826-2022-04023 |  | Injury | STAAR SURGICAL COMPANY | 2022-11-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -13.00/2.0/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE. THE LENS WAS EXPLANTED AND DISCARDED FOR AN UNKNOWN REASON. A REPLACEMENT LENS OF SHORTER LENGTH WAS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | -13.00/2.0/088 | NA | Other Health Care Professional | 10/24/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15835035&pc=QCB | 2023826-2022-03241 | 2022-06-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -12.0/2.5/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). | -12.0/2.5/089 | OD | Physician | 08/25/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15840982&pc=QCB | 2023826-2022-04238 | 2022-09-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.0/071 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.5/+1.0/071 | OD | Physician | 10/23/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15841006&pc=QCB | 2023826-2022-04245 | 2022-09-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/+3.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: DEVICE EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -14.5/+3.0/093 | OD | Physician | 10/25/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15841079&pc=QCB | 2023826-2022-04240 | 2022-08-04 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+2.0/085 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.5/+2.0/085 | OS | Physician | 10/24/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15841086&pc=QCB | 2023826-2022-04241 | 2022-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.5/094 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.5/+2.5/094 | OD | Physician | 10/24/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15841142&pc=QCB | 2023826-2022-04242 | 2022-08-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+5.5/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -7.0/+5.5/091 | OS | Physician | 10/25/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15845363&pc=QCB | 2023826-2022-04133 | 2022-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.5/2.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.5/2.5/087 | OS | Physician | 10/31/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15845368&pc=QCB | 2023826-2022-04153 | 2022-07-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -15.5/3.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -15.5/3.5/090 | OD | Physician | 11/03/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15845372&pc=QCB | 2023826-2022-04143 | 2022-09-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -12.0/2.0/011 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE- UNK. (B)(4). | -12.0/2.0/011 | OD | Physician | 11/02/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15845375&pc=QCB | 2023826-2022-04135 | 2022-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7, -10.0/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.0/1.0/089 | OD | Physician | 10/31/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15845394&pc=QCB | 2023826-2022-04152 | 2021-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/1.0/172 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. DEVICE PROBLEM CODE: (B)(4). CLAIM# (B)(4). | -10.5/1.0/172 | OS | Physician | 11/03/2022 | 2022 | 47 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15845395&pc=QCB | 2023826-2022-04147 | 1989-10-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2, -10.0/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTION DATA: B3: REPORTED AS 25-OCT-1989 - CORRECT TO 06-SEP-2022. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | -10.0/1.0/090 | OD | Physician | 11/03/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15845399&pc=QCB | 2023826-2022-04141 | 2022-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.5/3.5/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -7.5/3.5/095 | OS | Physician | 11/01/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15845402&pc=QCB | 2023826-2022-04137 | 2022-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -4.5/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -4.5/1.5/091 | OD | Physician | 11/01/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15845405&pc=QCB | 2023826-2022-04145 | 2022-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.0/1.0/132 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -13.0/1.0/132 | OS | Physician | 11/03/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15845407&pc=QCB | 2023826-2022-04138 | 2022-10-24 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.5/1.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.5/1.0/094 | OD | Physician | 11/01/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15845412&pc=QCB | 2023826-2022-04140 | 2022-02-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.0/1.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.0/1.5/090 | OS | Physician | 11/01/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15845423&pc=QCB | 2023826-2022-04142 | 2022-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.0/1.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: D4 - EXPIRATION DATE REPORTED AS 31-MAY-2022 - CORRECT TO 31-MAY-2025. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -10.0/1.5/091 | OD | Physician | 11/01/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15845426&pc=QCB | 2023826-2022-04134 | 2022-10-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.5/3.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.5/3.0/092 | OD | Physician | 10/31/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15845443&pc=QCB | 2023826-2022-04146 | 2022-06-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.5/2.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.5/2.0/094 | OD | Physician | 11/03/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15845652&pc=QCB | 2023826-2022-04256 | 2022-06-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -07.50/+2.5/094 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD), ON (B)(6) 2022 THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. D4- UNIQUE IDENTIFIER (UDI)#- (B)(4). CLAIM# (B)(4). | -07.50/+2.5/094 | OD | Physician | 10/24/2022 | 2022 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15845666&pc=QCB | 2023826-2022-04254 | 2022-10-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.00/+2.5/085 (SPHERE/CYLINDER/AXIS), TORE DURING LOADING AND THERE WAS NO PATIENT CONTACT. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION: UNK. PATIENT INFORMATION: UNK. PATIENT INFORMATION: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -14.00/+2.5/085 | NA | Physician | 10/28/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15845724&pc=QCB | 2023826-2022-04172 | 2022-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.00/+1.5/074 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -13.00/+1.5/074 | OS | Physician | 10/24/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15846634&pc=QCB | 2023826-2022-04171 | 2022-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -14.00/+2.0/112 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED BUT THE PROBLEM WAS NOT RESOLVED. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -14.00/+2.0/112 | OD | Physician | 10/24/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15848644&pc=QCB | 2023826-2022-04264 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -9.0/5.5/111 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS AN EXCHANGE LENS IN THE PATIENTS RIGHT EYE (OD) ON 31-JUL-2022 TO RESOLVE LOW VAULT. LOW VAULT IS STILL BEING OBSERVED. THE LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. SEE MFR # 2023826-2022-042643 FOR INITIAL LENS IMPLANT. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/5.5/111 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022 AS A REPLACEMENT LENS TO RESOLVE LOW VAULT. LOW VAULT IS STILL BEING OBSERVED. THE LENS REMAINS IMPLANTED. THE PROBLEM WAS NOT RESOLVED. REPORTEDLY "THE DOCTOR BELIEVED THAT AFTER REPLACING THE CRYSTAL, THE ARCH HEIGHT WAS NOT IDEAL, DID NOT SOLVE THE PROBLEM, AND CONTINUED CLINICAL OBSERVATION". CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -9.0/5.5/111 | OD | Physician | 10/26/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15848817&pc=QCB | 2023826-2022-04263 | 2021-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -9.0/5.5/110 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. LOW VAULT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS NOT RESOLVED. CAUSE REPORTED AS UNKNOWN. SEE MFR #2023826-2022-04264 FOR EXCHANGE LENS. Manufacturer Narrative: WEIGHT:UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). | -9.0/5.5/110 | OD | Physician | 10/26/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15848887&pc=QCB | 2023826-2022-04231 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.0/1.5/001 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS A REPLACEMENT INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022, DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE DID NOT RESOLVE THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. THE STATUS OF THE EYE IS REPORTED AS, "OK" AND "THE SURGEON DECIDED TO MAKE FOLLOW-UPS IN 3 MONTHS." Manufacturer Narrative: CORRECTED DATA: B1: ADVERSE EVENT SHOULD BE OMITTED AND REPLACED WITH PRODUCT PROBLEM FOR CORRECTION. B2: REQUIRED INTERVENTION TO PREVENT PERMEANT IMPAIRMENT/DAMAGE SHOULD BE OMITTED FOR CORRECTION. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4). | -9.0/1.5/001 | OS | Physician | 10/27/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15848891&pc=QCB | 2023826-2022-04228 | 2022-10-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -8.0/2.5/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AND REMOVED DUE TO THE LENS FLIPPED UPSIDE DOWN . THE REPLACEMENT LENS WAS IMPLANTED. Manufacturer Narrative: PATIENT INFORMATION: UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | -8.0/2.5/093 | NA | Physician | 10/24/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15849512&pc=QCB | 2023826-2022-04230 | 2022-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.5/013 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE EXCHANGE DID NOT RESOLVE THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. THE STATUS OF THE EYE IS REPORTED AS, "OK" AND "THE SURGEON DECIDED TO MAKE FOLLOW-UPS IN 3 MONTHS." Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). Manufacturer Narrative: H3: DHR REVIEW: DEVICE HISTORY RECORD REVIEW FOUND ASSOCIATED SOURCE LOTS WERE MANUFACTURED WITHIN ESTABLISHED PARAMETERS AND LIMITS. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: 20-DEC-2022. H3-DEVICE EVALUATION: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). | -9.0/1.5/013 | OS | Physician | 10/27/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15850078&pc=QCB | 2023826-2022-04259 | 2022-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -11.0/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM RESOLVED. CAUSE REPORTED AS UNKNOWN. STATUS OF THE EYE "PREOPERATIVE CONJUNCTIVA PLEASE HYPEREMIA, CORNEAL INCISION TO HE JIA, ANTERIOR CHAMBER AXIS DEPTH 3CT, ANTERIOR CHAMBER CLEAR, PUPIL ROUND, LENS POSITION POSITIVE, ANGLE 27.1 DEGREES." Manufacturer Narrative: (B)(4). | -11.0/2.0/089 | OS | Physician | 10/24/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15850090&pc=QCB | 2023826-2022-04227 | 2022-10-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF A -12.5/1.5/090 (SPHERE/CYLINDER/AXIS) WAS DAMAGED UPON INSERTION PARTIALLY IN THE EYE. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE: UNK. (B)(4). | -12.5/1.5/090 | NA | Physician | 10/24/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15850282&pc=QCB | 2023826-2022-04270 | 2022-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -8.0/1.5/104 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS OTHER " LENS IS OVERSIZED." Manufacturer Narrative: (B)(4). | -8.0/1.5/104 | OD | Physician | 10/31/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15852255&pc=QCB | 2023826-2022-04290 | 2022-07-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | NA | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/079 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS THAT WAS IMPLANTED VERTICALLY DUE ROTATION NOT ASSOCIATED TO A LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: (B)(6). CLAIM# (B)(4). | -12.0/1.5/079 | OS | Physician | 10/25/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15852882&pc=QCB | 2023826-2022-04291 | 2022-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF -10.00/2.5/94 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -10.00/2.5/94 | OS | Physician | 10/26/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15853893&pc=QCB | 2023826-2022-04292 | 2022-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -14.0/+3.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH LENS THAT WAS IMPLANTED VERTICALLY DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -14.0/+3.0/090 | OS | Physician | 10/26/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15854746&pc=QCB | 2023826-2022-04247 | 2022-10-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/+1.0/095 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -14.5/+1.0/095 | OS | Physician | 10/26/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15854765&pc=QCB | 2023826-2022-04251 | 2022-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+1.5/105 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -12.0/+1.5/105 | OD | Physician | 10/27/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15854777&pc=QCB | 2023826-2022-04252 | 2022-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.0/095 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/+1.0/095 | OS | Physician | 10/27/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15854783&pc=QCB | 2023826-2022-04248 | 2022-09-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+0.5/080 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -8.0/+0.5/080 | OD | Physician | 10/26/2022 | 2022 | 41 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15854788&pc=QCB | 2023826-2022-04250 | 2022-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.5/168 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LENS TWO SIZES SHORTER IN LENGTH DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.0/+2.5/168 | OS | Physician | 10/27/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15854789&pc=QCB | 2023826-2022-04246 | 2022-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+2.5/010 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LENS TWO SIZES LONGER IN LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.5/+2.5/010 | OD | Physician | 10/25/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15854821&pc=QCB | 2023826-2022-04249 | 2022-08-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+3.0/106 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.0/+3.0/106 | OD | Physician | 10/27/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15856891&pc=QCB | 2023826-2022-04297 | 2022-01-28 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+0.5/043 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SPHERICAL LENS LONGER IN LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -7.5/+0.5/043 | OS | Physician | 10/27/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15856896&pc=QCB | 2023826-2022-04301 | 2022-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/+2.0/086 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -14.5/+2.0/086 | OS | Physician | 10/27/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15856905&pc=QCB | 2023826-2022-04300 | 2022-07-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.0/085 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -11.0/+1.0/085 | OS | Physician | 10/27/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15856911&pc=QCB | 2023826-2022-04296 | 2022-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.0/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+1.0/095 | OD | Physician | 10/27/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15856946&pc=QCB | 2023826-2022-04223 | 2022-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -09.50/+3.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS ADJUSTED AXIALLY COUNTERCLOCKWISE 3° BEFORE REPLACEMENT. THE LENS WAS REPLACED WITH A LONGER LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). | -09.50/+3.0/092 | OD | Physician | 11/03/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15856949&pc=QCB | 2023826-2022-04299 | 2022-08-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.0/113 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.0/+1.0/113 | OD | Physician | 10/27/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15857614&pc=QCB | 2023826-2022-04298 | 2022-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+0.5/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+0.5/091 | OS | Physician | 10/27/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15857625&pc=QCB | 2023826-2022-04224 | 2022-06-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.00/+1.5/076 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). Manufacturer Narrative: CORRECTED DATA: H6: 4629 SHOULD BE OMITTED AND REPLACED WITH 4627: DEVICE EXPLANTATION FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 17-JUL-2023. G3: 14-AUG-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE/DEBRIS ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -11.00/+1.5/076 | OS | Physician | 11/03/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15857717&pc=QCB | 2023826-2022-04294 | 2022-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -10.0/+0.5/11 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO LOW VAULT. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: "ON (B)(6) 2022, THE LENS WAS REPLACED WITH A DIFFERENT MODEL/LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS REPORTED AS "ALL FINE" AND "PATIENT IS HAPPY". THE CAUSE OF THE EVENT WAS UNKNOWN." CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -10.0/+0.5/11 | OS | Physician | 10/28/2022 | 2022 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15874877&pc=QCB | 2023826-2022-04308 |  | Injury | STAAR SURGICAL COMPANY | 2022-11-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -11.00/3.0/085 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. THE SURGEON REPORTS THE LENS WAS EXPLANTED AND REPLACED WITH A SHORTER LENGTH LENS DUE TO THE LENGTH OF THE LENS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -11.00/3.0/085 | NA | Other Health Care Professional | 11/07/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15876562&pc=QCB | 2023826-2022-04309 | 2022-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.0/+1.5/096 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 . THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B7- PREEXISTING MEDICAL CONDITIONS: "IRITIS, SYNECHIA, PIGMENT DISPERSION, PSEUDOEXFOLIATION." CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -17.0/+1.5/096 | OD | Physician | 10/31/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15876563&pc=QCB | 2023826-2022-04311 | 2022-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/+3.0/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 . THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -14.5/+3.0/089 | OS | Physician | 10/31/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15876566&pc=QCB | 2023826-2022-04310 | 2022-08-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+3.5/071 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). (UNDER 21YRS OF AGE AT DATE OF IMPLANT). CLAIM# (B)(4). | -8.5/+3.5/071 | OS | Physician | 10/31/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15876568&pc=QCB | 2023826-2022-04312 | 2022-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.0/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 . THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -7.0/+1.0/094 | OS | Physician | 11/01/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15882890&pc=QCB | 2023826-2022-04335 | 2022-09-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -13.50/1.5/073 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS). IT WAS REPORTED THAT THE LENS WAS EXPLANTED AND REPLACED DUE TO LENGTH. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H3-DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -13.50/1.5/073 | OS | Other Health Care Professional | 11/07/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15882897&pc=QCB | 2023826-2022-04336 | 2022-09-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-11-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -13.50/1.5/073 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD). IT WAS REPORTED THAT THE LENS WAS EXPLANTED AND REPLACED DUE TO LENGTH. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B5-EVENT DESCRIPTION: THE LENS DIOPTER SHOULD BE CORRECTED TO: -13.50/2.50/088. ADDITIONAL INFORMATION: H3-DEVICE EVALUATION: RETURNED PRODUCT EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | -13.50/1.5/073 | OD | Other Health Care Professional | 11/07/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15891975&pc=QCB | 2023826-2022-04368 | 2022-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -4.50/1.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT, LENS REMAINS IMPLANTED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL DATA: B5- ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE SAME MODEL, SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. REPORTEDLY "PATIENT IS HAPPY". DEVICE EVALUATION: H3- TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -4.50/1.0/093 | OS | Physician | 11/03/2022 | 2022 | 45 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15895525&pc=QCB | 2023826-2022-04273 | 2022-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -15.0/1.0/178 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS WITH DIFFERENT AXIS AND WAS IMPLANTED VERTICALLY. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL;DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -15.0/1.0/178 | OD | Physician | 11/01/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15899044&pc=QCB | 2023826-2022-04380 | 2022-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-01 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -8.0/3.5/093 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -8.0/3.5/093 | OD | Physician | 11/28/2022 | 2022 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15901497&pc=QCB | 2023826-2022-04378 | 2022-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -10.0/2.0/086 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2021. THE PATIENT EXPERIENCED LENS ROTATION. ON (B)(6) 2021 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND IMPLANTED VERTICALLY. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. (B)(4). | -10.0/2.0/086 | OD | Physician | 11/28/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15901923&pc=QCB | 2023826-2022-04379 | 2022-09-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -9.5/1.5/089 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS; AND IMPLANTED VERTICALLY. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: (B)(4). | -9.5/1.5/089 | OS | Physician | 11/28/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15905755&pc=QCB | 2023826-2022-04369 | 2022-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -14.0/1.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL DATA: B5- ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE SAME MODEL, SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. REPORTEDLY "PATIENT IS HAPPY". CLAIM# (B)(4). Manufacturer Narrative: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: DISREGARD SUPPLEMENTAL MDR#1 AND MDR#2, IT WAS INTENDED FOR ANOTHER COMPLAINT, INFO IS NOT APPLICABLE. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -14.0/1.5/092 | OD | Physician | 11/03/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15909475&pc=QCB | 2023826-2022-04342 | 2022-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -18.00/+1.5/071 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE EYE IS QUIET AND IOP IS NORMAL. Manufacturer Narrative: CORRECTED DATA: D4: UNIQUE DEVICE IDENTIFIER (B)(4), SHOULD BE ADDED FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | -18.00/+1.5/071 | OD | Physician | 11/09/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15909566&pc=QCB | 2023826-2022-04340 | 2022-10-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -11.00/+2.0/079 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: WEIGHT: UNK; ETHNICITY: UNK; RACE: UNK. CLAIM # (B)(4). | -11.00/+2.0/079 | OS | Other Health Care Professional | 11/04/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15910389&pc=QCB | 2023826-2022-04347 | 2022-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.00/+3.0/066 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNKNOWN ,ETHINICITY: UNKNOWN ,RACE: UNKNOWN WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3: TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -07.00/+3.0/066 | OS | Physician | 11/06/2022 | 2022 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15910503&pc=QCB | 2023826-2022-04274 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -13.0/1.5/128 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. LOW VAULT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXPLANTED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF 13.0/1.5/128 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2020. LOW VAULT WAS OBSERVED. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS NOT RESOLVED. SEE MDR#: (B)(4) FOR EXCHANGE LENS. H3: PRODUCT EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | -13.0/1.5/128 | OD | Physician | 11/03/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15911188&pc=QCB | 2023826-2022-04275 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -14.0/1.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LOW VAULT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXPLANTED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -14.0/1.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LOW VAULT WAS OBSERVED WITH ROTATION. ON (B)(6) 2022. THE LENS WAS EXCHANGED FOR A LONGER LENGTH AND THE PROBLEM WAS RESOLVED. H3: PRODUCT EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL IN MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -14.0/1.5/092 | OS | Physician | 11/03/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15911280&pc=QCB | 2023826-2022-04276 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -11.5/2.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LOW VAULT WITH ROTATION WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS REMOVED AND ON THE SAME DAY DIFFERENT SURGERY A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -11.5/2.0/093 | OD | Physician | 11/03/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15911420&pc=QCB | 2023826-2022-04441 | 2022-10-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -14.0/1.0/140 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OD) ON (B)(6) 2022. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED, AND REPLACED FOR A SMALLER LENGTH LENS DUE TO A LENS TEAR DURING DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT, BUT NO PATIENT INJURY. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/1.0/140 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022 WHEN THE LENS TORE/BROKE DURING INJECTION/DELIVERY. THERE WAS NO PATIENT INJURY. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED, AND REPLACED FOR THE SAME MODEL; SIMILAR POWER; SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: PATIENT INFORMATION:UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | -14.0/1.0/140 | OS | Physician | 11/10/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15911536&pc=QCB | 2023826-2022-04277 | 2022-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT 12.1MM TMICL 12.1 IMPLANTABLE COLLAMER LENS OF -13.5/1.5/104 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6)-2021. LOW VAULT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. POST OPERATIVE INFORMATION IS NOT AVAILABLE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE REPORTER INDICATED THAT A 12.1MM TMICL 12.1 IMPLANTABLE COLLAMER LENS OF -13.5/1.5/104 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021. LOW VAULT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS.THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS OTHER. "LENS TOO SHORT." CLAIM# (B)(4). Manufacturer Narrative: DEVICE CODE (B)(4): OFF-LABEL USE (ACD < 3.0MM) .CLAIM# (B)(4). | -13.5/1.5/104 | OD | Physician | 11/04/2022 | 2022 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15911791&pc=QCB | 2023826-2022-04439 | 2022-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.0/+1.5/069 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING WHICH WAS STILL PRESENT AFTER A YEAR. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATE FACTOR. Manufacturer Narrative: CLAIM# (B)(4). | -10.0/+1.5/069 | OD | Physician | 11/11/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15911805&pc=QCB | 2023826-2022-04278 | 2022-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM TMICL 12.1 IMPLANTABLE COLLAMER LENS OF -9.0/3.0/102 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. LOW VAULT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. CAUSE REPORTED AS OTHER. "LENS TOO SHORT." POST OPERATIVE INFORMATION IS NOT AVAILABLE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5-THE REPORTER INDICATED THAT A 12.1MM TMICL 12.1 IMPLANTABLE COLLAMER LENS OF -9.0/3.0/102 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON 23-FEB-2022. LOW VAULT WAS OBSERVED. ON 30-NOV-2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE REPORTED AS OTHER. "LENS TOO SHORT." CORRECTED DATA: D4: CORRECT SERIAL NUMBER: T946635. EXPIRATION DATE: 31-JAN-2024. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -9.0/3.0/102 | OS | Physician | 11/04/2022 | 2022 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15911867&pc=QCB | 2023826-2022-04440 | 2022-06-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -9.0/+1.5/117 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/+1.5/117 | OS | Physician | 11/11/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15911891&pc=QCB | 2023826-2022-04442 | 2022-08-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-02 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -6.0/1.0/173 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6)2022, THE LENS WAS REMOVED AND THE REPLACEMENT LENS WAS IMPLANTED HORIZONTALLY DUE TO LOW VAULT WITHOUT ROTATION. REPORTEDLY, "THE 12.6MM/T1063214 CRYSTAL WAS IMPLANTED VERTICALLY AND THE 12.6MM/T1083410 CRYSTAL WAS IMPLANTED HORIZONTALLY." THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -6.0/1.0/173 | OD | Physician | 11/10/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15912686&pc=QCB | 2023826-2022-04279 | 2022-09-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Appropriate Term/Code Not Available | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -12.0/1.5/092 (SPHERE/CYLINDER/AXIS) TORE/ BROKE DURING INJECTION/ DELIVERY INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS FOR A SAME MODEL AND SIZE LENS AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -12.0/1.5/092 | OS | Other Health Care Professional | 11/02/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15915690&pc=QCB | 2023826-2022-04381 | 2022-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.0/1.5/078 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. IN A SECOND SURGERY, THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. (B)(4). CLAIM# (B)(4). | -13.0/1.5/078 | OD | Physician | 11/14/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15915692&pc=QCB | 2023826-2022-04382 | 2022-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM512.6, -7.50/2.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. CLAIM # (B)(4). | -7.50/2.5/088 | OD | Physician | 11/16/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15915693&pc=QCB | 2023826-2022-04384 | 2022-06-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -16.0/2.0/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM # (B)(4). | -16.0/2.0/104 | OS | Physician | 11/16/2022 | 2022 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15915701&pc=QCB | 2023826-2022-04386 | 2022-08-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED IN THE OBLIQUE DIRECTION A 13.7MM VTICMO13.7, -5.0/1.0/065 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4). | -5.0/1.0/065 | OS | Physician | 11/18/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15915707&pc=QCB | 2023826-2022-04385 | 2022-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.0/1.5/110 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/1.5/110 | OD | Physician | 11/18/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15917746&pc=QCB | 2023826-2022-03665 | 2022-03-03 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.0/+2.0/112 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -15.0/+2.0/112 | OD | Physician | 09/21/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15921152&pc=QCB | 2023826-2022-04456 | 2022-07-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.5/1.5/014 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE SURGEON REPORTS LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT SPHERICAL LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 17-MAY-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN, BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -11.5/1.5/014 | OD | Physician | 12/02/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15921189&pc=QCB | 2023826-2022-04349 | 2021-02-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -08.00/+1.5/110 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -08.00/+1.5/110 | OD | Physician | 11/11/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15921194&pc=QCB | 2023826-2022-04345 | 2022-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.00/+1.5/073 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL DATA: B5: THE LENS WAS REPLACED ON (B)(6) 2022 WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. CORRECTED DATA: B2: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM#: (B)(4). | -17.00/+1.5/073 | OD | Physician | 11/11/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15921198&pc=QCB | 2023826-2022-04344 | 2022-10-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -12.00/+4.0/174 (SPHERE/CYLINDER/AXIS) TORE DURING LOADING AND WAS STUCK IN THE CARTRIDGE OR INJECTION SYSTEM. THERE WAS NO PATIENT CONTACT OR INJURY. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. Manufacturer Narrative: CORRECTED DATA: D4: UNIQUE IDENTIFIER NUMBER: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -12.00/+4.0/174 | NA | Physician | 11/09/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15924581&pc=QCB | 2023826-2022-04467 | 2022-07-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.0/+2.0/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -17.0/+2.0/090 | OD | Physician | 11/07/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15924590&pc=QCB | 2023826-2022-04468 | 2022-06-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.5/040 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -7.0/+1.5/040 | OS | Physician | 11/07/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15927338&pc=QCB | 2023826-2022-04443 | 2022-10-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event; Patient Device Interaction Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.0/2.0/082 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED, AND REPLACED FOR A SMALLER LENGTH LENS DUE TO A LENS TEAR DURING INJECTION/DELIVERY INTO THE EYE. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. STATUS OF THE EYE IS REPORTED AS, "GOOD." Manufacturer Narrative: WEIGHT AND AGE:UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.00/2.0/082 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 WHEN THE LENS TORE/BROKE DURING INJECTION/DELIVERY. THERE WAS NO PATIENT INJURY. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED, AND REPLACED FOR THE SAME MODEL AND POWER BUT SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE WAS REPORTED AS "GOOD". THE REPORTER INDICATED THE CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR AND THE DEVICE FAILED TO PERFORM AS INTENDED. H6: MEDICAL DEVICE PROBLEM CODE 1494: OFF-LABEL USE (UNDER 21 YRS OF AGE AT DATE OF IMPLANTATION). CLAIM#: (B)(4). | -9.0/2.0/082 | OS | Physician | 11/10/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15929758&pc=QCB | 2023826-2022-04460 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 -7.5/1.5/073 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS' LEFT EYE (OS) ON (B)(6) 2022 AS A REPLACEMENT LENS TO ADDRESS EXCESSIVE VAULT BUT IT DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -7.5/1.5/073 | OS | Physician | 11/07/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15933673&pc=QCB | 2023826-2022-04346 | 2022-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -17.00/+2.0/086 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REPLACED ON 05-DEC-2022 WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM #(B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -17.00/+2.0/086 | OS | Physician | 11/11/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15936111&pc=QCB | 2023826-2022-04477 | 2022-08-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.0/157 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.5/+1.0/157 | OS | Physician | 11/14/2022 | 2022 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15936535&pc=QCB | 2023826-2022-04498 | 2022-11-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -10.50/+2.0/92 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE; INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SAME LENGTH LENS. THERE WAS NO PATIENT INJURY. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.50/+2.0/92 | OD | Physician | 12/01/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15936566&pc=QCB | 2023826-2022-04462 | 2021-04-21 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 -3.0/1.0/034 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED HORIZONTALLY INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2021 AS A REPLACEMENT LENS TO ADDRESS LOW VAULT. IT WAS REPORTED THAT AFTER IMPLANTATION OF THE LENS EXCESSIVE VAULT WAS OBSERVED. PER REPORTER THE LENS WAS IMPLANTED, REMOVED AND REPLACED WITH A SAME LENGTH LENS THAT WAS IMPLANTED VERTICALLY WITHIN THE SAME SURGERY AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE ON PRODUCT, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE OPTIC DEFORMED AND THE HAPTICS BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -3.0/1.0/034 | OD | Physician | 11/10/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15937238&pc=QCB | 2023826-2022-04283 | 2022-10-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-07 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -11.0/1.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS DEVICE. Manufacturer Narrative: CLAIM # (B)(4). | -11.0/1.0/092 | OD | Physician | 11/07/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15943836&pc=QCB | 2023826-2022-04473 | 2022-07-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+2.5/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.0/+2.5/090 | OD | Physician | 11/10/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15943880&pc=QCB | 2023826-2022-04476 | 2022-08-19 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+1.0/012 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.0/+1.0/012 | OD | Physician | 11/14/2022 | 2022 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15943885&pc=QCB | 2023826-2022-04470 | 2022-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.5/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+1.5/091 | OD | Physician | 11/10/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15943887&pc=QCB | 2023826-2022-04474 | 2022-06-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.0/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -10.0/+1.0/092 | OD | Physician | 11/10/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15946703&pc=QCB | 2023826-2022-04463 | 2022-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM, VTICMO12.1, -5.0/0.5/111 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXPLANTED DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTICS TORN/BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). | -5.0/0.5/111 | OS | Physician | 11/10/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15946737&pc=QCB | 2023826-2022-04500 | 2022-11-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11/+2.5/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS NOTED AS "USER ERROR; LENS ORDERED WAS TOO LONG". Manufacturer Narrative: (B)(4). | -11/+2.5/090 | OD | Physician | 11/21/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15946809&pc=QCB | 2023826-2022-04464 | 2022-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICMO12.6, -10.0/2.0/175 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED VERTICALLY INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS THAT WAS IMPLANTED HORIZONTALLY DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9-RETURN DATE: 12-JAN-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -10.0/2.0/175 | OS | Physician | 11/10/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15946915&pc=QCB | 2023826-2022-04403 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM TMICL 12.1 IMPLANTABLE COLLAMER LENS OF -12.50/2.5/094 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE ON AN UNKNOWN DATE. REPORTEDLY THE PATIENT EXPERIENCED A LOW VAULT AND THE LENS WAS EXCHANGED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CLAIM # (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3: DEVICE EVALUATION: LENS RETURNED IN A VIAL; DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC TORN WITH RESIDUE AND DEBRIS THROUGHOUT THE LENS. CLAIM# (B)(4). | -12.50/2.5/094 | NA | Physician | 11/08/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15946997&pc=QCB | 2023826-2022-04448 | 2022-10-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -13.5/1.0/175 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2022, THE LENS WAS REMOVED DUE TO LOW VAULT WITHOUT ROTATION. CAUSE OF THE EVENT WAS A PATIENT RELATED FACTOR. REPORTEDLY, "SUBTEMPORAL DEVIATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: DATE RETURNED: 20/DEC/2022 . H3- DEVICE EVALUATION: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, RACE, AND ETHNICITY:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -13.5/1.0/175 | OS | Physician | 11/10/2022 | 2022 | 42 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15947226&pc=QCB | 2023826-2022-04447 | 2022-08-12 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-08 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -12.5/2.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED AS A REPLACEMENT LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE INTIAL LENS WAS REPLACED DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. REPORTEDLY, STATUS OF THE EYE IS, "DISCOMFORT, GOOD VISION." Manufacturer Narrative: CORRECTED DATA: B1: ADVERSE EVENT OR PRODUCT PROBLEM: SHOULD BE CORRECTED TO PRODUCT PROBLEM. B2: OUTCOMES ATTRIBUTED TO ADVERSE EVENT: SHOULD BE OMITTED FOR CORRECTION. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM#: (B)(4). | -12.5/2.0/093 | OD | Physician | 11/10/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15951759&pc=QCB | 2023826-2022-04451 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS WITH A -9.5/2.5/117 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). LOW VAULT WAS OBSERVED. AN EXCHANGE IS PLANNED, AND THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: B1 :CORRECTED TO PRODUCT PROBLEM (E.G., DEFECTS/MALFUNCTIONS) IN INITIAL MDR. B2: REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) SHOULD BE REMOVED FROM THE INITIAL MDR B5 - CORRECTED TO: "THE REPORTER INDICATED THAT A 12.6 VTICM5_12.6 OF DIOPTER -9.50/2.5/077 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS THAT WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LOW VAULT WAS OBSERVED. AN EXCHANGE IS PLANNED, AND THE CAUSE OF THE EVENT IS UNKNOWN." IN THE INITIAL MDR. D4 :MODEL# VTICM5 12.6 SHOULD BE ADDED TO THE INITIAL MDR. D4 :SERIAL # (B)(6) SHOULD BE ADDED TO THE INITIAL MDR. D4 :EXPIRATION DATE: 30-NOV-2023 SHOULD BE ADDED TO THE INITIAL MDR. D6A :IMPLANT DATE (B)(6) 2021 SHOULD BE ADDED TO THE INITIAL MDR. . H4 : DEVICE MANUFACTURE DATE: 24-DEC-2020 SHOULD BE ADDED TO THE INITIAL MDR. H6 :INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: D4: EXPIRATION DATE: 31-JUL-2025 SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: B5 - LENS REPOSITIONED ON (B)(6) 2021 BUT DID NOT RESOLVE THE PROBLEM LENS WAS EXPLANTED ON (B)(6) 2023 AND REPLACED ON THE SAME DAY IN A DIFFERENT SURGERY WITH A LONGER LENGTH LENS. PROBLEM WAS RESOLVED. PATIENT HAPPY, GOOD VISION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -9.5/2.5/077 | OS | Physician | 11/11/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15952808&pc=QCB | 2023826-2022-04450 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.6 TORIC IMPLANTABLE COLLAMER LENS THAT WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). LOW VAULT WAS OBSERVED. AN EXCHANGE IS PLANNED, AND THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: B1: CORRECTED TO PRODUCT PROBLEM (E.G., DEFECTS/MALFUNCTIONS) IN INITIAL MDR. B2: REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) SHOULD BE REMOVED FROM THE INITIAL MDR B5: CORRECTED TO: "THE REPORTER INDICATED THAT A 12.6 VTICM5 OF DIOPTER -9.50/2.5/109 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS THAT WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD). LOW VAULT WAS OBSERVED. AN EXCHANGE IS PLANNED, AND THE CAUSE OF THE EVENT IS UNKNOWN." IN THE INITIAL MDR. D4: MODEL#: VTICM5 12.6 SHOULD BE ADDED TO THE INITIAL MDR. D4: SERIAL#: (B)(6) SHOULD BE ADDED TO THE INITIAL MDR. D4: EXPIRATION DATE: 31-AUG-2023 SHOULD BE ADDED TO THE INITIAL MDR. D6A: IMPLANT DATE: ON (B)(6) 2021 SHOULD BE ADDED TO THE INITIAL MDR. H4: DEVICE MANUFACTURE DATE: 01-OCT-2020 SHOULD BE ADDED TO THE INITIAL MDR. H6: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). Manufacturer Narrative: B5 - LENS REPOSITIONED ON (B)(6) 2021 BUT DID NOT RESOLVE THE PROBLEM LENS WAS EXCHANGED ON (B)(6) 2023 WITH A LONGER LENGTH LENS. PROBLEM WAS RESOLVED. PATIENT HAPPY, NO SYMPTOMS. CLAIM# (B)(4). Manufacturer Narrative: B5 - LENS WAS EXPLANTED ON (B)(6) 2023 AND REPLACED ON THE SAME DAY IN A DIFFERENT SURGERY WITH A LONGER LENGTH LENS. PROBLEM WAS RESOLVED. PATIENT HAPPY, GOOD VISION. CLAIM# (B)(4). | -9.50/2.5/109 | OD | Physician | 11/11/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15955223&pc=QCB | 2023826-2022-04523 | 2022-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+2.5/178 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LENS TWO SIZES LONGER IN LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.0/+2.5/178 | OS | Physician | 11/09/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15955249&pc=QCB | 2023826-2022-04406 | 2022-10-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -10.50/1.5/081 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). | -10.50/1.5/081 | OS | Physician | 11/09/2022 | 2022 | 28 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15955385&pc=QCB | 2023826-2022-04526 | 2022-05-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5_13.7; -8.0/2.0/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS. THE CAUSE OF THE EVENT WAS REPORTED AS "LENS TOO BIG; SURGEON PREFERENCE CHANGE". THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). | -8.0/2.0/094 | OS | Physician | 11/29/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15955509&pc=QCB | 2023826-2022-04524 | 2022-11-23 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -7.0/+2.5/078 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE PATIENT EXPERIENCED HIGH VAULTING. LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: B5 - REPORTEDLY, "THE PATIENT IS ACTUALLY QUITE HAPPY. A RE-LAUNCH IS PLANNED FOR THE END OF THE YEAR. I WOULD LIKE TO CLOSE THE CASE FIRST AS IT IS QUESTIONABLE WHETHER ANY ACTION WILL BE TAKEN." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM (B)(4). Manufacturer Narrative: (B)(4). | -7.0/+2.5/078 | OS | Physician | 11/24/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15955593&pc=QCB | 2023826-2022-04409 | 2022-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -9.0/1.0/085 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.0/1.0/085 | OS | Physician | 11/11/2022 | 2022 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15955634&pc=QCB | 2023826-2022-04407 | 2022-11-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -16.50/1.5/042 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LOW VAULT WITH ROTATION WAS OBSERVED. THE LENS REMAINS IMPLANTED. CAUSE REPORTED AS UNKNOWN. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: H10 IN PREVIOUS MDR SHOULD HAVE WORK ORDER SEARCH RESULTS AS FOLLOWS: H6-INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). Manufacturer Narrative: B2-REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE (DEVICES). B5-THE REPORTER INDICATED THAT A 13.1MM VTICMO 12.1 IMPLANTABLE COLLAMER LENS OF -16.50/1.5/042 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. LOW VAULT WITH ROTATION WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. D6B:(B)(6) 2023. CLAIM# (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | -16.50/1.5/042 | OS | Physician | 11/10/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15955683&pc=QCB | 2023826-2022-04405 | 2022-10-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -9.0/1.00/098 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/1.00/098 | OD | Physician | 11/09/2022 | 2022 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15955691&pc=QCB | 2023826-2022-04408 | 2022-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -8.50/2.5/099 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -8.50/2.5/099 | OD | Physician | 11/11/2022 | 2022 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15960125&pc=QCB | 2023826-2022-04525 | 2022-11-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6M VTICM5_12.6; -6.5/+5.5/84 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS IMPLANTED UPSIDE DOWN; INTRAOPERATIVELY THE LENS WAS REMOVED AND REPLACED WITH A SAME LENGTH LENS AND THE PROBLEM RESOLVED. THERE WAS NO PATIENT INJURY. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -6.5/+5.5/84 | OD | Physician | 11/16/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15965790&pc=QCB | 2023826-2022-04481 | 2022-09-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -7.5/1.5/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS INTRAOPERATIVELY IMPLANTED AND REMOVED DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. STATUS OF THE EYE IS REPORTED AS, "GOOD." ALSO REPORTED, "NEXT OPERATION DATE IS UNDECIDED." Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: H5: 3331- DEVICE HISTORY RECORD - BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDERS INCLUDING SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: PT INFO: UNK. WORK ORDER SEARCH FOUND ONE SIMILIAR COMPLAINT. (B)(4). | -7.5/1.5/90 | OD | Physician | 11/14/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15966117&pc=QCB | 2023826-2022-04482 | 2022-09-11 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -6.0/1.5/89 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS INTRAOPERATIVELY IMPLANTED AND REMOVED DUE TO EXCESSIVE VAULTING. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. STATUS OF THE EYE IS REPORTED AS, "GOOD.". ALSO REPORTED, "NEXT OPERATION DATE IS UNDECIDED.". Manufacturer Narrative: (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNKNOWN. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). | -6.0/1.5/89 | OS | Physician | 11/14/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15974300&pc=QCB | 2023826-2022-04585 | 2022-12-03 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -12.5/1.0/070 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. IT WAS INDICATED THAT THE LENS WAS STUCK IN THE CARTRIDGE SYSTEM AND THE SURGEON FOUND THE LENS COULD NOT BE PUSHED OUT SMOOTHLY. IT WAS NOTED "RETRACTION OCCURS WHEN THE LENS IS LOADED; CARTRIDGE TIP NARROW". THE LENS WAS STILL IMPLANTED WITH NO PATIENT INJURY. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). | -12.5/1.0/070 | OS | Physician | 12/01/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15975201&pc=QCB | 2023826-2022-04410 | 2022-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5 12.1 IMPLANTABLE COLLAMER LENS OF -13.00/3.0/095 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED ON (B)(6) 2022 AND THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL; DRY WITH RESIDUE AND DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -13.00/3.0/095 | OS | Physician | 11/14/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15975808&pc=QCB | 2023826-2022-04398 | 2022-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -14.50/1.0/036 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE- UNK. CLAIM# (B)(4). | -14.50/1.0/036 | OD | Physician | 11/24/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15975809&pc=QCB | 2023826-2022-04389 | 2022-10-01 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6, -6.50/1.50/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A LONGER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S RIGHT EYE (OD) DUE TO LOW VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -6.50/1.50/094 | OD | Physician | 11/21/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15975810&pc=QCB | 2023826-2022-04391 | 2022-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.0/2.5/072 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY, AND RACE - UNK. (B)(4). | -9.0/2.5/072 | OS | Physician | 11/22/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15975813&pc=QCB | 2023826-2022-04392 | 2022-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.0/1.0/128 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED VERTICALLY WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY - UNK, RACE - UNK. (B)(4). | -9.0/1.0/128 | OS | Physician | 11/23/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15975818&pc=QCB | 2023826-2022-04400 | 2022-05-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM512.6, -12.0/3.0/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY, AND RACE- UNK. CLAIM # (B)(4). | -12.0/3.0/096 | OD | Physician | 11/24/2022 | 2022 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15975835&pc=QCB | 2023826-2022-04402 | 2022-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -5.0/1.0/166 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. CLAIM# (B)(4). | -5.0/1.0/166 | OS | Physician | 11/25/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15975836&pc=QCB | 2023826-2022-04399 | 2022-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -15.00/1.0/121 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6)2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: (B)(4). | -15.00/1.0/121 | OS | Physician | 11/24/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15975837&pc=QCB | 2023826-2022-04390 | 2022-11-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.5/2.5/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. (B)(4). | -10.5/2.5/088 | OD | Physician | 11/22/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15975850&pc=QCB | 2023826-2022-04387 | 2022-08-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM5_12.1, -12.0/2.0/020 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S LEFT EYE (OS) DUE TO LOW VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -12.0/2.0/020 | OS | Physician | 11/21/2022 | 2022 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15975851&pc=QCB | 2023826-2022-04401 | 2022-04-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -5.0/1.0/026 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE LENS WAS REPLACED VERTICALLY WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -5.0/1.0/026 | OD | Physician | 11/25/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15975852&pc=QCB | 2023826-2022-04393 | 2022-09-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.5/1.0/107 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE- UNK. CLAIM# (B)(4). | -10.5/1.0/107 | OD | Physician | 11/23/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15980376&pc=QCB | 2023826-2022-04587 | 2022-11-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -12.00/5.5/098 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULT. PER UPDATED INFORMATION VA 6/6, EYE IS QUIET AND HEALTHY, VAULTING 100MICRON. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH - NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -12.00/5.5/098 | OD | Physician | 11/14/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15981105&pc=QCB | 2023826-2022-04588 | 2022-11-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -12.00/3.5/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULT. PER UPDATED INFORMATION VA 6/6, EYE IS QUIET AND HEALTHY, VAULTING 100MICRON. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -12.00/3.5/090 | OS | Physician | 11/14/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15982776&pc=QCB | 2023826-2022-04589 | 2022-11-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-14 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 13.2MM VTICM5_13.2 -11.50/1.5/086 IMPLANTABLE COLLAMER LENS TORE/BROKE BEFORE LOADING . FOR CAUSE OF EVENT THE REPORTER STATED " HAPTIC BROKEN WHILE REMOVING FROM VIAL". A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -11.50/1.5/086 | NA | Physician | 11/14/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15988554&pc=QCB | 2023826-2022-04506 | 2022-11-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+1.0/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -12.5/+1.0/094 | OS | Physician | 11/23/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15988567&pc=QCB | 2023826-2022-04509 | 2022-07-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -2.5/+1.5/103 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -2.5/+1.5/103 | OS | Physician | 11/23/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15988573&pc=QCB | 2023826-2022-04511 | 2022-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.0/+3.0/102 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -14.0/+3.0/102 | OS | Physician | 11/24/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15988576&pc=QCB | 2023826-2022-04501 | 2022-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/167 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN.Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/167 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. ON (B)(6) 2022 A SHORTER LENGTH LENS WITH A DIFFERENT AXIS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H6 HEALTH EFFECT - IMPACT CODE: CORRECTED TO 4627. CLAIM# (B)(4). | -9.5/+1.0/167 | OS | Physician | 11/20/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15988583&pc=QCB | 2023826-2022-04508 | 2022-10-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+2.0/087 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+2.0/087 | OD | Physician | 11/23/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15988606&pc=QCB | 2023826-2022-04504 | 2022-11-11 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS WITH DIOPTER -7.0/+0.5/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. IN A SEPARATE SURGERY THAT SAME DAY THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -7.0/+0.5/090 | OS | Physician | 11/22/2022 | 2022 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15988627&pc=QCB | 2023826-2022-04502 | 2022-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+1.5/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -6.5/+1.5/092 | OS | Physician | 11/20/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15989730&pc=QCB | 2023826-2022-04517 | 2022-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.5/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -7.0/+1.5/089 | OS | Physician | 11/25/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15989799&pc=QCB | 2023826-2022-04512 | 2022-08-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+1.5/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -12.5/+1.5/093 | OD | Physician | 11/24/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15989823&pc=QCB | 2023826-2022-04515 | 2022-07-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.0/078 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/+1.0/078 | OS | Physician | 11/24/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15989839&pc=QCB | 2023826-2022-04514 | 2022-07-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.0/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+1.0/094 | OS | Physician | 11/24/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15989860&pc=QCB | 2023826-2022-04520 | 2022-08-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.5/107 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+1.5/107 | OD | Physician | 11/25/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15990666&pc=QCB | 2023826-2022-04532 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/034 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6)2022 . THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). Manufacturer Narrative: CORRECTED DATA: D4: UNIQUE IDENTIFIER NUMBER: (B)(4). H6: OFF-LABEL-1494 PREVIOUS HISTORY OF OCULAR SURGERY INCLUDING REFRACTIVE CORNEAL SURGERY. CLAIM#: (B)(4). | -12.00/+2.0/034 | OD | Physician | 11/22/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15990670&pc=QCB | 2023826-2022-04531 | 2022-08-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -07.50/+3.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), IN (B)(6) 2022. THE LENS WAS RMOVED ON (B)(6) 2022 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR. Manufacturer Narrative: (B)(4). Manufacturer Narrative: D4 - UNIQUE IDENTIFIER (UDI)# (B)(4) SHOULD BE ADDED TO THE INITAL MDR. CLAIM #(B)(4). | -07.50/+3.5/088 | OD | Physician | 11/23/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15990697&pc=QCB | 2023826-2022-04527 | 2022-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.00/+4.0/093 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED DUE TO HIGH VAULTING. THE LENS WAS SUCCESSFULLY REPLACED WITH A SHORTER LENS WITH NO PATIENT HARM. THE LENS WAS DISCARDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -12.00/+4.0/093 | OD | Physician | 11/22/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15990720&pc=QCB | 2023826-2022-04533 | 2021-11-04 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/051 (SPHERE/CYLINDER/AXIS) WAS BROKEN DURING THE LOADING PROCESS. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50/+1.0/051 (SPHERE/CYLINDER/AXIS) WAS TORN DURING DELIVERY INTO THE PATIENTS RIGHT EYE (OD). THERE WAS NO PATIENT CONTACT OR INJURY. CLAIM # (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). | -10.50/+1.0/051 | OD | Physician | 11/21/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15990776&pc=QCB | 2023826-2022-04530 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -14.50/+1.5/108 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6)2022 . THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: D4 - UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INTIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #(B)(4). | -14.50/+1.5/108 | OS | Physician | 11/21/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15991309&pc=QCB | 2023826-2022-04592 | 2022-11-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF -15.5/2.0/045 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT, LENS REMAINS IMPLANTED. IN THE REPORTERS OPINION THE CAUSE OF THIS EVENT IS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER UPDATED INFORMATION THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6)2022 AND THE PROBLEM RESOLVED. CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -15.5/2.0/045 | OS | Physician | 11/15/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15992004&pc=QCB | 2023826-2022-04418 | 2022-11-13 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -14.50/2.50/025 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS). ON (B)(6) 2022 THE LENS WAS IMPLANTED AND REMOVED INTRA-OPERATIVELY DUE TO LENS BEING IMPLANTED UPSIDE DOWN. NO PATIENT INJURY REPORTED. A REPLACEMENT LENS OF THE SAME MODEL AND SIZE BUT DIFFERENT POWER WAS LATER IMPLANTED AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS USER ERROR. "LENS IMPLANTED UPSIDE DOWN AND WHILE REMOVING THE LENS THE HAPTIC TORN." Manufacturer Narrative: CLAIM # (B)(4). | -14.50/2.50/025 | OS | Physician | 11/15/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15992015&pc=QCB | 2023826-2022-04415 | 2022-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -10.0/1.0/097 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS ROTATION WAS OBSERVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3-DEVICE EVALUATION: LENS RETURNED IN A MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -10.0/1.0/097 | OS | Physician | 11/15/2022 | 2022 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15993051&pc=QCB | 2023826-2022-04422 | 2022-11-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 2.6 IMPLANTABLE COLLAMER LENS OF -14.5/3.5/110 (SPHERE/CYLINDER/AXIS) TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS NOT IMPLANTED. PATIENT CONTACT WAS REPORTED BUT NO PATIENT INJURY. A REPLACEMENT LENS WAS LATER IMPLANTED OF THE SAME MODEL/SIZE AND SIMILAR POWER. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -14.5/3.5/110 | OD | Other Health Care Professional | 11/16/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15993100&pc=QCB | 2023826-2022-04489 | 2021-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -10.5/1.5/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE LEFT EYE (OS) OF A PATIENT'S WITH PRE-EXISTING AMBLYOPIA ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. CAUSE IS REPORTED AS THE OTHER, AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. REPORTEDLY, "ICL ROTATE VERTICALLY AROUND ONE MONTH AFTER IMPLANTATION. ALSO, "KNOWN HISTORY OF BOTH EYES AMBLOYOPIA. BCVA PREOP: 0.6" THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: D4: VTICM5_12.1 IS THE CORRECT MODEL. ADDITIONAL INFORMATION: D9: 01-09-2023. H3-RETURNED PRODUCT EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE LENS HAPTIC TORN AND BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY AND RACE:UNK. OFF LABEL USE: PRE-EXISTING AMBLYOPIA. CLAIM#(B)(4). | -10.5/1.5/090 | OS | Physician | 11/16/2022 | 2022 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15993101&pc=QCB | 2023826-2022-04488 | 2021-05-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-15 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -7.5/2.0/101 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE RIGHT EYE (OD) OF A PATIENT'S WITH PRE-EXISTING AMBLYOPIA ON (B)(6) 2021. ON (B)(6) 2021, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. CAUSE IS REPORTED AS THE OTHER, AND THE DEVICE DID NOT FAIL TO PERFORM AS INTENDED. REPORTEDLY, "ICL ROTATE VERTICALLY AROUND ONE MONTH AFTER IMPLANTATON. ALSO, "KNOWN HISTORY OF BOTH EYES AMBLOYOPIA. BCVA PREOP: 0.6" THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: D4: VTICM5_12.1 IS THE CORRECT MODEL. D9: 01-09-2023. ADDITIONAL INFORMATION: H3-RETURNED PRODUCT EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -7.5/2.0/101 | OD | Physician | 11/16/2022 | 2022 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15993671&pc=QCB | 2023826-2022-04595 | 2022-10-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 13.7MM, VTICM5_13.7 -12.00/1.5/085 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS TORE/BROKE DURING LOADING. THERE WAS NO PATIENT CONTACT. ON (B)(6) 2022 A REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.00/1.5/085 | NA | Physician | 11/17/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=15993722&pc=QCB | 2023826-2022-04594 | 2022-04-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -6.00/2.5/069 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON 05-JUN-2021. ON (B)(6) 2022 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND A REPLACEMENT LENS OF THE SAME LENGTH WAS LATER IMPLANTED AND THE PROBLEM RESOLVED. Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT, IN MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND FIBERS ON THE LENS SURFACE, NO VISIBLE DAMAGE WAS OBSERVED . DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM # (B)(4). Manufacturer Narrative: CLAIM # (B)(4). | -6.00/2.5/069 | OS | Physician | 11/17/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16000210&pc=QCB | 2023826-2022-04528 | 2022-11-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.50/+2.0/072 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED DUE TO HIGH VAULTING. THE LENS WAS SUCCESSFULLY REPLACED WITH A SHORTER LENS WITH NO PATIENT HARM. THE LENS WAS DISCARDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. CLAIM# (B)(4). | -11.50/+2.0/072 | OS | Physician | 11/22/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16000294&pc=QCB | 2023826-2022-04534 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -13.50/+1.5/028 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING AND PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -13.50/+1.5/028 | OD | Physician | 11/23/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16000594&pc=QCB | 2023826-2022-04624 | 2022-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A VTICM5_12.6; -13.00/+4.5/2 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2022 THE LENS WAS REPOSITIONED BUT THIS DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXPLANTED; AND ON THIS SAME DATE IN A SEPARATE SURGERY A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID WITH RESIDUE/DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -13.00/+4.5/2 | OS | Physician | 12/15/2022 | 2022 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16000782&pc=QCB | 2023826-2022-04495 | 2022-11-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -9.5/1.5/91 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE IS REPORTED AS UNKNOWN. REPORTEDLY. "EYE QUIET LENS PLACED IN POSITION." Manufacturer Narrative: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -9.5/1.5/91 | OD | Physician | 11/19/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16001442&pc=QCB | 2023826-2022-04425 | 2022-11-11 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -11.5/1.5/095 (SPHERE/CYLINDER/AXIS) TORE/ BROKE DURING INJECTION/ DELIVERY INTO THE PATIENT'S LEFT EYE (OS) AND ALSO IMPLANTED UPSIDE DOWN. THE LENS WAS IMPLANTED; REMOVED ON (B)(6) 2022. A REPLACEMENT LENS WAS LATER IMPLANTED OF THE SAME MODEL/SIZE AND POWER. THE REPLACEMENT LENS RESOLVED THE PROBLEM. Manufacturer Narrative: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. (B)(4). | -11.5/1.5/095 | OS | Physician | 11/18/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16010277&pc=QCB | 2023826-2022-04650 | 2022-07-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -12.0/+1/084 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. IT WAS REPORTED THAT ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT DIFFERENT POWER LENS. THE CAUSE OF THE EVENT WAS REPORTER AS A USER ERROR AND NOTED "THE CUSTOMER CHOSE THE WRONG LENS BY ACCIDENT". Manufacturer Narrative: CLAIM# (B)(4). | -12.0/+1/084 | OS | Physician | 11/23/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16010408&pc=QCB | 2023826-2022-04647 | 2022-11-07 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -5.00/3.5/075 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE SURGEON REPORTS LOW VAULTING AND LENS ROTATION WITHIN 1 MONTH OF SURGERY AND 2 DIOPTER ASTIGMATISM. THE LENS REMAINS IMPLANTED. THE CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "MOST LIKELY RELATED TO THE ICL LENGTH; NEED LARGER SIZE BETTER". Manufacturer Narrative: (B)(4). | -5.00/3.5/075 | OS | Physician | 11/21/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16010515&pc=QCB | 2023826-2022-04541 | 2022-07-31 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.00/+5.5/091 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)# - (B)(4) SHOULD BE ADDED TO THE INITAL MDR. CLAIM # (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, AND RACE: UNK. DEVICE CODE: (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -07.00/+5.5/091 | OS | Physician | 11/23/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16010530&pc=QCB | 2023826-2022-04543 | 2020-04-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -07.50/+5.0/081 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2020 DUE TO PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5- "THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -07.50/+5.0/081 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE PATIENT EXPERIENCED REFRACTIVE SURPRISE. REPORTEDLY THE LENS WAS REPOSITIONED. ON 15/APR/2020 THE LENS WAS EXCHANGED WITH THE SAME MODEL AND LENGTH; STRONGER ASTIGMATIC POWERS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. " CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -07.50/+5.0/081 | OS | Physician | 11/22/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16010531&pc=QCB | 2023826-2022-04542 | 2021-01-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -17.00/+2.5/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2019. THE LENS WAS EXPLANTED ON (B)(6) 2021 DUE TO THE PATIENT EPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | -17.00/+2.5/088 | OS | Physician | 11/22/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16010543&pc=QCB | 2023826-2022-04536 | 2022-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -08.50/+4.0/084 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | -08.50/+4.0/084 | OS | Physician | 11/23/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16018260&pc=QCB | 2023826-2022-04656 | 2022-11-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -13.5/1.5/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022 AS AN EXCHANGE LENS. ICL HAS PERSISTENTLY LOW VAULT ~10% CENTRALLY WHICH DECREASES TOWARDS MIDPERIPHERY. THE LENS REMAINS IMPLANTED. NO FURTHER INFORMATION HAS BEEN PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. MFR# 2023826-2022-04428 FOR INITIAL LENS. Manufacturer Narrative: ADDITIONALINFORMATION: B5- RECEIVED INFORMATION THAT THIS EXCHANGE LENS DID RESOLVE THE ISSUE OF LOW VAULT. THIS CLAIM IS NOT A COMPLAINT DUE TO NO PRODUCT DEFICIENCY AGAINST THE LENS. PLEASE DISREGARD THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: INVESTIGATION TYPE: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM#: (B)(4). | -13.5/1.5/090 | OD | Other Health Care Professional | 11/30/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16019649&pc=QCB | 2023826-2022-04430 | 2022-11-21 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event; Ejection Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5 13.7 IMPLANTABLE COLLAMER LENS OF -8.5/1.0/071 (SPHERE/CYLINDER/AXIS) GOT STUCK IN CARTRIDGE OR INJECTION SYSTEM. THE LENS WAS NOT IMPLANTED INTO THE PATIENT BUT IT DID HAVE PATIENT CONTACT WITH NO PATIENT INJURY. REPORTEDLY "ON THE RIGHT EYE, AFTER SUCCESSFUL LOADING, THE LENS WAS VISUALIZED TO BE UPSIDE DOWN IN THE CARTRIDGE AFTER 1/4 OF THE LENS WAS INJECTED OUT. HE DIDN'T COMPLETE THE LENS INJECTION, AND INSTEAD REMOVED THE CARTRIDGE TIP FROM THE INCISION AND INJECTED THE LENS ONTO THE STERILE FIELD. ONE OF THE FOOT PLATES WAS VISUALIZED TO BE CAUGHT UNDER THE PLUNGER SILICONE TIP. THE BACK-UP LENS WAS OPENED AND LOADED INTO A NEW LIOLI AND CARTRIDGE, AND IMPLANTED INTO THE EYE WITHOUT COMPLICATION." CAUSE REPORTED AS DEVICE; "THREE LIOLI PLASTIC INJECTOR ARMS WERE VISIBLY WARPED FROM THE SAME BOX. DOCTOR PREVIOUSLY HAD AN INJECTOR EVENT ON (B)(6)2022. I WAS NOT PRESENT AT THE TIME OF THAT EVENT, BUT DR. (B)(6) CONFIRMED THE INJECTORS WERE FROM THE SAME BOX." Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM #: (B)(4). | -8.5/1.0/071 | OD | Physician | 11/21/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16022176&pc=QCB | 2023826-2022-04633 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -6.50/2.5/098 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF EVENT WAS USER ERROR. Manufacturer Narrative: CLAIM# (B)(4). | -6.50/2.5/098 | OD | Physician | 11/21/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16022258&pc=QCB | 2023826-2022-04630 | 2022-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS -7.00/2.0/107 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF EVENT WAS USER ERROR. Manufacturer Narrative: (B)(4). | -7.00/2.0/107 | OS | Physician | 11/21/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16022361&pc=QCB | 2023826-2022-04631 | 2022-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS -4.50/1.5/084 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF EVENT WAS USER ERROR. Manufacturer Narrative: (B)(4). | -4.50/1.5/084 | OD | Physician | 11/21/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16022415&pc=QCB | 2023826-2022-04632 | 2022-09-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS -4.50/1.5/114 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6)2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. IN THE REPORTERS OPINION THE CAUSE OF EVENT WAS USER ERROR. Manufacturer Narrative: CLAIM# (B)(4). | -4.50/1.5/114 | OS | Physician | 11/21/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16026092&pc=QCB | 2023826-2022-04607 | 2022-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.5/093 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2019. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.5/+1.5/093 | OS | Physician | 11/28/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16026093&pc=QCB | 2023826-2022-04605 | 2022-09-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.5/+3.0/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -15.5/+3.0/089 | OS | Physician | 11/28/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16026094&pc=QCB | 2023826-2022-04604 | 2022-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/+1.5/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -5.5/+1.5/090 | OD | Physician | 11/28/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16026095&pc=QCB | 2023826-2022-04609 | 2022-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.5/+3.0/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -15.5/+3.0/095 | OD | Physician | 11/28/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16026298&pc=QCB | 2023826-2022-04606 | 2022-07-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+1.5/108 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -10.5/+1.5/108 | OD | Physician | 11/28/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16026402&pc=QCB | 2023826-2022-03281 | 2022-06-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 12.6MM, VTICMO12.6, -14.5/1.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS TORE/BROKE DURING INJECTION/DELIVERY INTO THE RIGHT EYE (OD). THE LENS WAS IMPLANTED, REMOVED AND REPLACED INTRAOPERATIVELY. THERE WAS NO PATIENT INJURY. PROBLEM RESOLVED. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | -14.5/1.0/095 | OD | Physician | 06/23/2022 | 2022 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16028643&pc=QCB | 2023826-2022-04539 | 2022-11-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -05.00/+1.0/106 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: UNK. (B)(4). CLAIM # (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4). SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). | -05.00/+1.0/106 | OS | Physician | 11/24/2022 | 2022 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16028703&pc=QCB | 2023826-2022-04537 | 2022-10-07 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -08.50/+4.5/174 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE DUE TO THE WRONG DATA ENTRY (THE AXIS OF OD WAS ENTERED WRONG 80 DEGREES INSTEAD OF CORRECT AXIS 8 DEGREES, WHICH CAUSED A HUGE CHANGE IN REFRACTION). THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO USER ERROR, THE REPORTER INDICATED THERE WAS NOTHING WRONG WITH THE ICL LENS. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. EXPLANT DATE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -08.50/+4.5/174 | OD | Physician | 11/23/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16028719&pc=QCB | 2023826-2022-04538 | 2022-11-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -04.50/+0.5/085 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPORTED AS EXCESSIVE VAULTING AND REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT: UNK, ETHNICITY: UNK, RACE: UNK, AND EXPLANT DATE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -04.50/+0.5/085 | OD | Physician | 11/24/2022 | 2022 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16029482&pc=QCB | 2023826-2022-04433 | 2022-10-25 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -5.00/0.50/122 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. LOW VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY:UNK. RACE: UNK. INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM #: (B)(4). | -5.00/0.50/122 | OS | Physician | 11/21/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16029705&pc=QCB | 2023826-2022-04432 | 2022-09-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT 13.2 VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -10.0/2.5/113 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS DEVICE. "EXCESSIVE LOW VAULT ASSOCIATED TO LENS ROTATION." Manufacturer Narrative: (B)(4). | -10.0/2.5/113 | OD | Physician | 11/21/2022 | 2022 | 31 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16030625&pc=QCB | 2023826-2022-04615 | 2022-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.0/+1.0/064 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -16.0/+1.0/064 | OS | Physician | 11/28/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16030627&pc=QCB | 2023826-2022-04611 | 2022-07-16 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.5/+2.0/073 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -6.5/+2.0/073 | OS | Physician | 11/28/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16030631&pc=QCB | 2023826-2022-04616 | 2022-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+1.5/175 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LENS TWO SIZES SHORTER IN LENGTH DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -12.0/+1.5/175 | OS | Physician | 11/28/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16030632&pc=QCB | 2023826-2022-04617 | 2022-09-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+1.5/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -7.0/+1.5/091 | OS | Physician | 11/28/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16030634&pc=QCB | 2023826-2022-04618 | 2022-09-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -16.5/+6.0/102 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -16.5/+6.0/102 | OD | Physician | 11/28/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16030635&pc=QCB | 2023826-2022-04619 | 2022-10-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.5/+1.0/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -12.5/+1.0/092 | OD | Physician | 11/28/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16030636&pc=QCB | 2023826-2022-04620 | 2022-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-21 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.0/075 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.5/+1.0/075 | OS | Physician | 11/28/2022 | 2022 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16032224&pc=QCB | 2023826-2022-04634 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 -8.00/3.0/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) . THE LENS WAS EXPLANTED DUE TO LOW VAULT WITH ROTATION. ON (B)(6) 2022 A REPLACEMENT LENS OF LONGER LENGTH WAS IMPLANTED SUCCESSFULLY. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | -8.00/3.0/087 | OS | Other Health Care Professional | 11/22/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16032297&pc=QCB | 2023826-2022-04637 | 2022-10-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 -14.00/2.5/074 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 LENS REPOSITIONING WAS PERFORMED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. THE LENS WAS EXPLANTED AND A REPLACEMENT LENS OF LONGER LENGTH WAS LATER IMPLANTED. IT WAS REPORTED THAT THE PROBLEM RESOLVED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: H3- DEVICE EVALUATION: LENS WAS RETURNED DRY WITH RESIDUE/DEBRIS ON PRODUCT , INSIDE A SPECIMEN CUP. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM # (B)(4). Manufacturer Narrative: CORRECTED INFORMATION: B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.00/+2.50/074 (SPHERE/CYLINDER/AXIS) INTO THE LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED A ROTATION NOT ASSOCIATED WITH A LOW VAULT. ON (B)(6) 2022 THE LENS WAS REPOSITIONED BUT THAT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXPLANTED AND LATER REPLACED WITH THE SAME MODEL, LONGER LENGTH AND THE PROBLEM WAS RESOLVED. THE REPORTER INDICATED THAT THE CAUSE OF THE EVENT WAS "PATIENT RELATED FACTORS - LENS ROTATED FROM ORIGINAL POSITION, LENS WAS ROTATED MANUALLY TO CORRECT POSITION ON (B)(6) 2022 & SUBSEQUENTLY ROTATED BACK". CLAIM# (B)(4). | -14.00/2.5/074 | OS | Other Health Care Professional | 11/23/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16032444&pc=QCB | 2023826-2022-04722 | 2022-09-28 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT AN IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) AND THE PATIENT EXPERIENCED VAULT ISSUES. IT WAS REPORTED THAT THE VAULT RESOLVED ON ITS OWN, LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). Manufacturer Narrative: CORRECTED DATA: B5:THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 -11.5/+1.0/012 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) AND THE PATIENT EXPERIENCED VAULT ISSUES. IT WAS REPORTED THAT THE VAULT RESOLVED ON ITS OWN; LENS REMAINS IMPLANTED. REPORTER STATED CAUSE OF EVENT WAS OTHER STATING UNFORTUNATELY WE OPEN THE BOX OF THE LENSES. PROBLEM RESOLVED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. WORK ORDER SEARCH NOT REQUIRED. NO PATIENT INJURY REPORTED. CLAIM (B)(4). | -11.5/+1.0/012 | OD | Other Health Care Professional | 11/23/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16034295&pc=QCB | 2023826-2022-04642 | 2022-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -2.5/2.5/079 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -2.5/2.5/079 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. LENS ROTATION NOT ASSOCIATED TO A LOW VAULT WAS OBSERVED. ON (B)(6) 2023 THE LENS WAS REPOSITIONED. THIS RESOLVED THE PROBLEM. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -2.5/2.5/079 | OD | Physician | 11/25/2022 | 2022 | 45 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16035421&pc=QCB | 2023826-2022-04547 | 2022-10-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2, -2.0/1.0/096 DIOPTER, IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULT. A LONGER LENGTH LENS WAS IMPLANTED IN A SECOND SURGERY ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK. (B)(4). CLAIM# (B)(4). | -2.0/1.0/096 | OS | Physician | 11/28/2022 | 2022 | 44 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16035424&pc=QCB | 2023826-2022-04549 | 2022-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.0/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. IN A SECOND SURGERY ON (B)(6) 2022 A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK. (B)(4). CLAIM# (B)(4). | -15.0/2.5/091 | OD | Physician | 11/28/2022 | 2022 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16035430&pc=QCB | 2023826-2022-04551 | 2022-07-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.0/1.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED VERTICALLY WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/1.5/083 | OS | Physician | 11/28/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16035446&pc=QCB | 2023826-2022-04558 | 2022-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/2.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -11.0/2.0/087 | OS | Physician | 11/28/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16035455&pc=QCB | 2023826-2022-04557 | 2022-10-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.5/2.0/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -12.5/2.0/083 | OD | Physician | 11/28/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16035459&pc=QCB | 2023826-2022-04565 | 2021-11-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -12.0/1.5/096 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -12.0/1.5/096 | OS | Physician | 11/29/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16035463&pc=QCB | 2023826-2022-04568 | 2022-10-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -8.5/2.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.5/2.5/089 | OD | Physician | 11/29/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16035467&pc=QCB | 2023826-2022-04564 | 2022-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -7.5/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. (B)(4). | -7.5/2.0/090 | OS | Physician | 11/29/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16035471&pc=QCB | 2023826-2022-04575 | 2022-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.0/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -13.0/1.0/090 | OS | Physician | 12/01/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16035472&pc=QCB | 2023826-2022-04562 | 2022-09-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.0/1.0/106 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. IN A SECOND SURGERY ON (B)(6) 2022 A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY- UNK; RACE - UNK. (B)(4). | -9.0/1.0/106 | OD | Physician | 11/29/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16035475&pc=QCB | 2023826-2022-04566 | 2022-09-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -9.5/1.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED VERTICALLY WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. CLAIM# (B)(4). | -9.5/1.5/094 | OD | Physician | 11/29/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16035489&pc=QCB | 2023826-2022-04573 | 2022-07-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -14.0/4.5/082 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED VERTICALLY WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -14.0/4.5/082 | OD | Physician | 12/01/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16035498&pc=QCB | 2023826-2022-04572 | 2022-08-03 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED AND REMOVED A 12.6MM VTICMO 12.6, -13.5/1.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS FROM THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 DUE TO LOW VAULT . A LONGER LENGTH LENS WAS IMPLANTED ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.5/1.0/095 | OD | Physician | 12/01/2022 | 2022 | 17 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16035501&pc=QCB | 2023826-2022-04576 | 2022-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.0/2.0/101 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS IN A SECOND SURGERY ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. (B)(4). | -12.0/2.0/101 | OS | Physician | 12/01/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16035513&pc=QCB | 2023826-2022-04571 | 2022-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.5/1.5/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -11.5/1.5/086 | OS | Physician | 12/01/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16035524&pc=QCB | 2023826-2022-04579 | 2022-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/2.0/088 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -10.5/2.0/088 | OS | Physician | 12/02/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16035654&pc=QCB | 2023826-2022-04550 | 2022-11-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. IN A SECOND SURGERY ON (B)(6) 2022 A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -11.0/1.0/090 | OD | Physician | 11/28/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16035670&pc=QCB | 2023826-2022-04563 | 2022-09-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.0/1.0/070 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. IN A SECOND SURGERY ON (B)(6) 2022 A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY- UNK; RACE - UNK. CLAIM# (B)(4). | -9.0/1.0/070 | OS | Physician | 11/28/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16035679&pc=QCB | 2023826-2022-04574 | 2022-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -6.5/1.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTION: D4: UNIQUE IDENTIFIER (UDI): REPORTED AS N/A - CORRECT TO (B)(4). H4: REPORTED AS 31-DEC-2004, CORRECT TO 12-JAN-2022, CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -6.5/1.0/091 | OD | Physician | 12/01/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16039967&pc=QCB | 2023826-2022-04700 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -13.5/5.5/069 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, PRODUCT NON-COMFORMITY (NO PATIENT INVOLVEMENT), AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS THE DEVICE. REPORTEDLY, "ICL IS TOO BIG FOR THE ANTERIOR SEGMENT SIZE." ALSO, THE STATUS OF THE EYE WAS REPORTED AS, "PATIENT CURRENTLY ASYMPTOMATIC, NO PROBLEM REPORTED." Manufacturer Narrative: OFF LABEL USE- ACD <3.0. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: G2: (B)(6). CLAIM# (B)(4). | -13.5/5.5/069 | OS | Physician | 11/23/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16039977&pc=QCB | 2023826-2022-04699 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -12.5/5.5/82 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT, PRODUCT NON-COMFORMITY (NO PATIENT INVOLVEMENT), AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS THE DEVICE. REPORTEDLY, "ICL IS TOO BIG FOR THE ANTERIOR SEGMENT SIZE." ALSO, THE STATUS OF THE EYE WAS REPORTED AS, "PATIENT CURRENTLY ASYMPTOMATIC, NO PROBLEM REPORTED." Manufacturer Narrative: CORRECTED DATA: G2: COSTA RICA. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -12.5/5.5/82 | OD | Physician | 11/23/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16040282&pc=QCB | 2023826-2022-04710 | 2022-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.0/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LENS TWO SIZES LONGER IN LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -11.5/+1.0/094 | OS | Physician | 11/29/2022 | 2022 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16040283&pc=QCB | 2023826-2022-04712 | 2022-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+2.0/067 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -9.0/+2.0/067 | OS | Physician | 12/01/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16040284&pc=QCB | 2023826-2022-04714 | 2022-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+4.0/095 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: (B)(4). | -8.0/+4.0/095 | OS | Physician | 12/01/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16040285&pc=QCB | 2023826-2022-04713 | 2022-08-02 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.0/+1.0/091 | OD | Physician | 12/01/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16040286&pc=QCB | 2023826-2022-04711 | 2022-11-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.0/091 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.0/+2.0/091 | OS | Physician | 11/30/2022 | 2022 | 49 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16040287&pc=QCB | 2023826-2022-04716 | 2022-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.50/+2.5/083 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: (B)(4). | -12.50/+2.5/083 | OS | Physician | 12/01/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16040289&pc=QCB | 2023826-2022-04719 | 2022-10-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -14.5/+3.5/084 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -14.5/+3.5/084 | OS | Physician | 12/02/2022 | 2022 | 43 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16040290&pc=QCB | 2023826-2022-04721 | 2021-08-13 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+1.5/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. THE LENS WAS EXCHANGED ON (B)(6) 2021 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -13.0/+1.5/091 | OD | Physician | 12/02/2022 | 2022 | 18 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16040291&pc=QCB | 2023826-2022-04720 | 2022-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.0/102 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+1.0/102 | OD | Physician | 12/02/2022 | 2022 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16040539&pc=QCB | 2023826-2022-04715 | 2022-08-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-22 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.50/+2.5/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -12.50/+2.5/090 | OD | Physician | 12/01/2022 | 2022 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16041099&pc=QCB | 2023826-2022-04704 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -13.5/1.5/71 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2018. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). | -13.5/1.5/71 | OS | Physician | 11/25/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16041100&pc=QCB | 2023826-2022-04703 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -13.0/1.0/95 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2018. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION AND LENS OPACITY (CATARACT). THE PROBLEM WAS RESOLVED. Manufacturer Narrative: (B)(4). | -13.0/1.0/95 | OD | Physician | 11/25/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16045697&pc=QCB | 2023826-2022-04750 | 2022-11-19 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2; -10.00/1.50/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS THAT WAS INTENDED FOR THE PATIENT'S LEFT EYE (OS); HAD TORE DURING LOADING AND NOTED "LENS TEAR DUE TO SNAGGING". THIS OCCURRED ON (B)(6) 2022. ON THIS SAME DATE A REPLACEMENT LENS OF THE SAME MODEL/LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN AND NOTED "LENS TEAR DUE TO SNAGGING DURING INJECTOR LOADING". Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -10.00/1.50/092 | OS | Physician | 12/23/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16048686&pc=QCB | 2023826-2022-04680 | 2022-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.00/1.0/095 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS WAS TOO SMALL. IT WAS THE SURGEON'S PREFERENCE TO CHANGE THE LENS. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE IS GOOD. Manufacturer Narrative: CORRECTED DATA: D4: UNIQUE IDENTIFIER NUMBER: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -09.00/1.0/095 | OD | Physician | 11/29/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16048687&pc=QCB | 2023826-2022-04687 | 2022-11-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -11.50/+3.5/072 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPLACED WITH A LONGER LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.50/+3.5/072 | OS | Physician | 11/27/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16048696&pc=QCB | 2023826-2022-04681 | 2022-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -10.50/1.5/074 (SPHERE/CYLINDER/AXIS), INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS WAS TOO SMALL. IT WAS THE SURGEON'S PREFERENCE TO CHANGE THE LENS. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE STATUS OF THE EYE IS GOOD. Manufacturer Narrative: CORRECTED DATA: D4: UNIQUE IDENTIFIER NUMBER: (B)(4). H6:1494 OFF-LABEL: HISTORY OF PROGRESSIVE MYOPIA. CLAIM# (B)(4). Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -10.50/1.5/074 | OS | Physician | 11/29/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16050925&pc=QCB | 2023826-2022-04728 | 2022-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.5/2.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK; ETHNICITY: UNK; RACE: UNK. (B)(4). | -11.5/2.0/089 | OD | Physician | 12/07/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16050926&pc=QCB | 2023826-2022-04727 | 2022-10-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.5/3.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -12.5/3.0/090 | OS | Physician | 12/07/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16050929&pc=QCB | 2023826-2022-04730 | 2022-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.5/2.0/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO SIGNIFICANT REDUCTION OF IRIDO-CORNEAL ANGLES AND EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. (B)(4). CLAIM# (B)(4). | -8.5/2.0/091 | OD | Physician | 12/09/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16050932&pc=QCB | 2023826-2022-04729 | 2022-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED VERTICALLY A 12.1MM VTICMO12.1, -10.0/2.0/164 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED HORIZONTALLY WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. (B)(4). | -10.0/2.0/164 | OD | Physician | 12/09/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16057393&pc=QCB | 2023826-2022-04751 | 2022-10-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -16.0/2.5/106 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE PATIENT EXPERIENCED LENS ROTATION; NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH/MODEL LENS BUT IMPLANTED VERTICALLY. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -16.0/2.5/106 | OD | Physician | 12/23/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16057394&pc=QCB | 2023826-2022-04752 | 2022-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -8.5/2.0/090 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE SURGEON REPORTS LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND IMPLANTED VERTICALLY. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE AND RESIDUE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -8.5/2.0/090 | OS | Physician | 12/23/2022 | 2022 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16057492&pc=QCB | 2023826-2022-04735 | 2022-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -2.5/3.5/105 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. REPORTEDLY, "NORMOTENSIVE AQUEDUS MISDIRECTION OU WITH MYOPIC SHIFT OU.ATROPINE TID OU X 1 MONTH DIAMOX 25OMG TID X 1 WEEK." STATUS OF THE EYE IS SAID, "CURRENTLY ASYMPTOMATIC OU," AND "PATIENT RESPONDED VIEW FAST TO THE TREATMENT W ATROPINE AND DIAMOX, CURRENTLY IS ASYMPTOMATIC W/OUT TREATMENT. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -2.5/3.5/105 | OS | Physician | 11/28/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16057493&pc=QCB | 2023826-2022-04736 | 2022-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -10.0/3.5/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME MODEL, SIZE LENS IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT. REPORTEDLY, "THE SECOND LENS WAS IMPLANTED VERTICALLY." THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -10.0/3.5/088 | OS | Physician | 11/28/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16057500&pc=QCB | 2023826-2022-04734 | 2022-06-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -1.5/3.5/070 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REPORTEDLY, "NORMOTENSIVE AQUEDUS MISDIRECTION OU WITH MYOPIC SHIFT OU.ATROPINE TID OU X 1 MONTH DIAMOX 25OMG TID X 1 WEEK." STATUS OF THE EYE IS SAID, "CURRENTLY ASYMPTOMATIC OU," AND "PATIENT RESPONDED VIEW FAST TO THE TREATMENT W ATROPINE AND DIAMOX, CURRENTLY IS ASYMPTOMATIC W/OUT TREATMENT. Manufacturer Narrative: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -1.5/3.5/070 | OD | Physician | 11/28/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16064702&pc=QCB | 2023826-2022-04688 | 2019-06-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -12.00/+3.0/092 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2017. THE LENS WAS EXPLANTED ON (B)(6) 2019 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A SAME LENGTH BUT DIFFERENT MODEL (VICMO) LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CORRECTED DATA: H6: 4629 SHOULD BE OMITTED AND 4627 SHOULD BE ADDED FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.00/+3.0/092 | OD | Physician | 11/28/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16064711&pc=QCB | 2023826-2022-04689 | 2022-06-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -08.00/+2.0/161 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | -08.00/+2.0/161 | OD | Physician | 11/29/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16064723&pc=QCB | 2023826-2022-04690 | 2021-06-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -05.00/+3.5/103 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2021. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT. THE LENS WAS REPOSITIONED ON (B)(6) 2021 BUT THE PROBLEM WAS NOT RESOLVED, SO THE LENS WAS REMOVED. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND THE HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -05.00/+3.5/103 | OS | Physician | 11/30/2022 | 2022 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16064758&pc=QCB | 2023826-2022-04682 | 2022-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -09.50/+3.0/048 (SPHERE/CYLINDER/AXIS) IN THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6-HEALTH EFFECT- IMPACT CODE: "2199" SHOULD BE REMOVED AND CODE "4629" SHOULD BE ADDED. H6- MEDICAL DEVICE PROBLEM CODE: "2993" SHOULD BE ADDED. B5- THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICM5_ 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -09.50/+3.0/048 (SPHERE/CYLINDER/AXIS) IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE SAME MODEL, LENGTH AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE: "NO SIGNS". THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BENT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. FUNCTIONAL INSPECTION FOUND THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. ADDITONAL INFORMATION: H6- TYPE OF INVESTIGATION CODE: 3331- DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | -09.50/+3.0/048 | OS | Physician | 12/01/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16065464&pc=QCB | 2023826-2022-04738 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -11.0/2.0/074 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE. ON (B)(6) 2022, THE LENS WAS EXCHANGED DUE TO THE LENS BEING TOO SMALL AND LENS ROTATION. Manufacturer Narrative: CORRECTED DATA B5: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -11.0/1.5/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO THE LENS BEING TOO SMALL AND LENS ROTATION. REPORTEDLY, "IOP WAS GOOD." CLAIM# (B)(4). Manufacturer Narrative: B5: CORRECTED TO: THE REPORTER INDICATED THAT A 12.1MM VTICM512.1 IMPLANTABLE COLLAMER LENS OF -11.0/1.5/086 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXPLANTED AND REPLACED FOR A LONGER LENGTH LENS DUE TO THE LENS BEING TOO SMALL AND LENS ROTATION. REPORTEDLY, "IOP WAS GOOD." H6 - HEALTH EFFECT - IMPACT CODE: 4627 SHOULD BE ADDED TO INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4629 SHOULD BE REMOVED FROM THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -11.0/2.0/074 | NA | Physician | 11/29/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16065466&pc=QCB | 2023826-2022-04743 | 2022-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.5/2.0/072 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME MODEL, SAME SIZE LENS DUE TO LENS ROTATION NOT ASSOCIATED WITH LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: B5 - CORRECTED TO : "THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -12.5/2.0/072 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME MODEL, SAME SIZE LENS DUE TO LENS ROTATION NOT ASSOCIATED WITH LOW VAULT. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN." IN THE INTIAL MDR. D6A - CORRECTED TO: (B)(6) 2021 IN THE INITAL MDR. CLAIM # (B)(4). Manufacturer Narrative: D6A: CORRECTED TO: ON (B)(6) 2021 IN THE INITAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -12.5/2.0/072 | OS | Physician | 11/29/2022 | 2022 | 41 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16065468&pc=QCB | 2023826-2022-04742 | 2022-01-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-28 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -11.5/1.5/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2021. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME MODEL, SAME SIZE LENS DUE TO LENS ROTATION NOT ASSOCIATED WITH LOW VAULT. REPORTEDLY, "THE FIRST LENS WAS IMPLANTED AT CCQ 23." THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: D4 - UNIQUE IDENTIFER (UDI)# (B)(4). SHOULD BE ADDED TO THE INITAL MDR. CLAIM # (B)(4). Manufacturer Narrative: D4: CORRECTED TO: SERIAL#: (B)(6) IN THE INITIAL MDR. D4: CORRECTED TO: UNIQUE IDENTIFIER (UDI)#: N/A. CLAIM#: (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND OPTIC DEFORMED AND HAPTIC BENT/DEFORMED TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -11.5/1.5/092 | OD | Physician | 11/29/2022 | 2022 | 41 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16075211&pc=QCB | 2023826-2022-04765 | 2022-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -9.5/0.5/001 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. IT WAS REPORTED THAT THE EXCHANGE LENS DID NOT RESOLVE THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9-RETURN DATE: 12-JAN-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BROKEN. DIMENSIONAL INSPECTION WAS UNABLE TO BE PERFORMED DUE TO THE LENS RETURNED IS NOT THE MODEL STATED IN THE CLAIM. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -9.5/0.5/001 | OD | Physician | 11/29/2022 | 2022 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16075336&pc=QCB | 2023826-2022-04766 |  | Malfunction | STAAR SURGICAL COMPANY | 2022-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS -9.5/0.5/180 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 AS A REPLACEMENT LENS TO ADDRESS LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. LENS REMAINS IMPLANTED. THE REPORTER STATED " AFTER REPLACING THE LENS, THE PATIENT STILL HAD LOW VAULT BUT GOOD VISION, AND THE DOCTOR DECIDED TO KEEP THE PATIENT UNDER OBSERVATION." SEE MFR # 2023826-2022-04765 FOR CLAIM AGAINST THE INITIAL LENS. Manufacturer Narrative: (B)(4). | -9.5/0.5/180 | OD | Physician | 11/29/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16075538&pc=QCB | 2023826-2022-04767 |  | Injury | STAAR SURGICAL COMPANY | 2022-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF -11.5/3.5/004 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT AND SIGNIFICANT REDUCTION OF IRIDOCORNEAL ANGLES. THIS EXCHANGED RESOLVED THE PROBLEM. Manufacturer Narrative: CLAIM# (B)(4). | -11.5/3.5/004 | OS | Physician | 11/29/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16075649&pc=QCB | 2023826-2022-04768 | 2022-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -7.0/2.0/087 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -7.0/2.0/087 | OS | Physician | 11/29/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16075700&pc=QCB | 2023826-2022-04769 | 2022-08-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-29 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -8.5/1.5/109 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -8.5/1.5/109 | OD | Physician | 11/29/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16077730&pc=QCB | 2023826-2022-04793 | 2022-10-10 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON (B)(6) 2022 A 12.6MM, VTICMO12.6 -13.5/4.0/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED UPSIDE DOWN INTO THE PATIENTS RIGHT EYE (OD). THE LENS WAS IMPLANTED, REMOVED AND REPLACED INTRAOPERATIVELY, WITH NO PATIENT INJURY. CAUSE OF EVENT WAS UNKNOWN. PROBLEM RESOLVED. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | -13.5/4.0/087 | OD | Physician | 12/02/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16078383&pc=QCB | 2023826-2022-04758 | 2022-12-05 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -15.00/+2.0/094, TORE DURING DELIVERY INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT CONTACT. THE LENS WAS DISCARDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: B5: THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS ON (B)(6) 2022 AND THE PROBLEM WAS RESOLVED. THE STATUS IS OK. CAUSE OF THE EVENT WAS UNKNOWN. CLAIM # (B)(4). Manufacturer Narrative: D4: UNIQUE IDENTIFIER (UDI)#: (B)(4) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | -15.00/+2.0/094 | OS | Physician | 12/08/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16079371&pc=QCB | 2023826-2022-04737 | 2022-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -5.0/1.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6)2022, THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. REPORTEDLY, "SECOND LENS ELECTIVE SURGERY." Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 23-MAR-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT WITH RESIDUE/DEBRIS ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: H3-SHOULD BE CORRECTED TO- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT WITH RESIDUE/DEBRIS ON THE LENS. DIMENSIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4). | -5.0/1.0/092 | OD | Physician | 12/02/2022 | 2022 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16080960&pc=QCB | 2023826-2022-04658 | 2022-08-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2022-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF -12.5/1.0/94 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON THE SAME DAY DIFFERENT SURGERY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE REPLACEMENT LENS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - DEVICE CODE: 1494 - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE LESS THAN THAT OF 21 YEARS SHOULD BE ADDED TO THE INITIAL MDR. CLAIM #(B)(4). Manufacturer Narrative: (B)(4). | -12.5/1.0/94 | OS | Physician | 12/01/2022 | 2022 | 17 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16080981&pc=QCB | 2023826-2022-04660 | 2022-12-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2022-12-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -12.00/1.0/037 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED. THE LENS REMAINS IMPLANTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: "REPORTEDLY, THE PATIENT DOES NOT WANT AN EXCHANGE AND WILL CONTINUE TO BE CHECKED. THE PROBLEM IS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. CLAIM#: (B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM # (B)(4). | -12.00/1.0/037 | OS | Physician | 12/02/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16085502&pc=QCB | 2023826-2023-00003 | 2022-10-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -17.5/4.0/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. CLAIM # (B)(4). | -17.5/4.0/103 | OS | Physician | 12/12/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16085503&pc=QCB | 2023826-2023-00001 | 2022-07-31 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.0/1.0/098 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -12.0/1.0/098 | OD | Physician | 12/12/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16085504&pc=QCB | 2023826-2023-00002 | 2022-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.5/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. ON (B)(6) 2022 A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.5/1.5/092 | OD | Physician | 12/12/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16085507&pc=QCB | 2023826-2023-00009 | 2022-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 137MM VTICMO13.7, -10.5/2.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. (B)(4). | -10.5/2.0/092 | OS | Physician | 12/12/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16085511&pc=QCB | 2023826-2023-00005 | 2022-03-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM512.1, -8.0/2.5/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT WITH ROTATION ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT - UNK; ETHNICITY - UNK; RACE - UNK. CLAIM# (B)(4). | -8.0/2.5/087 | OS | Physician | 12/13/2022 | 2022 | 26 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16085512&pc=QCB | 2023826-2023-00004 | 2022-11-17 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6 VTICM5_12.6, -10.5/3.0/103 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/3.0/103 | OS | Physician | 12/12/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16085516&pc=QCB | 2023826-2023-00010 | 2022-12-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1, -7.5/2.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT, ETHNICITY AND RACE - UNK. (B)(4). | -7.5/2.5/090 | OS | Physician | 12/13/2022 | 2022 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16087494&pc=QCB | 2023826-2022-04796 | 2022-10-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -13.0/3.5/091 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXPLANTED WITH NO PLANS TO REPLACE THE LENS. PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BROKEN/BENT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -13.0/3.5/091 | OS | Physician | 12/29/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16087504&pc=QCB | 2023826-2022-04795 | 2022-10-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-03 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.5/3.0/094 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. THE PATIENT EXPERIENCED LENS ROTATION NOT ASSOCIATED TO LOW VAULTING. ON (B)(6) 2022 THE LENS WAS EXPLANTED WITH NO PLANS TO REPLACE THE LENS. PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY WITH RESIDUE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#(B)(4). | -11.5/3.0/094 | OD | Physician | 12/29/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16098680&pc=QCB | 2023826-2023-00076 | 2022-11-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF -8.0/4.0/94 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS, THE REPORTER STATED " LENS FOR OS WAS IMPLANTED BY MISTAKE". IT WAS REPORTED THAT THE PROBLEM RESOLVED AFTER EXCHANGE. Manufacturer Narrative: CLAIM# (B)(4). | -8.0/4.0/94 | OD | Physician | 12/05/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16099315&pc=QCB | 2023826-2023-00062 | 2022-05-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -5.5/+1.5/079 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -5.5/+1.5/079 | OS | Physician | 12/07/2022 | 2022 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16099318&pc=QCB | 2023826-2023-00060 | 2022-11-29 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+3.0/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. WITHIN THE INITIAL SURGERY THE LENS WAS REMOVED AND REPLACED INTRA-OPERATIVELY FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.0/+3.0/092 | OD | Physician | 12/07/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16100023&pc=QCB | 2023826-2023-00065 | 2022-10-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-04 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.50/+1.0/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CORRECTED DATA: H6-DEVICE CODE 1494: OFF-LABEL USE (UNDER 21YRS OF AGE AT DATE OF IMPLANT) SHOULD HAVE BEEN INCLUDED IN THE INITIAL MDR . CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -11.50/+1.0/095 | OD | Physician | 12/12/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16109354&pc=QCB | 2023826-2023-00088 | 2022-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2; -11.50/2.5/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS HAD CRINKLED UP DURING INSERTION. THE LENS TOUCHED THE PATIENT'S EYE WITH NO INJURY. Manufacturer Narrative: TYPE OF INVESTIGATION: (B)(4): NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: B1: CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2: CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. B5: "THE LENS WAS EXPLANTED AND RETURNED FOR CREDIT." SHOULD BE ADDED TO THE INITIAL MDR. H1: CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6: CORRECTED TO HEALTH EFFECT - IMPACT CODE: 4629. CLAIM#: (B)(4). | -11.50/2.5/088 | NA | Physician | 12/14/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16109392&pc=QCB | 2023826-2023-00084 | 2022-12-28 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -0.50/+6.0/081 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE. INTRAOPERATIVELY THE LENS WAS REMOVED WITH NO PATIENT INJURY. ON (B)(6) 2022 A REPLACEMENT LENS OF THE SAME LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -0.50/+6.0/081 | OS | Physician | 01/03/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16110233&pc=QCB | 2023826-2023-00047 |  | Injury | STAAR SURGICAL COMPANY | 2023-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -9.5/1.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE. LOW VAULT WAS OBSERVED. Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4). Manufacturer Narrative: CORRECTED DATA: B5 A 13.2MM TMICL13.2 IMPLANTABLE COLLAMER LENS OF A -10.0/1.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE. H3: DEVICE EVALUATION: PRODUCT WAS RETURNED WITH LIQUID IN A VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. H6: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM#: (B)(4). Manufacturer Narrative: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.00/1.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE. THE PATIENT EXPERIENCED A LOW VAULT. THE LENS WAS EXPLANTED. H6- REMOVE "HEALTH EFFECT - IMPACT CODE: 2199" CLAIM\# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM TMICL 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.00/1.0/089 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. THE LENS WAS EXPLANTED. CLAIM# (B)(4). | -10.00/1.0/089 | OS | Physician | 12/06/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16110234&pc=QCB | 2023826-2023-00050 | 2022-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -9.5/2.0/95 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON 14-JUN-2022. ON 06-DEC-2022, THE LENS WAS EXCHANGED FOR A SAME SIZE LENS IMPLANTED VERTICALLY AND THE PROBLEM WAS RESOLVED. REPORTEDLY, "CHANGED FROM EAR-SIDE FIXED TO VERTICAL FIXED." CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT SHOULD BE ADDED TO THE INITAL MDR. CLAIM # (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -9.5/2.0/95 | OS | Physician | 12/06/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16110237&pc=QCB | 2023826-2023-00049 | 2022-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -8.5/1.0/85 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME SIZE LENS IMPLANTED VERTICALLY AND THE PROBLEM WAS RESOLVED. REPORTEDLY, "CHANGED FROM EAR-SIDE FIXED TO VERTICAL FIXED." CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: H6 - TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: PATIENT INFORMATION: UNK. CLAIM# (B)(4). | -8.5/1.0/85 | OD | Physician | 12/06/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16110238&pc=QCB | 2023826-2023-00048 | 2022-11-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-05 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -5.0/1.0/054 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED AND REPLACED DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(6). (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 12-JAN-2023. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN AND RESIDUE/DEBRIS ON THE PRODUCT. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -5.0/1.0/054 | OD | Physician | 12/06/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16113230&pc=QCB | 2023826-2023-00080 | 2022-11-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.7MM VTICM5_13.7 IMPLANTABLE COLLAMER LENS OF -12.5/1.5/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULT WAS OBSERVED, LENS REMAINS IMPLANTED . ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5: PER UPDATED INFORMATION THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS ON (B)(6)2023 AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -12.5/1.5/90 | OS | Physician | 12/08/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16114299&pc=QCB | 2023826-2023-00101 | 2022-12-26 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -11.50/+1.0/095 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. IT WAS REPORTED THAT THE LENS WAS SURPRISINGLY THICK. IT WAS NOTED THE "VISUAL ACUITY AND VAULT IS GOOD; BUT THE SURGEON IS WONDERING WHY THIS TICL LENS IS SURPRISINGLY THICK; COULD IT BE SOMETHING WRONG?". PROBLEM WAS NOT RESOLVED. CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE AND NOTED "CAUSED LOADING AND MANIPULATING DIFFICULTIES". LENS REMAINS IMPLANTED. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 TYPE OF INVESTIGATION 3331 DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. (B)(4). Manufacturer Narrative: (B)(4). | -11.50/+1.0/095 | OD | Physician | 01/02/2023 | 2023 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16116774&pc=QCB | 2023826-2023-00051 | 2022-09-23 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -7.0/2.0/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS INTRAOPERATIVELY IMPLANTED, REMOVED, AND REPLACED FOR A LONGER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -7.0/2.0/089 | OS | Physician | 12/07/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16116793&pc=QCB | 2023826-2023-00052 |  | Injury | STAAR SURGICAL COMPANY | 2023-01-06 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -9.5/1.5/147 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE ON (B)(6) 2022. LENS ROTATION WAS OBSERVED AND THE LENS WAS EXCHANGED FOR A SPHERICAL LENS. Manufacturer Narrative: WEIGHT, ETHNICITY, RACE:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B5- "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/1.5/147 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE ON (B)(6) 2022. THE PATIENT EXPERIENCED LENS ROTATION. REPORTEDLY " LENS ROTATION (TWICE)". ON (B)(6) 2022 THE LENS WAS EXPLANTED. THE LENS WAS REPLACED WITH A SPHERICAL LENS. CLAIM# (B)(4). | -9.5/1.5/147 | NA | Physician | 12/09/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16123141&pc=QCB | 2023826-2023-00016 | 2022-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.0/2.5/113 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM.CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -12.0/2.5/113 | OD | Physician | 12/20/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16123160&pc=QCB | 2023826-2023-00027 | 2022-11-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2, -5.0/1.0/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM (B)(4). | -5.0/1.0/085 | OD | Physician | 12/23/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16124552&pc=QCB | 2023826-2023-00021 | 2022-08-27 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6, -5.0/2.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S RIGHT EYE (OD) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -5.0/2.0/094 | OD | Physician | 08/27/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16124776&pc=QCB | 2023826-2023-00017 | 2022-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.5/3.0/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.5/3.0/074 | OS | Physician | 12/20/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16124957&pc=QCB | 2023826-2023-00028 | 2022-11-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2, -04.00/1.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT.THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PT INFO: UNK. CLAIM # (B)(4). | -04.00/1.0/094 | OS | Physician | 12/23/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16126617&pc=QCB | 2023826-2023-00015 | 2022-08-11 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2, -10.0/3.0/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. CLAIM# (B)(4). | -10.0/3.0/092 | OS | Physician | 12/19/2022 | 2022 | 35 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16127556&pc=QCB | 2023826-2023-00117 | 2022-08-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -11.50/2.5/087 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. THE SURGEON REPORTS THE LENS IS TOO SMALL AND HAS ROTATED. LENS REMAINS IMPLANTED. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED IN A MICROCENTRIFUGE VIAL WITH LIQUID. VISUAL INSPECTION FOUND OPTIC/HAPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. ADDIITONAL INFORMATION: B5- THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.50/2.5/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS). THE PATIENT EXPERIENCED LOW VAUT WITH ROTATION. REPORTEDLY " LENS IS TOO SMALL AND HAS ROTATED, PLANNED EXCHANGE WITH 13.2MM LENS LENGTH." THE LENS WAS EXPLANTED. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -11.50/2.5/087 | OS | Physician | 12/16/2022 | 2022 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16127574&pc=QCB | 2023826-2023-00025 | 2022-11-05 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-09 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6, -5.50/1.0/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED, REMOVED, AND REPLACED INTRAOPERATIVELY WITH A SHORTER LENGTH LENS ON (B)(6) 2022 FROM PATIENT'S LEFT EYE (OS) DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -5.50/1.0/094 | OS | Physician | 12/23/2022 | 2022 | 34 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16129918&pc=QCB | 2023826-2023-00090 | 2022-12-15 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.50/+1.0/086 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS WAS INSERTED UPSIDE DOWN WHEN INJECTED INTO THE PATIENTS LEFT EYE (OS). THE LENS WAS TORN UPON REMOVAL AND WAS DISCARDED. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS AND THERE WAS NO PATIENT HARM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.50/+1.0/086 | OS | Physician | 12/15/2022 | 2022 | 31
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16129936&pc=QCB | 2023826-2023-00097 | 2022-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00/+1.0/052 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH A SHORTER LENS AND THE PROBLEM WAS RESOLVED. THE PATIENT IS HAPPY. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). | -09.00/+1.0/052 | OS | Physician | 12/15/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16129938&pc=QCB | 2023826-2023-00096 | 2022-11-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -06.50/+3.0/001 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6)2022. THE LENS WAS REPORTED AS LOW VAULTING AND LENS ROTATION. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION: B1: ADVERSE EVENT. B2: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. B5: THE LENS WAS REPLACED ON (B)(6) 2023 WITH A LONGER LENGTH LENS. D4: UNIQUE IDENTIFIER NUMBER: (B)(4). D6B: ON (B)(6) 2023. H1: SERIOUS INJURY. CLAIM#: (B)(4). Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -06.50/+3.0/001 | OS | Physician | 12/14/2022 | 2022 | 42 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16129942&pc=QCB | 2023826-2023-00091 | 2022-12-15 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -12.50/+1.0/087 (SPHERE/CYLINDER/AXIS), INTO PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE HAPTIC WAS STUCK BY THE FOAM TIP PLUNGER TIP UPON INSERTION AND A SMALL PIECE WAS RIPPED. THE LENS REMAINED IMPLANTED WITH NO PATIENT HARM. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: (B)(4). | -12.50/+1.0/087 | OS | Physician | 12/15/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16129951&pc=QCB | 2023826-2023-00095 | 2022-11-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -08.00/+3.0/004 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPORTED AS LOW VAULTING AND LENS ROTATION. THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS DUE TO THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: ADDITIONAL INFORMATION: B1: ADVERSE EVENT. B2: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. B5: THE LENS WAS REPLACED ON (B)(6) 2023 WITH A LONGER LENGTH LENS. D4: UNIQUE IDENTIFIER NUMBER: (B)(4). D6B: (B)(6) 2023. H1: SERIOUS INJURY. CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -08.00/+3.0/004 | OD | Physician | 12/14/2022 | 2022 | 42 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16129968&pc=QCB | 2023826-2023-00093 | 2022-09-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -03.50/+1.5/071 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REMOVED ON (B)(6) 2022 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH A DIFFERENT MODEL BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CORRECTED DATA: D4- UNIQUE IDENTIFIER (UDI) #: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -03.50/+1.5/071 | OS | Physician | 12/12/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16134902&pc=QCB | 2023826-2023-00055 |  | Injury | STAAR SURGICAL COMPANY | 2023-01-10 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions |  | Event Description: THE REPORTER INDICATED THAT A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS WITH A -15.0/3.5/090 (SPHERE/CYCLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE. THE LENS WAS REMOVED AND REPLACED FOR A VTICMO LENS. REPORTEDLY, "THE FIRST LENS WAS KINKED AND AFFECTED THE OUTCOME OF PATIENT'S VISION SO WE HAD TO EXCHANGE THE LENS." Manufacturer Narrative: WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: D6A.- "(B)(6) 2022." B5- "THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL 12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -15.00/3.5/090 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S EYE IN (B)(6) 2022. REPORTEDLY "THE FIRST LENS WAS KINKED AND AFFECTED THE OUTCOME OF PATIENT'S VISION SO WE HAD TO EXCHANGE THE LENS". THE LENS WAS EXPLANTED AND REPLACED WITH A DIFFERENT MODEL/ SAME LENGTH AND THE PROBLEM WAS RESOLVED." CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3: DEVICE EVALUATION: THE LENS WAS RETURNED IN A VIAL WITH LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. CLAIM#: (B)(4). | -15.0/3.5/090 | NA | Physician | 12/12/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16140022&pc=QCB | 2023826-2023-00068 | 2022-11-30 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.5/090 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. IN A SEPARATE SURGERY THAT DAY THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.0/+1.5/090 | OD | Physician | 12/12/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16140025&pc=QCB | 2023826-2023-00070 | 2022-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.0/087 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -11.0/+2.0/087 | OS | Physician | 12/12/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16141732&pc=QCB | 2023826-2023-00067 | 2022-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-11 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.5/+0.5/085 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -4.5/+0.5/085 | OS | Physician | 12/12/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16148130&pc=QCB | 2023826-2023-00071 | 2022-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -6.0/+2.0/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -6.0/+2.0/093 | OD | Physician | 12/14/2022 | 2022 | 22 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16148131&pc=QCB | 2023826-2023-00073 | 2022-12-09 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.0/+2.5/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -13.0/+2.5/090 | OS | Physician | 12/15/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16148132&pc=QCB | 2023826-2023-00075 | 2021-08-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM TMICL12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+1.0/089 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2020. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+1.0/089 | OD | Physician | 12/15/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16148135&pc=QCB | 2023826-2023-00074 | 2022-11-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+1.5/094 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT INDICATE THE CAUSE OF THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -8.0/+1.5/094 | OS | Physician | 12/15/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16151004&pc=QCB | 2023826-2023-00121 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -4.00/+1/105 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 AS A REPLACEMENT LENS. IT WAS REPORTED THAT THE UCVA IS VERY GOOD; BUT VAULTING IS STILL VERY LOW. THE SURGEON HAS DECIDED TO KEEP THE LENS IN THE EYE AND MONITOR THE PATIENT. LENS REMAINS IMPLANTED. PROBLEM IS NOTED AS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.00/1.0/105 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 AS A REPLACEMENT LENS. REPORTEDLY "UCVA IS VERY GOOD - VAULTING STILL VERY LOW!" AND THAT "THE SURGEON DECIDED TO KEEP THE LENS IN THE EYE AND DO A CLOSE MONITORING". IT WAS LATER REPORTED THAT "THE LOW VAULT IS STABLE AND THE PATIENT IS HAPPY WITH THE VISUAL OUTCOME - SO NO EXCHANGE PLANNED". THE LENS REMAINS IMPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -4.00/+1/105 | OD | Physician | 01/09/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16151452&pc=QCB | 2023826-2023-00122 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -4.50/+1.00/75 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022 AS A REPLACEMENT LENS. IT WAS REPORTED THAT THE UCVA IS VERY GOOD; BUT VAULTING IS STILL VERY LOW. THE SURGEON HAS DECIDED TO KEEP THE LENS IN THE EYE AND MONITOR THE PATIENT. LENS REMAINS IMPLANTED. PROBLEM IS NOTED AS RESOLVED. Manufacturer Narrative: ADDITIONAL DATA: B5: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -4.50/1.0/075 (SPHERE/CYLINDER/AXIS) INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 AS A REPLACEMENT LENS. REPORTEDLY "UCVA IS VERY GOOD - VAULTING STILL VERY LOW!" AND THAT "THE SURGEON DECIDED TO KEEP THE LENS IN THE EYE AND DO A CLOSE MONITORING". IT WAS LATER REPORTED THAT "THE LOW VAULT IS STABLE AND THE PATIENT IS HAPPY WITH THE VISUAL OUTCOME - SO NO EXCHANGE PLANNED". THE LENS REMAINS IMPLANTED. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -4.50/+1.00/75 | OS | Physician | 01/09/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16151777&pc=QCB | 2023826-2023-00141 | 2022-07-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER STATED THAT A 12.6MM VTICM5_12.6 -14.00/3.5/092 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2020. ON (B)(6) 2020 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON 20-OCT-2022 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS AND IT WAS REPORTED THAT LRI WAS PERFORMED FOR ASTIGMATISM, PROBLEM RESOLVED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: ONE SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). Manufacturer Narrative: CORRECTED DATA: H6-1494: OFF-LABEL: AGE<21 YEARS OLD. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9:19-DEC-2022. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT AND RESIDUE ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -14.00/3.5/092 | OD | Physician | 12/13/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16152652&pc=QCB | 2023826-2023-00142 | 2022-09-02 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER STATED THAT A 12.1MM VTICM5_12.1 -14.00/1.5/69 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS REPOSITIONED DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT BUT IT DID NOT RESOLVE THE PROBLEM. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS AND IT WAS REPORTED THAT LRI WAS PERFORMED FOR ASTIGMATISM, PROBLEM RESOLVED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: 19-DEC-2022. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE AND RESIDUE ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM#: (B)(4). | -14.00/1.5/69 | OS | Physician | 12/13/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16152962&pc=QCB | 2023826-2023-00145 | 2022-07-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER STATED THAT A 12.6MM VTICM5_12.6 -12.00/1.00/91 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | -12.00/1.00/91 | OS | Physician | 12/13/2022 | 2022 | 23 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16153270&pc=QCB | 2023826-2023-00144 | 2022-11-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER STATED THAT A 12.1MM VTICM5_12.1 -7.00/1.00/97 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A LONGER LENGTH LENS DUE TO LENS ROTATION NOT ASSOCIATED TO A LOW VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -7.00/1.00/97 | OS | Physician | 12/13/2022 | 2022 | 40 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16153508&pc=QCB | 2023826-2023-00171 | 2023-01-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -10.5/1.5/097 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2023. THE SURGEON REPORTS THE LENS WAS "SURPRISINGLY STIFF; WHICH CAUSED LOADING AND MANIPULATING DIFFICULTIES". THE PROBLEM IS REPORTED AS NOT RESOLVED. LENS REMAINS IMPLANTED. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). | -10.5/1.5/097 | OD | Physician | 01/11/2023 | 2023 | 48 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16153566&pc=QCB | 2023826-2023-00172 | 2023-01-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -11.0/2.5/149 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS THE LENS WAS "SURPRISINGLY STIFF; WHICH CAUSED LOADING AND MANIPULATING DIFFICULTIES". THE PROBLEM IS REPORTED AS NOT RESOLVED. LENS REMAINS IMPLANTED. CAUSE OF THE EVENT WAS REPORTED AS THE DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION: H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM#: (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. MEDICAL DEVICE PROBLEM CODE - 1494: OFF-LABEL USE (OVER 45YRS OF AGE AT DATE OF IMPLANT). (B)(4). | -11.0/2.5/149 | OS | Physician | 01/11/2023 | 2023 | 48 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16153693&pc=QCB | 2023826-2023-00108 | 2022-12-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Improper or Incorrect Procedure or Method; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.5/075 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022. REPORTEDLY, "WRONG SIGN IN THE CYLINDER IN THE CALCULATION." CAUSE OF THE EVENT WAS USER ERROR. LENS REMAINS IMPLANTED. Manufacturer Narrative: CORRECTED DATA: B1: SHOULD BE CORRECTED TO MALFUNCTION. B2: SHOULD BE OMITTED FOR CORRECTION. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: THE PATIENT DOES NOT WANT A LENS EXCHANGE. A LASER ENHANCEMENT IS PLANNED FOR THE SUMMER. CLAIM# 727633. Manufacturer Narrative: B1: SHOULD BE CORRECTED TO ADVERSE EVENT. B2: SHOULD BE CORRECTED TO REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. B5: SHOULD BE CORRECTED TO: THE REPORTER INDICATED THAT A 12.6MM VTICM512.6 IMPLANTABLE COLLAMER LENS OF -9.0/1.5/075 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED IN THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. REPORTEDLY, "WRONG SIGN IN THE CYLINDER IN THE CALCULATION." CAUSE OF THE EVENT WAS USER ERROR. ON (B)(6) 2022 THE LENS WAS REPOSITIONED AND THIS DID NOT RESOLVE THE PROBLEM. REPORTEDLY, "THE PATIENT DOES NOT WANT A LENS EXCHANGE. A LASER ENHANCEMENT IS PLANNED FOR THE SUMMER". LENS REMAINS IMPLANTED. H1: SHOULD BE CORRECTED TO SERIOUS INJURY. (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -9.0/1.5/075 | OD | Physician | 12/16/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16155173&pc=QCB | 2023826-2023-00135 | 2022-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-12 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -9.5/3.5/006 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. CAUSE REPORTED AS UNKNOWN Manufacturer Narrative: CLAIM#: (B)(4). | -9.5/3.5/006 | OS | Physician | 12/14/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16155878&pc=QCB | 2023826-2023-00149 | 2022-12-14 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -5.5/1.50/104 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULT, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: B5 - REPORTEDLY, PROBLEM WAS RESOLVED. STATUS OF EYE IS "ALL FINE! CURRENTLY NO TREATMENT REQUIRED; SURGEON PLANS A CLOSED MONITORING. PATIENT IS HAPPY." CLAIM# (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: ZM0009705220. TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -5.5/1.50/104 | OD | Physician | 12/16/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16155879&pc=QCB | 2023826-2023-00150 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -7.5/0.50/99 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022 AS A REPLACEMENT LENS TO ADDRESS EXCESSIVE VAULT AND REFRACTIVE SURPRISE. IT WAS REPORTED THAT THE LENS DID NOT REDUCE THE VAULTING, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: H6 - TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -7.5/0.50/99 | OS | Physician | 12/16/2022 | 2022 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16155880&pc=QCB | 2023826-2023-00147 | 2022-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM, VTICM5_13.2, -12.5/4.0/73 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022 AND THE PATIENT EXPERIENCED EXCESSIVE VAULT, LENS REMAINS IMPLANTED. CAUSE OF EVENT WAS UNKNOWN. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5- "ON (B)(6) 2023 THE LENS WAS EXCHANGED FOR THE SAME MODEL/SHORTER LENGTH LENS AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE WAS "ALL FINE". ADDITIONAL INFORMATION PROVIDED WAS "PATIENT IS HAPPY". CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -12.5/4.0/73 | OS | Physician | 12/14/2022 | 2022 | 42 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16155885&pc=QCB | 2023826-2023-00148 | 2022-12-08 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-13 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM, VTICM5_12.6, -7.0/-2.5/89(SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME LENGTH LENS THAT WAS IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT AND THE PROBLEM RESOLVED. CAUSE OF EVENT WAS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS CLAIM# (B)(4). Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENT REPORTED FOR UNITS WITHIN THE SAME LOT. (B)(4). | -7.0/-2.5/89 | OS | Physician | 12/14/2022 | 2022 | 29 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16178190&pc=QCB | 2023826-2023-00157 | 2022-12-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -09.00/+0.50/087 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6)2022. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. THE CAUSE OF THE EVENT WAS UNKNOWN. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: B1 - CORRECTED TO ADVERSE EVENT IN INITIAL MDR. B2 - CORRECTED TO REQUIRED INTERVENTION TO PREVENT IMPAIRMENT/DAMAGE (DEVICES) IN INITIAL MDR. D4 - UNIQUE IDENTIFIER (UDI)# : (B)(4), SHOULD BE ADDED TO THE INITAL MDR. B5 - "ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SAME LENGTH/ MODEL LENS DUE TO REFRACTIVE SURPRISE. THE PROBLEM IS RESOLVED." SHOULD BE ADDED TO THE INITIAL MDR. D6B - EXPLANT DATE (B)(6) 2023 SHOULD BE ADDED TO INITIAL MDR. H1 - CORRECTED TO SERIOUS INJURY IN INITIAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 4629 SHOULD BE ADDED TO THE INITAL MDR. H6 - HEALTH EFFECT - IMPACT CODE: 2199 SHOULD BE REMOVED FROM THE INITAL MDR. CLAIM # (B)(4). | -09.00/+0.50/087 | OD | Physician | 12/20/2022 | 2022 | 47 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16178211&pc=QCB | 2023826-2023-00158 | 2022-10-18 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.00/+2.0/073 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE LENS REMAINED IMPLANTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NONE HAS BEEN FORTHCOMING. Manufacturer Narrative: PATIENT WEIGHT: UNK. PATIENT ETHNICITY: UNK. PATIENT RACE: UNK. EXPLANT DATE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -09.00/+2.0/073 | OD | Physician | 12/20/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16178322&pc=QCB | 2023826-2023-00162 | 2022-11-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -12.00/+2.0/095 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING AND LENS ROTATION. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: D4 - UNIQUE IDENTIFIER (UDI)# (B)(4), SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: (B)(4). | -12.00/+2.0/095 | OD | Physician | 12/19/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16178334&pc=QCB | 2023826-2023-00163 | 2022-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50/+1.0/088 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: D4 - UNIQUE IDENTIFIER (UDI)# (B)(4), SHOULD BE ADDED TO THE INITIAL MDR. CLAIM # (B)(4). Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -09.50/+1.0/088 | OS | Physician | 12/23/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16178430&pc=QCB | 2023826-2023-00164 | 2022-12-15 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing; Off-Label Use | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/105 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO PATIENT EXPERIENCING A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE: 23-MAR-2023. H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE/DEBRIS ON THE LENS. VISUAL INSPECTION FOUND THE HAPTIC BENT WITH RESIDUE/DEBRIS ON THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -10.00/+1.0/105 | OD | Physician | 12/23/2022 | 2022 | 19 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16178463&pc=QCB | 2023826-2023-00160 | 2022-11-30 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -06.50/+4.5/080 (SPHERE/CYLINDER/AXIS), WAS TORN DURING LOADING (THE LENS SNAGGED WHILE LOADING THE INJECTOR) AND THERE WAS NO PATIENT CONTACT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT WEIGHT: UNK. PATIENT ETHNICITY: UNK. PATIENT RACE: UNK. (B)(4). WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # : (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: "H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE." CLAIM# (B)(4). | -06.50/+4.5/080 | NA | Physician | 12/22/2022 | 2022 | 49 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16178483&pc=QCB | 2023826-2023-00161 | 2022-11-30 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-16 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS, -04.50/+2.0/164 (SPHERE/CYLINDER/AXIS), WAS TORN DURING LOADING (THE LENS SNAGGED WHILE LOADING THE INJECTOR) AND THERE WAS NO PATIENT CONTACT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: "H6 - TYPE OF INVESTIGATION 3331 - DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE." CLAIM#(B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. DEVICE CODE: (B)(4) - THIS LENS MODEL IS CONTRAINDICATED FOR PATIENTS WITH AGE MORE THAN THAT OF 45 YEARS. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # 727683 | -04.50/+2.0/164 | NA | Physician | 12/22/2022 | 2022 | 49 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16189409&pc=QCB | 2023826-2023-00201 | 2023-01-08 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -9.5/2.0/073 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2023. THE LENS TORE DURING INJECTION/DELIVERY INTO THE EYE. INTRAOPERATIVELY THE LENS WAS REMOVED. THERE WAS NO PATIENT INJURY. IT WAS INDICATED THAT "PATIENTS NEED TO WAIT UNTIL THEY HAVE TIME BEFORE CHOOSING TO RE-IMPLANT". CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -9.5/2.0/073 | OS | Physician | 01/09/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16190016&pc=QCB | 2023826-2023-00187 | 2022-11-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.0/+4.0/090 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -7.0/+4.0/090 | OS | Physician | 12/20/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16190017&pc=QCB | 2023826-2023-00186 | 2022-11-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-17 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICM5_12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+4.50/089 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/+4.50/089 | OD | Physician | 12/20/2022 | 2022 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16196247&pc=QCB | 2023826-2023-00029 | 2022-12-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.0/1.0/087 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH SPHERICAL LENS DUE TO LOW VAULT WITH ROTATION ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -18.0/1.0/087 | OS | Physician | 12/27/2022 | 2022 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16196250&pc=QCB | 2023826-2023-00031 | 2022-11-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -8.5/6.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -8.5/6.0/086 | OS | Physician | 12/26/2022 | 2022 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16196253&pc=QCB | 2023826-2023-00030 | 2022-11-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -18.0/1.5/112 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH SPHERICAL LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -18.0/1.5/112 | OS | Physician | 12/26/2022 | 2022 | 39 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16196262&pc=QCB | 2023826-2023-00043 | 2022-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.0/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO EXCESSIVE VAULT. ON (B)(6) 2022 A SHORTER LENGTH LENS WAS IMPLANTED. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.0/2.0/093 | OS | Physician | 12/30/2022 | 2022 | 20 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16196265&pc=QCB | 2023826-2023-00036 | 2022-10-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -6.5/2.0/099 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED VERTICALLY WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -6.5/2.0/099 | OD | Physician | 12/29/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16196267&pc=QCB | 2023826-2023-00042 | 2022-07-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -15.5/2.5/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. (B)(4). CLAIM# (B)(4). | -15.5/2.5/090 | OD | Physician | 12/30/2022 | 2022 | 20 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16196269&pc=QCB | 2023826-2023-00037 | 2022-11-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -13.0/1.5/074 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -13.0/1.5/074 | OS | Physician | 12/29/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16196270&pc=QCB | 2023826-2023-00039 | 2022-11-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -13.0/1.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -13.0/1.0/090 | OS | Physician | 12/29/2022 | 2022 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16196273&pc=QCB | 2023826-2023-00038 | 2022-11-29 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -11.0/2.5/069 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNK. ETHNICITY - UNK. RACE - UNK. CLAIM# (B)(4). | -11.0/2.5/069 | OS | Physician | 12/29/2022 | 2022 | 34 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16196274&pc=QCB | 2023826-2023-00040 | 2022-10-17 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -11.0/2.0/086 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). CLAIM # (B)(4). | -11.0/2.0/086 | OS | Physician | 12/29/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16196275&pc=QCB | 2023826-2023-00044 | 2022-10-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -6.0/1.5/092 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -6.0/1.5/092 | OS | Physician | 12/29/2022 | 2022 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16199835&pc=QCB | 2023826-2023-00170 | 2022-11-13 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -06.50/+1.0/171 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH A DIFFERENT MODEL/DIOPTER BUT SAME LENGTH LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: ADDITIONAL INFORMATION: D9: LENS RETURNED 16-FEB-2023. H3: DEVICE EVALUATION: THE LENS WAS RETURNED WITH MOISTURE IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO THE LENS. DIMENSIONAL AND FUNCTIONAL INSPECTIONS FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -06.50/+1.0/171 | OD | Physician | 12/29/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16199869&pc=QCB | 2023826-2023-00168 | 2022-12-21 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -10.00/+1.0/105 (SPHERE/CYLINDER/AXIS), INTO THE PATIENTS RIGHT EYE (OD) ON (B)(6)2022. THE SURGEON NOTED THE LENS FELT THICK, WHICH CAUSED LOADING AND MANIPULATING DIFFICULTIES. THE LENS REMAINED IMPLANTED. THE POST-OP VISUAL ACUITY AND VAULT ARE GOOD. CAUSE OF THE EVENT WAS THE DEVICE. Manufacturer Narrative: WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CLAIM# (B)(4). Manufacturer Narrative: H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER, INCLUDING THE SUSPECTED DEVICE, HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM # (B)(4). | -10.00/+1.0/105 | OD | Physician | 12/26/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16200096&pc=QCB | 2023826-2023-00233 | 2022-07-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -14.5/1.0/003 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A SAME LENGTH BUT IMPLANTED VERTICALLY. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -14.5/1.0/003 | OS | Physician | 01/12/2023 | 2023 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16200178&pc=QCB | 2023826-2023-00137 | 2022-11-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -10.5/2.0/179 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS MODEL OF THE SAME SIZE. LRI PERFORMED FOR ASTIGMATISM. THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: INVESTIGATION TYPE 4110: LENS WORK ORDER SEARCH- NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -10.5/2.0/179 | OS | Physician | 12/19/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16200367&pc=QCB | 2023826-2023-00110 | 2022-12-19 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-18 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF -7.0/1.0/096 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. EXCESSIVE VAULTING WAS OBSERVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: B5 - LENS WAS REPLACED ON (B)(6) 2023 WITH A SHORTER LENGTH LENS AND PROBLEM WAS RESOLVED. PATIENT EYE IS QUIET AND GOOD. CLAIM# (B)(4). Manufacturer Narrative: PT INFO.:UNK. WORK ORDER SEARCH FOUND NO SIMILAR COMPLAINTS. (B)(4). | -7.0/1.0/096 | OS | Physician | 12/19/2022 | 2022 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16206773&pc=QCB | 2023826-2023-00112 | 2022-12-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.5/1.5/93 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A DIFFERENT MODEL, SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "ALL FINE," AND "PATIENT IS HAPPY!". Manufacturer Narrative: PATIENT INFORMATION:UNK. (B)(4). | -9.5/1.5/93 | OD | Physician | 12/20/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16211325&pc=QCB | 2023826-2023-00245 |  | Injury | STAAR SURGICAL COMPANY | 2023-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6; -14.5/+2.5/88 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED LOW VAULTING WITH ROTATION. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH A LONGER LENGTH LENS AND THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS "OTHER" AND NOTED "ICL SMALL SIZE?". Manufacturer Narrative: (B)(4). CLAIM#: (B)(4). | -14.5/+2.5/88 | OS | Physician | 01/17/2023 | 2023 |  | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16212577&pc=QCB | 2023826-2023-00248 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-01-19 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6; -7.50/1.5/113 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS HAD TORN AND COULD NOT BE USED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: CORRECTED DATA: B1: ADVERSE EVENT ADDED AND PRODUCT PROBLEM OMITTED FOR CORRECTION. B2: REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE ADDED FOR CORRECTION. H1: SERIOUS INJURY ADDED AND MALFUNCTION OMITTED FOR CORRECTION. H6: 4627 ADDED AND 2199 OMITTED FOR CORRECTION. ADDITIONAL INFORMATION: D9: 10-JAN-2023. G3: 02/24/2023. H3: DEVICE EVALUATION: RETURNED PRODUCT EVALUATION FOUND THE LENS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND THE HAPTIC TORN. CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -7.50/1.5/113 | NA | Physician | 12/27/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16214807&pc=QCB | 2023826-2023-00113 | 2022-11-16 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -7.5/2.0/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS INTRAOPERATIVELY IMPLANTED AND REMOVED. REPORTEDLY, "OVERSIZED LENS." ON (B)(6) 2022, A DIFFERENT MODEL, SMALLER LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: PATIENT IDENTIFIER: UNK. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -7.5/2.0/088 | OD | Physician | 12/22/2022 | 2022 | 35 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16216469&pc=QCB | 2023826-2023-00236 | 2022-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -9.5/3.0/093 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED AND THE A SAME SIZE LENS WAS IMPLANTED VERTICALLY. REPORTEDLY, "THE SECOND LENS WAS IMPLANTED VERTICALLY." THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: H6- INVESTIGATION TYPE CODE: 4110- LENS WORK ORDER SEARCH-NO SIMILAR COMPLAINT EVENT(S) WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT AND FIBER ON THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# : (B)(4). Manufacturer Narrative: PATIENT INFO: UNK. CLAIM# (B)(4). | -9.5/3.0/093 | OD | Physician | 12/22/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16218347&pc=QCB | 2023826-2023-00250 | 2023-01-12 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -11.5/+2.5/130 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2023. THE SURGEON REPORTS EXCESSIVE VAULTING. IT WAS NOTED "NO EXCHANGE IS PLANNED YET; THE PATIENT IS CLOSELY MONITORED". THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. LENS REMAINS IMPLANTED. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -11.5/+2.5/130 | OS | Physician | 01/19/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16218361&pc=QCB | 2023826-2023-00249 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6; -7.50/1.5/113 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S EYE. IT WAS REPORTED THAT THE LENS FLIPPED WHEN INSERTED INTO THE PATIENT'S EYE. IT WAS REMOVED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). Manufacturer Narrative: CORRECTED DATA: B1: SHOULD BE CORRECTED TO ADVERSE EVENT. B2: SHOULD BE CORRECTED TO REQUIRED INTERVENTION TO PREVENT PERMANENT IMPAIRMENT/DAMAGE. H1: SHOULD BE CORRECTED TO SERIOUS INJURY. H6: CODES 2199 AND 3189 SHOULD BE OMITTED FOR CORRECTION. H6: CODEE 4627 SHOULD BE ADDED FOR CORRECTION. CLAIM# (B)(4). | -7.50/1.5/113 | NA | Physician | 12/27/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16219630&pc=QCB | 2023826-2023-00197 | 2022-10-28 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.5/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -8.5/+1.5/091 | OD | Physician | 12/23/2022 | 2022 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16219631&pc=QCB | 2023826-2023-00194 | 2022-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.0/109 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.5/+1.0/109 | OD | Physician | 12/23/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16219633&pc=QCB | 2023826-2023-00198 | 2022-12-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.5/074 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: H6-DEVICE CODE 1494: OFF-LABEL USE (ACD < 3.0MM) SHOULD BE ADDED TO THE INITIAL MDR. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -11.5/+1.5/074 | OS | Physician | 12/23/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16219635&pc=QCB | 2023826-2023-00191 | 2022-12-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.5/113 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.0/+1.5/113 | OS | Physician | 12/22/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16219636&pc=QCB | 2023826-2023-00192 | 2022-12-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-20 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+0.5/092 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -9.5/+0.5/092 | OD | Physician | 12/22/2022 | 2022 | 31 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16230096&pc=QCB | 2023826-2023-00273 | 2023-01-09 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-23 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1; -16.0/2.0/112 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2023. THE LENS HAD TORE DURING INJECTION/DELIVERY INTO THE EYE; INTRAOPERATIVELY THE LENS WAS REMOVED. THERE WAS NO PATIENT INJURY. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -16.0/2.0/112 | OD | Physician | 01/19/2023 | 2023 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16234186&pc=QCB | 2023826-2023-00210 | 2022-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -5.5/2.0/093 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION: UNK. (B)(4). | -5.5/2.0/093 | OD | Physician | 01/01/2023 | 2023 | 29 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16234191&pc=QCB | 2023826-2023-00213 | 2022-12-07 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -9.0/1.0/108 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK, ETHNICITY- UNK, RACE- UNK. CLAIM# (B)(4). | -9.0/1.0/108 | OS | Physician | 01/01/2023 | 2023 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16234196&pc=QCB | 2023826-2023-00211 | 2022-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED VERTICALLY A 12.6MM VTICMO12.6, -16.0/3.0/177 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED HORIZONTALLY WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION: UNK. (B)(4). | -16.0/3.0/177 | OD | Physician | 01/01/2023 | 2023 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16234197&pc=QCB | 2023826-2023-00212 | 2022-10-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED VERTICALLY A 12.6MM VTICMO12.6, -16.0/3.0/178 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED HORIZONTALLY WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT : UNK. ETHNICITY : UNK. RACE : UNK. CLAIM # (B)(4). | -16.0/3.0/178 | OS | Physician | 01/01/2023 | 2023 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16234198&pc=QCB | 2023826-2023-00216 | 2022-05-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -11.0/1.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION : UNK. (B)(4). | -11.0/1.0/095 | OD | Physician | 01/01/2023 | 2023 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16234201&pc=QCB | 2023826-2023-00227 | 2022-12-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -12.5/1.0/174 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT - UNKNOWN. ETHNICITY - UNKNOWN. RACE - UNKNOWN. (B)(4). | -12.5/1.0/174 | OD | Physician | 01/06/2023 | 2023 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16234202&pc=QCB | 2023826-2023-00215 | 2022-09-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED VERTICALLY A 12.6MM VTICMO12.6, -13.0/1.0/010 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT : UNK. ETHNICITY : UNK. RACE : UNK .(B)(4). CLAIM # (B)(4). | -13.0/1.0/010 | OD | Physician | 01/01/2023 | 2023 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16234203&pc=QCB | 2023826-2023-00226 | 2022-08-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6, -9.50/2.5/062 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION: UNK. (B)(4). | -9.50/2.5/062 | OS | Physician | 01/05/2023 | 2023 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16234215&pc=QCB | 2023826-2023-00230 | 2022-10-01 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2, -10.5/1.5/083 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION: UNK. (B)(4). | -10.5/1.5/083 | OS | Physician | 01/05/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16234218&pc=QCB | 2023826-2023-00222 | 2022-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.7MM VTICM5 13.7, -11.50/4.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT IN A SECOND SURGERY ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PT INFO: UNK. CLAIM# (B)(4). | -11.50/4.0/089 | OD | Physician | 01/05/2023 | 2023 | 33 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16234222&pc=QCB | 2023826-2023-00223 | 2022-12-26 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5 12.6, -11.0/2.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. Manufacturer Narrative: PATIENT INFORMATION- UNK. (B)(4). | -11.0/2.5/091 | OS | Physician | 01/05/2023 | 2023 | 26 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16234231&pc=QCB | 2023826-2023-00228 | 2022-11-18 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1, -10.0/2.5/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6)2022 . THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON(B)(6)2022 . THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE PROBLEM CODE: (B)(4). CLAIM #(B)(4). | -10.0/2.5/089 | OD | Physician | 01/05/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16237329&pc=QCB | 2023826-2023-00307 | 2022-11-20 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: HOLD FOR RH 1.24 THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -13.5/+3.5/172 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2023 THE LENS WAS EXCHANGED WITH A SHORTER LENGTH LENS. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS A PATIENT RELATED FACTOR AND UNKNOWN AND NOTED "ANATOMICALLY NARROW SULCUS". ADDITIONAL IT WAS UNKNOWN IF THE DEVICE FAILED TO PERFORM AS INTENDED. Manufacturer Narrative: MEDICAL DEVICE PROBLEM CODE: 1494. CLAIM#: (B)(4). | -13.5/+3.5/172 | OD | Physician | 01/19/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16237499&pc=QCB | 2023826-2023-00195 | 2022-12-02 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.0/072 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -11.5/+1.0/072 | OS | Physician | 12/23/2022 | 2022 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16237759&pc=QCB | 2023826-2023-00239 | 2022-12-02 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM513.2 IMPLANTABLE COLLAMER LENS OF A -13.5/1.5/87 (SPHERE/CYLINDER/AXIS), TORE/BROKE DURING INJECTION/DELIVERY INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS INTRAOPERATIVELY IMPLANTED AND REMOVED. THERE WERE NO PATIENT INJURIES. CAUSE IS REPORTED AS USER ERROR. ON 22-DEC-2022, THE REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: WEIGHT, ETHNICITY, AND RACE:UNK. CLAIM# (B)(4). | -13.5/1.5/87 | OS | Physician | 12/27/2022 | 2022 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16238118&pc=QCB | 2023826-2023-00243 | 2022-11-29 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-24 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -17.5/1.0/088 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS INTRAOPERATIVELY IMPLANTED AND REMOVED DUE TO A LENS TEAR/BREAK DURING INJECTION/DELIVERY INTO THE EYE. THERE WAS PATIENT CONTACT, BUT NO PATIENT INJURY. THE PROBLEM WAS RESOLVED. CAUSE OF THE EVENT IS UNKNOWN. REPORTEDLY, "THE SECOND PIECE; ELECTIVE SURGERY." Manufacturer Narrative: PATIENT IDENTIFIER: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -17.5/1.0/088 | OS | Physician | 12/29/2022 | 2022 | 25 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16247113&pc=QCB | 2023826-2023-00295 | 2022-12-24 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.0/+1.5/093 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.0/+1.5/093 | OD | Physician | 12/27/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16247114&pc=QCB | 2023826-2023-00293 | 2022-11-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+3.5/178 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+3.5/178 | OS | Physician | 12/27/2022 | 2022 | 35
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16247115&pc=QCB | 2023826-2023-00294 | 2022-11-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -17.0/+2.5/109 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -17.0/+2.5/109 | OS | Physician | 12/27/2022 | 2022 | 25 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16247116&pc=QCB | 2023826-2023-00296 | 2022-10-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+2.0/098 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -11.0/+2.0/098 | OS | Physician | 12/27/2022 | 2022 | 23
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16247117&pc=QCB | 2023826-2023-00297 | 2022-10-21 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -13.5/+3.0/089 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2019. DUE TO LOW VAULT THE LENS WAS REMOVED ON (B)(6) 2022. LATER ON (B)(6) 2022 A LENS OF LONGER LENGTH WAS IMPLANTED AND THIS RESOLVED THE PROBLEM. THE REPORTER DID NOT GIVE A CAUSE FOR THIS EVENT. Manufacturer Narrative: CLAIM# (B)(4). | -13.5/+3.0/089 | OD | Physician | 12/27/2022 | 2022 | 37 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16247118&pc=QCB | 2023826-2023-00304 | 2022-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+1.5/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.5/+1.5/095 | OD | Physician | 12/29/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16247119&pc=QCB | 2023826-2023-00302 | 2022-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+1.0/112 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -9.0/+1.0/112 | OD | Physician | 12/29/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16247120&pc=QCB | 2023826-2023-00299 | 2022-10-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+1.5/092 INTO THE PATIENT'S RIGHT EYE (OD) (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.0/+1.5/092 | OD | Physician | 12/29/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16247121&pc=QCB | 2023826-2023-00300 | 2022-10-14 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+2.0/086 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM#: (B)(4). | -9.0/+2.0/086 | OS | Physician | 12/29/2022 | 2022 | 36 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16247122&pc=QCB | 2023826-2023-00301 | 2022-10-27 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.0/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -11.0/+1.0/091 | OD | Physician | 12/29/2022 | 2022 | 19 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16247123&pc=QCB | 2023826-2023-00303 | 2022-07-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+1.0/070 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.0/+1.0/070 | OS | Physician | 12/29/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16247126&pc=QCB | 2023826-2023-00305 | 2022-09-06 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+2.0/095 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5- "THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO 13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -12.0/+2.0/93 (SPHERE/CYLINDER/AXIS) INTO THE PATIENTS LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED AN EXCESSIVE VAULT. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SHORTER LENGTH AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -12.0/+2.0/095 | OS | Physician | 12/29/2022 | 2022 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16247136&pc=QCB | 2023826-2023-00306 | 2022-11-11 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -7.5/+2.5/88 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -7.5/+2.5/88 | OD | Physician | 12/29/2022 | 2022 | 39 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16248039&pc=QCB | 2023826-2023-00334 | 2022-12-17 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT ON 17-DEC-2022 A 13.2MM, VTICMO13.2, -17.5/1.5/094 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS GOT STUCK IN THE CARTRIDGE AND TORE/BROKE DURING INJECTION/DELIVERY INTO THE RIGHT EYE (OD). THE LENS WAS IMPLANTED AND REMOVED INTRAOPERATIVELY, WITH NO PATIENT INJURY. A REPLACEMENT LENS WAS IMPLANTED AT A LATER DATE AND THE PROBLEM RESOLVED. Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). | -17.5/1.5/094 | OD | Physician | 12/29/2022 | 2022 |  | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16248061&pc=QCB | 2023826-2023-00332 | 2022-12-05 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+2.0/079 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.0/+2.0/079 | OS | Physician | 12/30/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16248064&pc=QCB | 2023826-2023-00333 | 2022-12-25 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/+2.0/159 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -11.5/+2.0/159 | OD | Physician | 12/30/2022 | 2022 | 23 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16248416&pc=QCB | 2023826-2023-00335 | 2022-10-25 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS -8.0/1.0/90 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL/ LENGTH LENS THAT WAS IMPLANTED VERTICALLY DUE TO EXCESSIVE VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID WITH DEBRIS IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND AN OPTIC TORN AND FIBRE ON THE LENS SURFACE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: TYPE OF INVESTIGATION - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -8.0/1.0/90 | OS | Physician | 12/29/2022 | 2022 | 21 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16248777&pc=QCB | 2023826-2023-00336 | 2022-12-04 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS -9.5/1.5/18 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED VERTICALLY INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SAME MODEL/ LENGTH LENS THAT WAS IMPLANTED HORIZONTALLY DUE TO LOW VAULT. THE EXCHANGE RESOLVED THE PROBLEM. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -9.5/1.5/18 | OD | Physician | 12/30/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16251266&pc=QCB | 2023826-2023-00298 | 2022-11-21 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.5/+1.0/095 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SPHERICAL LENS TWO SIZES LONGER IN LENGTH DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -8.5/+1.0/095 | OD | Physician | 12/29/2022 | 2022 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16251275&pc=QCB | 2023826-2023-00241 | 2022-06-15 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -9.0/2.0/082 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SMALLER LENGTH LENS DUE TO LOW VAULT WITHOUT ROTATION. THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT IS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). Manufacturer Narrative: ORIGINALLY THE CLAIM WAS DETERMINED TO BE REPORTABLE, BUT UPON FURTHER REVIEW WAS DETERMINED TO NOT MEET DEFINITION OF A COMPLAINT. HOWEVER, THE CLAIM CANNOT BE VOIDED AS AN MDR HAS ALREADY BEEN SUBMITTED. CLAIM# (B)(4). | -9.0/2.0/082 | OS | Physician | 12/28/2022 | 2022 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16251424&pc=QCB | 2023826-2023-00321 | 2021-01-12 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5 13.2 IMPLANTABLE COLLAMER LENS OF -10.50/2.0/82 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2020. LENS ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED AND THE PROBLEM WAS NOT RESOLVED. THE LENS REMAINS IMPLANTED. Manufacturer Narrative: B5 - THE LENS WAS EXCHANGED ON (B)(6) 2023, WITH A LONGER LENGTH LENS AND PROBLEM WAS RESOLVED. REPORTEDLY, "BINOCULAR VISUS SC 20/20. ALL FINE! PATIENT IS HAPPY!" CLAIM# (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND HAPTIC BROKEN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -10.50/2.0/82 | OS | Physician | 12/27/2022 | 2022 | 22 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16251449&pc=QCB | 2023826-2023-00322 | 2022-06-18 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-26 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICM5 12.6 IMPLANTABLE COLLAMER LENS OF -7.00/3.00/089 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. LENS ROTATION WAS OBSERVED. THE LENS WAS REPOSITIONED AND THE PROBLEM WAS NOT RESOLVED. ON (B)(6) 2022 THE LENS WAS EXCHANGED FOR A SPHERICAL LENS AND THE PROBLEM WAS RESOLVED. Manufacturer Narrative: H3: DEVICE EVALUATION LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: CLAIM# (B)(4). | -7.00/3.00/089 | OS | Physician | 12/27/2022 | 2022 | 30 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16254747&pc=QCB | 2023826-2023-00174 |  | Malfunction | STAAR SURGICAL COMPANY | 2023-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS, -12.50/+1.5/090 (SPHERE/CYLINDER/AXIS), WAS TORN AND THERE WAS NO PATIENT CONTACT. ANOTHER SAME MODEL/LENGTH LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. THE PATIENT IS DOING WELL. CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: CORRECTED DATA: B5: THE REPORTER INDICATED THAT A 12.6MM VTICM5_12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.5/1.5/090 (SPHERE/CYLINDER/AXIS) TORE. THERE WAS NO PATIENT CONTACT. A BACKUP LENS OF THE SAME MODEL/LENGTH WAS IMPLANTED INTO THE LEFT EYE (OS) AND THE PROBLEM WAS RESOLVED. REPORTEDLY "PATIENT IS DOING WELL". THE REPORTER INDICATED THE CAUSE OF THE EVENT IS UNKNOWN. CLAIM#: (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: UNK. DATE OF EVENT: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -12.50/+1.5/090 | OS | Other Health Care Professional | 01/04/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16254788&pc=QCB | 2023826-2023-00180 | 2022-08-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -13.00/+2.0/004 (SPHERE/CYLINDER/AXIS), (IMPLANTED VERTICALLY) IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS (IMPLANTED HORIZONTALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND OPTIC TORN. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM # (B)(4). | -13.00/+2.0/004 | OD | Physician | 01/01/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16254797&pc=QCB | 2023826-2023-00179 | 2022-12-06 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -16.00/+4.0/087 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY. VISUAL INSPECTION FOUND HAPTIC BENT/DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. FUNCTIONAL INSPECTION FOUND THAT THE RETURNED LENS DID NOT MEET ORIGINAL VALUES MEASURED AT THE TIME OF MANUFACTURING. H6: DEVICE HISTORY RECORD (DHR) REVIEW: BASED ON THE RESULTS OF THE INVESTIGATION, ALL RELEASED DEVICES FROM THE ASSOCIATED WORK ORDER(S), INCLUDING THE SUSPECTED DEVICE; HAVE BEEN MANUFACTURED WITHIN ESTABLISHED PROCESS PARAMETERS; AND THERE IS NO INDICATION THAT THE MANUFACTURING AND PROCESSING OF THE DEVICE CONTRIBUTED TO THE COMPLAINT ISSUE. CLAIM# (B)(4). Manufacturer Narrative: PATIENT WEIGHT: UNK. PATIENT ETHNICITY: UNK. PATIENT RACE: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM #: (B)(4). | -16.00/+4.0/087 | OS | Physician | 01/04/2023 | 2023 | 24 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16254806&pc=QCB | 2023826-2023-00178 | 2022-09-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -14.00/+2.0/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS REMOVED LATER THAT SAME DAY DUE TO EXCESSIVE VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH LENS (IMPLANTED VERTICALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. (B)(4). CLAIM # (B)(4). | -14.00/+2.0/090 | OS | Physician | 01/01/2023 | 2023 | 18 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16254886&pc=QCB | 2023826-2023-00177 | 2022-08-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -16.00/+1.5/090 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE. THE LENS WAS REPLACED WITH A SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM # (B)(4). | -16.00/+1.5/090 | OD | Physician | 01/01/2023 | 2023 | 32 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16254887&pc=QCB | 2023826-2023-00176 | 2022-08-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Misfocusing | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -18.00/+2.0/073 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO THE PATIENT EXPERIENCED A REFRACTIVE SURPRISE AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). | -18.00/+2.0/073 | OD | Physician | 01/01/2023 | 2023 | 33 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16254919&pc=QCB | 2023826-2023-00185 | 2022-09-08 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.7MM VTICMO13.7 IMPLANTABLE COLLAMER LENS, -09.50/+2.0/178 (SPHERE/CYLINDER/AXIS) (IMPLANTED VERTICALLY), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING. THIS WAS THE REPLACEMENT LENS FOR MFR. REPORT # 2023826-2023-00184. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL DRY WITH RESIDUE/DEBRIS ON PRODUCT AND RETURNED IN PIECES. VISUAL INSPECTION FOUND OPTIC SPLIT AND HAPTIC BROKEN. UNABLE TO IDENTIFY LENS OR COMPLETE DIMENSIONAL INSPECTION DUE TO DAMAGED STATE OF LENS. CLAIM# (B)(4). Manufacturer Narrative: WEIGHT: UNK; ETHNICITY: UNK; RACE: UNK. (B)(4). | -09.50/+2.0/178 | OS | Physician | 01/05/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16254933&pc=QCB | 2023826-2023-00184 | 2022-07-26 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS, -09.50/+2.0/089 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS (IMPLANTED VERTICALLY) BUT THE PROBLEM WAS NOT RESOLVED. SEE MFR. REPORT # 2023826-2023-00185 FOR REPLACEMENT LENS. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: CLAIM#: (B)(4). Manufacturer Narrative: (B)(4). | -09.50/+2.0/089 | OS | Physician | 01/05/2023 | 2023 | 36 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16255992&pc=QCB | 2023826-2023-00323 |  | Injury | STAAR SURGICAL COMPANY | 2023-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -16.0/1.0/090 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S EYE ON AN UNKNOWN DATE. ON AN UNKNOWN DATE THE LENS WAS IMPLANTED AND REMOVED DUE TO THE LENS FLIPPED IN THE EYE. A BACKUP LENS WAS IMPLANTED. Manufacturer Narrative: (B)(4). | -16.0/1.0/090 | NA | Other Health Care Professional | 12/28/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16257852&pc=QCB | 2023826-2023-00181 | 2022-08-12 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS, -11.50/+1.0/177 (SPHERE/CYLINDER/AXIS), (IMPLANTED VERTICALLY) IN THE PATIENTS LEFT EYE (OS), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH ANOTHER SAME MODEL/LENGTH BUT DIFFERENT DIOPTER LENS (IMPLANTED HORIZONTALLY) AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL WITH MOISTURE AND RESIDUE ON PRODUCT. VISUAL INSPECTION FOUND NO VISIBLE DAMAGE TO LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). | -11.50/+1.0/177 | OS | Physician | 01/01/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16257913&pc=QCB | 2023826-2023-00325 | 2022-11-16 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-27 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Appropriate Term/Code Not Available | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO 12.6 IMPLANTABLE COLLAMER LENS OF -16.0/3.5/110 (SPHERE/CYLINDER/AXIS) TORE/ BROKE DURING INJECTION/ DELIVERY INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS EXCHANGED FOR A SAME MODEL AND SIZE LENS AND THE PROBLEM WAS RESOLVED. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -16.0/3.5/110 | OS | Other Health Care Professional | 12/29/2022 | 2022 | 27 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16269214&pc=QCB | 2023826-2023-00356 | 2022-09-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+1.5/091 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE EVALUATION: H3 - TYPE OF INVESTIGATION CODE: 10- DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL WITH RESIDUE ON THE PRODUCT. VISUAL INSPECTION FOUND HAPTIC DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -8.0/+1.5/091 | OD | Physician | 01/01/2023 | 2023 | 43 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16269218&pc=QCB | 2023826-2023-00358 | 2022-09-30 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.0/+2.0/092 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -9.0/+2.0/092 | OS | Physician | 01/01/2023 | 2023 | 37 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16269220&pc=QCB | 2023826-2023-00357 | 2022-09-20 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -8.0/+1.5/089 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT. THIS RESOLVED THE PROBLEM. CAUSE REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -8.0/+1.5/089 | OS | Physician | 01/01/2023 | 2023 | 43 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16269221&pc=QCB | 2023826-2023-00352 | 2022-09-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -9.5/+3.0/110 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -9.5/+3.0/110 | OD | Physician | 01/01/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16269230&pc=QCB | 2023826-2023-00353 | 2022-09-29 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size; Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 IMPLANTABLE COLLAMER LENS OF DIOPTER -10.5/+3.5/071 INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A LONGER LENGTH LENS DUE TO LOW VAULT AND LENS ROTATION. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM# (B)(4). | -10.5/+3.5/071 | OS | Physician | 01/01/2023 | 2023 | 21 | Male
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16269231&pc=QCB | 2023826-2023-00355 | 2022-01-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF DIOPTER -11.0/+1.0/170 INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS EXCHANGED ON (B)(6) 2022 FOR A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT. THIS RESOLVED THE PROBLEM. CAUSE OF EVENT REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -11.0/+1.0/170 | OD | Physician | 01/01/2023 | 2023 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16269397&pc=QCB | 2023826-2023-00367 | 2021-07-26 04:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS OF -11.0/0.5/087 (SPHERE/CYLINDER/AXIS) TORE DURING LOADING AFTER OPENING THE VIAL. THE LENS WAS NOT IMPLANTED. CAUSE OF THE EVENT WAS THE DEVICE. REPORTEDLY, "A SMALL SCRATCH/CRACK WAS NOTED ON THE LENS UNDER MICROSCOPE WHEN ATTEMPTED TO LOAD THE LENS INTO THE CARTRIDGE." ON (B)(6) 2021, THE REPLACEMENT LENS WAS IMPLANTED AND THE PROBLEM WAS RESOLVED. STATUS OF THE EYE IS REPORTED AS, "ANTERIOR CHAMBER CELL/PIGMENTS 1+." Manufacturer Narrative: CORRECTED DATA: H6-MEDICAL DEVICE PROBLEM CODE: 1494 OFF-LABEL USE (ACD<3.0MM) SHOULD BE REMOVED. CLAIM# (B)(4). Manufacturer Narrative: PATIENT IDENTIFIER: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINTS WITHIN ASSOCIATED LOTS WERE FOUND. CLAIM# (B)(4). | -11.0/0.5/087 | NA | Physician | 11/26/2022 | 2022 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16269429&pc=QCB | 2023826-2023-00313 | 2022-10-10 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-30 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICMO13.2 IMPLANTABLE COLLAMER LENS OF -4.0/1.0/092 (SPHERE/CYLINDER/AXIS) WAS IMPLANTED INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. ON (B)(6) 2022, THE LENS WAS REMOVED DUE TO EXCESSIVE VAULT. REPORTEDLY, "THE SECOND LENS" ELECTIVE SURGERY." THE PROBLEM WAS RESOLVED. CAUSE IS REPORTED AS UNKNOWN. Manufacturer Narrative: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND HAPTIC BENT TO THE LENS. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -4.0/1.0/092 | OS | Physician | 01/01/2023 | 2023 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16273247&pc=QCB | 2023826-2023-00276 | 2022-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.5/1.5/105 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: PATIENT INFORMATION: UNK. CLAIM#: (B)(4). | -8.5/1.5/105 | OS | Physician | 01/09/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16273248&pc=QCB | 2023826-2023-00275 | 2021-09-09 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -9.0/1.5/073 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2021. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. (B)(4). CLAIM# (B)(4). | -9.0/1.5/073 | OS | Physician | 01/09/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16273254&pc=QCB | 2023826-2023-00278 | 2022-09-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.0/2.5/068 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 . THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022 . THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -7.0/2.5/068 | OD | Physician | 01/09/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16273255&pc=QCB | 2023826-2023-00279 | 2022-09-23 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -5.0/2.0/095 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022 . THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022 . THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). | -5.0/2.0/095 | OS | Physician | 01/09/2023 | 2023 | 30 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16273259&pc=QCB | 2023826-2023-00282 | 2022-11-03 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -8.5/1.0/089 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: CLAIM # (B)(4). | -8.5/1.0/089 | OS | Physician | 01/10/2023 | 2023 | 32 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16273267&pc=QCB | 2023826-2023-00285 | 2022-11-14 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -10.0/0.5/085 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT: UNK. ETHNICITY: UNK. RACE: UNK. CLAIM # (B)(4). | -10.0/0.5/085 | OD | Physician | 01/12/2023 | 2023 | 40 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16273272&pc=QCB | 2023826-2023-00288 | 2022-07-07 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -18.0/3.5/091 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022 . THE LENS WAS REPLACED WITH A SHORTER LENGTH LENS DUE TO EXCESSIVE VAULT ON (B)(6) 2022 . THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM# (B)(4). | -18.0/3.5/091 | OD | Physician | 01/12/2023 | 2023 | 27 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16273617&pc=QCB | 2023826-2023-00277 | 2022-08-05 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.6MM VTICMO12.6, -7.5/2.0/063 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: DEVICE PROBLEM CODE: (B)(4). CLAIM # (B)(4). | -7.5/2.0/063 | OD | Physician | 01/09/2023 | 2023 | 24 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16273630&pc=QCB | 2023826-2023-00289 | 2022-09-19 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 12.1MM VTICMO12.1 -12.5/2.0/090 (SPHERE/CYLINDER/AXIS), IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S RIGHT EYE (OD) ON (B)(6) 2022. THE LENS WAS REPLACED WITH A LONGER LENGTH LENS DUE TO LOW VAULT ON (B)(6) 2022. THIS RESOLVED THE PROBLEM. CAUSE OF THE EVENT IS REPORTED AS UNKNOWN. Manufacturer Narrative: WEIGHT- UNK. ETHNICITY- UNK. RACE- UNK. (B)(4). CLAIM# (B)(4). | -12.5/2.0/090 | OD | Physician | 01/12/2023 | 2023 | 38 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16273842&pc=QCB | 2023826-2023-00369 |  | Injury | STAAR SURGICAL COMPANY | 2023-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Device Dislodged or Dislocated | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -7.50/4.0/084 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE SURGEON REPORTS LENS ROTATION WITH A PLANNED REPLACEMENT. LENS REMAINS IMPLANTED. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION: B5: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICM5_13.2; -7.50/4.0/084 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6)2022. REPOSITIONING WAS ATTEMPTED ON (B)(6) 2022 AND DID NOT RESOLVE THE PROBLEM. ON (B)(6)2023 THE LENS WAS EXPLANTED. IN A SEPARATE SURGERY LATER THAT DAY A LONGER LENGTH LENS OF THE SAME POWER WAS IMPLANTED. ADDITIONAL INFORMATION STATES THE PATIENT RUBBED EYE ON (B)(6) 2023 CAUSING A WOUND LEAK. THE SURGEON ATTEMPTED TO SUTURE THE WOUND AND SENT THE PATIENT TO A DIFFERENT DOCTOR FOR POSSIBLE GLUE TO CLOSE THE INCISION; SEIDEL AND HYPOTONY. CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. THE PROBLEM WAS NOT RESOLVED. A LENS WORK ORDER SEARCH WAS PERFORMED AND NO SIMILAR COMPLAINT TYPES WERE FOUND IN THE SAME LENS LOT. D9: RETURNED TO MANUFACTURER 05/01/2023. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: H3 - DEVICE EVALUATION: THE LENS WAS RETURNED IN LIQUID IN A VIAL. VISUAL INSPECTION FOUND THE OPTIC TORN. CLAIM# (B)(4). Manufacturer Narrative: LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM# (B)(4). | -7.50/4.0/084 | OS | Physician | 01/03/2022 | 2022
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16273846&pc=QCB | 2023826-2023-00368 | 2022-12-10 05:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | No Apparent Adverse Event | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT THE SURGEON IMPLANTED A 13.2MM VTICMO13.2; -10.0/2.5/097 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS INTO THE PATIENT'S LEFT EYE (OS) ON (B)(6) 2022. THE PATIENT EXPERIENCED EXCESSIVE VAULTING. ON (B)(6) 2022 THE LENS WAS EXCHANGED WITH THE SAME LENGTH LENS BUT IMPLANTED VERTICALLY; AS OPPOSED TO THE INITIAL LENS BEING IMPLANTED HORIZONTALLY. THIS RESOLVED THE PROBLEM. THE CAUSE OF THE EVENT WAS REPORTED AS UNKNOWN. Manufacturer Narrative: TYPE OF INVESTIGATION: 4110 - LENS WORK ORDER SEARCH: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). Manufacturer Narrative: ADDITIONAL INFORMATION: D9: RETURN DATE- (B)(6) 2023. H3: DEVICE EVALUATION: THE LENS WAS RETURNED DRY IN A MICROCENTRIFUGE VIAL. VISUAL INSPECTION FOUND THE HAPTIC BENT AND DEFORMED. DIMENSIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATIONS. CLAIM# (B)(4). | -10.0/2.5/097 | OS | Physician | 01/03/2023 | 2023 | 28 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16273935&pc=QCB | 2023826-2023-00372 |  | Injury | STAAR SURGICAL COMPANY | 2023-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Adverse Event Without Identified Device or Use Problem | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THAT A 13.2MM VTICM5_13.2; -12.50/2.0/088 (SPHERE/CYLINDER/AXIS) IMPLANTABLE COLLAMER LENS HAD BEEN EXPLANTED FROM A PATIENT'S RIGHT EYE (OD). ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE NOT BEEN SUCCESSFUL. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL MEDWATCH REPORT WILL BE SUBMITTED. Manufacturer Narrative: NO SIMILAR COMPLAINT TYPE EVENTS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM#: (B)(4). | -12.50/2.0/088 | OD | Physician | 01/04/2023 | 2023
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16275250&pc=QCB | 2023826-2023-00256 | 2023-01-10 05:00 | Malfunction | STAAR SURGICAL COMPANY | 2023-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Unintended Movement | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON INSERTED AND REMOVED A 13.2MM VTICM5_13.2 IMPLANTABLE COLLAMER LENS, -11.00/+1.0/015 (SPHERE/CYLINDER/AXIS), WITHIN THE SAME SURGERY DUE TO THE LENS WAS INSERTED UPSIDE DOWN WHEN INJECTED INTO THE PATIENTS LEFT EYE (OS). THERE WAS NO PATIENT INJURY AND THE LENS WAS TORN UPON REMOVAL. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: PATIENT IDENTIFIER: UNK. WORK ORDER SEARCH: NO SIMILAR COMPLAINT WAS REPORTED FOR UNITS WITHIN THE SAME LOT. CLAIM # (B)(4). | -11.00/+1.0/015 | OS | Physician | 01/10/2023 | 2023 | 42 | Female
https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=16275289&pc=QCB | 2023826-2023-00269 | 2022-10-22 04:00 | Injury | STAAR SURGICAL COMPANY | 2023-01-31 05:00 | QCB | IMPLANTABLE COLLAMER LENS (ICL) | Off-Label Use; Inadequacy of Device Shape and/or Size | No Clinical Signs, Symptoms or Conditions | P030016 | Event Description: THE REPORTER INDICATED THE SURGEON IMPLANTED A 12.6MM VTICMO12.6 IMPLANTABLE COLLAMER LENS, -09.00/+1.5/113 (SPHERE/CYLINDER/AXIS), IN THE PATIENTS RIGHT EYE (OD), ON (B)(6) 2022. THE LENS WAS EXPLANTED ON (B)(6) 2022 DUE TO LOW VAULTING. THE LENS WAS REPLACED WITH A LONGER LENS AND THE PROBLEM WAS RESOLVED. THE CAUSE OF THE EVENT WAS UNKNOWN. Manufacturer Narrative: (B)(4). CLAIM # (B)(4). Manufacturer Narrative: (B)(4). Manufacturer Narrative: H3: DEVICE EVALUATION - LENS WAS RETURNED IN MICROCENTRIFUGE VIAL IN LIQUID. VISUAL INSPECTION FOUND HAPTIC TORN. DIMENSIONAL AND FUNCTIONAL INSPECTION FOUND THE LENS TO BE WITHIN SPECIFICATION. CLAIM# (B)(4). | -09.00/+1.5/113 | OD | Physician | 01/10/2023 | 2023 | 33 | Female